## Table_S1_S5_DescriptiveMeasure
Table S1. Invasive and non-invasive hemodynamic levels obtained during the records obtained with brachial oscillometry (Mobil-O-Graph) (n=34) | [9 empty] | Table S2. Invasive and non-invasive hemodynamic levels obtained during the records obtained with radial applanation tonometry (SphygmoCor) (n=34) | [9 empty] | Table S3. Hemodynamic data obtained from carotid artery applanation tonometry (SphygmoCor) and from aortic and brachial invasive measurements (n=34) | [9 empty] | Table S4. Hemodynamic data obtained from brachial applanation tonometry (SphygmoCor) and from aortic and brachial invasive measurements (n=34) | [9 empty] | Table S5. Hemodynamic data obtained from carotid ultrasound (SonoSite - Hemodyn-4M) and from aortic and brachial invasive measurements (n=34)
Brachial oscillometry (Mobil-O-Graph) | [9 empty] | Radial Applanation Tonometry | [9 empty] | Carotid Applanation Tonometry | [9 empty] | Brachial Applanation Tonometry | [9 empty] | Carotid Ultrasonography
Variable | MV | SEM | SD | Min | p25th | p50th | p75th | Max | Range | Variable | MV | SEM | SD | Min | p25th | p50th | p75th | Max | Range | Variable | MV | SEM | SD | Min | p25th | p50th | p75th | Max | Range | Variable | MV | SEM | SD | Min | p25th | p50th | p75th | Max | Range | Variable | MV | SEM | SD | Min | p25th | p50th | p75th | Max | Range
Invasive records | [9 empty] | Invasive records | [9 empty] | Invasive records | [9 empty] | Invasive records | [9 empty] | Invasive records
Aortic SBP Invasive [mmHg] | 131.5445 | 4.1486 | 24.1904 | 83 | 116 | 127.5069 | 151.7273 | 178.4444 | 95.4444 | Aortic SBP Invasive [mmHg] | 134.2108 | 4.1148 | 23.9935 | 82 | 117.5 | 132.9583 | 153.6667 | 177.3333 | 95.3333 | Aortic SBP Invasive [mmHg] | 130.9162 | 4.368 | 25.0924 | 77.5 | 114.3333 | 124.5 | 151 | 189.5 | 112 | Aortic SBP Invasive [mmHg] | 133.6562 | 4.337 | 24.5335 | 77 | 120.5 | 130 | 154.25 | 179 | 102 | Aortic SBP Invasive [mmHg] | 132.4768 | 4.3751 | 25.5109 | 78.75 | 115.6 | 131.05 | 151.75 | 188 | 109.25
Aortic MBP Invasive [mmHg] | 92.8861 | 2.5506 | 14.8722 | 66.697 | 81 | 90.487 | 102.25 | 120.8889 | 54.1919 | Aortic MBP Invasive [mmHg] | 94.2451 | 2.4612 | 14.3509 | 66.6667 | 83 | 92.1667 | 99.6667 | 120.6667 | 54 | Aortic MBP Invasive [mmHg] | 92.6616 | 2.6742 | 15.3623 | 65 | 80 | 90 | 102 | 130.5 | 65.5 | Aortic MBP Invasive [mmHg] | 93.6979 | 2.6117 | 14.7741 | 65 | 85 | 90.5 | 106 | 121 | 56 | Aortic MBP Invasive [mmHg] | 93.5746 | 2.6569 | 15.4925 | 63 | 84 | 91.125 | 103.5 | 128 | 65
Aortic DBP Invasive [mmHg] | 67.6918 | 1.7749 | 10.3496 | 51.5 | 61 | 65.9667 | 76 | 93 | 41.5 | Aortic DBP Invasive [mmHg] | 68.3897 | 1.7034 | 9.9323 | 51.6667 | 62 | 66.1667 | 76.5 | 93 | 41.3333 | Aortic DBP Invasive [mmHg] | 67.5793 | 1.8322 | 10.525 | 51.6667 | 61 | 65 | 73.5 | 93 | 41.3333 | Aortic DBP Invasive [mmHg] | 67.8906 | 1.8167 | 10.2766 | 52 | 61.5 | 64.5 | 74.5 | 92 | 40 | Aortic DBP Invasive [mmHg] | 68.1016 | 1.8263 | 10.6489 | 50.75 | 61.6667 | 65.625 | 76 | 93.25 | 42.5
Aortic PP Invasive [mmHg] | 63.8527 | 3.3929 | 19.7841 | 24.4545 | 50.2222 | 59.8542 | 82.0909 | 98.3333 | 73.8788 | Aortic PP Invasive [mmHg] | 65.8211 | 3.4662 | 20.2112 | 23 | 53.25 | 62.25 | 83 | 97.6667 | 74.6667 | Aortic PP Invasive [mmHg] | 63.3369 | 3.5879 | 20.611 | 23.5 | 50.3333 | 59 | 82.6667 | 102 | 78.5 | Aortic PP Invasive [mmHg] | 65.7656 | 3.6608 | 20.7088 | 22 | 53.5 | 64 | 82 | 103 | 81 | Aortic PP Invasive [mmHg] | 64.3751 | 3.5459 | 20.6762 | 24 | 53.6 | 60.5833 | 83.3333 | 102 | 78
HR Invasive [beat/minute] | 70.4359 | 2.2892 | 12.9497 | 44.8333 | 59.5545 | 71.5195 | 80.9444 | 104.2727 | 59.4394 | HR Invasive [beat/minute] | 70.4245 | 2.2651 | 12.8135 | 45 | 60 | 71 | 78 | 104 | 59 | HR Invasive [beat/minute] | 70.0349 | 2.2192 | 12.3558 | 45 | 60.6667 | 71.25 | 79 | 104 | 59 | HR Invasive [beat/minute] | 70.1437 | 2.7821 | 14.9818 | 50 | 57.5 | 68 | 78 | 104 | 54 | HR Invasive [beat/minute] | 69.7158 | 2.2459 | 12.705 | 45 | 58.8333 | 71.1 | 78.125 | 104.25 | 59.25
Brachial SBP Invasive [mmHg] | 147.3752 | 4.5698 | 25.8504 | 83 | 132 | 146 | 167 | 191 | 108 | Brachial SBP Invasive [mmHg] | 147.8594 | 4.5877 | 25.952 | 83 | 132 | 146 | 167 | 191 | 108 | Brachial SBP Invasive [mmHg] | 147.871 | 4.776 | 26.5917 | 83 | 131 | 146 | 169 | 193 | 110 | Brachial SBP Invasive [mmHg] | 145.0286 | 4.8506 | 27.4392 | 85 | 133.25 | 143.5 | 165.6667 | 189.3333 | 104.3333 | Brachial SBP Invasive [mmHg] | 147.5156 | 4.636 | 26.2253 | 83 | 132 | 145.5 | 167 | 193 | 110
Brachial MBP Invasive [mmHg] | 100.1721 | 2.5811 | 14.6012 | 67 | 92.1429 | 101 | 111 | 123 | 56 | Brachial MBP Invasive [mmHg] | 100.5313 | 2.5708 | 14.5425 | 67 | 92 | 101.5 | 111 | 123 | 56 | Brachial MBP Invasive [mmHg] | 100.6452 | 2.7402 | 15.2568 | 67 | 93 | 101 | 111 | 131 | 64 | Brachial MBP Invasive [mmHg] | 97.7969 | 2.6118 | 14.7745 | 68 | 88.25 | 97.75 | 109.75 | 120.3333 | 52.3333 | Brachial MBP Invasive [mmHg] | 100.375 | 2.6733 | 15.1226 | 67 | 92 | 101 | 111 | 131 | 64
Brachial DBP Invasive [mmHg] | 72.2728 | 1.9221 | 10.873 | 56 | 63.1429 | 72 | 81.5 | 94 | 38 | Brachial DBP Invasive [mmHg] | 72.4375 | 1.9124 | 10.818 | 56 | 63 | 72 | 81 | 94 | 38 | Brachial DBP Invasive [mmHg] | 72.6129 | 1.9907 | 11.0836 | 56 | 65 | 72 | 82 | 94 | 38 | Brachial DBP Invasive [mmHg] | 69.7839 | 1.7427 | 9.858 | 53 | 62.75 | 68.5 | 76.8333 | 91.3333 | 38.3333 | Brachial DBP Invasive [mmHg] | 72.4062 | 1.9623 | 11.1003 | 56 | 63 | 72 | 81.5 | 94 | 38
Brachial PP Invasive [mmHg] | 75.1024 | 4.2848 | 24.2386 | 27 | 60 | 77.75 | 85.5 | 132 | 105 | Brachial PP Invasive [mmHg] | 75.4219 | 4.3054 | 24.3553 | 27 | 60 | 77.75 | 88 | 132 | 105 | Brachial PP Invasive [mmHg] | 75.2581 | 4.4451 | 24.7494 | 27 | 59 | 78 | 86 | 132 | 105 | Brachial PP Invasive [mmHg] | 75.2448 | 4.3306 | 24.4976 | 25 | 61.1667 | 75.25 | 89.25 | 136.3333 | 111.3333 | Brachial PP Invasive [mmHg] | 75.1094 | 4.2943 | 24.2925 | 27 | 60 | 77.5 | 85.5 | 132 | 105
Mobil-O-Graph Non-invasive records: sd calibration | [9 empty] | Mobil-O-Graph Non-invasive records: sd calibration | [9 empty] | Mobil-O-Graph Non-invasive records: sd calibration | [9 empty] | Mobil-O-Graph Non-invasive records: sd calibration | [9 empty] | Mobil-O-Graph Non-invasive records: sd calibration
bSBP (MOG_GTF_sd) [mmHg] | 132.8506 | 3.5336 | 20.6044 | 90.2727 | 119.0909 | 133.9545 | 147.5455 | 175.7778 | 85.5051 | bSBP (MOG_GTF_sd) [mmHg] | 134.2819 | 3.5747 | 20.844 | 87 | 127 | 133.8333 | 150 | 180 | 93 | bSBP (MOG_GTF_sd) [mmHg] | 132.598 | 3.928 | 22.5647 | 88.5 | 122.3333 | 131.3333 | 148 | 174.3333 | 85.8333 | bSBP (MOG_GTF_sd) [mmHg] | 135.651 | 3.7202 | 21.0447 | 78 | 126.1667 | 136.0833 | 150 | 166.6667 | 88.6667 | bSBP (MOG_GTF_sd) [mmHg] | 132.9127 | 3.699 | 21.5689 | 84.75 | 118 | 135.25 | 145.5 | 175 | 90.25
bMBP (MOG_GTF_sd) [mmHg] | 97.769 | 2.5509 | 14.8744 | 61.6364 | 88.5 | 100.0985 | 107.3333 | 125.7778 | 64.1414 | bMBP (MOG_GTF_sd) [mmHg] | 98.4461 | 2.5219 | 14.7049 | 58 | 90 | 101.75 | 109.3333 | 122.6667 | 64.6667 | bMBP (MOG_GTF_sd) [mmHg] | 97.099 | 2.8384 | 16.3054 | 58 | 89 | 94 | 107.5 | 128 | 70 | bMBP (MOG_GTF_sd) [mmHg] | 98.7292 | 2.783 | 15.7429 | 54 | 91.8333 | 100.1667 | 105.75 | 124.3333 | 70.3333 | bMBP (MOG_GTF_sd) [mmHg] | 97.8147 | 2.7312 | 15.9253 | 57 | 87.4 | 99.25 | 105.25 | 127 | 70
bDBP (MOG_GTF_sd) [mmHg] | 80.6807 | 2.2586 | 13.1696 | 48 | 72.1111 | 80.1528 | 89.8571 | 103.6923 | 55.6923 | bDBP (MOG_GTF_sd) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | bDBP (MOG_GTF_sd) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | bDBP (MOG_GTF_sd) [mmHg] | 80.7708 | 2.5852 | 14.6244 | 43 | 73.5 | 79.25 | 92.25 | 104.6667 | 61.6667 | bDBP (MOG_GTF_sd) [mmHg] | 80.7299 | 2.4606 | 14.3476 | 43.75 | 71 | 80.625 | 90 | 104 | 60.25
bPP (MOG_GTF_sd) [mmHg] | 52.1698 | 2.1872 | 12.7532 | 30.4 | 41.7273 | 49 | 63 | 82.4545 | 52.0545 | bPP (MOG_GTF_sd) [mmHg] | 53.3113 | 2.3952 | 13.9666 | 34 | 44 | 49.5833 | 59.5 | 86 | 52 | bPP (MOG_GTF_sd) [mmHg] | 52.7909 | 2.3644 | 13.5822 | 31.5 | 43 | 51 | 62 | 82.6667 | 51.1667 | bPP (MOG_GTF_sd) [mmHg] | 54.8802 | 2.4872 | 14.0698 | 28 | 45.1667 | 52.75 | 66 | 82.6667 | 54.6667 | bPP (MOG_GTF_sd) [mmHg] | 52.1828 | 2.2124 | 12.9004 | 31.6667 | 41.25 | 49 | 63.5 | 82 | 50.3333
HR (MOG_GTF_sd) [beat/minute] | 71.4016 | 2.2863 | 13.3314 | 42.5 | 61.5556 | 72 | 82.0588 | 105 | 62.5 | HR (MOG_GTF_sd) [beat/minute] | 71.9216 | 2.3023 | 13.4244 | 44 | 60 | 73.4167 | 79.75 | 105 | 61 | HR (MOG_GTF_sd) [beat/minute] | 71.9586 | 2.3074 | 13.2548 | 44 | 64.6667 | 72.5 | 80.5 | 105 | 61 | HR (MOG_GTF_sd) [beat/minute] | 70.5469 | 2.4989 | 14.1357 | 44 | 60.5 | 70.9167 | 81 | 104.3333 | 60.3333 | HR (MOG_GTF_sd) [beat/minute] | 71.2304 | 2.246 | 13.0964 | 45 | 63 | 72.375 | 82.75 | 105 | 60
aoSBP (MOG_GTF_sd) [mmHg] | 123.009 | 3.4583 | 20.1654 | 84.0909 | 108.5455 | 122.6234 | 135 | 168.2727 | 84.1818 | aoSBP (MOG_GTF_sd) [mmHg] | 124.1338 | 3.5984 | 20.6713 | 83 | 111 | 123.25 | 138.6667 | 170.6667 | 87.6667 | aoSBP (MOG_GTF_sd) [mmHg] | 123.7273 | 3.8591 | 22.1687 | 81 | 111.3333 | 122 | 137 | 166 | 85 | aoSBP (MOG_GTF_sd) [mmHg] | 129.9414 | 3.2996 | 17.1453 | 79 | 120.6667 | 129 | 141 | 158.3333 | 79.3333 | aoSBP (MOG_GTF_sd) [mmHg] | 123.0938 | 3.8067 | 21.5339 | 83 | 107 | 123.9583 | 136.5 | 172 | 89
aoDBP (MOG_GTF_sd) [mmHg] | 82.1371 | 2.2576 | 13.1639 | 48.3636 | 73.2222 | 81.5278 | 90.8571 | 104.3077 | 55.9441 | aoDBP (MOG_GTF_sd) [mmHg] | 82.4672 | 2.339 | 13.4366 | 44 | 73 | 84 | 92 | 104.25 | 60.25 | aoDBP (MOG_GTF_sd) [mmHg] | 81.1298 | 2.5057 | 14.3941 | 42 | 70 | 80 | 90.5 | 109 | 67 | aoDBP (MOG_GTF_sd) [mmHg] | 85.9414 | 2.3935 | 12.4368 | 57 | 76.5 | 86 | 95.5 | 106.3333 | 49.3333 | aoDBP (MOG_GTF_sd) [mmHg] | 82.1036 | 2.6074 | 14.7497 | 44 | 71.8333 | 82.5 | 92.5 | 106 | 62
aoPP (MOG_GTF_sd) [mmHg] | 40.8705 | 2.1269 | 12.4021 | 23.4545 | 31.9091 | 37.3273 | 44.6667 | 75.0909 | 51.6364 | aoPP (MOG_GTF_sd) [mmHg] | 41.6591 | 2.3756 | 13.6467 | 23.3333 | 31.75 | 38.75 | 47 | 76 | 52.6667 | aoPP (MOG_GTF_sd) [mmHg] | 42.6561 | 2.3194 | 13.3237 | 23 | 32.75 | 40 | 47.5 | 77 | 54 | aoPP (MOG_GTF_sd) [mmHg] | 43.9383 | 2.4591 | 12.7776 | 22 | 35 | 46 | 53 | 67.6667 | 45.6667 | aoPP (MOG_GTF_sd) [mmHg] | 40.9865 | 2.2992 | 13.0062 | 24 | 32.4583 | 37 | 47.25 | 78 | 54
Mobil-O-Graph Non-invasive records: 033 calibration | [9 empty] | Mobil-O-Graph Non-invasive records: 033 calibration | [9 empty] | Mobil-O-Graph Non-invasive records: 033 calibration | [9 empty] | Mobil-O-Graph Non-invasive records: 033 calibration | [9 empty] | Mobil-O-Graph Non-invasive records: 033 calibration
bSBP (MOG_GTF_033) [mmHg] | 132.8506 | 3.5336 | 20.6044 | 90.2727 | 119.0909 | 133.9545 | 147.5455 | 175.7778 | 85.5051 | bSBP (MOG_GTF_033) [mmHg] | 134.2819 | 3.5747 | 20.844 | 87 | 127 | 133.8333 | 150 | 180 | 93 | bSBP (MOG_GTF_033) [mmHg] | 132.598 | 3.928 | 22.5647 | 88.5 | 122.3333 | 131.3333 | 148 | 174.3333 | 85.8333 | bSBP (MOG_GTF_033) [mmHg] | 135.651 | 3.7202 | 21.0447 | 78 | 126.1667 | 136.0833 | 150 | 166.6667 | 88.6667 | bSBP (MOG_GTF_033) [mmHg] | 132.9127 | 3.699 | 21.5689 | 84.75 | 118 | 135.25 | 145.5 | 175 | 90.25
bMBP (MOG_GTF_033) [mmHg] | 97.769 | 2.5509 | 14.8744 | 61.6364 | 88.5 | 100.0985 | 107.3333 | 125.7778 | 64.1414 | bMBP (MOG_GTF_033) [mmHg] | 98.4461 | 2.5219 | 14.7049 | 58 | 90 | 101.75 | 109.3333 | 122.6667 | 64.6667 | bMBP (MOG_GTF_033) [mmHg] | 97.099 | 2.8384 | 16.3054 | 58 | 89 | 94 | 107.5 | 128 | 70 | bMBP (MOG_GTF_033) [mmHg] | 98.7292 | 2.783 | 15.7429 | 54 | 91.8333 | 100.1667 | 105.75 | 124.3333 | 70.3333 | bMBP (MOG_GTF_033) [mmHg] | 97.8147 | 2.7312 | 15.9253 | 57 | 87.4 | 99.25 | 105.25 | 127 | 70
bDBP (MOG_GTF_033) [mmHg] | 80.6807 | 2.2586 | 13.1696 | 48 | 72.1111 | 80.1528 | 89.8571 | 103.6923 | 55.6923 | bDBP (MOG_GTF_033) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | bDBP (MOG_GTF_033) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | bDBP (MOG_GTF_033) [mmHg] | 80.7708 | 2.5852 | 14.6244 | 43 | 73.5 | 79.25 | 92.25 | 104.6667 | 61.6667 | bDBP (MOG_GTF_033) [mmHg] | 80.7299 | 2.4606 | 14.3476 | 43.75 | 71 | 80.625 | 90 | 104 | 60.25
bPP (MOG_GTF_033) [mmHg] | 52.1698 | 2.1872 | 12.7532 | 30.4 | 41.7273 | 49 | 63 | 82.4545 | 52.0545 | bPP (MOG_GTF_033) [mmHg] | 53.3113 | 2.3952 | 13.9666 | 34 | 44 | 49.5833 | 59.5 | 86 | 52 | bPP (MOG_GTF_033) [mmHg] | 52.7909 | 2.3644 | 13.5822 | 31.5 | 43 | 51 | 62 | 82.6667 | 51.1667 | bPP (MOG_GTF_033) [mmHg] | 54.8802 | 2.4872 | 14.0698 | 28 | 45.1667 | 52.75 | 66 | 82.6667 | 54.6667 | bPP (MOG_GTF_033) [mmHg] | 52.1828 | 2.2124 | 12.9004 | 31.6667 | 41.25 | 49 | 63.5 | 82 | 50.3333
HR (MOG_GTF_033) [beat/minute] | 71.4016 | 2.2863 | 13.3314 | 42.5 | 61.5556 | 72 | 82.0588 | 105 | 62.5 | HR (MOG_GTF_033) [beat/minute] | 71.9216 | 2.3023 | 13.4244 | 44 | 60 | 73.4167 | 79.75 | 105 | 61 | HR (MOG_GTF_033) [beat/minute] | 71.9586 | 2.3074 | 13.2548 | 44 | 64.6667 | 72.5 | 80.5 | 105 | 61 | HR (MOG_GTF_033) [beat/minute] | 70.5469 | 2.4989 | 14.1357 | 44 | 60.5 | 70.9167 | 81 | 104.3333 | 60.3333 | HR (MOG_GTF_033) [beat/minute] | 71.2304 | 2.246 | 13.0964 | 45 | 63 | 72.375 | 82.75 | 105 | 60
aoSBP (MOG_GTF_033) [mmHg] | 123.0134 | 3.3082 | 19.29 | 84.5294 | 110.3 | 121.1667 | 140.8462 | 158 | 73.4706 | aoSBP (MOG_GTF_033) [mmHg] | 124.4975 | 3.4026 | 19.5464 | 86 | 116.25 | 125 | 137 | 158 | 72 | aoSBP (MOG_GTF_033) [mmHg] | 121.9056 | 3.5701 | 20.5085 | 77 | 109.75 | 120 | 137 | 159 | 82 | aoSBP (MOG_GTF_033) [mmHg] | 127.0772 | 3.378 | 17.5525 | 74 | 120 | 124.3333 | 142.6667 | 156.25 | 82.25 | aoSBP (MOG_GTF_033) [mmHg] | 122.6354 | 3.5266 | 19.9496 | 81 | 109.5 | 120.6 | 139 | 154 | 73
aoDBP (MOG_GTF_033) [mmHg] | 82.3315 | 2.3285 | 13.5774 | 48.0909 | 73 | 81.9167 | 92.1818 | 104.0909 | 56 | aoDBP (MOG_GTF_033) [mmHg] | 82.7197 | 2.4184 | 13.8928 | 43 | 74 | 84 | 92.5 | 104.6667 | 61.6667 | aoDBP (MOG_GTF_033) [mmHg] | 80.7394 | 2.4855 | 14.2784 | 41 | 68 | 80 | 90 | 108 | 67 | aoDBP (MOG_GTF_033) [mmHg] | 85.5247 | 2.39 | 12.419 | 56 | 76.5 | 85 | 95 | 105.3333 | 49.3333 | aoDBP (MOG_GTF_033) [mmHg] | 82.2161 | 2.6898 | 15.2156 | 43 | 71.3333 | 83 | 94.125 | 105.25 | 62.25
aoPP (MOG_GTF_033) [mmHg] | 40.7341 | 2.1478 | 12.5235 | 23.7 | 30.6667 | 39.1307 | 49 | 76.5 | 52.8 | aoPP (MOG_GTF_033) [mmHg] | 41.8081 | 2.3696 | 13.6125 | 23.3333 | 30.3333 | 38 | 53 | 73.5 | 50.1667 | aoPP (MOG_GTF_033) [mmHg] | 41.0889 | 2.2754 | 13.0713 | 20 | 31 | 39.3333 | 46 | 73.5 | 53.5 | aoPP (MOG_GTF_033) [mmHg] | 41.534 | 2.5869 | 13.4419 | 18 | 34 | 38.3333 | 49.3333 | 77.25 | 59.25 | aoPP (MOG_GTF_033) [mmHg] | 40.3714 | 2.0589 | 11.647 | 24 | 30.2667 | 40.75 | 47.25 | 69 | 45
Mobil-O-Graph Non-invasive records: osc calibration | [9 empty] | Mobil-O-Graph Non-invasive records: osc calibration | [9 empty] | Mobil-O-Graph Non-invasive records: osc calibration | [9 empty] | Mobil-O-Graph Non-invasive records: osc calibration | [9 empty] | Mobil-O-Graph Non-invasive records: osc calibration
bSBP (MOG_GTF_osc) [mmHg] | 132.8506 | 3.5336 | 20.6044 | 90.2727 | 119.0909 | 133.9545 | 147.5455 | 175.7778 | 85.5051 | bSBP (MOG_GTF_osc) [mmHg] | 134.2819 | 3.5747 | 20.844 | 87 | 127 | 133.8333 | 150 | 180 | 93 | bSBP (MOG_GTF_osc) [mmHg] | 132.598 | 3.928 | 22.5647 | 88.5 | 122.3333 | 131.3333 | 148 | 174.3333 | 85.8333 | bSBP (MOG_GTF_osc) [mmHg] | 135.651 | 3.7202 | 21.0447 | 78 | 126.1667 | 136.0833 | 150 | 166.6667 | 88.6667 | bSBP (MOG_GTF_osc) [mmHg] | 132.9127 | 3.699 | 21.5689 | 84.75 | 118 | 135.25 | 145.5 | 175 | 90.25
bMBP (MOG_GTF_osc) [mmHg] | 104.5229 | 2.7146 | 15.8287 | 67 | 96.2727 | 106.5779 | 113.6 | 135.1111 | 68.1111 | bMBP (MOG_GTF_osc) [mmHg] | 105.3284 | 2.6866 | 15.6654 | 63 | 97 | 107.4583 | 116 | 133.3333 | 70.3333 | bMBP (MOG_GTF_osc) [mmHg] | 103.8995 | 3.0208 | 17.3533 | 64 | 94.6667 | 105 | 114 | 136 | 72 | bMBP (MOG_GTF_osc) [mmHg] | 105.9219 | 2.898 | 16.3935 | 59 | 98.6667 | 107.5 | 113.6667 | 132.6667 | 73.6667 | bMBP (MOG_GTF_osc) [mmHg] | 104.6412 | 2.9061 | 16.9452 | 62 | 95.6 | 106.5 | 114 | 136 | 74
bDBP (MOG_GTF_osc) [mmHg] | 80.6807 | 2.2586 | 13.1696 | 48 | 72.1111 | 80.1528 | 89.8571 | 103.6923 | 55.6923 | bDBP (MOG_GTF_osc) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | bDBP (MOG_GTF_osc) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | bDBP (MOG_GTF_osc) [mmHg] | 80.7708 | 2.5852 | 14.6244 | 43 | 73.5 | 79.25 | 92.25 | 104.6667 | 61.6667 | bDBP (MOG_GTF_osc) [mmHg] | 80.7299 | 2.4606 | 14.3476 | 43.75 | 71 | 80.625 | 90 | 104 | 60.25
bPP (MOG_GTF_osc) [mmHg] | 52.1698 | 2.1872 | 12.7532 | 30.4 | 41.7273 | 49 | 63 | 82.4545 | 52.0545 | bPP (MOG_GTF_osc) [mmHg] | 53.3113 | 2.3952 | 13.9666 | 34 | 44 | 49.5833 | 59.5 | 86 | 52 | bPP (MOG_GTF_osc) [mmHg] | 52.7909 | 2.3644 | 13.5822 | 31.5 | 43 | 51 | 62 | 82.6667 | 51.1667 | bPP (MOG_GTF_osc) [mmHg] | 54.8802 | 2.4872 | 14.0698 | 28 | 45.1667 | 52.75 | 66 | 82.6667 | 54.6667 | bPP (MOG_GTF_osc) [mmHg] | 52.1828 | 2.2124 | 12.9004 | 31.6667 | 41.25 | 49 | 63.5 | 82 | 50.3333
HR (MOG_GTF_osc) [beat/minute] | 71.4016 | 2.2863 | 13.3314 | 42.5 | 61.5556 | 72 | 82.0588 | 105 | 62.5 | HR (MOG_GTF_osc) [beat/minute] | 71.9216 | 2.3023 | 13.4244 | 44 | 60 | 73.4167 | 79.75 | 105 | 61 | HR (MOG_GTF_osc) [beat/minute] | 71.9586 | 2.3074 | 13.2548 | 44 | 64.6667 | 72.5 | 80.5 | 105 | 61 | HR (MOG_GTF_osc) [beat/minute] | 70.5469 | 2.4989 | 14.1357 | 44 | 60.5 | 70.9167 | 81 | 104.3333 | 60.3333 | HR (MOG_GTF_osc) [beat/minute] | 71.2304 | 2.246 | 13.0964 | 45 | 63 | 72.375 | 82.75 | 105 | 60
aoSBP (MOG_GTF_osc) [mmHg] | 139.4676 | 3.9441 | 22.9976 | 95.5294 | 126.1429 | 136.2583 | 157.6667 | 188.5 | 92.9706 | aoSBP (MOG_GTF_osc) [mmHg] | 141.2702 | 4.0998 | 23.5514 | 100 | 127.5 | 144 | 159 | 185.5 | 85.5 | aoSBP (MOG_GTF_osc) [mmHg] | 138.2707 | 4.2291 | 24.2941 | 85.5 | 124.6667 | 137 | 155 | 185.5 | 100 | aoSBP (MOG_GTF_osc) [mmHg] | 144.1543 | 4.196 | 21.8033 | 82 | 130 | 141 | 164 | 187 | 105 | aoSBP (MOG_GTF_osc) [mmHg] | 138.9281 | 4.1754 | 23.6199 | 92.4 | 126.5 | 135 | 158.5 | 181 | 88.6
aoDBP (MOG_GTF_osc) [mmHg] | 82.9403 | 2.3843 | 13.9026 | 48.0909 | 73.1111 | 82.8889 | 92.4167 | 109.3636 | 61.2727 | aoDBP (MOG_GTF_osc) [mmHg] | 83.3586 | 2.4714 | 14.1972 | 43 | 76 | 84 | 92.75 | 110 | 67 | aoDBP (MOG_GTF_osc) [mmHg] | 81.2136 | 2.5014 | 14.3695 | 41 | 68.3333 | 80.6667 | 90 | 108 | 67 | aoDBP (MOG_GTF_osc) [mmHg] | 85.9784 | 2.3996 | 12.4689 | 56 | 77 | 85 | 96 | 105.6667 | 49.6667 | aoDBP (MOG_GTF_osc) [mmHg] | 82.849 | 2.7557 | 15.5885 | 43 | 71.8333 | 83.5 | 94.375 | 106.5 | 63.5
aoPP (MOG_GTF_osc) [mmHg] | 56.4993 | 2.9073 | 16.9525 | 32.1 | 42.9286 | 53.8636 | 67.7778 | 104.1667 | 72.0667 | aoPP (MOG_GTF_osc) [mmHg] | 57.9596 | 3.1783 | 18.2577 | 33.3333 | 43.3333 | 52.5 | 74 | 100.5 | 67.1667 | aoPP (MOG_GTF_osc) [mmHg] | 56.998 | 3.0616 | 17.5874 | 28.5 | 44 | 56 | 64.5 | 100.5 | 72 | aoPP (MOG_GTF_osc) [mmHg] | 58.1574 | 3.6298 | 18.861 | 26 | 45 | 54 | 69.5 | 106.75 | 80.75 | aoPP (MOG_GTF_osc) [mmHg] | 56.1109 | 2.871 | 16.2407 | 32 | 42 | 55.5 | 66.5 | 95 | 63
MV: mean value. SEM: standard error of the mean. SD: standard deviation. p25th, p50th and p75th: 25th percentile (first quartile), 50th percentile (median or second quartile) and 75th percentile (third quartile). SBP, MBP, DBP and PP: systolic, mean, diastolic and pulse blood pressure, respectively. HR: heart rate. Prefix "p", "b" and "ao" indicate "peripheral", "brachial" artery and "aortic" root, respectively. Brachial MBP was obtained using oscillometry (osc) and six different equations (see text). osc: oscillometry. inv: invasive. | [9 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, sd calibration | [9 empty] | Carotid Applanation Tonometry Non-invasive records: GTF, sd calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, sd calibration | [9 empty] | Common carotid artery ultrasonographic records
[10 empty] | Radial SBP (SCOR_RT_GTF_sd) [mmHg] | 134.2819 | 3.5747 | 20.844 | 87 | 127 | 133.8333 | 150 | 180 | 93 | cSBP (SCOR_CT_GTF_sd) [mmHg] | 136.3343 | 3.8925 | 22.3608 | 93 | 127 | 135.3333 | 151 | 179.6667 | 86.6667 | bSBP (SCOR_BT_GTF_sd) [mmHg] | 135.5573 | 3.7138 | 21.0084 | 78 | 126.1667 | 136.0833 | 150 | 166.6667 | 88.6667 | CCA Systolic Diameter [mm] | 7.8097 | 0.15408 | 0.89841 | 5.47 | 7.13 | 7.92 | 8.36 | 9.82 | 4.35
[10 empty] | Radial DBP (SCOR_RT_GTF_sd) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | cDBP (SCOR_CT_GTF_sd) [mmHg] | 79.6187 | 2.4579 | 14.1197 | 42 | 68.3333 | 79 | 89 | 107 | 65 | bDBP (SCOR_BT_GTF_sd) [mmHg] | 80.6615 | 2.5879 | 14.6395 | 43 | 73.5 | 78.75 | 92.25 | 104.6667 | 61.6667 | CCA Diastolic Diameter [mm] | 7.4324 | 0.15032 | 0.87648 | 5 | 6.79 | 7.575 | 7.94 | 9.41 | 4.41
[10 empty] | Radial MBP (SCOR_RT_GTF_sd) [mmHg] | 100.4387 | 2.8518 | 16.629 | 58 | 91.5 | 102.8333 | 110.3333 | 128 | 70 | cMBP (SCOR_CT_GTF_sd) [mmHg] | 103.2389 | 3.0399 | 17.4628 | 64 | 93.6667 | 104 | 116 | 136 | 72 | bMBP (SCOR_BT_GTF_sd) [mmHg] | 104.0781 | 3.1625 | 17.8896 | 52 | 95.8333 | 104.25 | 115.9167 | 132.6667 | 80.6667 | CCA Mean Diameter [mm] | 7.6212 | 0.15192 | 0.88584 | 5.22 | 6.96 | 7.755 | 8.18 | 9.6 | 4.38
[10 empty] | aoSBP (SCOR_RT_P2_sd) [mmHg] | 126.9297 | 4.0446 | 22.8795 | 76 | 113.625 | 123 | 140.25 | 178 | 102 | aoSBP (SCOR_CT_GTF_sd) [mmHg] | 132.598 | 3.928 | 22.5647 | 88.5 | 122.3333 | 131.3333 | 148 | 174.3333 | 85.8333 | b_aoSBP (SCOR_BT_GTF_sd) [mmHg] | 128.5182 | 3.8727 | 21.9072 | 65 | 120.6667 | 130 | 143.75 | 160.6667 | 95.6667 | CCA IMT [mm] | 0.7943 | 0.041038 | 0.22477 | 0.441 | 0.619 | 0.7975 | 1.002 | 1.242 | 0.801
[10 empty] | aoPP (SCOR_RT_P2_sd) [mmHg] | 45.3255 | 2.7263 | 15.4225 | 22.75 | 33.6667 | 41.5 | 55 | 84 | 61.25 | aoDBP (SCOR_CT_GTF_sd) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | b_aoDBP (SCOR_BT_GTF_sd) [mmHg] | 82.4844 | 2.6227 | 14.8365 | 44 | 75.3333 | 80.8333 | 95 | 107 | 63 | CCA Peak Systolic Velocity [cm/s] | 65.4818 | 4.0447 | 23.2349 | 33.5 | 47.6 | 61.3 | 79.3 | 141.8 | 108.3
[10 empty] | aoSBP (SCOR_RT_GTF_sd) [mmHg] | 124.3873 | 3.6673 | 21.384 | 75 | 114 | 123.75 | 139.5 | 170.6667 | 95.6667 | aoPP (SCOR_CT_GTF_sd) [mmHg] | 52.7909 | 2.3644 | 13.5822 | 31.5 | 43 | 51 | 62 | 82.6667 | 51.1667 | b_aoPP (SCOR_BT_GTF_sd) [mmHg] | 46.1406 | 2.3765 | 13.4438 | 21 | 37.5 | 45.25 | 57.75 | 70.3333 | 49.3333 | CCA End Diastolic Velocity [cm/s] | 14.7448 | 0.91063 | 5.2312 | 0.72 | 12.3 | 14.4 | 17.3 | 28.8 | 28.08
[10 empty] | aoDBP (SCOR_RT_GTF_sd) [mmHg] | 82.1176 | 2.3039 | 13.4339 | 45 | 72.75 | 84 | 91.5 | 105.75 | 60.75 | Carotid Applanation Tonometry Non-invasive records: GTF, 033 calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, 033 calibration | [9 empty] | CCA Resistive Index | 0.76576 | 0.013078 | 0.075126 | 0.66 | 0.72 | 0.76 | 0.81 | 1.01 | 0.35
[10 empty] | aoPP (SCOR_RT_GTF_sd) [mmHg] | 42.3333 | 2.3085 | 13.4606 | 22.3333 | 31.3333 | 40.25 | 51.3333 | 75.3333 | 53 | cSBP (SCOR_CT_GTF_033) [mmHg] | 122.2803 | 3.5618 | 20.4608 | 82 | 114.5 | 120 | 137 | 155.5 | 73.5 | bSBP (SCOR_BT_GTF_033) [mmHg] | 124.237 | 3.1004 | 17.5383 | 73 | 116.1667 | 124.75 | 134.625 | 157.5 | 84.5 | CCA Pulsatile Index | 1.9091 | 0.10467 | 0.60128 | 1.14 | 1.52 | 1.79 | 2.15 | 4.07 | 2.93
[10 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, 033 calibration | [9 empty] | cDBP (SCOR_CT_GTF_033) [mmHg] | 79.7096 | 2.4638 | 14.1532 | 42 | 68 | 79 | 89 | 107 | 65 | bDBP (SCOR_BT_GTF_033) [mmHg] | 80.6615 | 2.5879 | 14.6395 | 43 | 73.5 | 78.75 | 92.25 | 104.6667 | 61.6667 | CCA Systo-diastolic Index | 6.6906 | 2.2167 | 12.734 | 2.95 | 3.52 | 4.2 | 5.27 | 77.1 | 74.15
[10 empty] | Radial SBP (SCOR_RT_GTF_033) [mmHg] | 130.2745 | 3.0555 | 17.8166 | 87 | 123.5 | 132.875 | 141.6667 | 166 | 79 | cMBP (SCOR_CT_GTF_033) [mmHg] | 97.1596 | 2.8365 | 16.2946 | 58 | 89 | 95.3333 | 107.5 | 128 | 70 | bMBP (SCOR_BT_GTF_033) [mmHg] | 98.6198 | 2.7816 | 15.7352 | 54 | 91.8333 | 99.75 | 105.75 | 124.3333 | 70.3333 | Common carotid artery diameter re-calibration: 033 calibration
[10 empty] | Radial DBP (SCOR_RT_GTF_033) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | aoSBP (SCOR_CT_GTF_033) [mmHg] | 119.5152 | 3.5966 | 20.661 | 77 | 112.3333 | 118 | 134.6667 | 153 | 76 | aoSBP (SCOR_BT_GTF_033) [mmHg] | 118.2031 | 3.1814 | 17.9966 | 70 | 111.5 | 119 | 130.3333 | 148 | 78 | aoSBP (Echo_ExpAdj_033) [mmHg] | 120.1076 | 3.4104 | 19.8857 | 84.8329 | 109.2649 | 118.9319 | 131.1063 | 164.2197 | 79.3868
[10 empty] | Radial MBP (SCOR_RT_GTF_033) [mmHg] | 98.4461 | 2.5219 | 14.7049 | 58 | 90 | 101.75 | 109.3333 | 122.6667 | 64.6667 | aoDBP (SCOR_CT_GTF_033) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoDBP (SCOR_BT_GTF_033) [mmHg] | 82.0417 | 2.6009 | 14.713 | 45 | 74.8333 | 79.8333 | 93.9167 | 106.6667 | 61.6667 | aoPP (Echo_ExpAdj_033) [mmHg] | 39.3777 | 2.2071 | 12.8693 | 22.8442 | 29.4355 | 35.099 | 47.957 | 71.2197 | 48.3756
[10 empty] | aoSBP (SCOR_RT_P2_033) [mmHg] | 122.3385 | 3.5024 | 19.8124 | 76 | 113.25 | 121.1667 | 134.5 | 164.6667 | 88.6667 | aoPP (SCOR_CT_GTF_033) [mmHg] | 39.7081 | 2.0799 | 11.948 | 21 | 30 | 37 | 47.6667 | 65.3333 | 44.3333 | aoPP (SCOR_BT_GTF_033) [mmHg] | 36.2083 | 1.8792 | 10.6303 | 13 | 28.5 | 35.8333 | 46.5 | 57 | 44 | Common carotid artery diameter re-calibration: 033HR calibration
[10 empty] | aoPP (SCOR_RT_P2_033) [mmHg] | 40.7344 | 2.3915 | 13.5282 | 24 | 32 | 36 | 47.125 | 80 | 56 | Carotid Applanation Tonometry Non-invasive records: GTF, 033HR calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, 033HR calibration | [9 empty] | aoSBP (Echo_ExpAdj_033HR) [mmHg] | 132.2105 | 3.9945 | 23.2918 | 96.1959 | 117.8056 | 129.1803 | 144.6711 | 187.1569 | 90.961
[10 empty] | aoSBP (SCOR_RT_GTF_033) [mmHg] | 120.6103 | 3.0799 | 17.9588 | 75 | 111 | 122.2083 | 130 | 158.3333 | 83.3333 | cSBP (SCOR_CT_GTF_033HR) [mmHg] | 133.2116 | 3.885 | 22.3178 | 91 | 122 | 131.4 | 146.3333 | 168.5 | 77.5 | bSBP (SCOR_BT_GTF_033HR) [mmHg] | 135.7344 | 3.4755 | 19.6604 | 79 | 125 | 136.25 | 149.5 | 174.5 | 95.5 | aoPP (Echo_ExpAdj_033HR) [mmHg] | 51.4806 | 2.9842 | 17.4006 | 28.7933 | 38.061 | 46.353 | 61.1711 | 99.4812 | 70.688
[10 empty] | aoDBP (SCOR_RT_GTF_033) [mmHg] | 81.8382 | 2.268 | 13.2248 | 45 | 72.75 | 84 | 91.5 | 105 | 60 | cDBP (SCOR_CT_GTF_033HR) [mmHg] | 79.6288 | 2.4575 | 14.1171 | 42 | 68.3333 | 79 | 89 | 107 | 65 | bDBP (SCOR_BT_GTF_033HR) [mmHg] | 80.7708 | 2.5852 | 14.6244 | 43 | 73.5 | 79.25 | 92.25 | 104.6667 | 61.6667 | Common carotid artery diameter re-calibration: 0412 calibration
[10 empty] | aoPP (SCOR_RT_GTF_033) [mmHg] | 38.6765 | 2.0106 | 11.7237 | 22.6667 | 30 | 34.8333 | 46 | 71.3333 | 48.6667 | cMBP (SCOR_CT_GTF_033HR) [mmHg] | 101.7237 | 2.9486 | 16.9387 | 62 | 92.6667 | 102 | 112 | 133.5 | 71.5 | bMBP (SCOR_BT_GTF_033HR) [mmHg] | 103.4714 | 2.8983 | 16.3953 | 57 | 95.6667 | 104.75 | 111.3333 | 131 | 74 | aoSBP (Echo_ExpAdj_0412) [mmHg] | 131.7847 | 3.926 | 22.8923 | 95.9238 | 117.6911 | 130.087 | 144.318 | 185.6535 | 89.7297
[10 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, 033HR calibration | [9 empty] | aoSBP (SCOR_CT_GTF_033HR) [mmHg] | 129.9727 | 3.832 | 22.0133 | 86 | 119 | 128.5 | 143.3333 | 165 | 79 | aoSBP (SCOR_BT_GTF_033HR) [mmHg] | 128.0599 | 3.4711 | 19.6357 | 75 | 119.6667 | 128.5 | 141.625 | 162 | 87 | aoPP (Echo_ExpAdj_0412) [mmHg] | 51.0547 | 2.9415 | 17.1517 | 28.5583 | 37.3471 | 45.2522 | 62.3878 | 92.6535 | 64.0952
[10 empty] | Radial SBP (SCOR_RT_GTF_033HR) [mmHg] | 143.3824 | 3.4705 | 20.2365 | 99 | 133.3333 | 143.875 | 158 | 185 | 86 | aoDBP (SCOR_CT_GTF_033HR) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoDBP (SCOR_BT_GTF_033HR) [mmHg] | 82.4453 | 2.597 | 14.6906 | 45 | 75 | 80.5 | 94.6667 | 106.6667 | 61.6667 | Common carotid artery diameter re-calibration: osc calibration
[10 empty] | Radial DBP (SCOR_RT_GTF_033HR) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | aoPP (SCOR_CT_GTF_033HR) [mmHg] | 50.1657 | 2.4127 | 13.8602 | 27 | 39.5 | 48.5 | 61.3333 | 75.3333 | 48.3333 | aoPP (SCOR_BT_GTF_033HR) [mmHg] | 45.5859 | 2.2645 | 12.8101 | 18 | 37 | 45.1667 | 58.25 | 67 | 49 | aoSBP (Echo_ExpAdj_osc) [mmHg] | 138.4232 | 4.319 | 25.1841 | 100.7828 | 122.3932 | 136.5615 | 151.8688 | 199.708 | 98.9252
[10 empty] | Radial MBP (SCOR_RT_GTF_033HR) [mmHg] | 103.1201 | 2.6411 | 15.4002 | 62 | 94 | 105.8333 | 115 | 130 | 68 | Carotid Applanation Tonometry Non-invasive records: GTF, 0412 calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, 0412 calibration | [9 empty] | aoPP (Echo_ExpAdj_osc) [mmHg] | 57.6933 | 3.4249 | 19.9702 | 31.5378 | 41.843 | 50.8453 | 69.9857 | 106.708 | 75.1702
[10 empty] | aoSBP (SCOR_RT_P2_033HR) [mmHg] | 132.9948 | 3.8456 | 21.7541 | 85 | 121 | 134 | 145.25 | 183 | 98 | cSBP (SCOR_CT_GTF_0412) [mmHg] | 132.8985 | 3.9494 | 22.6874 | 89.5 | 120.6667 | 129 | 147 | 168 | 78.5 | bSBP (SCOR_BT_GTF_0412) [mmHg] | 135.8385 | 3.4216 | 19.3552 | 79 | 126.1667 | 136.25 | 149.75 | 172 | 93 | Common carotid artery diameter re-calibration: inv calibration
[10 empty] | aoPP (SCOR_RT_P2_033HR) [mmHg] | 51.3906 | 2.7904 | 15.7849 | 31.5 | 41.5 | 46 | 57.25 | 92.3333 | 60.8333 | cDBP (SCOR_CT_GTF_0412) [mmHg] | 79.6348 | 2.4571 | 14.1151 | 42 | 68.3333 | 79 | 89 | 107 | 65 | bDBP (SCOR_BT_GTF_0412) [mmHg] | 80.7708 | 2.5852 | 14.6244 | 43 | 73.5 | 79.25 | 92.25 | 104.6667 | 61.6667 | aoSBP (Echo_ExpAdj_inv) [mmHg] | 130.7744 | 4.8337 | 28.185 | 80.57 | 108.2718 | 126.9682 | 150.3815 | 196.4296 | 115.8596
[10 empty] | aoSBP (SCOR_RT_GTF_033HR) [mmHg] | 130.9975 | 3.3594 | 19.5883 | 84 | 119.5 | 132.375 | 142.3333 | 174.6667 | 90.6667 | cMBP (SCOR_CT_GTF_0412) [mmHg] | 101.4823 | 2.939 | 16.8833 | 62 | 92.25 | 101 | 112 | 133 | 71 | bMBP (SCOR_BT_GTF_0412) [mmHg] | 103.4583 | 2.8561 | 16.1563 | 57 | 96.1667 | 104.8333 | 110.3333 | 129.3333 | 72.3333 | aoPP (Echo_ExpAdj_inv) [mmHg] | 61.6978 | 4.147 | 23.8226 | 24.57 | 39.8525 | 59.9355 | 77.3931 | 113.7876 | 89.2176
[10 empty] | aoDBP (SCOR_RT_GTF_033HR) [mmHg] | 82.1299 | 2.2831 | 13.3127 | 45 | 73 | 84 | 91.5 | 105.5 | 60.5 | aoSBP (SCOR_CT_GTF_0412) [mmHg] | 129.3929 | 3.9597 | 22.7466 | 85 | 117.6667 | 127 | 146 | 164.5 | 79.5 | aoSBP (SCOR_BT_GTF_0412) [mmHg] | 128.1458 | 3.4436 | 19.4798 | 75 | 120.6667 | 128 | 144.4167 | 160 | 85 | MV: mean value. SEM: standard error of the mean. SD: standard deviation. p25th, p50th and p75th: 25th percentile (first quartile), 50th percentile (median or second quartile) and 75th percentile (third quartile). SBP, MBP, DBP and PP: systolic, mean, diastolic and pulse blood pressure, respectively. HR: heart rate. Prefix "p", "b" and "ao" indicate "peripheral", "brachial" artery and "aortic" root, respectively. Brachial MBP was obtained using oscillometry (osc) and six different equations (see text). osc: oscillometry. inv: invasive. ExpAdj: use of an exponential fit for the diameter-pressure transformation.
[10 empty] | aoPP (SCOR_RT_GTF_033HR) [mmHg] | 48.9069 | 2.3629 | 13.7779 | 29.6667 | 38 | 44.6667 | 57 | 82.3333 | 52.6667 | aoDBP (SCOR_CT_GTF_0412) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoDBP (SCOR_BT_GTF_0412) [mmHg] | 82.4089 | 2.5898 | 14.6503 | 45 | 75 | 80.5 | 94.5 | 106.6667 | 61.6667
[10 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, 0412 calibration | [9 empty] | aoPP (SCOR_CT_GTF_0412) [mmHg] | 49.5859 | 2.5632 | 14.7243 | 26 | 37 | 46.5 | 59.5 | 81 | 55 | aoPP (SCOR_BT_GTF_0412) [mmHg] | 45.7526 | 2.3269 | 13.1627 | 18 | 36.25 | 45.8333 | 58.25 | 70.6667 | 52.6667
[10 empty] | Radial SBP (SCOR_RT_GTF_0412) [mmHg] | 142.7819 | 3.4149 | 19.9122 | 99 | 133.3333 | 144.125 | 155 | 183 | 84 | Carotid Applanation Tonometry Non-invasive records: GTF, osc calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, osc calibration
[10 empty] | Radial DBP (SCOR_RT_GTF_0412) [mmHg] | 80.9706 | 2.2501 | 13.1202 | 44 | 73 | 83.25 | 90.5 | 103.5 | 59.5 | cSBP (SCOR_CT_GTF_osc) [mmHg] | 138.6298 | 4.2474 | 24.3993 | 89 | 124.3333 | 134.3333 | 152.6667 | 175.5 | 86.5 | bSBP (SCOR_BT_GTF_osc) [mmHg] | 141.8203 | 3.6547 | 20.6742 | 81 | 131 | 140.25 | 156.625 | 179.5 | 98.5
[10 empty] | Radial MBP (SCOR_RT_GTF_0412) [mmHg] | 102.8946 | 2.6142 | 15.2433 | 62 | 95 | 105.375 | 113.6667 | 129.3333 | 67.3333 | cDBP (SCOR_CT_GTF_osc) [mmHg] | 79.65 | 2.4527 | 14.0897 | 42 | 68.3333 | 79 | 89 | 107 | 65 | bDBP (SCOR_BT_GTF_osc) [mmHg] | 80.6615 | 2.5879 | 14.6395 | 43 | 73.5 | 78.75 | 92.25 | 104.6667 | 61.6667
[10 empty] | aoSBP (SCOR_RT_P2_0412) [mmHg] | 132.7656 | 3.9152 | 22.1479 | 85 | 121 | 132.25 | 145.5 | 181.3333 | 96.3333 | cMBP (SCOR_CT_GTF_osc) [mmHg] | 103.8692 | 3.0313 | 17.4134 | 64 | 94.6667 | 105 | 114 | 136 | 72 | bMBP (SCOR_BT_GTF_osc) [mmHg] | 105.8281 | 2.8953 | 16.378 | 59 | 98.6667 | 106.5 | 113.6667 | 132.6667 | 73.6667
[10 empty] | aoPP (SCOR_RT_P2_0412) [mmHg] | 51.1615 | 2.9263 | 16.5538 | 30.5 | 40.75 | 44.875 | 58.625 | 99 | 68.5 | aoSBP (SCOR_CT_GTF_osc) [mmHg] | 134.7914 | 4.2419 | 24.3677 | 85 | 121.3333 | 132.3333 | 150 | 171.5 | 86.5 | aoSBP (SCOR_BT_GTF_osc) [mmHg] | 133.2734 | 3.6148 | 20.4487 | 77 | 123.6667 | 132.5 | 150.25 | 166.5 | 89.5
[10 empty] | aoSBP (SCOR_RT_GTF_0412) [mmHg] | 130.652 | 3.3936 | 19.788 | 84 | 121 | 132.4167 | 142 | 173.3333 | 89.3333 | aoDBP (SCOR_CT_GTF_osc) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoDBP (SCOR_BT_GTF_osc) [mmHg] | 82.5703 | 2.5867 | 14.6328 | 45 | 75.4167 | 80.8333 | 94.6667 | 107 | 62
[10 empty] | aoDBP (SCOR_RT_GTF_0412) [mmHg] | 82.1005 | 2.2836 | 13.3153 | 45 | 73 | 84 | 91.5 | 105.5 | 60.5 | aoPP (SCOR_CT_GTF_osc) [mmHg] | 54.9843 | 2.8849 | 16.5725 | 28 | 43 | 50.5 | 66.6667 | 91 | 63 | aoPP (SCOR_BT_GTF_osc) [mmHg] | 50.7344 | 2.638 | 14.9228 | 19 | 40 | 50.1667 | 64.75 | 80 | 61
[10 empty] | aoPP (SCOR_RT_GTF_0412) [mmHg] | 48.549 | 2.461 | 14.3503 | 29 | 38 | 43.5833 | 56.3333 | 88 | 59 | Carotid Applanation Tonometry Non-invasive records: GTF, inv calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: GTF, inv calibration
[10 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, osc calibration | [9 empty] | cSBP (SCOR_CT_GTF_inv) [mmHg] | 128.7242 | 4.1546 | 23.8665 | 81.5 | 114 | 129.4 | 140.6667 | 187.5 | 106 | bSBP (SCOR_BT_GTF_inv) [mmHg] | 136.224 | 4.1084 | 23.2404 | 80 | 123 | 142.5 | 151.25 | 180.5 | 100.5
[10 empty] | Radial SBP (SCOR_RT_GTF_osc) [mmHg] | 149.5147 | 3.7196 | 21.6886 | 102 | 139.3333 | 150.0417 | 163 | 193.3333 | 91.3333 | cDBP (SCOR_CT_GTF_inv) [mmHg] | 67.546 | 1.7532 | 10.0713 | 53.6667 | 61 | 64.6667 | 73 | 91 | 37.3333 | bDBP (SCOR_BT_GTF_inv) [mmHg] | 69.8151 | 1.7396 | 9.8406 | 53 | 62.75 | 68.5 | 76.8333 | 91.3333 | 38.3333
[10 empty] | Radial DBP (SCOR_RT_GTF_osc) [mmHg] | 80.6765 | 2.2415 | 13.0699 | 44 | 73 | 82 | 90.5 | 103.5 | 59.5 | cMBP (SCOR_CT_GTF_inv) [mmHg] | 92.8333 | 2.6555 | 15.2548 | 65 | 82 | 90 | 102 | 130.5 | 65.5 | bMBP (SCOR_BT_GTF_inv) [mmHg] | 97.7708 | 2.616 | 14.7981 | 68 | 88.25 | 97.75 | 109.75 | 120.3333 | 52.3333
[10 empty] | Radial MBP (SCOR_RT_GTF_osc) [mmHg] | 105.0784 | 2.6907 | 15.6893 | 63 | 97 | 107.125 | 116 | 133.3333 | 70.3333 | aoSBP (SCOR_CT_GTF_inv) [mmHg] | 124.9803 | 4.1229 | 23.6842 | 78 | 104.3333 | 126.5 | 137.3333 | 182 | 104 | aoSBP (SCOR_BT_GTF_inv) [mmHg] | 127.9297 | 3.9716 | 22.4667 | 76 | 116.5 | 134.3333 | 142.75 | 164.5 | 88.5
[10 empty] | aoSBP (SCOR_RT_P2_osc) [mmHg] | 138.224 | 4.264 | 24.1206 | 87 | 125 | 137.375 | 151.75 | 191.3333 | 104.3333 | aoDBP (SCOR_CT_GTF_inv) [mmHg] | 67.7308 | 1.7977 | 10.3268 | 54 | 61 | 65 | 73.5 | 93 | 39 | aoDBP (SCOR_BT_GTF_inv) [mmHg] | 71.8828 | 1.8077 | 10.226 | 57 | 64.5 | 70.5833 | 79.75 | 97 | 40
[10 empty] | aoPP (SCOR_RT_P2_osc) [mmHg] | 56.9323 | 3.3108 | 18.7287 | 33.5 | 45 | 50.0833 | 66.125 | 111.6667 | 78.1667 | aoPP (SCOR_CT_GTF_inv) [mmHg] | 57.2495 | 3.3213 | 19.0792 | 17 | 42 | 55 | 73 | 94.5 | 77.5 | aoPP (SCOR_BT_GTF_inv) [mmHg] | 56.0052 | 3.1839 | 18.0111 | 19 | 40.3333 | 61.9167 | 69.5 | 81.5 | 62.5
[10 empty] | aoSBP (SCOR_RT_GTF_osc) [mmHg] | 136.0049 | 3.6499 | 21.2822 | 86 | 125.3333 | 137.8333 | 148.75 | 182.3333 | 96.3333 | Carotid Applanation Tonometry Non-invasive records: NPROC, sd calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: NPMA, sd calibration
[10 empty] | aoDBP (SCOR_RT_GTF_osc) [mmHg] | 81.9706 | 2.2862 | 13.331 | 45 | 73.5 | 84 | 91.5 | 105.5 | 60.5 | aoSBP (SCOR_CT_NPROC_sd) [mmHg] | 132.598 | 3.928 | 22.5647 | 88.5 | 122.3333 | 131.3333 | 148 | 174.3333 | 85.8333 | aoSBP (SCOR_BT_NPMA_6.0_sd) [mmHg] | 132.0766 | 3.8456 | 21.754 | 70.34 | 123.465 | 133.2075 | 146.5192 | 163.53 | 93.19
[10 empty] | aoPP (SCOR_RT_GTF_osc) [mmHg] | 54.0294 | 2.7889 | 16.262 | 31.6667 | 41.6667 | 49.4167 | 64 | 99.3333 | 67.6667 | aoDBP (SCOR_CT_NPROC_sd) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoPP (SCOR_BT_NPMA_6.0_sd) [mmHg] | 51.4151 | 2.4655 | 13.9471 | 25.69 | 42.4958 | 50.095 | 63.155 | 77.4333 | 51.7433
[10 empty] | Radial Applanation Tonometry Non-invasive records: P2 and GTF, inv calibration | [9 empty] | aoPP (SCOR_CT_NPROC_sd) [mmHg] | 52.7909 | 2.3644 | 13.5822 | 31.5 | 43 | 51 | 62 | 82.6667 | 51.1667 | Brachial Applanation Tonometry Non-invasive records: NPMA, 033 calibration
[10 empty] | Radial SBP (SCOR_RT_GTF_inv) [mmHg] | 140.1397 | 3.7179 | 21.6786 | 83 | 130 | 140 | 154 | 183.3333 | 100.3333 | Carotid Applanation Tonometry Non-invasive records: NPROC, 033 calibration | [9 empty] | aoSBP (SCOR_BT_NPMA_6.0_033) [mmHg] | 121.1662 | 3.1472 | 17.8033 | 71.68 | 113.5442 | 120.5925 | 132.7188 | 153.55 | 81.87
[10 empty] | Radial DBP (SCOR_RT_GTF_inv) [mmHg] | 68.3897 | 1.7034 | 9.9323 | 51.6667 | 62 | 66.1667 | 76.5 | 93 | 41.3333 | aoSBP (SCOR_CT_NPROC_033) [mmHg] | 121.7869 | 3.5573 | 20.4353 | 81 | 111.3333 | 120 | 137.6667 | 155.5 | 74.5 | aoPP (SCOR_BT_NPMA_6.0_033) [mmHg] | 40.5048 | 2.0032 | 11.3317 | 15.68 | 33.3625 | 39.6183 | 51.7375 | 61.85 | 46.17
[10 empty] | Radial MBP (SCOR_RT_GTF_inv) [mmHg] | 94.2255 | 2.4536 | 14.307 | 67 | 83 | 92.1667 | 99.6667 | 120.6667 | 53.6667 | aoDBP (SCOR_CT_NPROC_033) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | Brachial Applanation Tonometry Non-invasive records: NPMA, 033HR calibration
[10 empty] | aoSBP (SCOR_RT_P2_inv) [mmHg] | 129.1927 | 3.9403 | 22.2897 | 77 | 116.125 | 126.4167 | 142.5 | 179 | 102 | aoPP (SCOR_CT_NPROC_033) [mmHg] | 41.9798 | 2.1883 | 12.5705 | 23 | 31.5 | 39 | 51 | 67.6667 | 44.6667 | aoSBP (SCOR_BT_NPMA_6.0_033HR) [mmHg] | 131.8191 | 3.4699 | 19.6289 | 76.91 | 123.1933 | 130.7825 | 144.6675 | 169.4 | 92.49
[10 empty] | aoPP (SCOR_RT_P2_inv) [mmHg] | 60.138 | 3.3282 | 18.827 | 18 | 50.1667 | 61 | 72 | 99 | 81 | Carotid Applanation Tonometry Non-invasive records: NPROC, 033HR calibration | [9 empty] | aoPP (SCOR_BT_NPMA_6.0_033HR) [mmHg] | 51.0482 | 2.4363 | 13.7816 | 20.91 | 42.7325 | 48.8483 | 64.265 | 75.29 | 54.38
[10 empty] | aoSBP (SCOR_RT_GTF_inv) [mmHg] | 126.4681 | 3.4949 | 20.3788 | 77 | 115.3333 | 126.7083 | 140.3333 | 169.6667 | 92.6667 | aoSBP (SCOR_CT_NPROC_033HR) [mmHg] | 132.8813 | 3.8364 | 22.0386 | 90 | 121.3333 | 131.3333 | 147 | 168 | 78 | Brachial Applanation Tonometry Non-invasive records: NPMA, 0412 calibration
[10 empty] | aoDBP (SCOR_RT_GTF_inv) [mmHg] | 69.7917 | 1.7766 | 10.3594 | 52 | 63.5 | 67.4167 | 78 | 95 | 43 | aoDBP (SCOR_CT_NPROC_033HR) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoSBP (SCOR_BT_NPMA_6.0_0412) [mmHg] | 131.7154 | 3.4616 | 19.582 | 76.91 | 123.395 | 130.7475 | 146.9 | 166.95 | 90.04
[10 empty] | aoPP (SCOR_RT_GTF_inv) [mmHg] | 56.7304 | 2.7806 | 16.2133 | 17 | 48.6667 | 57.5 | 69 | 88 | 71 | aoPP (SCOR_CT_NPROC_033HR) [mmHg] | 53.0742 | 2.5383 | 14.5815 | 30 | 41.6667 | 49.5 | 65.5 | 78.6667 | 48.6667 | aoPP (SCOR_BT_NPMA_6.0_0412) [mmHg] | 50.9446 | 2.4767 | 14.0106 | 20.91 | 42.11 | 49.75 | 65.0075 | 76.42 | 55.51
[10 empty] | Radial Applanation Tonometry Non-invasive records: NPMA_4.0_sd calibration | [9 empty] | Carotid Applanation Tonometry Non-invasive records: NPROC, 0412 calibration | [9 empty] | Brachial Applanation Tonometry Non-invasive records: NPMA, osc calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_sd) [mmHg] | 124.0849 | 3.6345 | 21.1927 | 74.32 | 115.105 | 123.6525 | 139.115 | 170.08 | 95.76 | aoSBP (SCOR_CT_NPROC_0412) [mmHg] | 132.3874 | 3.9132 | 22.4799 | 88.5 | 120 | 129.5 | 147 | 166.5 | 78 | aoSBP (SCOR_BT_NPMA_6.0_osc) [mmHg] | 137.432 | 3.6146 | 20.4473 | 78.65 | 127.4758 | 135.6 | 152.4425 | 174.295 | 95.645
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_sd) [mmHg] | 43.1143 | 2.3207 | 13.532 | 23.7533 | 32.16 | 41.82 | 52.91 | 76.08 | 52.3267 | aoDBP (SCOR_CT_NPROC_0412) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | aoPP (SCOR_BT_NPMA_6.0_osc) [mmHg] | 56.7705 | 2.7933 | 15.8013 | 22.65 | 46.79 | 55.3125 | 72.045 | 86.4133 | 63.7633
[10 empty] | Radial Applanation Tonometry Non-invasive records: NPMA_4.4_sd calibration | [9 empty] | aoPP (SCOR_CT_NPROC_0412) [mmHg] | 52.5803 | 2.6876 | 15.4391 | 28.5 | 40 | 50 | 63 | 84 | 55.5 | Brachial Applanation Tonometry Non-invasive records: NPMA, inv calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_sd) [mmHg] | 125.6469 | 3.6598 | 21.3402 | 76.11 | 116.405 | 125.0737 | 140.47 | 172.5567 | 96.4467 | Carotid Applanation Tonometry Non-invasive records: NPROC, osc calibration | [9 empty] | aoSBP (SCOR_BT_NPMA_6.0_inv) [mmHg] | 132.1462 | 4.1151 | 23.2785 | 77.65 | 120.4175 | 138.1217 | 147.3 | 173.69 | 96.04
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_sd) [mmHg] | 44.6763 | 2.3597 | 13.7594 | 24.9633 | 34.3533 | 42.87 | 54.76 | 78.5567 | 53.5933 | aoSBP (SCOR_CT_NPROC_osc) [mmHg] | 138.2859 | 4.1918 | 24.08 | 91 | 123.3333 | 134.3333 | 152.3333 | 175 | 84 | aoPP (SCOR_BT_NPMA_6.0_inv) [mmHg] | 62.3311 | 3.3284 | 18.8283 | 22.65 | 49.8183 | 66.9642 | 73.5737 | 94.9533 | 72.3033
[10 empty] | Radial Applanation Tonometry Non-invasive records: NPMA_4.0_033 calibration | [9 empty] | aoDBP (SCOR_CT_NPROC_osc) [mmHg] | 79.8071 | 2.4841 | 14.2699 | 42 | 68 | 79 | 89 | 107 | 65 | MV: mean value. SEM: standard error of the mean. SD: standard deviation. p25th, p50th and p75th: 25th percentile (first quartile), 50th percentile (median or second quartile) and 75th percentile (third quartile). SBP, MBP, DBP and PP: systolic, mean, diastolic and pulse blood pressure, respectively. HR: heart rate. Prefix "p", "b" and "ao" indicate "peripheral", "brachial" artery and "aortic" root, respectively. Brachial MBP was obtained using oscillometry (osc) and six different equations (see text). osc: oscillometry. inv: invasive. BT: brachial artery applanation tonometry.
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_033) [mmHg] | 120.2731 | 3.0478 | 17.7717 | 74.11 | 112.15 | 122.8504 | 129.0033 | 157.8733 | 83.7633 | aoPP (SCOR_CT_NPROC_osc) [mmHg] | 58.4788 | 3.0056 | 17.2661 | 32 | 45 | 55 | 71 | 94.6667 | 62.6667
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_033) [mmHg] | 39.3025 | 1.9649 | 11.457 | 23.9533 | 30.11 | 35.6275 | 47.4067 | 69.1633 | 45.21 | Carotid Applanation Tonometry Non-invasive records: NPROC, inv calibration
[10 empty] | Radial Applanation Tonometry Non-invasive records: NPMA_4.4_033 calibration | [9 empty] | aoSBP (SCOR_CT_NPROC_inv) [mmHg] | 127.9545 | 4.0922 | 23.508 | 80.5 | 114 | 128.6667 | 140.6667 | 186.5 | 106
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_033) [mmHg] | 121.7684 | 3.0667 | 17.8819 | 75.89 | 113.355 | 123.89 | 130.84 | 159.97 | 84.08 | aoDBP (SCOR_CT_NPROC_inv) [mmHg] | 67.5591 | 1.8365 | 10.5498 | 51.6667 | 61 | 65 | 73.5 | 93 | 41.3333
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_033) [mmHg] | 40.7978 | 2.0108 | 11.7246 | 25.1967 | 31.8625 | 37.08 | 48.88 | 71.2367 | 46.04 | aoPP (SCOR_CT_NPROC_inv) [mmHg] | 60.3955 | 3.3529 | 19.2607 | 20 | 49 | 57.25 | 76.3333 | 99 | 79
[10 empty] | Radial Applanation Tonometry Non-invasive records: NPMA_4.0_033HR calibration | [9 empty] | MV: mean value. SEM: standard error of the mean. SD: standard deviation. p25th, p50th and p75th: 25th percentile (first quartile), 50th percentile (median or second quartile) and 75th percentile (third quartile). SBP, MBP, DBP and PP: systolic, mean, diastolic and pulse blood pressure, respectively. HR: heart rate. Prefix "p", "b" and "ao" indicate "peripheral", "brachial" artery and "aortic" root, respectively. Brachial MBP was obtained using oscillometry (osc) and six different equations (see text). osc: oscillometry. inv: invasive. CT: common carotid artery applanation tonometry.
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_033HR) [mmHg] | 130.6479 | 3.3324 | 19.4308 | 82.71 | 121.075 | 131.9737 | 142.8333 | 174.2033 | 91.4933
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_033HR) [mmHg] | 49.6773 | 2.3239 | 13.5504 | 31.2433 | 38.71 | 45.2496 | 58.195 | 80.2033 | 48.96
Radial Applanation Tonometry Non-invasive records: NPMA_4.4_033HR calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_033HR) [mmHg] | 132.5435 | 3.3658 | 19.626 | 85 | 122.55 | 134.0025 | 144.7567 | 176.84 | 91.84
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_033HR) [mmHg] | 51.5729 | 2.3908 | 13.9407 | 32.8733 | 40.975 | 46.945 | 59.745 | 83.7467 | 50.8733
Radial Applanation Tonometry Non-invasive records: NPMA_4.0_0412 calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_0412) [mmHg] | 130.0889 | 3.3402 | 19.4764 | 82.71 | 121.1567 | 132.295 | 142.54 | 172.7567 | 90.0467
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_0412) [mmHg] | 49.1183 | 2.3811 | 13.8838 | 30.5033 | 38.69 | 44.3863 | 54.97 | 85.69 | 55.1867
Radial Applanation Tonometry Non-invasive records: NPMA_4.4_0412 calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_0412) [mmHg] | 131.949 | 3.3651 | 19.6219 | 85 | 122.86 | 134.31 | 144.3 | 175.3467 | 90.3467
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_0412) [mmHg] | 50.9784 | 2.4368 | 14.2088 | 32.09 | 40.395 | 46.2442 | 56.9133 | 88.2533 | 56.1633
Radial Applanation Tonometry Non-invasive records: NPMA_4.0_Osc calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_osc) [mmHg] | 135.5832 | 3.6066 | 21.0301 | 84.86 | 125.6267 | 137.5533 | 146.9025 | 181.67 | 96.81
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_osc) [mmHg] | 54.9067 | 2.6953 | 15.7161 | 33.4033 | 42.82 | 49.6525 | 63.8933 | 96.51 | 63.1067
Radial Applanation Tonometry Non-invasive records: NPMA_4.4_Osc calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_osc) [mmHg] | 137.6558 | 3.6447 | 21.2523 | 87.28 | 127.3033 | 139.2733 | 148.33 | 184.55 | 97.27
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_osc) [mmHg] | 56.9793 | 2.7661 | 16.1292 | 35.14 | 44.3675 | 52.0125 | 65.8833 | 99.4067 | 64.2667
Radial Applanation Tonometry Non-invasive records: NPMA_4.0_inv calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.0_inv) [mmHg] | 126.1364 | 3.4672 | 20.2172 | 76.2 | 115.4633 | 125.7875 | 140.33 | 169.12 | 92.92
[10 empty] | aoPP (SCOR_RT_NPMA_4.0_inv) [mmHg] | 57.7467 | 2.777 | 16.1927 | 17.2 | 48.13 | 58.5 | 68.73 | 89.12 | 71.92
Radial Applanation Tonometry Non-invasive records: NPMA_4.4_inv calibration
[10 empty] | aoSBP (SCOR_RT_NPMA_4.4_inv) [mmHg] | 128.267 | 3.5238 | 20.5469 | 77.22 | 118 | 128.1242 | 142.11 | 172.0233 | 94.8033
[10 empty] | aoPP (SCOR_RT_NPMA_4.4_inv) [mmHg] | 59.8773 | 2.826 | 16.4781 | 18.22 | 51.0375 | 61.4583 | 70.54 | 92.0233 | 73.8033
MV: mean value. SEM: standard error of the mean. SD: standard deviation. p25th, p50th and p75th: 25th percentile (first quartile), 50th percentile (median or second quartile) and 75th percentile (third quartile). SBP, MBP, DBP and PP: systolic, mean, diastolic and pulse blood pressure, respectively. HR: heart rate. Prefix "p", "b" and "ao" indicate "peripheral", "brachial" artery and "aortic" root, respectively. Brachial MBP was obtained using oscillometry (osc) and six different equations (see text). osc: oscillometry. inv: invasive. RT: radial artery applanation tonometry.

## Table_S6_S7_CCC
Table S6. Lin´s Concordance Correlation Coefficient (CCC) (Non-invasive: x, Invasive: y): aoSBP | [6 empty] | Table S7. Lin´s Concordance Correlation Coefficient (CCC) (Non-invasive: x, Invasive: y): aoPP
| Cal.: bSBP/bDBP | Cal.: bMBP0.33/bDBP | Cal.: bMBP0.33HR/bDBP | Cal.: bMBP0.412/bDBP | Cal.: bMBPosc/bDBP | Cal.: bMBPinv/bDBPinv |  | Cal.: bSBP/bDBP | Cal.: bMBP0.33/bDBP | Cal.: bMBP0.33HR/bDBP | Cal.: bMBP0.412/bDBP | Cal.: bMBPosc/bDBP | Cal.: bMBPinv/bDBPinv
BA Oscillometry: Brachial-Aortic GTF | [6 empty] | BA Oscillometry: Brachial-Aortic GTF
| aoSBP (MOG_GTF_sd) | aoSBP (MOG_GTF_033) | --------- | --------- | aoSBP (MOG_GTF_osc) | --------- |  | aoPP (MOG_GTF_sd) | aoPP (MOG_GTF_033) | --------- | --------- | aoPP (MOG_GTF_osc) | ---------
CCC | 0.8111 | 0.7839 | --------- | --------- | 0.7986 | --------- | CCC | 0.3006 | 0.3039 | --------- | --------- | 0.6166 | ---------
CCC, 95% CI | 0.6790 to 0.8923 | 0.6393 to 0.8750 | --------- | --------- | 0.6463 to 0.8897 | --------- | CCC, 95% CI | 0.1459 to 0.4409 | 0.1493 to 0.4440 | --------- | --------- | 0.3823 to 0.7764 | ---------
r (Pearson coefficient; "precision") | 0.887 | 0.8671 | --------- | --------- | 0.846 | --------- | r (Pearson coefficient; "precision") | 0.6672 | 0.674 | --------- | --------- | 0.6752 | ---------
Bias correction factor Cb ("accuracy") | 0.9145 | 0.9041 | --------- | --------- | 0.9439 | --------- | Bias correction factor Cb ("accuracy") | 0.4505 | 0.4509 | --------- | --------- | 0.9132 | ---------
RA Tonometry: Radial-Aortic GTF | [6 empty] | RA Tonometry: Radial-Aortic GTF
| aoSBP (SCOR_RT_GTF_sd) | aoSBP (SCOR_RT_GTF_033) | aoSBP (SCOR_RT_GTF_033HR) | aoSBP (SCOR_RT_GTF_0412) | aoSBP (SCOR_RT_GTF_osc) | aoSBP (SCOR_RT_GTF_inv) |  | aoPP (SCOR_RT_GTF_sd) | aoPP (SCOR_RT_GTF_033) | aoPP (SCOR_RT_GTF_033HR) | aoPP (SCOR_RT_GTF_0412) | aoPP (SCOR_RT_GTF_osc) | aoPP (SCOR_RT_GTF_inv)
CCC | 0.8346 | 0.6994 | 0.8548 | 0.8386 | 0.8476 | 0.8856 | CCC | 0.3712 | 0.2307 | 0.4048 | 0.3995 | 0.5074 | 0.7973
CCC, 95% CI | 0.7194 to 0.9051 | 0.5438 to 0.8085 | 0.7429 to 0.9202 | 0.7141 to 0.9117 | 0.7216 to 0.9193 | 0.8045 to 0.9343 | CCC, 95% CI | 0.2180 to 0.5065 | 0.1008 to 0.3529 | 0.2066 to 0.5711 | 0.1979 to 0.5686 | 0.2709 to 0.6861 | 0.6728 to 0.8779
r (Pearson coefficient; "precision") | 0.921 | 0.8836 | 0.8821 | 0.8657 | 0.8565 | 0.9534 | r (Pearson coefficient; "precision") | 0.79 | 0.6354 | 0.6491 | 0.6348 | 0.63 | 0.9203
Bias correction factor Cb ("accuracy") | 0.9062 | 0.7915 | 0.969 | 0.9686 | 0.9897 | 0.9289 | Bias correction factor Cb ("accuracy") | 0.4698 | 0.3631 | 0.6236 | 0.6293 | 0.8054 | 0.8663
RA Tonometry: 2nd Radial Shoulder | [6 empty] | RA Tonometry: 2nd Radial Shoulder
| aoSBP (SCOR_RT_P2_sd) | aoSBP (SCOR_RT_P2_033) | aoSBP (SCOR_RT_P2_033HR) | aoSBP (SCOR_RT_P2_0412) | aoSBP (SCOR_RT_P2_osc) | aoSBP (SCOR_RT_P2_inv) |  | aoPP (SCOR_RT_P2_sd) | aoPP (SCOR_RT_P2_033) | aoPP (SCOR_RT_P2_033HR) | aoPP (SCOR_RT_P2_0412) | aoPP (SCOR_RT_P2_osc) | aoPP (SCOR_RT_P2_inv)
CCC | 0.8304 | 0.7027 | 0.8373 | 0.8171 | 0.8085 | 0.883 | CCC | 0.4298 | 0.2807 | 0.4726 | 0.4678 | 0.5729 | 0.8399
CCC, 95% CI | 0.6969 to 0.9083 | 0.5270 to 0.8207 | 0.6975 to 0.9157 | 0.6620 to 0.9051 | 0.6451 to 0.9012 | 0.7833 to 0.9384 | CCC, 95% CI | 0.2486 to 0.5820 | 0.1256 to 0.4224 | 0.2493 to 0.6482 | 0.2395 to 0.6471 | 0.3201 to 0.7496 | 0.7108 to 0.9142
r (Pearson coefficient; "precision") | 0.8897 | 0.8493 | 0.8472 | 0.8263 | 0.8136 | 0.9202 | r (Pearson coefficient; "precision") | 0.7619 | 0.6518 | 0.6614 | 0.6462 | 0.6461 | 0.8889
Bias correction factor Cb ("accuracy") | 0.9334 | 0.8273 | 0.9884 | 0.989 | 0.9938 | 0.9595 | Bias correction factor Cb ("accuracy") | 0.5642 | 0.4307 | 0.7146 | 0.7239 | 0.8867 | 0.9448
RA Tonometry: Radial NPMA (fs/4.0) | [6 empty] | RA Tonometry: Radial NPMA (fs/4.0)
| aoSBP (SCOR_RT_NPMA_40_sd) | aoSBP (SCOR_RT_NPMA_40_033) | aoSBP (SCOR_RT_NPMA_40_033HR) | aoSBP (SCOR_RT_NPMA_40_0412) | aoSBP (SCOR_RT_NPMA_40_osc) | aoSBP (SCOR_RT_NPMA_40_inv) |  | aoPP (SCOR_RT_NPMA_40_sd) | aoPP (SCOR_RT_NPMA_40_033) | aoPP (SCOR_RT_NPMA_40_033HR) | aoPP (SCOR_RT_NPMA_40_0412) | aoPP (SCOR_RT_NPMA_40_osc) | aoPP (SCOR_RT_NPMA_40_inv)
CCC | 0.8293 | 0.6963 | 0.8585 | 0.8375 | 0.8568 | 0.8822 | CCC | 0.3911 | 0.2427 | 0.4267 | 0.4098 | 0.5408 | 0.8321
CCC, 95% CI | 0.7127 to 0.9013 | 0.5431 to 0.8047 | 0.7508 to 0.9218 | 0.7147 to 0.9102 | 0.7382 to 0.9240 | 0.8008 to 0.9316 | CCC, 95% CI | 0.2354 to 0.5272 | 0.1122 to 0.3650 | 0.2287 to 0.5908 | 0.2099 to 0.5769 | 0.3113 to 0.7107 | 0.7255 to 0.8997
r (Pearson coefficient; "precision") | 0.9218 | 0.8914 | 0.8897 | 0.8714 | 0.8659 | 0.9562 | r (Pearson coefficient; "precision") | 0.8028 | 0.6626 | 0.6704 | 0.649 | 0.6625 | 0.938
Bias correction factor Cb ("accuracy") | 0.8996 | 0.7812 | 0.9649 | 0.961 | 0.9895 | 0.9226 | Bias correction factor Cb ("accuracy") | 0.4872 | 0.3663 | 0.6364 | 0.6315 | 0.8163 | 0.8871
RA Tonometry: Radial NPMA (fs/4.4) | [6 empty] | RA Tonometry: Radial NPMA (fs/4.4)
| aoSBP (SCOR_RT_NPMA_44_sd) | aoSBP (SCOR_RT_NPMA_44_033) | aoSBP (SCOR_RT_NPMA_44_033HR) | aoSBP (SCOR_RT_NPMA_44_0412) | aoSBP (SCOR_RT_NPMA_44_osc) | aoSBP (SCOR_RT_NPMA_44_inv) |  | aoPP (SCOR_RT_NPMA_44_sd) | aoPP (SCOR_RT_NPMA_44_033) | aoPP (SCOR_RT_NPMA_44_033HR) | aoPP (SCOR_RT_NPMA_44_0412) | aoPP (SCOR_RT_NPMA_44_osc) | aoPP (SCOR_RT_NPMA_44_inv)
CCC | 0.8537 | 0.7221 | 0.8625 | 0.8439 | 0.8424 | 0.9099 | CCC | 0.42 | 0.2586 | 0.4532 | 0.4356 | 0.5646 | 0.8699
CCC, 95% CI | 0.7465 to 0.9177 | 0.5717 to 0.8255 | 0.7553 to 0.9248 | 0.7230 to 0.9146 | 0.7139 to 0.9160 | 0.8414 to 0.9496 | CCC, 95% CI | 0.2568 to 0.5599 | 0.1186 to 0.3886 | 0.2418 to 0.6235 | 0.2228 to 0.6089 | 0.3258 to 0.7356 | 0.7778 to 0.9254
r (Pearson coefficient; "precision") | 0.9226 | 0.8877 | 0.8826 | 0.8657 | 0.8587 | 0.9545 | r (Pearson coefficient; "precision") | 0.7994 | 0.65 | 0.653 | 0.6351 | 0.6487 | 0.9357
Bias correction factor Cb ("accuracy") | 0.9254 | 0.8134 | 0.9772 | 0.9748 | 0.981 | 0.9533 | Bias correction factor Cb ("accuracy") | 0.5255 | 0.3979 | 0.694 | 0.6859 | 0.8702 | 0.9297
CA Tonometry: Carotid-Aortic GTF | [6 empty] | CA Tonometry: Carotid-Aortic GTF
| aoSBP (SCOR_CT_GTF_sd) | aoSBP (SCOR_CT_GTF_033) | aoSBP (SCOR_CT_GTF_033HR) | aoSBP (SCOR_CT_GTF_0412) | aoSBP (SCOR_CT_GTF_osc) | aoSBP (SCOR_CT_GTF_inv) |  | aoPP (SCOR_CT_GTF_sd) | aoPP (SCOR_CT_GTF_033) | aoPP (SCOR_CT_GTF_033HR) | aoPP (SCOR_CT_GTF_0412) | aoPP (SCOR_CT_GTF_osc) | aoPP (SCOR_CT_GTF_inv)
CCC | 0.9022 | 0.7649 | 0.8868 | 0.8713 | 0.868 | 0.9428 | CCC | 0.6087 | 0.3361 | 0.5581 | 0.562 | 0.6969 | 0.9172
CCC, 95% CI | 0.8155 to 0.9493 | 0.6139 to 0.8619 | 0.7891 to 0.9407 | 0.7592 to 0.9333 | 0.7521 to 0.9319 | 0.8950 to 0.9692 | CCC, 95% CI | 0.4191 to 0.7474 | 0.1895 to 0.4680 | 0.3656 to 0.7047 | 0.3665 to 0.7100 | 0.5043 to 0.8234 | 0.8519 to 0.9544
r (Pearson coefficient; "precision") | 0.9096 | 0.8783 | 0.8951 | 0.8774 | 0.8794 | 0.9732 | r (Pearson coefficient; "precision") | 0.7871 | 0.7802 | 0.7773 | 0.7747 | 0.7869 | 0.9645
Bias correction factor Cb ("accuracy") | 0.9918 | 0.8709 | 0.9907 | 0.9931 | 0.9871 | 0.9688 | Bias correction factor Cb ("accuracy") | 0.7733 | 0.4308 | 0.7179 | 0.7255 | 0.8856 | 0.951
CA Tonometry: Carotid calibration | [6 empty] | CA Tonometry: Carotid calibration
| aoSBP (SCOR_CT_NPROC_sd) | aoSBP (SCOR_CT_NPROC_033) | aoSBP (SCOR_CT_NPROC_033HR) | aoSBP (SCOR_CT_NPROC_0412) | aoSBP (SCOR_CT_NPROC_osc) | aoSBP (CT_NPROC_inv) |  | aoPP (SCOR_CT_NPROC_sd) | aoPP (SCOR_CT_NPROC_033) | aoPP (SCOR_CT_NPROC_033HR) | aoPP (SCOR_CT_NPROC_0412) | aoPP (SCOR_CT_NPROC_osc) | aoPP (CT_NPROC_inv)
CCC | 0.9022 | 0.7915 | 0.8719 | 0.8552 | 0.8234 | 0.9715 | CCC | 0.6087 | 0.3618 | 0.6009 | 0.5892 | 0.7042 | 0.9582
CCC, 95% CI | 0.8155 to 0.9493 | 0.6475 to 0.8809 | 0.7625 to 0.9328 | 0.7313 to 0.9244 | 0.6818 to 0.9056 | 0.9456 to 0.9852 | CCC, 95% CI | 0.4191 to 0.7474 | 0.1980 to 0.5059 | 0.3949 to 0.7494 | 0.3741 to 0.7442 | 0.4961 to 0.8358 | 0.9202 to 0.9783
r (Pearson coefficient; "precision") | 0.9096 | 0.8746 | 0.8823 | 0.862 | 0.8623 | 0.981 | r (Pearson coefficient; "precision") | 0.7871 | 0.7353 | 0.7458 | 0.7244 | 0.7394 | 0.9712
Bias correction factor Cb ("accuracy") | 0.9918 | 0.905 | 0.9881 | 0.992 | 0.9549 | 0.9903 | Bias correction factor Cb ("accuracy") | 0.7733 | 0.492 | 0.8057 | 0.8133 | 0.9525 | 0.9866
BA Tonometry: Radial-Aortic GTF | [6 empty] | BA Tonometry: Radial-Aortic GTF
| aoSBP (SCOR_BT_GTF_sd) | aoSBP (SCOR_BT_GTF_033) | aoSBP (SCOR_BT_GTF_033HR) | aoSBP (SCOR_BT_GTF_0412) | aoSBP (SCOR_BT_GTF_osc) | aoSBP (SCOR_BT_GTF_inv) |  | aoPP (SCOR_BT_GTF_sd) | aoPP (SCOR_BT_GTF_033) | aoPP (SCOR_BT_GTF_033HR) | aoPP (SCOR_BT_GTF_0412) | aoPP (SCOR_BT_GTF_osc) | aoPP (SCOR_BT_GTF_inv)
CCC | 0.8479 | 0.5873 | 0.7878 | 0.7782 | 0.8011 | 0.8028 | CCC | 0.4556 | 0.2095 | 0.3761 | 0.3752 | 0.4867 | 0.6908
CCC, 95% CI | 0.7221 to 0.9194 | 0.3917 to 0.7321 | 0.6283 to 0.8838 | 0.6136 to 0.8780 | 0.6379 to 0.8954 | 0.6356 to 0.8980 | CCC, 95% CI | 0.2843 to 0.5987 | 0.09263 to 0.3206 | 0.1960 to 0.5316 | 0.1899 to 0.5346 | 0.2649 to 0.6596 | 0.4860 to 0.8237
r (Pearson coefficient; "precision") | 0.8832 | 0.8113 | 0.8414 | 0.8327 | 0.8193 | 0.8318 | r (Pearson coefficient; "precision") | 0.8442 | 0.7047 | 0.7268 | 0.7068 | 0.6991 | 0.7963
Bias correction factor Cb ("accuracy") | 0.96 | 0.7239 | 0.9363 | 0.9345 | 0.9777 | 0.9651 | Bias correction factor Cb ("accuracy") | 0.5397 | 0.2973 | 0.5174 | 0.5308 | 0.6962 | 0.8675
BA Tonometry: Brachial NPMA (fs/6.0) | [6 empty] | BA Tonometry: Brachial NPMA (fs/6.0)
| aoSBP (SCOR_BT_NPMA_60_sd | aoSBP (SCOR_BT_NPMA_60_033) | aoSBP (SCOR_BT_NPMA_60_033HR) | aoSBP (SCOR_BT_NPMA_60_0412) | aoSBP (SCOR_BT_NPMA_60_osc) | aoSBP (SCOR_BT_NPMA_60_inv) |  | aoPP (SCOR_BT_NPMA_60_sd | aoPP (SCOR_BT_NPMA_60_033) | aoPP (SCOR_BT_NPMA_60_033HR) | aoPP (SCOR_BT_NPMA_60_0412) | aoPP (SCOR_BT_NPMA_60_osc) | aoPP (SCOR_BT_NPMA_60_inv)
CCC | 0.8725 | 0.6322 | 0.7989 | 0.7881 | 0.7733 | 0.8198 | CCC | 0.5759 | 0.2535 | 0.464 | 0.4612 | 0.5754 | 0.7768
CCC, 95% CI | 0.7637 to 0.9331 | 0.4337 to 0.7722 | 0.6384 to 0.8928 | 0.6204 to 0.8869 | 0.5913 to 0.8804 | 0.6572 to 0.9095 | CCC, 95% CI | 0.3902 to 0.7166 | 0.1099 to 0.3867 | 0.2454 to 0.6376 | 0.2405 to 0.6365 | 0.3245 to 0.7507 | 0.5890 to 0.8850
r (Pearson coefficient; "precision") | 0.8908 | 0.8021 | 0.8244 | 0.8141 | 0.8013 | 0.8224 | r (Pearson coefficient; "precision") | 0.8409 | 0.6697 | 0.6879 | 0.6805 | 0.6691 | 0.7934
Bias correction factor Cb ("accuracy") | 0.9795 | 0.7882 | 0.9691 | 0.9682 | 0.9651 | 0.9969 | Bias correction factor Cb ("accuracy") | 0.6849 | 0.3785 | 0.6745 | 0.6777 | 0.8601 | 0.9791
CA Distension waveform re-calibration | [6 empty] | CA Distension waveform re-calibration
| aoSBP (Echo_ExpAdj_sd) | aoSBP (Echo_ExpAdj_033) | aoSBP (Echo_ExpAdj_033HR) | aoSBP (Echo_ExpAdj_0412) | aoSBP (Echo_ExpAdj_osc) | aoSBP (Echo_ExpAdj_inv) |  | aoPP (Echo_ExpAdj_sd) | aoPP (Echo_ExpAdj_033) | aoPP (Echo_ExpAdj_033HR) | aoPP (Echo_ExpAdj_0412) | aoPP (Echo_ExpAdj_osc) | aoPP (Echo_ExpAdj_inv)
CCC | 0.8004 | 0.7291 | 0.826 | 0.8254 | 0.7893 | 0.9424 | CCC | 0.5469 | 0.2617 | 0.4631 | 0.4613 | 0.5574 | 0.7352
CCC, 95% CI | 0.6540 to 0.8891 | 0.5711 to 0.8350 | 0.6823 to 0.9082 | 0.6823 to 0.9075 | 0.6241 to 0.8869 | 0.8915 to 0.9698 | CCC, 95% CI | 0.3584 to 0.6926 | 0.1111 to 0.4006 | 0.2174 to 0.6536 | 0.2198 to 0.6495 | 0.2905 to 0.7438 | 0.5382 to 0.8561
r (Pearson coefficient; "precision") | 0.869 | 0.8651 | 0.8294 | 0.8306 | 0.8117 | 0.9491 | r (Pearson coefficient; "precision") | 0.766 | 0.6082 | 0.5802 | 0.5883 | 0.5888 | 0.7547
Bias correction factor Cb ("accuracy") | 0.9211 | 0.8428 | 0.9958 | 0.9937 | 0.9724 | 0.993 | Bias correction factor Cb ("accuracy") | 0.7141 | 0.4303 | 0.7981 | 0.7841 | 0.9467 | 0.9742
CCC: Concordance correlation coefficient. 95% CI: 95% Confidence Interval. Association between aoSBP obtained invasively and non-invasively (using several techniques and calibration methods (see text). | [6 empty] | CCC: Concordance correlation coefficient. 95% CI: 95% Confidence Interval. Association between aoPP obtained invasively and non-invasively (using several techniques and calibration methods (see text).

## Table_S8_ByA_aoSBP
Table S8. Comparison of invasive (reference) and non-invasive methods used to quantify aoSBP: Bland-Altman analysis
| Systematic differences (mmHg) |  |  |  | Limits of agreement (mmHg) | [5 empty] | Regression equation | [6 empty] | Absolute % error
BA Oscillometry: Brachial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
bSBP (MOG_Osc) | 1.306 | -1.7091 | 4.3211 | 8.6413 | -15.6309 | -20.8349 | -10.4269 | 18.2429 | 13.038 | 23.4469 | 27.7539 | 13.54 | 41.9674 | -0.2011 | -0.3074 | -0.09474 | 0.00053 | 5% | 15.25%
aoSBP (MOG_Osc_GTF_sd) | -8.5355 | -12.459 | -4.6114 | 11.2467 | -30.579 | -37.3521 | -23.806 | 13.508 | 6.735 | 20.281 | 25.7479 | 7.2145 | 44.2814 | -0.2606 | -0.3993 | -0.122 | 0.000564 | 8.47% | 17.7%
aoSBP (MOG_Osc_GTF_033) | -8.5311 | -12.777 | -4.2853 | 12.1687 | -32.3817 | -39.71 | -25.0535 | 15.3195 | 7.9912 | 22.6477 | 32.0609 | 12.937 | 51.1847 | -0.3086 | -0.4516 | -0.1655 | 0.000114 | 7.24% | 21.33%
aoSBP (MOG_Osc_GTF_osc) | 7.923 | 3.3374 | 12.5086 | 13.1424 | -17.8361 | -25.7508 | -9.9215 | 33.6822 | 25.767 | 41.5969 | 33.6632 | 9.2654 | 58.0609 | -0.1957 | -0.3782 | -0.01318 | 0.0364 | 8.52% | 22.28%
RA Tonometry: Radial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_RT_GTF_sd) | -9.8235 | -13.093 | -6.5533 | 9.3725 | -28.1936 | -33.8379 | -22.5493 | 8.5465 | 2.9022 | 14.1908 | 14.2187 | -2.815 | 31.2526 | -0.1791 | -0.3041 | -0.05414 | 0.0064 | 7.56% | 14.92%
aoSBP (SCOR_RT_GTF_033) | -13.6005 | -17.68 | -9.5208 | 11.6924 | -36.5177 | -43.5591 | -29.4762 | 9.3167 | 2.2752 | 16.3582 | 31.8462 | 14.649 | 49.0429 | -0.3386 | -0.4648 | -0.2124 | 5.12e-06 | 9.89% | 20.82%
aoSBP (SCOR_RT_GTF_033HR) | -3.2132 | -7.1949 | 0.7684 | 11.4114 | -25.5797 | -32.4519 | -18.7074 | 19.1532 | 12.281 | 26.0254 | 34.345 | 15.474 | 53.2158 | -0.2798 | -0.4183 | -0.1414 | 0.000252 | 5.98% | 13.95%
aoSBP (SCOR_RT_GTF_0412) | -3.5588 | -7.7628 | 0.6451 | 12.0485 | -27.1739 | -34.4298 | -19.918 | 20.0563 | 12.8 | 27.3122 | 34.825 | 14.57 | 55.0793 | -0.286 | -0.4346 | -0.1374 | 0.000439 | 6.34% | 15.69%
aoSBP (SCOR_RT_GTF_osc) | 1.7941 | -2.5345 | 6.1228 | 12.406 | -22.5217 | -29.9929 | -15.0505 | 26.1099 | 18.638 | 33.5811 | 34.0436 | 11.577 | 56.5098 | -0.2403 | -0.4051 | -0.07543 | 0.0056 | 7.36% | 15.95%
aoSBP (SCOR_RT_GTF_inv) | -7.7426 | -10.414 | -5.0712 | 7.6563 | -22.7489 | -27.3597 | -18.1382 | 7.2636 | 2.6529 | 11.8744 | 17.7864 | 5.2142 | 30.3586 | -0.1902 | -0.2825 | -0.09796 | 0.000199 | 5.48% | 13.83%
RA Tonometry: 2nd Radial Shoulder | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_RT_P2_sd) | -8.6484 | -12.647 | -4.6496 | 11.0912 | -30.3872 | -37.2934 | -23.481 | 13.0903 | 6.1841 | 19.9966 | 10.4266 | -9.529 | 30.3827 | -0.1455 | -0.2991 | 0.008138 | 0.0626 | 7.34% | 19.39%
aoSBP (SCOR_RT_P2_033) | -13.2396 | -17.828 | -8.6504 | 12.7286 | -38.1876 | -46.1133 | -30.2618 | 11.7084 | 3.7827 | 19.6342 | 23.3395 | 3.1833 | 43.4957 | -0.2806 | -0.4357 | -0.1254 | 0.000883 | 9.6% | 20.83%
aoSBP (SCOR_RT_P2_033HR) | -2.5833 | -7.2224 | 2.0557 | 12.867 | -27.8026 | -35.8145 | -19.7907 | 22.6359 | 14.624 | 30.6478 | 24.7227 | 2.466 | 46.9795 | -0.2128 | -0.3842 | -0.04145 | 0.0166 | 8.4% | 15.13%
aoSBP (SCOR_RT_P2_0412) | -2.8125 | -7.77 | 2.145 | 13.7502 | -29.7629 | -38.3248 | -21.201 | 24.1379 | 15.576 | 32.6998 | 25.1437 | 1.1488 | 49.1386 | -0.2178 | -0.4025 | -0.03306 | 0.0224 | 7.49% | 17.36%
aoSBP (SCOR_RT_P2_osc) | 2.6458 | -2.6602 | 7.9518 | 14.7169 | -26.1992 | -35.363 | -17.0354 | 31.4909 | 22.327 | 40.6547 | 23.4581 | -3.442 | 50.3587 | -0.1643 | -0.3714 | 0.04278 | 0.1156 | 6% | 23.53%
aoSBP (SCOR_RT_P2_inv) | -6.3854 | -9.7839 | -2.9869 | 9.4262 | -24.8607 | -30.7302 | -18.9913 | 12.0899 | 6.2205 | 17.9593 | 11.5967 | -5.097 | 28.2906 | -0.138 | -0.2665 | 0.009454 | 0.0362 | 5.87% | 16.84%
RA Tonometry: Radial NPMA (fs/4.0) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_RT_NPMA_40_sd) | -10.1259 | -13.386 | -6.8652 | 9.3451 | -28.4423 | -34.0701 | -22.8144 | 8.1905 | 2.5627 | 13.8183 | 14.806 | -1.992 | 31.6043 | -0.1858 | -0.309 | -0.0625 | 0.0043 | 7.47% | 16.31%
aoSBP (SCOR_RT_NPMA_40_033) | -13.9377 | -17.936 | -9.9386 | 11.4616 | -36.4024 | -43.3049 | -29.5 | 8.527 | 1.6246 | 15.4295 | 31.6654 | 15.189 | 48.1416 | -0.3398 | -0.4607 | -0.2189 | 2.41e-06 | 9.72% | 20.67%
aoSBP (SCOR_RT_NPMA_40_033HR) | -3.5629 | -7.4426 | 0.3168 | 11.1193 | -25.3566 | -32.0529 | -18.6604 | 18.2308 | 11.534 | 24.9271 | 33.9442 | 15.803 | 52.0848 | -0.2795 | -0.4126 | -0.1463 | 0.00016 | 5.23% | 13.78%
aoSBP (SCOR_RT_NPMA_40_0412) | -4.1219 | -8.2587 | 0.01492 | 11.8561 | -27.3599 | -34.5 | -20.2199 | 19.1162 | 11.976 | 26.2562 | 35.152 | 15.613 | 54.6907 | -0.2926 | -0.436 | -0.1493 | 0.000225 | 5.77% | 14.36%
aoSBP (SCOR_RT_NPMA_40_osc) | 1.3724 | -2.8168 | 5.5616 | 12.0064 | -22.1601 | -29.3906 | -14.9296 | 24.9049 | 17.674 | 32.1354 | 33.7281 | 12.21 | 55.2455 | -0.2411 | -0.399 | -0.08319 | 0.0039 | 7.38% | 15.9%
aoSBP (SCOR_RT_NPMA_40_inv) | -8.0744 | -10.702 | -5.4467 | 7.531 | -22.8352 | -27.3705 | -18.2998 | 6.6865 | 2.1511 | 11.2218 | 17.9979 | 5.9 | 30.0957 | -0.1943 | -0.283 | -0.1055 | 9.54e-05 | 5.46% | 13.33%
RA Tonometry: Radial NPMA (fs/4.4) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_RT_NPMA_44_sd) | -8.5639 | -11.805 | -5.3227 | 9.2893 | -26.7709 | -32.3651 | -21.1767 | 9.6431 | 4.0489 | 15.2373 | 15.5154 | -1.319 | 32.3505 | -0.1794 | -0.303 | -0.05588 | 0.0058 | 6.75% | 15.22%
aoSBP (SCOR_RT_NPMA_44_033) | -12.4424 | -16.476 | -8.4083 | 11.5617 | -35.1034 | -42.0661 | -28.1406 | 10.2186 | 3.2558 | 17.1813 | 32.9712 | 16.136 | 49.8061 | -0.3384 | -0.4619 | -0.2148 | 3.68e-06 | 8.7% | 19.87%
aoSBP (SCOR_RT_NPMA_44_033HR) | -1.6673 | -5.6395 | 2.3049 | 11.3844 | -23.9807 | -30.8367 | -17.1248 | 20.6461 | 13.79 | 27.5021 | 35.6467 | 16.779 | 54.5139 | -0.278 | -0.4165 | -0.1396 | 0.000271 | 6.53% | 14.28%
aoSBP (SCOR_RT_NPMA_44_0412) | -2.2618 | -6.4714 | 1.9479 | 12.065 | -25.9092 | -33.175 | -18.6433 | 21.3856 | 14.119 | 28.6514 | 36.9312 | 16.843 | 57.0187 | -0.292 | -0.4394 | -0.1446 | 0.000317 | 6.35% | 15.42%
aoSBP (SCOR_RT_NPMA_44_osc) | 3.445 | -0.8519 | 7.7419 | 12.3148 | -20.6921 | -28.1084 | -13.2758 | 27.5821 | 20.165 | 34.9984 | 35.58 | 13.302 | 57.8574 | -0.2394 | -0.4029 | -0.07596 | 0.0054 | 7.39% | 18.07%
aoSBP (SCOR_RT_NPMA_44_inv) | -5.9438 | -8.5731 | -3.3144 | 7.5358 | -20.714 | -25.2522 | -16.1757 | 8.8264 | 4.2882 | 13.3646 | 18.5699 | 6.0321 | 31.1078 | -0.1827 | -0.2747 | -0.09065 | 0.00031 | 4.83% | 11.71%
CA Tonometry: Carotid-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_CT_GTF_sd) | 1.6818 | -2.0163 | 5.3799 | 10.4294 | -18.7597 | -25.1445 | -12.3749 | 22.1234 | 15.738 | 28.5081 | 25.5136 | 7.2785 | 43.7488 | -0.182 | -0.3189 | -0.04517 | 0.0108 | 5.47% | 17.45%
aoSBP (SCOR_CT_GTF_033) | -11.401 | -15.682 | -7.1191 | 12.0759 | -35.0698 | -42.4626 | -27.677 | 12.2678 | 4.875 | 19.6605 | 24.8382 | 5.6257 | 44.0507 | -0.2768 | -0.421 | -0.1326 | 0.000463 | 9.01% | 22.24%
aoSBP (SCOR_CT_GTF_033HR) | -0.9434 | -4.9134 | 3.0265 | 11.196 | -22.8876 | -29.7417 | -16.0335 | 21.0007 | 14.146 | 27.8548 | 27.1674 | 8.0795 | 46.2553 | -0.2147 | -0.358 | -0.07145 | 0.0046 | 6.41% | 14.4%
aoSBP (SCOR_CT_GTF_0412) | -1.5232 | -5.8002 | 2.7537 | 12.0619 | -25.1646 | -32.5489 | -17.7804 | 22.1182 | 14.733 | 29.5024 | 25.2691 | 4.042 | 46.4962 | -0.2047 | -0.364 | -0.04532 | 0.0135 | 6.29% | 16.21%
aoSBP (SCOR_CT_GTF_osc) | 3.8753 | -0.4385 | 8.189 | 12.1655 | -19.9692 | -27.4168 | -12.5215 | 27.7197 | 20.272 | 35.1673 | 22.988 | 0.4254 | 45.5506 | -0.146 | -0.3153 | 0.02336 | 0.0886 | 7.29% | 20.66%
aoSBP (SCOR_CT_GTF_inv) | -5.9359 | -7.9975 | -3.8743 | 5.8141 | -17.3315 | -20.8909 | -13.7722 | 5.4598 | 1.9004 | 9.0191 | 4.7192 | -5.868 | 15.3065 | -0.08139 | -0.1609 | 0.00192 | 0.045 | 4.97% | 12.56%
CA Tonometry: Carotid calibration | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_CT_NPROC_sd) | 1.6818 | -2.0163 | 5.3799 | 10.4294 | -18.7597 | -25.1445 | -12.3749 | 22.1234 | 15.738 | 28.5081 | 25.5136 | 7.2785 | 43.7488 | -0.182 | -0.3189 | -0.04517 | 0.0108 | 5.47% | 17.45%
aoSBP (SCOR_CT_NPROC_033) | -9.1293 | -13.476 | -4.7817 | 12.2611 | -33.1611 | -40.6672 | -25.6549 | 14.9025 | 7.3963 | 22.4086 | 28.5429 | 9.2699 | 47.816 | -0.2878 | -0.4324 | -0.1431 | 0.000312 | 7.95% | 20.82%
aoSBP (SCOR_CT_NPROC_033HR) | 1.9652 | -2.2221 | 6.1524 | 11.8089 | -21.1804 | -28.4097 | -13.951 | 25.1107 | 17.881 | 32.34 | 31.4262 | 11.254 | 51.5975 | -0.225 | -0.3764 | -0.07363 | 0.0049 | 6.12% | 18.38%
aoSBP (SCOR_CT_NPROC_0412) | 1.4712 | -3.0483 | 5.9907 | 12.7459 | -23.5108 | -31.3138 | -15.7078 | 26.4532 | 18.65 | 34.2562 | 31.2817 | 9.1224 | 53.441 | -0.2277 | -0.394 | -0.06138 | 0.0089 | 5.77% | 18.55%
aoSBP (SCOR_CT_NPROC_osc) | 7.3697 | 2.7811 | 11.9583 | 12.9407 | -17.994 | -25.9162 | -10.0718 | 32.7334 | 24.811 | 40.6556 | 29.9546 | 6.2362 | 53.6731 | -0.1725 | -0.3505 | 0.005515 | 0.0571 | 7.53% | 23.59%
aoSBP (SCOR_CT_NPROC_inv) | -2.9616 | -4.7306 | -1.1926 | 4.989 | -12.7401 | -15.7943 | -9.6858 | 6.8169 | 3.7626 | 9.8711 | 7.6315 | -1.231 | 16.4948 | -0.08092 | -0.1474 | -0.01439 | 0.0188 | 2.49% | 6.94%
BA Tonometry: Radial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_BT_GTF_sd) | -4.2806 | -8.4257 | -0.1354 | 11.1009 | -26.0382 | -33.2023 | -18.8742 | 17.4771 | 10.313 | 24.6412 | 33.1903 | 12.944 | 53.4366 | -0.2785 | -0.4268 | -0.1302 | 0.000634 | 6.13% | 15.57%
aoSBP (SCOR_BT_GTF_033) | -15.1667 | -20.388 | -9.945 | 13.9838 | -42.575 | -51.5996 | -33.5504 | 12.2417 | 3.2171 | 21.2663 | 43.8298 | 22.445 | 65.2137 | -0.4385 | -0.5951 | -0.2818 | 3.76e-06 | 9.32% | 24.34%
aoSBP (SCOR_BT_GTF_033HR) | -5.1861 | -9.9586 | -0.4136 | 12.7809 | -30.2367 | -38.4849 | -21.9884 | 19.8645 | 11.616 | 28.1127 | 42.4143 | 20.483 | 64.3453 | -0.3538 | -0.5144 | -0.1931 | 0.000106 | 4.63% | 18.2%
aoSBP (SCOR_BT_GTF_0412) | -5.0944 | -9.9815 | -0.2074 | 13.0877 | -30.7463 | -39.1925 | -22.3 | 20.5574 | 12.111 | 29.0036 | 44.3067 | 22.072 | 66.541 | -0.3672 | -0.53 | -0.2043 | 7.89e-05 | 4.32% | 20.39%
aoSBP (SCOR_BT_GTF_osc) | 0.075 | -4.9436 | 5.0936 | 13.44 | -26.2674 | -34.9411 | -17.5938 | 26.4174 | 17.743 | 35.0911 | 45.5347 | 21.048 | 70.0206 | -0.3379 | -0.5172 | -0.1585 | 0.000614 | 4.01% | 19.7%
aoSBP (SCOR_BT_GTF_inv) | -5.575 | -10.485 | -0.6647 | 13.15 | -31.3489 | -39.8354 | -22.8625 | 20.1989 | 11.712 | 28.6854 | 26.7389 | -0.057 | 53.5355 | -0.2402 | -0.4365 | -0.04386 | 0.0183 | 6.8% | 20.91%
BA Tonometry: Brachial NPMA (fs/6.0) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (SCOR_BT_NPMA_60_sd) | -0.7678 | -4.7958 | 3.2601 | 10.787 | -21.9104 | -28.8719 | -14.9489 | 20.3747 | 13.413 | 27.3362 | 35.7373 | 16.09 | 55.3844 | -0.2713 | -0.4152 | -0.1274 | 0.000609 | 5.84% | 13.27%
aoSBP (SCOR_BT_NPMA_60_033) | -12.3265 | -17.628 | -7.0247 | 14.1987 | -40.1559 | -49.3192 | -30.9927 | 15.5029 | 6.3396 | 24.6661 | 46.6798 | 24.591 | 68.7676 | -0.4385 | -0.6004 | -0.2767 | 6.15e-06 | 7.73% | 22.59%
aoSBP (SCOR_BT_NPMA_60_033HR) | -1.5722 | -6.5369 | 3.3925 | 13.2957 | -27.6318 | -36.2123 | -19.0512 | 24.4874 | 15.906 | 33.0679 | 45.706 | 22.17 | 69.2418 | -0.3514 | -0.5238 | -0.179 | 0.000263 | 5.43% | 15.79%
aoSBP (SCOR_BT_NPMA_60_0412) | -1.6827 | -6.7717 | 3.4062 | 13.6285 | -28.3946 | -37.1898 | -19.5993 | 25.0291 | 16.233 | 33.8244 | 46.9016 | 22.814 | 70.9883 | -0.3611 | -0.5375 | -0.1846 | 0.000251 | 4.8% | 18.18%
aoSBP (SCOR_BT_NPMA_60_osc) | 4.0573 | -1.1842 | 9.2988 | 14.037 | -23.4553 | -32.5143 | -14.3964 | 31.5699 | 22.511 | 40.6289 | 49.1503 | 22.918 | 75.3816 | -0.3351 | -0.5273 | -0.143 | 0.0013 | 5.81% | 23.55%
aoSBP (SCOR_BT_NPMA_60_inv) | -1.3949 | -6.5023 | 3.7126 | 13.6779 | -28.2036 | -37.0307 | -19.3764 | 25.4138 | 16.586 | 34.241 | 27.7747 | -0.858 | 56.4083 | -0.2168 | -0.4266 | -0.007037 | 0.0433 | 6.11% | 20.55%
CA Distension waveform re-calibration | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p50th | p 95th
aoSBP (Echo_CCA_ExpAdj_033) | -12.3685 | -16.898 | -7.839 | 12.9817 | -37.8127 | -45.6306 | -29.9948 | 13.0756 | 5.2577 | 20.8935 | 30.772 | 11.786 | 49.7578 | -0.3256 | -0.4665 | -0.1848 | 4.59e-05 | 9.7% | 21.11%
aoSBP (Echo_CCA_ExpAdj_033HR) | -0.2659 | -5.2935 | 4.7618 | 14.4094 | -28.5082 | -37.1859 | -19.8306 | 27.9764 | 19.298 | 36.6541 | 31.8847 | 7.1207 | 56.6488 | -0.2427 | -0.4263 | -0.05903 | 0.0112 | 8.55% | 21.72%
aoSBP (Echo_CCA_ExpAdj_0412) | -0.6915 | -5.6843 | 4.3013 | 14.3095 | -28.738 | -37.3555 | -20.1205 | 27.3551 | 18.737 | 35.9726 | 33.0488 | 8.7837 | 57.3139 | -0.2547 | -0.4346 | -0.07473 | 0.007 | 8.97% | 20.78%
aoSBP (Echo_CCA_ExpAdj_osc) | 5.9462 | 0.5179 | 11.3744 | 15.5574 | -24.5463 | -33.9153 | -15.1773 | 36.4387 | 27.069 | 45.8077 | 32.2648 | 4.2705 | 60.2591 | -0.1987 | -0.4063 | 0.008946 | 0.0601 | 9.1% | 26.42%
aoSBP (Echo_CCA_ExpAdj_inv) | -1.7021 | -4.8302 | 1.426 | 8.9652 | -19.2738 | -24.6728 | -13.8748 | 15.8697 | 10.47 | 21.2687 | -8.1379 | -25.03 | 8.7612 | 0.04858 | -0.0767 | 0.1739 | 0.4356 | 4.33% | 15.48%
"Mean error" (average error considering 10 methods) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | Median | p 95th
Average Calibration: sd | -5.2647 | -8.9605 | -1.5686 | 10.3435 | -25.5379 | -31.9202 | -19.1556 | 15.0086 | 8.6261 | 21.3908 | 22.2966 | 3.9056 | 40.6878 | -0.20713 | -0.34381 | -0.07048 | 0.01139 | 0.067222 | 0.16364
Average Calibration: 033 | -12.2143 | -16.6769 | -7.7513 | 12.5114 | -36.7367 | -44.4433 | -29.0302 | 12.3081 | 4.6015 | 20.0146 | 32.6546 | 13.5812 | 51.7277 | -0.33732 | -0.47901 | -0.1956 | 0.0001839 | 0.08886 | 0.21462
Average Calibration: 033HR | -1.8921 | -6.2693 | 2.485 | 12.2526 | -25.9071 | -33.4656 | -18.3487 | 22.1229 | 14.564 | 29.6813 | 34.1397 | 13.2921 | 54.987 | -0.27086 | -0.42561 | -0.11611 | 0.0042613 | 0.063644 | 0.16181
Average Calibration: 0412 | -2.2528 | -6.8388 | 2.333 | 12.8393 | -27.4178 | -35.3367 | -19.4989 | 22.9121 | 14.9928 | 30.831 | 34.7622 | 12.7788 | 56.7453 | -0.2782 | -0.4414 | -0.11497 | 0.0059012 | 0.062333 | 0.17438
Average Calibration: osc | 3.3979 | -1.3479 | 8.1436 | 13.2872 | -22.645 | -30.8398 | -14.4502 | 29.4408 | 21.2456 | 37.6356 | 34.078 | 9.8383 | 58.3174 | -0.23059 | -0.41056 | -0.050609 | 0.037579 | 0.068744 | 0.20819
Average Calibration: inv | -5.0795 | -8.2243 | -1.9347 | 8.7495 | -22.2285 | -27.6596 | -16.7974 | 12.0695 | 6.6382 | 17.5006 | 13.853 | -2.3327 | 30.0399 | -0.14177 | -0.26164 | -0.019347 | 0.066423 | 0.051489 | 0.14683
SD: Standard deviation. ME: mean error. 95% CI: 95% Confidence Interval. P: p value. LL and UP: lower and upper limit, respectively. fs: sampling frequency. Prefix "c" and "ao" indicate central and aortic root, respectively. SBP: systolic blood pressure. MOG, SCOR and Echo: values obtained using signals recorded with Mobil-O-Graph (Oscillometry), SphygmoCor (Applanation Tonometry) or Echography (vascular ultrasound), respectively. Osc: oscillometry. RT, BT and CT: radial, brachial and carotid artery applanation tonometry, respectively. P2: Radial artery 2nd-shoulder algorithm. NPMA: N-point moving average algorithm. GTF: general transfer function. ExpAdj: Exponential adjustement. RA, BA and CCA: radial, brachial and Common carotid artery, respectively. NPROC: non-processed (without application of GTF; direct calibration of the carotid wave). n: number of subjects.

## Table_S9_ByA_aoPP
Table S9. Comparison of invasive (reference) and non-invasive methods used to quantify aoPP: Bland-Altman analysis
| Systematic differences (mmHg) |  |  |  | Limits of agreement (mmHg) | [5 empty] | Regression equation | [6 empty] | Absolute % error
BA Oscillometry: Brachial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
bPP (MOG_Osc) | -11.6829 | -15.784 | -7.5812 | 11.7554 | -34.7235 | -41.802 | -27.6441 | 11.3578 | 4.2784 | 18.4372 | 18.2488 | 9.4712 | 27.0264 | -0.4688 | -0.6002 | -0.3373 | 2.98e-08 | 22.3% | 40.48%
aoPP (MOG_Osc_GTF_sd) | -22.9822 | -28.131 | -17.8329 | 14.758 | -51.9078 | -60.795 | -43.0202 | 5.9435 | -2.9441 | 14.8311 | 14.1634 | 2.9386 | 25.3883 | -0.5817 | -0.7499 | -0.4136 | 5.44e-08 | 36.26% | 54.95%
aoPP (MOG_Osc_GTF_033) | -23.1186 | -28.226 | -18.0111 | 14.6381 | -51.8092 | -60.624 | -42.9938 | 5.572 | -3.2433 | 14.3874 | 13.4925 | 2.2506 | 24.7344 | -0.5734 | -0.7418 | -0.405 | 7.45e-08 | 39.97% | 54.65%
aoPP (MOG_Osc_GTF_osc) | -7.3534 | -12.597 | -2.1092 | 15.0301 | -36.8124 | -45.863 | -27.761 | 22.1056 | 13.0541 | 31.157 | 19.5579 | 4.3629 | 34.7528 | -0.4215 | -0.6491 | -0.1939 | 0.000662 | 23.71% | 38%
RA Tonometry: Radial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_RT_GTF_sd) | -23.4877 | -27.899 | -19.0764 | 12.643 | -48.2681 | -55.882 | -40.6542 | 1.2926 | -6.3213 | 8.9065 | 7.7034 | -2.408 | 17.8148 | -0.4739 | -0.6209 | -0.3268 | 2.14e-07 | 37.23% | 53.99%
aoPP (SCOR_RT_GTF_033) | -27.1446 | -32.603 | -21.6854 | 15.6461 | -57.8109 | -67.233 | -48.3885 | 3.5217 | -5.9008 | 12.9441 | 14.415 | 3.3245 | 25.5055 | -0.6314 | -0.7927 | -0.4701 | 4.2e-09 | 38.75% | 62.55%
aoPP (SCOR_RT_GTF_033HR) | -16.9142 | -22.283 | -11.5448 | 15.3888 | -47.0762 | -56.343 | -37.8087 | 13.2477 | 3.9803 | 22.5152 | 19.7816 | 6.941 | 32.6222 | -0.5575 | -0.7442 | -0.3708 | 8.57e-07 | 21.55% | 56.42%
aoPP (SCOR_RT_GTF_0412) | -17.2721 | -22.747 | -11.7971 | 15.6912 | -48.0268 | -57.476 | -38.5772 | 13.4827 | 4.0331 | 22.9322 | 18.8839 | 5.2988 | 32.469 | -0.5493 | -0.7469 | -0.3518 | 2.88e-06 | 24.12% | 56.42%
aoPP (SCOR_RT_GTF_osc) | -11.7917 | -17.405 | -6.1783 | 16.0881 | -43.3244 | -53.013 | -33.6358 | 19.7411 | 10.0524 | 29.4297 | 20.6658 | 5.1931 | 36.1385 | -0.4931 | -0.7181 | -0.2681 | 9.36e-05 | 16.48% | 53.08%
aoPP (SCOR_RT_GTF_inv) | -9.0907 | -11.972 | -6.2091 | 8.2586 | -25.2775 | -30.251 | -20.304 | 7.0961 | 2.1226 | 12.0697 | 8.1364 | 0.3664 | 15.9065 | -0.2617 | -0.3747 | -0.1487 | 4.5e-05 | 15.76% | 27.33%
RA Tonometry: 2nd Radial Shoulder | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_RT_P2_sd) | -21.2031 | -26.014 | -16.392 | 13.3444 | -47.3581 | -55.667 | -39.0489 | 4.9519 | -3.3573 | 13.2611 | 6.6917 | -4.9901 | 18.3735 | -0.4256 | -0.6033 | -0.2479 | 3.17e-05 | 35.7% | 52.55%
aoPP (SCOR_RT_P2_033) | -25.7943 | -31.427 | -20.1614 | 15.6236 | -56.4165 | -66.145 | -46.6881 | 4.828 | -4.9004 | 14.5564 | 11.1231 | -0.8945 | 23.1406 | -0.5654 | -0.7482 | -0.3826 | 5.78e-07 | 39.23% | 58.94%
aoPP (SCOR_RT_P2_033HR) | -15.138 | -20.762 | -9.5134 | 15.6006 | -45.7151 | -55.429 | -36.0011 | 15.4391 | 5.725 | 25.1531 | 16.0357 | 2.1497 | 29.9217 | -0.4836 | -0.6949 | -0.2724 | 5.8e-05 | 23.29% | 55.4%
aoPP (SCOR_RT_P2_0412) | -15.3672 | -21.154 | -9.5803 | 16.0508 | -46.8267 | -56.821 | -36.8323 | 16.0923 | 6.0979 | 26.0867 | 15.0728 | 0.2688 | 29.8768 | -0.4717 | -0.6969 | -0.2465 | 0.000177 | 24.1% | 55.4%
aoPP (SCOR_RT_P2_osc) | -9.5964 | -15.592 | -3.6004 | 16.6305 | -42.1921 | -52.547 | -31.8367 | 22.9994 | 12.644 | 33.3547 | 16.1983 | -0.5486 | 32.9453 | -0.4029 | -0.6577 | -0.1482 | 0.003 | 18.45% | 51.73%
aoPP (SCOR_RT_P2_inv) | -6.3906 | -9.7979 | -2.9833 | 9.4506 | -24.9138 | -30.798 | -19.0292 | 12.1325 | 6.2479 | 18.0172 | 5.525 | -4.5606 | 15.6106 | -0.1843 | -0.3377 | -0.03088 | 0.0202 | 13.84% | 28.34%
RA Tonometry: Radial NPMA (fs/4.0) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_RT_NPMA_40_sd) | -22.7068 | -27.015 | -18.3985 | 12.3475 | -46.9078 | -54.343 | -39.4719 | 1.4943 | -5.9417 | 8.9302 | 7.734 | -2.1503 | 17.6182 | -0.4625 | -0.6062 | -0.3187 | 2.21e-07 | 35.31% | 52%
aoPP (SCOR_RT_NPMA_40_033) | -26.5186 | -31.843 | -21.194 | 15.2603 | -56.4288 | -65.618 | -47.2387 | 3.3916 | -5.7985 | 12.5817 | 14.5785 | 4.0659 | 25.0911 | -0.6244 | -0.7772 | -0.4715 | 1.67e-09 | 36.29% | 61.69%
aoPP (SCOR_RT_NPMA_40_033HR) | -16.1438 | -21.376 | -10.9115 | 14.9959 | -45.5357 | -54.566 | -36.5049 | 13.2482 | 4.2173 | 22.2791 | 20.0912 | 7.7738 | 32.4087 | -0.5505 | -0.7296 | -0.3714 | 5.12e-07 | 20.65% | 55.03%
aoPP (SCOR_RT_NPMA_40_0412) | -16.7028 | -22.074 | -11.331 | 15.3956 | -46.8781 | -56.149 | -37.6066 | 13.4726 | 4.201 | 22.7442 | 19.7739 | 6.833 | 32.7148 | -0.5542 | -0.7424 | -0.366 | 1.09e-06 | 24.9% | 55.03%
aoPP (SCOR_RT_NPMA_40_osc) | -10.9144 | -16.258 | -5.5702 | 15.3166 | -40.9348 | -50.158 | -31.7108 | 19.1061 | 9.8821 | 28.3301 | 20.9972 | 6.5747 | 35.4197 | -0.4848 | -0.6946 | -0.2751 | 4.63e-05 | 16.16% | 50.48%
aoPP (SCOR_RT_NPMA_40_inv) | -8.0744 | -10.702 | -5.4467 | 7.531 | -22.8352 | -27.37 | -18.2998 | 6.6865 | 2.1511 | 11.2218 | 8.281 | 1.4055 | 15.1566 | -0.2485 | -0.3485 | -0.1485 | 1.66e-05 | 11.86% | 27.79%
RA Tonometry: Radial NPMA (fs/4.4) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_RT_NPMA_44_sd) | -21.1448 | -25.463 | -16.8259 | 12.3781 | -45.4059 | -52.86 | -37.9515 | 3.1163 | -4.3381 | 10.5706 | 8.8571 | -1.2716 | 18.9857 | -0.4558 | -0.6031 | -0.3085 | 4.52e-07 | 33.89% | 50.37%
aoPP (SCOR_RT_NPMA_44_033) | -25.0233 | -30.405 | -19.6416 | 15.424 | -55.2544 | -64.543 | -45.9657 | 5.2078 | -4.0809 | 14.4966 | 15.9789 | 5.0628 | 26.8949 | -0.6229 | -0.7817 | -0.4642 | 4e-09 | 35.39% | 60.3%
aoPP (SCOR_RT_NPMA_44_033HR) | -14.2482 | -19.595 | -8.9011 | 15.3249 | -44.285 | -53.514 | -35.056 | 15.7887 | 6.5596 | 25.0177 | 21.9261 | 8.991 | 34.8612 | -0.5496 | -0.7377 | -0.3615 | 1.25e-06 | 16.77% | 53.91%
aoPP (SCOR_RT_NPMA_44_0412) | -14.8426 | -20.31 | -9.3747 | 15.6712 | -45.5583 | -54.995 | -36.1207 | 15.873 | 6.4354 | 25.3105 | 21.5883 | 8.1425 | 35.0342 | -0.5535 | -0.749 | -0.358 | 2.14e-06 | 21.39% | 53.91%
aoPP (SCOR_RT_NPMA_44_osc) | -8.8418 | -14.31 | -3.3728 | 15.6743 | -39.5633 | -49.002 | -30.1239 | 21.8798 | 12.4404 | 31.3192 | 22.903 | 7.8648 | 37.9412 | -0.4823 | -0.701 | -0.2636 | 8.63e-05 | 11.99% | 51.44%
aoPP (SCOR_RT_NPMA_44_inv) | -5.9438 | -8.5731 | -3.3144 | 7.5358 | -20.714 | -25.252 | -16.1757 | 8.8264 | 4.2882 | 13.3646 | 9.666 | 2.5418 | 16.7902 | -0.2372 | -0.3408 | -0.1336 | 5.28e-05 | 9.4% | 22.63%
CA Tonometry: Carotid-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_CT_GTF_sd) | -10.546 | -15.15 | -5.9414 | 12.9858 | -35.9982 | -43.948 | -28.0484 | 14.9063 | 6.9564 | 22.8561 | 19.9414 | 10.0375 | 29.8452 | -0.4814 | -0.6303 | -0.3324 | 2.3e-07 | 20.7% | 42.66%
aoPP (SCOR_CT_GTF_033) | -23.6288 | -28.429 | -18.8281 | 13.5389 | -50.165 | -58.453 | -41.8766 | 2.9074 | -5.381 | 11.1959 | 11.0622 | 2.228 | 19.8965 | -0.5477 | -0.6805 | -0.4149 | 1.69e-09 | 40% | 54.21%
aoPP (SCOR_CT_GTF_033HR) | -13.1712 | -17.832 | -8.5102 | 13.1451 | -38.9356 | -46.982 | -30.8883 | 12.5932 | 4.5458 | 20.6405 | 17.0576 | 6.7511 | 27.3641 | -0.4773 | -0.6322 | -0.3223 | 5.52e-07 | 24.3% | 46.47%
aoPP (SCOR_CT_GTF_0412) | -13.751 | -18.393 | -9.1085 | 13.0929 | -39.4131 | -47.428 | -31.3977 | 11.9111 | 3.8957 | 19.9264 | 14.5345 | 3.5284 | 25.5406 | -0.4466 | -0.6121 | -0.2811 | 5.06e-06 | 26.41% | 46.62%
aoPP (SCOR_CT_GTF_osc) | -8.3525 | -12.863 | -3.8412 | 12.7229 | -33.2893 | -41.078 | -25.5005 | 16.5843 | 8.7954 | 24.3731 | 14.9087 | 2.8212 | 26.9962 | -0.3673 | -0.549 | -0.1855 | 0.00026 | 19.15% | 41.66%
aoPP (SCOR_CT_GTF_inv) | -6.0874 | -8.0378 | -4.137 | 5.5005 | -16.8683 | -20.235 | -13.5009 | 4.6935 | 1.3262 | 8.0609 | 0.7003 | -5.2541 | o 6.6548 | -0.1072 | -0.1967 | -0.01764 | 0.0205 | 8.7% | 27.43%
CA Tonometry: Carotid calibration | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_CT_NPROC_sd) | -10.546 | -15.15 | -5.9414 | 12.9858 | -35.9982 | -43.948 | -28.0484 | 14.9063 | 6.9564 | 22.8561 | 19.9414 | 10.0375 | 29.8452 | -0.4814 | -0.6303 | -0.3324 | 2.3e-07 | 20.7% | 42.66%
aoPP (SCOR_CT_NPROC_033) | -21.3571 | -26.394 | -16.3199 | 14.2059 | -49.2007 | -57.897 | -40.5039 | 6.4865 | -2.2102 | 15.1833 | 13.5757 | 3.5054 | 23.646 | -0.5515 | -0.703 | -0.4001 | 2.28e-08 | 36.41% | 55.87%
aoPP (SCOR_CT_NPROC_033HR) | -10.2626 | -15.139 | -5.3863 | 13.7523 | -37.2172 | -45.636 | -28.7981 | 16.692 | 8.2729 | 25.111 | 19.6545 | 8.1733 | 31.1357 | -0.4723 | -0.645 | -0.2997 | 4.06e-06 | 21.4% | 48.54%
aoPP (SCOR_CT_NPROC_0412) | -10.7566 | -15.797 | -5.7153 | 14.2173 | -38.6225 | -47.326 | -29.9188 | 17.1094 | 8.4056 | 25.8131 | 18.2107 | 5.6304 | 30.7911 | -0.4574 | -0.6465 | -0.2682 | 2.61e-05 | 23.39% | 48.54%
aoPP (SCOR_CT_NPROC_osc) | -4.8581 | -9.8308 | 0.1146 | 14.0241 | -32.3452 | -40.93 | -23.7598 | 22.6291 | 14.0436 | 31.2145 | 19.2481 | 5.5072 | 32.989 | -0.3806 | -0.5872 | -0.174 | 0.000714 | 13.73% | 60.74%
aoPP (SCOR_CT_NPROC_inv) | -2.9414 | -4.7034 | -1.1794 | 4.9693 | -12.6812 | -15.723 | -9.6391 | 6.7984 | 3.7562 | 9.8406 | 2.9131 | -2.5119 | o 8.3381 | -0.09243 | -0.174 | -0.01087 | 0.0276 | 4.37% | 18.79%
BA Tonometry: Radial-Aortic GTF | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_BT_GTF_sd) | -19.45 | -23.673 | -15.2263 | 11.3112 | -41.6199 | -48.919 | -34.3201 | 2.7199 | -4.5799 | 10.0197 | 10.5278 | 1.1223 | 19.9333 | -0.4494 | -0.5848 | -0.3141 | 2.17e-07 | 31.5% | 43.89%
aoPP (SCOR_BT_GTF_033) | -29.9444 | -35.344 | -24.5441 | 14.4625 | -58.2909 | -67.624 | -48.9574 | -1.598 | -10.931 | 7.7355 | 11.298 | 0.9584 | 21.6377 | -0.6183 | -0.7671 | -0.4696 | 2.96e-09 | 46.66% | 63.89%
aoPP (SCOR_BT_GTF_033HR) | -20.425 | -25.55 | -15.2995 | 13.7264 | -47.3287 | -56.187 | -38.4703 | 6.4787 | -2.3797 | 15.3372 | 14.6968 | 2.6595 | 26.7342 | -0.5266 | -0.6998 | -0.3534 | 9.94e-07 | 33.01% | 50.56%
aoPP (SCOR_BT_GTF_0412) | -20.2472 | -25.502 | -14.9922 | 14.0733 | -47.8308 | -56.913 | -38.7485 | 7.3364 | -1.7459 | 16.4187 | 14.8837 | 2.1369 | 27.6305 | -0.5267 | -0.7101 | -0.3433 | 2.51e-06 | 32% | 55%
aoPP (SCOR_BT_GTF_osc) | -15.2333 | -20.557 | -9.9089 | 14.2592 | -43.1814 | -52.383 | -33.9791 | 12.7147 | 3.5124 | 21.917 | 15.9356 | 1.2767 | 30.5945 | -0.4673 | -0.6782 | -0.2564 | 9.79e-05 | 25.05% | 50.56%
aoPP (SCOR_BT_GTF_inv) | -9.7111 | -14.228 | -5.1939 | 12.0972 | -33.4217 | -41.228 | -25.6146 | 13.9994 | 6.1924 | 21.8065 | 10.9397 | -3.1687 | 25.048 | -0.3096 | -0.5126 | -0.1066 | 0.0041 | 15.34% | 48.04%
BA Tonometry: Brachial NPMA (fs/6.0) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (SCOR_BT_NPMA_60_sd) | -14.1901 | -18.372 | -10.008 | 11.1997 | -36.1415 | -43.369 | -28.9136 | 7.7614 | 0.5335 | 14.9892 | 14.2056 | 4.2614 | 24.1498 | -0.4257 | -0.5688 | -0.2826 | 1.42e-06 | 25.42% | 35.68%
aoPP (SCOR_BT_NPMA_60_033) | -25.7488 | -31.295 | -20.2019 | 14.8548 | -54.8641 | -64.45 | -45.2774 | 3.3666 | -6.2201 | 12.9532 | 14.9375 | 3.3005 | 26.5744 | -0.61 | -0.7774 | -0.4426 | 3.97e-08 | 40.29% | 54.49%
aoPP (SCOR_BT_NPMA_60_033HR) | -15.1111 | -20.497 | -9.7248 | 14.4246 | -43.3832 | -52.692 | -34.0742 | 13.1611 | 3.8521 | 22.4701 | 19.0901 | 5.2577 | 32.9224 | -0.5128 | -0.7118 | -0.3137 | 1.3e-05 | 25.54% | 42.43%
aoPP (SCOR_BT_NPMA_60_0412) | -15.2216 | -20.665 | -9.7781 | 14.5779 | -43.7942 | -53.202 | -34.3863 | 13.351 | 3.943 | 22.7589 | 18.77 | 4.5639 | 32.9761 | -0.5096 | -0.714 | -0.3052 | 2.07e-05 | 26.79% | 43.27%
aoPP (SCOR_BT_NPMA_60_osc) | -9.3649 | -14.98 | -3.7489 | 15.0399 | -38.8431 | -48.549 | -29.137 | 20.1133 | 10.4072 | 29.8194 | 21.0337 | 4.7685 | 37.2989 | -0.4558 | -0.6898 | -0.2217 | 0.000433 | 19.51% | 46.35%
aoPP (SCOR_BT_NPMA_60_inv) | -3.4476 | -8.0485 | 1.1532 | 12.3213 | -27.5974 | -35.549 | -19.6458 | 20.7022 | 12.7505 | 28.6539 | 14.9805 | 0.04012 | 29.9209 | -0.2763 | -0.4913 | -0.06132 | 0.0136 | 10.69% | 36.39%
CA Distension waveform re-calibration | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
aoPP (Echo_CCA_ExpAdj_033) | -24.9979 | -30.725 | -19.2706 | 16.4148 | -57.1709 | -67.056 | -47.2856 | 7.175 | -2.7103 | 17.0604 | 15.005 | 2.9922 | 27.0179 | -0.6214 | -0.7993 | -0.4435 | 4.51e-08 | 43.87% | 60.29%
aoPP (Echo_CCA_ExpAdj_033HR) | -12.895 | -19.066 | -6.7239 | 17.6864 | -47.5603 | -58.211 | -36.9092 | 21.7703 | 11.1192 | 32.4214 | 20.0464 | 3.3808 | 36.712 | -0.5117 | -0.7585 | -0.2649 | 0.000186 | 30.59% | 62.39%
aoPP (Echo_CCA_ExpAdj_0412) | -13.3206 | -19.408 | -7.2328 | 17.4478 | -47.5182 | -58.025 | -37.0108 | 20.8771 | 10.3696 | 31.3845 | 19.639 | 3.3301 | 35.9479 | -0.512 | -0.7535 | -0.2705 | 0.000142 | 32.2% | 55.47%
aoPP (Echo_CCA_ExpAdj_osc) | -6.6826 | -13.117 | -0.2481 | 18.4416 | -42.8282 | -53.934 | -31.7222 | 29.4629 | 18.3569 | 40.5689 | 21.084 | 2.1029 | 40.065 | -0.4313 | -0.7124 | -0.1502 | 0.0038 | 24.08% | 75.01%
aoPP (Echo_CCA_ExpAdj_inv) | -3.583 | -9.1855 | 2.0194 | 15.8 | -34.5511 | -44.223 | -24.8784 | 27.385 | 17.7124 | 37.0577 | 3.7326 | -14.9 | 22.3658 | -0.1138 | -0.3945 | 0.1669 | 0.4147 | 12.25% | 46.92%
"Mean error" (average error considering 10 methods) | Mean error | 95 % CI ME |  | SD | Lower limit | 95% CI LL |  | Upper limit | 95% CI UL |  | Intercept | 95% CI Intercept |  | Slope | 95% CI Slope |  | P | p 50th | p 95th
Average Calibration: sd | -18.473 | -22.9852 | -13.9603 | 12.6615 | -43.2895 | -51.0812 | -35.4975 | 6.3436 | -1.4485 | 14.1356 | 12.1962 | 1.953 | 22.4393 | -0.47082 | -0.62196 | -0.31967 | 3.8598e-06 | 0.30746 | 0.47639
Average Calibration: 033 | -25.3276 | -30.6691 | -19.9858 | 15.0069 | -54.7411 | -63.9643 | -45.5176 | 4.0859 | -5.1377 | 13.3094 | 13.5466 | 2.6794 | 24.4139 | -0.59664 | -0.75689 | -0.43641 | 7.7462e-08 | 0.39686 | 0.58688
Average Calibration: 033HR | -14.9232 | -20.2333 | -9.6128 | 14.8939 | -44.1152 | -53.2844 | -34.9456 | 14.2688 | 5.0992 | 23.4384 | 18.7089 | 5.7864 | 31.6314 | -0.51577 | -0.70597 | -0.32557 | 2.9469e-05 | 0.24122 | 0.5235
Average Calibration: 0412 | -15.2757 | -20.6722 | -9.8789 | 15.1353 | -44.941 | -54.2594 | -35.6221 | 14.3895 | 5.0706 | 23.7084 | 17.9285 | 4.4148 | 31.4423 | -0.509 | -0.70793 | -0.31007 | 4.2164e-05 | 0.26144 | 0.52184
Average Calibration: osc | -9.5151 | -14.9903 | -4.0394 | 15.3552 | -39.6113 | -49.066 | -30.1562 | 20.5812 | 11.126 | 30.0363 | 19.2194 | 3.9512 | 34.4876 | -0.4406 | -0.66533 | -0.21587 | 0.0009479 | 0.18289 | 0.5345
Average Calibration: inv | -6.1411 | -9.472 | -2.8101 | 9.2738 | -24.3178 | -30.0699 | -18.5653 | 12.0356 | 6.2831 | 17.7881 | 7.2083 | -2.8935 | 20.1141 | -0.20345 | -0.35231 | -0.054579 | 0.055646 | 0.11357 | 0.31518
SD: Standard deviation. ME: mean error. 95% CI: 95% Confidence Interval. LL and UP: lower and upper limit, respectively. P: p value. fs: sampling frequency. Prefix "c" and "ao" indicate central and aortic root, respectively. PP: pulse blood pressure. MOG, SCOR and Echo: values obtained using signals recorded with Mobil-O-Graph (Oscillometry), SphygmoCor (Applanation tonometry) or Echography (vascular ultrasound), respectively. Osc: oscillometry. RT, BT and CT: radial, brachial and carotid artery applanation tonometry, respectively. P2: Radial artery 2nd-shoulder algorithm. NPMA: N-point moving average algorithm. GTF: general transfer function. ExpAdj: Exponential adjustement. RA, BA and CCA: radial, brachial and Common carotid artery, respectively. NPROC: non-processed (without application of GTF; direct calibration of the carotid wave). n: number of subjects.