Table 2: Extraction methods for GC-MS analysis breath collected in sampling bags.

| Extraction method | Adsorbent phase | Protocol | Desorption method | Sample storage conditions | Analysis method | GC-MS system | GC-column | Sensibility | Ref. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| NTD | PDMS, Carbopack-X, Carboxen 1000 | air pump, 30mL/min | Thermal Desorption at 260°C (2min) | not specified | GC-MS | 6890 A gas chromatograph coupled with a 5975 Inert XL MSD (Agilent) | DB-624 (Agilent) 60m × 0.32mm × 1.8µm | ppb | [1] |
| NTD | 80 mg Carbotrap-B and 260 mg Carbopack-X | air pump, 30mL/min | Thermal Desorption Unit TD-100 (Markes International). Desorption at 320 °C for 10 min using helium as an inert carrier gas with a flow rate of 50 mL/min. Cryo-focusing onto a cold trap kept at −20 °C filled with graphitized carbon blacks. The temperature of the cold trap is then rapidly raised to 340 °C hold for 2 min | samples are extracted immediately after sampling and analyzed on the same day (within a few hours) | GC-ToF-MS | 7890B Gas Chromatograph (Agilent) coupled with a Bench-TOFdx time-of-flight mass spectrometer (Five Technologies) | apolar column Restek-Q-Bond 30m x 0.25mm x 8μm | ppb | [2] |
| SPME | PDMS | 60min at 37°C | Thermal desorption at 250°C in a splitless mode (3min) | not specified | GC-MS | Gas Chromatograph Mass Spectrometer GCMS-QP2010/ PLUS (Shimadzu) | Rtx-1 (Restek) 30m × 0.25mm× 0.25µm | not specified | [3] |
| SPME | Carboxen, PDMS | 10min at 40°C | Thermal desorption at 290°C in a splitless mode (1 min) | not specified | GC-MS | 7890 Gas Chromatograph (Agilent) coupled with a mass selective detector MSD type 5975C (Agilent) | PoraBond Q column (Varian) 25m × 0.32mm x 5μm | ppb | [4] |
| SPME | DVB, Carboxen, PDMS | 60min at 22°C | Thermal desorption at 250°C in a splitless mode (2min) | storage at 22°C for no longer than 6 hours | GCxGC-ToF-MS | GC 7890A (Agilent) with a dual stage jet cryogenic modulator (licensed from Zoex) and a secondary oven coupled to a high-speed ToF mass spectrometer (LECO) | 1D column: HP-5 column (J&W Scientific Inc.) 30m × 0.32mm x 0.25µm 2D column: DB-FFAP (J&W Scientific Inc.) 0.79m × 0.25mm x 0.25µ | ppq (pg/L) | [5] |
| SPME | Carboxen, PDMS | 25min at room temperature | Thermal desorption at 270°C in a splitless mode (5min) | samples were analyzed within 6h | GC-MS | Finnigan Trace GC Ultra coupled to a Polaris Q Quadrupole Ion Trap GC-MS system | DB-624 (Agilent) 30m × 0.25mm × 1.4µm | ppt | [6] |
| TF-SPME | PDMS Thin-Film, 5% Carboxen | 3h at 25°C | Thermal Desorption at 270°C for 5 min in tapered glass TD tube (5 mm I.D) at 270°C with a helium flow rate of 60mL/min. Cryo-focusing at -120°C in a cooled injection system (CIS). The CIS is then heated to 270°C at 12°C/s | Samples are analyzed immediately after the 3h of extraction | GC-MS | GC-MS 6890/5973(Hewlett Packard) | SLBTM-5MB (Sigma-Aldrich) 30m × 0.25mm × 0.25μm | ppb | [7] |
| Trapping in TD tube | Carbotrap C, Carbopack C | air pump | Thermal Desorption at 250°C for 8min with helium flowing at 10mL/min. Cryo-focusing at -150°C. Injection at 250°C over 1.5min | not specified | GC-MS | HP 5890 Series II Plus GC coupled to 5972 MSD (Hewlett-Packard) | HP5MS 30m x 0.25mm x 0.25µm | not specified | [8] |
| Trapping in TD tube | Tenax | air pump, 250mL/min | Thermal Desorption at 300°C for 3 minutes. Cryo-focusing into a Tenax TA cold trap at −150°C, which was heated after 2 minutes to 280°C at 20°C/s | storage at 4°C for no longer than 2 weeks | GC-MS | 6890 N GC (Agilent) coupled to a quadrupole mass spectrometer 5975 MSD (Agilent) | VF1-MS 30m × 0.25mm x 1 µm | ppb | [9] |
| Trapping in TD tube | Tenax TA | air pump, 100mL/min | Thermal Desorption at 270°C for 5 min under 60mL/min flow. Cryo-focusing on a Tenax cold trap at 0°C; injection in splitless mode at constant 3mL/min total flow | stored at 4 °C until analysis | GC-MS and electronic nose | GCMS-QP2010 (Shimadzu) | SLB-5 MS (Sigma-Aldrich) 30m x 0.25mm x 0.5 | ppb | [10] |
| Trapping in TD tube | Tenax, Carbograph, Carboxen | air pump, 200mL/min | Thermal Desorption at 270°C for 10min. Cryo-focusing at 10°C. The cold trap is subsequently heated to 300°C | storage at 4°C for no longer than 2 weeks | GCxGC BenchTOF-MS | 7890B GC (Agilent) fitted with a flow modulator and a three-way splitter plate coupled to a flame ionization detector and a time-of-flight mass spectrometer with electron ionization (SepSolve). | 1D column: Stabilwax (Restek) 30m × 0.25mm × 0.25μm 2D column: Rtx-200 MS (Restek) 5m × 0.25mm × 0.1μm | ppq (pg/L) | [11] |
| Headspace | no adsorbent phase | Gas Syringe 50 mL/min and flow path temperature of 150 ℃ | Thermal Desorption; a multiple channel thermal desorption system (UNITYxrTM) with an auto-sampler CIA Advantage-xrTM (Markes) is used to sample 100 mL of exhaled breath from each of the Tedlar bags at a flow rate of 50 mL/min and flow path temperature of 150 ℃ | room temperature no longer than 24h | GC-MS | Trace GC-Ultra gas chromatograph attached to an ISQ Mass Spectrometer (GC-MS, Thermo Scientific) | Rtx-VMS (Restek) 30m × 0.25mm x 1.40μm | ppm | [12] |

[1]	F. Monedeiro, M. Monedeiro-Milanowski, I. A. Ratiu, B. Brożek, T. Ligor, and B. Buszewski, “Needle trap device-gc-ms for characterization of lung diseases based on breath voc profiles,” Molecules, vol. 26, no. 6, Mar. 2021, doi: 10.3390/molecules26061789.

[2]	A. Pizzini et al., “Analysis of volatile organic compounds in the breath of patients with stable or acute exacerbation of chronic obstructive pulmonary disease,” J Breath Res, vol. 12, no. 3, Mar. 2018, doi: 10.1088/1752-7163/aaa4c5.

[3]	Y. Wang et al., “The analysis of volatile organic compounds biomarkers for lung cancer in exhaled breath, tissues and cell lines,” Cancer Biomarkers, vol. 11, pp. 129–137, 2012, doi: 10.3233/CBM-2012-0270.

[4]	P. Mochalski, J. King, M. Haas, K. Unterkofler, A. Amann, and G. Mayer, “Blood and breath profiles of volatile organic compounds in patients with end-stage renal disease,” BMC Nephrol, vol. 15, no. 1, pp. 1–14, Mar. 2014, doi: 10.1186/1471-2369-15-43.

[5]	M. Caldeira et al., “Allergic asthma exhaled breath metabolome: A challenge for comprehensive two-dimensional gas chromatography,” J Chromatogr A, vol. 1254, pp. 87–97, Sep. 2012, doi: 10.1016/j.chroma.2012.07.023.

[6]	M. Koureas, P. Kirgou, G. Amoutzias, C. Hadjichristodoulou, K. Gourgoulianis, and A. Tsakalof, “Target analysis of volatile organic compounds in exhaled breath for lung cancer discrimination from other pulmonary diseases and healthy persons,” Metabolites, vol. 10, no. 8, pp. 1–18, Aug. 2020, doi: 10.3390/metabo10080317.

[7]	K. Murtada, V. Galpin, J. J. Grandy, V. Singh, F. Sanchez, and J. Pawliszyn, “Development of porous carbon/polydimethylsiloxane thin-film solid-phase microextraction membranes to facilitate on-site sampling of volatile organic compounds,” Sustain Chem Pharm, vol. 21, Jun. 2021, doi: 10.1016/j.scp.2021.100435.

[8]	M. Phillips et al., “Prediction of breast cancer risk with volatile biomarkers in breath,” Breast Cancer Res Treat, vol. 170, no. 2, pp. 343–350, Jul. 2018, doi: 10.1007/s10549-018-4764-4.

[9]	P. Brinkman et al., “Exhaled breath profiles in the monitoring of loss of control and clinical recovery in asthma,” Clinical and Experimental Allergy, vol. 47, no. 9, pp. 1159–1169, Sep. 2017, doi: 10.1111/cea.12965.

[10]	Y. Y. Broza et al., “Exhaled Breath Markers for Nonimaging and Noninvasive Measures for Detection of Multiple Sclerosis,” ACS Chem Neurosci, vol. 8, no. 11, pp. 2402–2413, Nov. 2017, doi: 10.1021/acschemneuro.7b00181.

[11]	A. Z. Berna et al., “Reproducible Breath Metabolite Changes in Children with SARS-CoV-2 Infection,” ACS Infect Dis, vol. 7, no. 9, pp. 2596–2603, Sep. 2021, doi: 10.1021/acsinfecdis.1c00248.

[12]	E. Smirnova et al., “Predictive performance of selected breath volatile organic carbon compounds in stage 1 lung cancer,” Transl Lung Cancer Res, vol. 11, no. 6, pp. 1009–1018, Jun. 2022, doi: 10.21037/tlcr-21-953.