## Lens-related Issue-1
Web Address...26 | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8197958&pc=MTA | 2023826-2018-01925 | 2018-09-05 | Injury | STAAR SURGICAL COMPANY | 2018-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Capsular Bag Tear; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR/ DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 IMPLANTABLE COLLAMRE LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THIS COMPLAINT LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 M WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER INDICATED THAT THE INITIAL LENS WAS 12.6MM BUT PROVIDED THE SERIAL NUMBER OF A 13.2MM LENS. NO ADDITIONAL INFORMATION HAS BEEN FORTHCOMING. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973350&pc=MTA | 2023826-2021-04716 | 2021-09-20 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | N/A | NA | 1 | Event Description: REPORTEDLY A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS), HOWEVER THE SURGEON REPORTS A CAPSULAR TEAR AND PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. ATTEMPTS TO OBTAIN CLARIFICATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: PER RESPONSE-THERE WAS NO PATIENT CONTACT WITH THE LENS. THE EVENT WAS CAUSED BY A SURGICAL BLADE THAT WASN'T SHARP ENOUGH. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481825&pc=MTA | 2023826-2018-00636 | 2018-02-23 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Capsular Bag Tear; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 12.6 MM VICMO12.6, -13 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO CAPSULAR TEAR. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN; FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS); AND THE HAPTIC TORN PIECE IS MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266114&pc=MTA | 2023826-2015-01601 | 2014-06-04 | Injury | STAAR SURGICAL COMPANY | 2015-12-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Capsular Contracture; Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -12.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. A ANTERIOR SUBCAPSULAR CATARACT OPACITY WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: DIOPTER: -12.00. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158221&pc=MTA | 2023826-2020-01234 | 2020-03-23 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY ON (B)(6) 2020. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256324&pc=MTA | 2023826-2022-00148 | 2020-10-08 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC/PSC. PHACO-EMULSIFICATION (CCATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC/PSC. PHACO-EMULSIFICATION (CCATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION : UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION : UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980707&pc=MTA | 2023826-2020-02950 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN 2003. ASC LEN OPACITY WAS OBSERVED ON (B)(6) 2020, 17 YEARS AFTER ICL IMPLANTATION. THE REPORTER STATED " NO FURTHER STEPS ARE NEEDED WITH OD, BECAUSE THEY INTERNALLY SWITCHED OD <-> OS. SO WITH OD ALL IS FINE." . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER REPORTER, THERE IS NO COMPLAINT ASSOCIATED TO THE RIGHT EYE (LENS SERIAL # (B)(6). CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN 2003. ASC LENS OPACITY WAS OBSERVED ON (B)(6) 2020, 17 YEARS AFTER ICL IMPLANTATION. LENS REMAINS IMPLANTED. PER NEW INFORMATION REPORTER MEANT TO SAY THAT THERE IS A MILD ASC LENS OPACITY BUT THE REFRACTIVE SURPRISE IS ONLY AT LEFT EYE (OS) AND NOT AT THE REPORTED RIGHT EYE (OD). FOR RIGHT EYE (OD) NO LENS EXCHANGE IS PLANNED RIGHT NOW. CLAIM # (B)(4). Manufacturer Narrative: 2003. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366535&pc=MTA | 2023826-2021-00476 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE IN (B)(6) 2007. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-UPDATED INFO: THE PATIENT DEVELOPED A CATARACT ASC IN THE LEFT (OS) EYE 13-14 YEARS AFTER ICL IMPLANTATION. THE ICL AND CATARACT WERE REMOVED AND A MULTIFOCAL IOL WAS IMPLANTED WITHOUT COMPLICATION. D6B-UNK. H6-HEALTH IMPACT CODE UPDATED. (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366550&pc=MTA | 2023826-2021-00475 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S EYE IN (B)(6) 2007. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-UPDATED INFO: IN (B)(6) 2021 THE LENS WAS EXPLANTED FROM THE RIGHT (OD) EYE DUE TO LENS OPACITY ASC. CATARACT WAS REMOVED AND A MULTIFOCAL LENS WAS IMPLANTED WITH NO COMPLICATIONS. H6-HEALTH IMPACT CODE UPDATED. (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523570&pc=MTA | 2023826-2021-03689 | 2021-05-26 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO A CATARACT FORMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN GAUZE. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS THROUGHOUT THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872783&pc=MTA | 2023826-2021-04489 | 2021-06-18 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A NS (NUCLEAR SCLEROTIC) CATARACT WAS OBSERVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12875660&pc=MTA | 2023826-2021-04490 | 2021-06-18 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A NS (NUCLEAR SCLEROTIC) CATARACT WAS OBSERVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783787&pc=MTA | 2023826-2021-04216 | 2021-06-21 | Injury | STAAR SURGICAL COMPANY | 2021-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR CATARACT) WAS OBSERVED. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13022168&pc=MTA | 2023826-2021-04813 | 2021-10-06 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC) WAS OBSERVED. PHACO-EMULSIFICATION (CATARACT) SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722040&pc=MTA | 2023826-2022-01933 | 2022-05-09 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT HAS LENS OPACITY (CATARACT ASC). IT WAS ALSO NOTED "THE INVESTIGATOR DISCUSSED THE DEVELOPMENT OF ANTERIOR SUBCAPSULAR OPACITIES WITH THE PATIENT AND THE DECISION WAS MADE TO MONITOR THE VISION FOR NOW". THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14881811&pc=MTA | 2023826-2022-02099 | 2022-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -14.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITIES) WHICH WAS OBSERVED ON (B)(6) 2022 AND NOTICED IN THE LENS PERIPHERY AT APPROXIMATELY 4 O'CLOCK WHICH WAS DISCOVERED ON SLIT LAMP EXAMINATION. IT WAS NOTED THE ASC WAS DISCUSSED THE WITH PATIENT AND A DECISION WAS MADE TO MONITOR THE VISION CLOSELY. LENS REMAINS IMPLANTED. IT IS REPORTED THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637143&pc=MTA | 2023826-2023-03752 | 2023-07-19 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2015. IT WAS INDICATED THE PATIENT DEVELOPED A CATARACT AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048071&pc=MTA | 2023826-2023-04719 | 2023-09-25 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. CATARACTS WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ON (B)(6) 2023 LENS OPACITY(CATARACT) PSC, NS, WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. CAUSE OF EVENT PATIENT RELATED FACTOR: ADL COMPLAINTS, DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ENTERED INTO FORM CORRECTION: G2 - REPORTED AS FOREIGN - CORRECT TO REMOVE FOREIGN. CLAIM# (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223150&pc=MTA | 2023826-2023-05380 | 2023-11-15 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -3.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT HAD LENS OPACITY, ASC/NS WHICH WAS OBSERVED ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED AND CATARACT WITH A POSTERIOR CHAMBER LENS WAS PERFORMED WAS ADDITIONALLY REPORTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176027&pc=MTA | 2023826-2016-01771 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED A CATARACT HAD DEVELOPED BUT THERE WAS NO VISION LOSS. THE ICL REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: A2: (B)(6) 1980. FACILITY CONFIRMED PATIENT REPORT AND THE DATE OF LAST VISIT WAS (B)(6) 2016. THE PATIENT WILL BE MONITORED AND VA POST-OP WAS 20/20. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441180&pc=MTA | 2023826-2019-02346 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. IT IS UNKNOWN AS TO WHETHER THE LENS REMAINS IMPLANTED; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441184&pc=MTA | 2023826-2019-02345 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. IT IS UNKNOWN AS TO WHETHER THE LENS REMAINS IMPLANTED; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158238&pc=MTA | 2023826-2020-01236 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY 2 - 3 WEEKS AFTER SURGERY. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G4 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 14-MAY-2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158240&pc=MTA | 2023826-2020-01237 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY 2 - 3 WEEKS AFTER INITIAL SURGERY. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G4 - DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12641181&pc=MTA | 2023826-2021-04028 | NA | Injury | STAAR SURGICAL COMPANY | 2021-10-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD). POST-OPERATIVE OBSERVATION SHOWED THE DEVELOPMENT OF CATARACT. THE REFRACTIVE RESULT IS STABLE. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-03029. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EVENT DATE: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14121142&pc=MTA | 2023826-2022-01174 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, -21.00 DIOPTER IN THE PATIENTS LEFT EYE (OS). AT ONE YEAR POST-OP, THE EYE DEVELOPED A VISUALLY SIGNIFICANT ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT PHACO ASPIRATION OF THE CATARACT WITH ICL IN SITU AND A STANDARD POWER PCIOL WAS IMPLANTED. ONE YEAR AFTER THE LEFT EYE CATARACT SURGERY, THE PATIENT HAD UDVA OF 20/120, CDVA OF 20/40 AND SUBJECTIVE REFRACTION OF +3.00 DS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138558&pc=MTA | 2023826-2022-01175 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, -21.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). EIGHT YEARS POST-OP, THE EYE DEVELOPED AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT PHACOASPIRATION OF THE CATARACT WITH ICL IN SITU AND A SINGLE PIECE HYDROPHOBIC PCIOL WAS IMPLANTED. ONE YEAR AFTER THE RIGHT EYE CATARACT SURGERY, THE PATIENT HAD CDVA OF 20/40 AND SUBJECTIVE REFRACTION OF -0.75 DC@10 DEGREES. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621351&pc=MTA | 2023826-2022-03696 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621428&pc=MTA | 2023826-2022-03697 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT LEFT (OS) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674548&pc=MTA | 2023826-2023-03841 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO THE PATIENT DEVELOPING A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899883&pc=MTA | 2023826-2020-02746 | NA | Injury | STAAR SURGICAL COMPANY AG | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINED IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE REPORTER INDICATED THE CATARACT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6), BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM #: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601473&pc=MTA | 2023826-2015-00297 | 2013-11-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) AND WAS NOT REPLACED. SEE MFR REPORT #2023826-2015-00296 FOR THE OTHER EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW-REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601598&pc=MTA | 2023826-2015-00296 | 2013-11-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) AND WAS NOT REPLACED. Manufacturer Narrative: (B)(4) - CATARACT, INDUCED; (LENS EXPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598784&pc=MTA | 2023826-2022-01733 | 2022-03-25 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LENS OPACITY (CATARACT); ANTERIOR SUBCAPSULAR CATARACTS (ASC). THE PROBLEM WAS RESOLVED. PATIENT HAD A CATARACT SURGERY WITH PSEUDOPHAKIC IOL IMPLANTATION. CAUSE REPORTED AS USER ERROR. Manufacturer Narrative: WEIGHT-RACE: UNK. (B)(4). LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6303780&pc=MTA | 2023826-2017-00133 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY ASCO (ANTERIOR SUBCAPSULAR OPACITY), WHICH WAS OBSERVED ON (B)(6) 2016. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/30 WHICH IS THE SAME AS PREOPERATIVE DATA AS OF (B)(6) 2015. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648707&pc=MTA | 2023826-2018-00976 | NA | Injury | STAAR SURGICAL COMPANY AG | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WILL BE REMOVING A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS BECAUSE THE PATIENT HAS A CATARACT. ON (B)(6) 2018 THE REPORTER STATED THAT THE LENS REMAINS IMPLANTED, LENS REMOVAL HAS NOT BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12213000&pc=MTA | 2023826-2021-02697 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED LOW VAULTING AND LENS OPACITY (ASC). THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT". Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE OVERALL LENGTH WITHIN SPECIFICATIONS, WIDE INSPECTION COULD NOT BE PERFORMED DUE TO HAPTIC SIGNIFICANTLY DAMAGED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14283191&pc=MTA | 2023826-2022-01401 | 2022-02-11 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND LENS OPACITY (ASC). ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS BUT IT WAS NOTED "AFTER LENS REPLACEMENT; THERE IS A DECREASE IN VISUAL ACUITY; BUT THE PATIENT IS UNDER OBSERVATION. IT IS EXPECTED THAT VISION WILL RECOVER IN THE FUTURE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-01402 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189499&pc=MTA | 2023826-2022-01284 | 2022-03-16 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT NO INTERVETNION WAS PERFORMED FOR THE LENS OPACITY AND THE PATIENT IS UNDER FOLLOW-UP OBSERVATION. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS. REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189674&pc=MTA | 2023826-2022-01285 | 2022-03-16 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT NO INTERVENTION WAS PERFORMED FOR THE LENS OPACITY AND THE PATIENT IS UNDER FOLLOW-UP OBSERVATION. Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649701&pc=MTA | 2023826-2023-01087 | 2023-02-28 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, A CENTRAL ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED (VA WAS 20/25). THE ICL VAULT WAS HIGH (300%). THE LENS REMAINED IMPLANTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REPORTEDLY, ENLARGEMENT OF INCISION AND ADDITION OF SUTURE PERFORMED" SHOULD BE ADDED FOR CORRECTION. H6: 4625-ADDITIONAL SURGERY: ADDITIONAL SUTURE SHOULD BE ADDED FOR ENLARGEMENT OF INCISION SHOULD BE ADDED FOR CORRECTION. (B)(4) | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190574&pc=MTA | 2023826-2022-01287 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT AND ANTERIOR CAPSULE CATARACT FORMATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190575&pc=MTA | 2023826-2022-01288 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT AND ANTERIOR CAPSULE CATARACT FORMATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT- UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14656722&pc=MTA | 2023826-2022-01787 | NA | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS PATIENT-RELATED FACTOR. Manufacturer Narrative: AGE-RACE:UNK. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584913&pc=MTA | 2023826-2023-00976 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2 -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT AND IT WAS REPORTED THAT ON (B)(6) 2023 LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. "OTHER/ LOW VAULT" WAS REPORTED FOR CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539252&pc=MTA | 2023826-2017-00718 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO DEVELOPMENT OF A CATARACT. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY CONFIRMED THE PATIENT REPORT OF A CATARACT IN THE LEFT EYE AND WILL MONITOR THE PATIENT FOR ANY CHANGES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED LOW ICL VAULTING WAS OBSERVED, AS WELL AS BEGINNING OF ANTERIOR SUBCAPSULAR CATARACT (ASC). UCVA WAS 20/25-. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513713&pc=MTA | 2023826-2015-00141 | NA | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract | P030016 | NA | 1 | Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON INDICATED THE CATARACT WAS RELATED TO THE DEVICE AND WAS EITHER A SURGICAL TRAUMA OR PROGRESSIVE CATARACT. THE PATIENT'S VISUAL ACUITY WAS 20/20. Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.7MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED A CATARACT WAS FORMING ON UPPER LENS. THE ICL WAS IMPLANTED ON (B)(6) 2010 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) - CATARACT, INDUCED. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - SEVERAL FACTORS COULD CAUSE THE LATE ONSET OF CATARACT WHICH INCLUDES AGE, CONCOMMITANT DISEASES SUCH AS DIABETES, TRAUMA, SMOKING, ALCOHOL USE OR SYSTEMIC DRUGS SUCH AS STEROIDS. AGING IS A MAJOR CONTRIBUTING FACTOR FOR PROGRESSIVE CATARACT WHICH COULD NOT BE RULED OUT ESPECIALLY IN THIS CASE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262921&pc=MTA | 2023826-2023-02665 | 2023-04-09 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, ASC LENS OPACITY WAS OBSERVED AND THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283881&pc=MTA | 2023826-2023-02672 | 2023-05-31 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P880091 | NA | 1 | Event Description: THE REPORTER, INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF -23.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. ON (B)(6) 2023, THE LENS WAS REMOVED, DUE TO SECONDARY CATARACT AND SIGNIFICANT VISIONS DECREASE. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10877667&pc=MTA | 2023826-2020-02718 | 2020-01-21 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS OPACITY (ASC). CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED WITHIN THE SAME SURGERY AS REMOVAL OF INITIAL LENS. PROBLEM WAS RESOLVED. THE REPORTER STATED " WHEN THE ICL LENS LOADING WAS FINISHED, THE OPERATING ROOM NURSE INFORMED DOCTOR THAT IT WOULD TAKE ANOTHER 10 MINUTES TO DISINFECT THE SURGICAL INSTRUMENTS. WHEN THE LENS WAS IMPLANTED THE EYE, IT TOUCHED THE ANTERIOR CAPSULE OF THE CRYSTAL, CAUSING CATARACCT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511797&pc=MTA | 2023826-2020-01990 | 2020-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY WAS OBSERVED ON (B)(6) 2020. CATARACT SURGERY WAS NOT PERFORMED AND A DIFFERENT TYPE OF LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS DEBRIS ON LENS. CLAIM # (B)(6). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511799&pc=MTA | 2023826-2020-01989 | 2020-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY WAS OBSERVED ON (B)(6) 2020. CATARACT SURGERY WAS NOT PERFORMED AND A DIFFERENT TYPE OF LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS . CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302644&pc=MTA | 2023826-2023-02839 | 2023-01-10 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. LENS OPACITY WAS OBSERVED ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302731&pc=MTA | 2023826-2023-02840 | 2023-01-10 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS OPACITY WAS OBSERVED ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998309&pc=MTA | 2023826-2023-01941 | 2023-01-26 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A CATARACT WAS OBSERVED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED ON 04-APR-2023 AND THE PROBLEM WAS RESOLVED. CATARACT SURGERY WAS PERFORMED AND AN IOL LENS WAS IMPLANTED DURING THE SAME PROCEDURE AS ICL EXPLANT. THE PATIENT IS PSEUDOPHAKIC. THE CAUSE OF THE EVENT WAS DUE TO DEVICE AND PATIENT-RELATED FACTOR (SHALLOW CHAMBER, AGE AND HIGH MYOPIA). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725719&pc=MTA | 2023826-2023-01328 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VISIAN ICL IMPLANTABLE COLLAMER LENS, INTO HER LEFT EYE (OS), IN 2011. THE REPORTER INDICATED A CATARACT HAD DEVELOPED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT AGE/ DOB, WEIGHT, ETHNICITY & RACE UNKNOWN. EVENT DATE: UNKNOWN. DEVICE EXPIRATION DATE UNKNOWN. DEVICE EXPLANT DATE: UNKNOWN. INITIAL REPORTER OCCUPATION: UNKNOWN. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, INTO HER LEFT EYE (OS) ON (B)(6) 2009. THE REPORTER INDICATED A CATARACT HAD DEVELOPED AND THE LENS REMAINED IMPLANTED. THE SURGEON INDICATED THE PATIENT WAS LAST SEEN ON (B)(6) 2013. D6A: (B)(6) 2009. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587376&pc=MTA | 2023826-2023-00949 | 2023-02-14 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | UNK | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2013. THE LENS WAS REPORTED AS LOW VAULTING AND AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587382&pc=MTA | 2023826-2023-00948 | 2023-02-14 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | UNK | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS REPORTED AS LOW VAULTING AND AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16891772&pc=MTA | 2023826-2023-01749 | 2023-04-13 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS LOW VAULTING WITHOUT ROTATION AND LENS CATARACT (CORTICAL OBSERVED ON (B)(6) 2023). LENS REMAINS IMPLANTED. PROBLEM IS NOT RESOLVED. IT WAS ADDITIONALLY NOTED "GOOD. NO COMPLAINTS PER PATIENT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15891259&pc=MTA | 2023826-2022-04271 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM MICL 12.1 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. NARROW VAULT AND PATIENT NEEDED CE (CATARACT EXTRACTION) OD DUE TO CONTACT WITH ICL. FORMING CATARACT. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: HEALTH IMPACT:(B)(4) - CATARACT EXTRACTION INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM MICL 12.1-7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. NARROW VAULT AND PATIENT NEEDED CE (CATARACT EXTRACTION) OD DUE TO CONTACT WITH ICL FORMING CATARACTS. GLARES AND HALOS AND BLURRED VISION REPORTED. THE LENS WAS REMOVED ON 14-NOV-2022. PROBLEM RESOLVED. CAUSE WAS REPORTED AS DEVICE. PANOPTIX AND IOL IMPLANTATION IN PLACE. YAG CAP SCHEDULED. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: D6B: LENS EXPLANTED (B)(6) 2022. H6: 1494- ACD<3.00, AGE >45 AT TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687738&pc=MTA | 2023826-2020-02387 | 2020-07-31 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, THE TIP OF THE KNIFE TOUCHED THE ANTERIOR LENS CAPSULE. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED INTRA-OPERATIVELY WITHIN INITIAL SURGERY; DUE TO LENS OPACITY ASC. PHACOEMULSIFICATION & IOL IMPLANTATION ALSO REPORTED. THE REPORTER STATES THE PROBLEM WAS RESOLVED. H6 - HEALTH IMPACT CODE 4625-ADDITIONAL SURGERY (PHACOEMULSIFICATION & IOL IMPLANTATION). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934940&pc=MTA | 2023826-2016-01143 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934943&pc=MTA | 2023826-2016-01144 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/12.5. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645517&pc=MTA | 2023826-2017-00956 | 2017-03-22 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE REPORTER INDICATED LENS OPACITY (ANTERIOR) WAS FIRST OBSERVED ON (B)(6) 2017 AND THE LENS WAS EXPLANTED DUE TO LOW VAULTING ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT HAD ADDITIONAL SYMPTOMS - MILD BLURRING. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652095&pc=MTA | 2023826-2016-00641 | 2013-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW/NO VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LOW VAULT OF THE LENS WAS FIRST OBSERVED ON (B)(6) 2013. LENS OPACITY WAS OBSERVED (1 TR ASC) AND PATIENT EXPERIENCED BLURRY VISION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2, WITH NO DISCOMFORT OR COMPLAINTS. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054105&pc=MTA | 2023826-2016-01523 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Corneal Edema; Hyphema; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 LENS, -3.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULTING, RESULTING IN ELEVATED INTRAOCULAR PRESSURE (IOP), SHALLOW ANTERIOR CHAMBER, CORNEAL EDEMA, HYPHEMA, AND DEVELOPMENT OF A CATARACT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A SMALL TEAR ON THE HAPTIC. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218705&pc=MTA | 2023826-2015-01478 | 2015-08-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY, PUPIL BLOCK, ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS OCCURRED WITH EXPLANT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE: FIBERS. (B)(4): ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS. CONCLUSION: THE PATIENT IS UNDER THE AGE OF 21 YEARS. THA PATIENT'S EYE HAS ANTERIOR CHAMBER DEPTH (ACD), LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328083&pc=MTA | 2023826-2015-01750 | NA | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HIS VISION HAS CHANGED TO -0.5. THE FACILITY REPORTED THE PATIENT HAS AN EARLY CATARACT, DUE TO THE AGING PROCESS AND NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH RESULTS: PER MEDICAL REVIEW: ON (B)(6) 2015 THE PATIENT COMPLETED 36 MONTH POST-TREATMENT PATIENT FOLLOW-UP FORM AND REPORTED DECREASE OF VISUAL ACUITY FOLLOWING ICL IMPLANTATION. ACCORDING TO DCR FORM, RECEIVED ON 11/19/15, THE CAUSE OF THE EVENT WAS A CATARACT DEVELOPMENT DUE TO AGING PROCESS AND NOT RELATED TO THE DEVICE. ICL REMAINED IMPLANTED AND UCVA WAS 20/25 WITH GOOD PROGNOSIS. SERIOUSNESS IS BASED ON NO REPORTED LOSS OF BCVA AND NO REPORT OF ANY SECONDARY INTERVENTION PLANNED OR SCHEDULED. IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE", THEREFORE THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS IMPLANTED. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760961&pc=MTA | 2023826-2015-00534 | 2014-02-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device remains implanted | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED PATIENT HAD A MICL 13.6 MM IMPLANTABLE COLLAMER LENS IMPLANTED IN THE PATIENT'S LEFT EYE ON (B)(6) 2012. AS PART OF THE POST MARKET STUDY, THE PATIENT REPORTED EXPERIENCING HAVING LOST ACUITY IN HIS NEAR VISION AFTER THE SURGERY AND HAVE DIFFICULTY FOCUSING WITHOUT THE USE OF CONTACTS OR GLASSES TO BALANCE HIS EYESIGHT. ALSO BEING TOLD BY HIS OPTOMETRIST THAT SHE CAN SEE THE BEGINNINGS OF CATARACTS. THE PHYSICIAN'S OFFICE WAS CONTACTED FOR FURTHER INFORMATION. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2014. THE PATIENT IS NOW HYPEROPIC. STATED ADVERSE EVENT IS RELATED TO THE DEVICE. PROGNOSIS IS GOOD. VISUAL ACUITY POST-TREATMENT BCDVA 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). ICL REMAINS IMPLANTED. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536601&pc=MTA | 2023826-2016-00303 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER RIGHT EYE (OD) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY CONFIRMED THE PATIENT REPORT AND INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015. THE TYPE OF CATARACT WAS CORTICAL PERIPHERAL. THE LENS HAD A LOW VAULT AND THIS PROBABLY WAS THE CAUSE OF THE SUBCAPSULAR HAZE. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328002&pc=MTA | 2023826-2015-01751 | 2014-04-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR AND POST SUBCAPSULAR). THE LENS WAS EXPLANTED ON (B)(6) 2014, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S VA WAS 20/25 AND THE PROGNOSIS IS GOOD. Manufacturer Narrative: PER MEDICAL REVIEW: REPORTEDLY, ICL WAS EXPLANTED FROM A (B)(6) - YEAR OLD PATIENT, TWO YEARS POSTOPERATIVELY, TO ADDRESS CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST (DCR), RECEIVED ON 11/19/15, CATARACT WAS EXTRACTED AND IOL WAS IMPLANTED AT THE SAME SURGERY DATE. THE SAME REPORTED ATTRIBUTED CATARACT (ANTERIOR SUBCAPSULAR COMBINED WITH POSTERIOR SUBCAPSULAR) TO THE DEVICE. THE PROGNOSIS AFTER SECONDARY SURGICAL INTERVENTION WAS GOOD (VA 20/25). A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084245&pc=MTA | 2023826-2015-01194 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE IS STILL SEDIMENT (VISCOELASTIC) ON THE CRYSTALLINE LENS CENTER. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION: IMPLANTED DATE - (B)(6) 2015. ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION ((B)(4)) AND OCUCOAT ((B)(4)) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536606&pc=MTA | 2023826-2016-00302 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT REPORTED THE EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON REPORTED THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROGNOSIS IS EXCELLENT. THE VA POST-OP WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536972&pc=MTA | 2023826-2016-00300 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS REMOVED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2015. THE REPORTER INDICATED THE CATARACT WAS A CORTICAL CENTRAL AND THE EVENT WAS NOT DEVICE RELATED. AN IOL WAS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327964&pc=MTA | 2023826-2015-01752 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE LENS WAS EXPLANTED ON (B)(6) 2015, THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: PER MEDICAL REVIEW: REPORTEDLY ICL WAS EXPLANTED 3 YEARS POSTOPERATIVELY TO ADDRESS LOW VAULT AND A CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST, DATED ON (B)(6) 2015, THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR AND WAS ATTRIBUTED TO THE DEVICE. FOLLOWING ICL REMOVAL, CATARACT WAS EXTRACTED, AND IOL WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE PATIENT PROGNOSIS WAS GOOD (VA:20/20). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILIAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5139024&pc=MTA | 2023826-2015-01342 | NA | Injury | STAAR SURGICAL COMPANY AG | 2015-10-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT, THE DEVELOPMENT OF A CATARACT AND DECREASED VISUAL ACUITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER MEDICAL REVIEW - CENTRAL AND/OR PERIPHERAL CONTACT BETWEEN THE ICL AND THE CRYSTALLINE LENS IS ASSOCIATED WITH ANTERIOR SUBCAPSULAR CATARACT FORMATION. HOWEVER, THIS REPORT PROVIDES NO INFORMATION ON THE DEGREE OF VAULTING OR THE POSITION OF THE LENS OPACITIES. ADDITIONALLY, THERE IS NO INFORMATION ABOUT THE CRYSTALLINE LENS PRIOR TO THE INITIAL SURGERY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284955&pc=MTA | 2023826-2017-00114 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(6) 1966. NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284956&pc=MTA | 2023826-2017-00115 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(4). DATE OF BIRTH - (B)(6) 1966. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047775&pc=MTA | 2023826-2015-01116 | 2015-02-13 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (CORTICAL) AND A CAPSULAR TEAR. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293145&pc=MTA | 2023826-2015-01653 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -22.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2008. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/40. Manufacturer Narrative: AGE AT TIME OF EVENT IS (B)(6) WHICH IS INCORRECT, CORRECTED AGE AT TIME OF EVENT IS (B)(6). (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THERE WAS NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: DIOPTER: -22.00. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826213&pc=MTA | 2023826-2016-00983 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. THE PATIENT'S NATURAL LENS WAS ALSO REMOVED AND REPLACED WITH A DIFFERENT INTRAOCULAR LENS (IOL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. EXPLANT DATE: UNK. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680358&pc=MTA | 2023826-2016-00764 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER STATED THE PATIENT IS DOING WELL AND THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680364&pc=MTA | 2023826-2016-00765 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER STATED THE PATIENT IS DOING WELL AND THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533731&pc=MTA | 2023826-2016-00354 | 2014-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -06.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2014. THE SURGEON PERFORMED CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293227&pc=MTA | 2023826-2015-01652 | 2015-05-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -17.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015.THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015 . PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DIOPTER: -17.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553204&pc=MTA | 2023826-2016-00408 | 2015-11-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE; MEASUREMENT WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873542&pc=MTA | 2023826-2016-01065 | 2016-06-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT AND REFRACTIVE SURPRISE. THE LENS WAS REMOVED, CATARACT SURGERY WAS PERFORMED AND A TORIC IOL WAS IMPLANTED WITH GOOD RESULTS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE DIOPTER WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS REPORT SOURCE 'USER FACILITY' IS INCORRECT. THE CORRECT REPORT SOURCE IS 'DISTRIBUTOR'. DEVICE CODE: (B)(4) (EXPLANTED) TO (B)(4)- PREVIOUS CODE NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).DEVICE NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533727&pc=MTA | 2023826-2016-00353 | 2012-05-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2015. THE SURGEON PERFORMED CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084270&pc=MTA | 2023826-2015-01197 | 2015-05-15 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -20.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975337&pc=MTA | 2023826-2016-01317 | 2016-08-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSION OF INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7427120&pc=MTA | 2023826-2018-00550 | NA | Injury | STAAR SURGICAL COMPANY | 2018-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract, Induced; No Known Impact Or Consequence To Patient | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM MICL12.1, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2017. THE REPORTER STATED THAT THE SURGEON WILL BE REMOVING THE LENS BECAUSE IT IS TOO FLAT AND IT IS CAUSING A CATARACT. A LONGER LENS IN THE SAME POWER WAS ORDERED TO REPLACE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE SURGEON HAS DECIDED NOT TO IMPLANT A LARGER LENS SINCE "IS LIKELY TO CAUSE NARROW ANGLES" AND WILL MONITOR AND CONSIDER CATARACT SURGERY AS NEEDED IN THE FUTURE. (B)(4) Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327956&pc=MTA | 2023826-2015-01753 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Retinal Tear; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE PATIENT HAD A SMALL RETINAL TEAR AND IT WAS REPAIRED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, (B)(6) YEAR OLD PATIENT RECEIVED ICL IN 2012 AND WAS COMPLAINING ABOUT DECREASE IN VA DUE TO CATARACT DEVELOPMENT BY COMPLETING 36 MONTH FOLLOW-UP FORM. THE PATIENT ALSO REPORTED SMALL RETINAL TEAR THAT WAS REPAIRED. DATA CLARIFICATION REQUEST FORM WAS RECEIVED ON 11/19/2015, STATING THAT ANTERIOR SUBCAPSULAR CATARACT WAS CAUSED BY LOW VAULT OF ICL DECREASING THE VA FROM 20/15 TO 20/20. ICL REMAINED IMPLANTED AND NO SECONDARY OR SURGICALLY INTERVENTION WAS PERFORMED. HOWEVER, THE SMALL RETINAL TEAR, THAT WAS TREATED AND RESOLVED BY CRYOTHERAPY, WAS NOT CAUSED BY DEVICE BUT BY PATIENT RELATED FACTOR (VITREOUS DEGENERATION). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) YEARS OLD AND PERF DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE "THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4) LENS IMPLANTED. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12449121&pc=MTA | 2023826-2021-03482 | NA | Injury | STAAR SURGICAL COMPANY | 2021-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT #2023826-2021-03481 AND THE PROBLEM WAS NOT RESOLVED. THE PATIENT NOW HAS A SMALL ANTERIOR SUBCAPSULAR CATARACT AND VISION IS WORSE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900067&pc=MTA | 2023826-2020-02747 | NA | Injury | STAAR SURGICAL COMPANY AG | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE REPORTER INDICATED THE CATARACT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: IT WAS ALSO REPORTED THAT PHACOEMULSIFICATION WAS PERFORMED ON THE PATIENT. SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PHACOEMULSIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848263&pc=MTA | 2023826-2022-04295 | 2022-10-06 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -14.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS LENS OPACITY OF ASC AND PSC WHICH WAS OBSERVED ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. THE PROBLEM AS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "PATIENT REPORTED BEING HIT HARD IN THE LEFT EYE AND HER VISION BEGAN TO BLUR/GET HAZY". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547267&pc=MTA | 2023826-2023-03522 | 2023-03-22 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158233&pc=MTA | 2023826-2020-01235 | 2020-03-23 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS EYE ON (B)(6) 2020. THE PATIENT COMPLAINED OF BLURRED VISION. EXPLANTATION OF THE LENS AND CLEAR LENS SURGERY IS PLANNED FOR (B)(6) 2020. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4 - CORRECTED TO (B)(6) 2020. B5 - "IT WAS REPORTED THAT CLEAR LENS SURGERY AND IMPLANTATION OF NON-STAAR LENS WAS PERFORMED, WHICH RESOLVED THE ISSUE." SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CORRECTED DATA: H6 - DEVICE CODE:1401 SHOULD HAVE BEEN INCLUDED. B3 - CORRECT DATE SHOULD BE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECT DATE SHOULD BE REVERTED BACK TO (B)(6) 2020. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED AS IT IS N/A. H6 - PATIENT CODE: 2137 SHOULD BE REMOVED AS IT IS N/A. B5 - "THE PATIENT COMPLAINED OF BLURRED VISION", SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS N/A. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914025&pc=MTA | 2023826-2021-04571 | NA | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN HIS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14252557&pc=MTA | 2023826-2022-01313 | 2022-02-15 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022 WITH PAINLESS PROGRESSIVE VISION LOSS IN EYE FOR HALF A YEAR. LATER, ON (B)(6) 2022, THE PATIENT RECEIVED ICL IOL EXTRACTION FOR PHACOEMULSIFICATION EXTRACTION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14253045&pc=MTA | 2023826-2022-01314 | 2022-02-15 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT HAD PAINLESS PROGRESSIVE VISION LOSS, AND WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022. ON (B)(6) 2022, THE PATIENT UNDERWENT ICL IOL EXTRACTION, PHACOEMULSIFICATION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13374207&pc=MTA | 2023826-2022-00351 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). IT WAS INDICATED THAT THE LENS WAS TOO FLAT AND RESULTED IN MINIMAL CLEARANCE OVER THE NATURAL LENS, AND THERE WAS A SLIGHT CATARACT WITH A "SOMEWHAT" HYPEROPIC OUTCOME. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PATIENT IS DOING QUITE WELL. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12904968&pc=MTA | 2023826-2021-04570 | NA | Injury | STAAR SURGICAL COMPANY | 2021-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A (B)(4) IMPLANTABLE (B)(6) LENS, -12.00 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE PATIENT SEES HALOES IN HIS RIGHT EYE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, EXPLANT DATE: UNKNOWN, INITIAL REPORTER OCCUPATION: UNKNOWN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191994&pc=MTA | 2023826-2016-01814 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT COMPLAINED OF BLURRY VISION AT NIGHT. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING FINE. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN PIECES AND UNABLE TO EVALUATE. THE LENS WAS RETURNED DRY. METHOD CODE(S): LENS WORK ORDER SEARCH. RESULTS CODE(S): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4959171&pc=MTA | 2023826-2015-00960 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY | 2015-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY) AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED BY PHACOEMULSIFICATION AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140284&pc=MTA | 2023826-2016-01711 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. AT 3 MONTHS POST-OP, PATIENT WAS 20/25 UNCORRECTED AND THE PRESSURES WERE NORMAL. THE REPORTER INDICATED THE BEGINNING OF AN ANTERIOR SUBCAPSULAR CATARACT WAS NOTED AT THAT VISIT. THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140273&pc=MTA | 2023826-2016-01712 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINS IMPLANTED BUT MAY BE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY (ANTERIOR) WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. THE LENS DIOPTER WAS -13.0. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978092&pc=MTA | 2023826-2017-01620 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -3.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS OPACITY-ASCO. SURGEON ALSO PERFORMED PHACO-EMULSIFICATION AND IOL IMPLANTATION. SURGEON STATES THAT THE PATIENT "DEVELOPED ASCO AFTER 5 MONTHS SURGERY." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/25 AT THE LAST EXAM WHICH WAS A DECREASE COMPARED TO BASELINE BCVA WHICH WAS 20/22. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021462&pc=MTA | 2023826-2017-01676 | 2017-09-06 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LOSS OF BCVA, FROM 20/40 TO 20/100, AND NS AND CORTICAL LENS OPACITY WAS NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. ON (B)(6) 2017, THE PATIENT'S POST-OP UCVA IS 20/30 AND THE STATUS OF THE STATUS IS PERFECT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH ANY CATARACT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438479&pc=MTA | 2023826-2018-00544 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT THE PROBLEMS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM 12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS OF SIMILAR MODEL ON (B)(6) 2018 AND THE "VAULT WAS NORMAL AROUND 400UM APROX.". THE FINAL BCVA VALUE FOR PATIENT IS 0.8. EXPLANTED DATE: IT IS CORRECTED TO (B)(6) 2018. (B)(4). DEVICE MANUFACTURE DATE: IT IS CORRECTED TO 08-MAR-2011. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190944&pc=MTA | 2023826-2020-01388 | 2019-05-12 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: A CASE REPORT WAS RECEIVED AND INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -22.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND THE PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT FEMTOSECOND LASER ASSISTED CATARACT SURGERY (FLACS) COMBINED WITH ICL EXTRACTION AND AN IOL WAS IMPLANTED. VISUAL ACUITY WAS REDUCED FOR THE LAST 2 YEARS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252472&pc=MTA | 2023826-2015-01553 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS HAD A LOW VAULT, REFRACTIVE SURPRISE AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE: (B)(4) PREVIOUS CODE INCORRECT. NO INDUCED CATARACT WAS REPORTED. DEVICE CODE: (B)(4) (IMPLANTED)-NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. ADDITIONAL DATA: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTION TORN, HAPTIC BROKEN, AND HAIR ON LENS SURFACE. LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS, DATE UNKNOWN. PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/25. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074711&pc=MTA | 2023826-2016-01569 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR AND POSTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2016. THE PATIENT EXPERIENCED BLURRED VISION. THE PATIENT WAS SCHEDULED FOR YAG CAPSULOTOMY. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WERE FIBERS ON THE LENS. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT AND THE LENS WAS FOUND TO BE IN SPECIFICATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6742631&pc=MTA | 2023826-2017-01242 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE SURGEON ALSO NOTED THAT THE PATIENT'S VISUAL ACUITY WAS 20/60 ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) AND EXPLANTED ON (B)(6) 2016. THE CATARACT WAS FIRST OBSERVED ON (B)(6) 2016 AND WAS LOCATED IN THE NUCLEAR PART OF THE EYE. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. (B)(4). DATE OF EVENT - (B)(6) 2016. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870372&pc=MTA | 2023826-2018-01382 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC); PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. ACCORDING TO THE SURGEON, THE EXACT CAUSE OF THE CATARACT IS UNKNOWN BUT MAY HAVE OCCURRED DUE TO HIGH MYOPIA. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFGUE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL DATA RECEIVED ON 15-FEB-2019: " THE PATIENT ALREADY HAD IOL IN THE LEFT EYE DUE TO CATARACT FORMATION. PATIENT CAME BACK TO THE CLINIC LAST SUMMER COMPLAINING BLURRY VISION AND DOCOTR FOUND NSC, NOT ASCO. HE ADVISED TO REMOVE ICL AND IMPLANTED IOL, (B)(6) 2018" (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438480&pc=MTA | 2023826-2018-00543 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT PROBLEMS. Manufacturer Narrative: THE BEST CORRECTED VISUAL ACUITY (BCVA) FOR THE PATIENT IS CURRENTLY 1.0. (B)(4) Manufacturer Narrative: "SURGERY IS PLANNED TO CORRECT PROBLEMS." IN INITIAL MDR SHOULD BE UPDATED TO "THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. NO INFORMATION REGARDING THE DETAILS OF THE SECONDARY SURGERY HAS BEEN PROVIDED DESPITE THE REQUESTS FOR DETAILS." PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938159&pc=MTA | 2023826-2020-00798 | NA | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE VAULT IS MINIMAL. LASIK ENHANCEMENT WAS PERFORMED 9 MONTHS LATER. THE PATIENT HAS NOW DEVELOPED AN ANTERIOR CORTICAL CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER ON (B)(6) 2019. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D6: IMPLANTED DATE - (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611974&pc=MTA | 2023826-2019-00792 | 2019-02-23 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -1.75 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS FOLLOWING ICL IMPLANTATION. PHACOEMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION WAS PERFORMED ON THE RIGHT EYE 3 WEEKS AFTER EXPLANTATION. THE PATIENT PASSED AWAY ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722066&pc=MTA | 2023826-2022-01932 | 2022-05-09 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT HAS LENS OPACITY (CATARACT ASC). IT WAS ALSO NOTED "THE INVESTIGATOR DISCUSSED THE DEVELOPMENT OF ANTERIOR SUBCAPSULAR OPACITIES WITH THE PATIENT AND THE DECISION WAS MADE TO MONITOR THE VISION FOR NOW". THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4614" AND "2199" SHOULD BE REMOVED AND CODE "4627" AND "4625" SHOULD BE ADDED. H6-HEALTH EFFECT- CLINICAL CODE: "2140" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S LEFT EYE (OS) ON 03/JUNE/2020 AS A BACKUP LENS. ON 09/MAY/2022 LENS OPACITY CATARACT (ASC) WAS OBSERVED. THE PATIENT ALSO EXPERIENCED BLURRY/HAZY VISION. ON 29/AUG/2023 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "SUBJ INITIALLY REPORTED BLURRY & HAZY VA OS AT V7 (09MAY22). PI NOTED AN ASC OS. PI & SUBJ DECIDED TO MONITOR. AT V8 (28JUL23), SUBJ REPORTED VA HAD WORSENED, ESP IN BRIGHT LIGHTS. PI & SUBJ DECIDED TO PROCEED W/ CATARACT SX OS.SAE EVENT SUMMARY 2 : SUBJ HAD UNEVENTFUL ICL EXPLANT & SIMULTANEOUS CATARACT EXTRACTION SURGERY ON (B)(6) 2023. UCVA OS IMPROVED TO 20/12.5 AT 1 DAY PO. SITE WILL CONT TO FOLLOW SUBJ TO ENSURE HEALING AND STABILITY". THE CAUSE OF THE EVENT WAS THE DEVICE. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069167&pc=MTA | 2023826-2017-01792 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069168&pc=MTA | 2023826-2017-01791 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877264&pc=MTA | 2023826-2017-01463 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM12.0 IMPLANTABLE COLLAMER LENS, -21.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2016 THE LENS WAS EXPLANTED. THE SURGEON NOTED LOW VAULT, LENS ROTATION, LENS OPACITY ASCO, AND "ARTIPHAKIA" AS THE STATUS OF THE EYE. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION IS ALSO REPORTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT'S DATE OF BIRTH IS CORRECTED FROM "(B)(6) 1986" TO "(B)(6) 1983". BLURRED VISION AND PRESENCE OF GLARES/ HALOES EXPERIENCED BY THE PATIENT, SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.9MM)]. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816274&pc=MTA | 2023826-2016-00974 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, LENS OPACITY (CATARACT), AND LOSS OF VISION. THE PATIENT EXPERIENCED AN ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, BLURRED VISION, IRIS ATROPHY, AND CORNEAL EDEMA. THE ANTERIOR CHAMBER IRRIGATION OF REMAINING VISCO FLUIDS WAS PERFORMED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/200. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND SPOT ON LENS SURFACE. DIMENSIONAL INSPECTION CONLUDED THAT THE LENS IS IN SPECIFICATION. (B)(4). Manufacturer Narrative: NO REPORTED LOSS OF VISION. DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661678&pc=MTA | 2023826-2016-00727 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553032&pc=MTA | 2023826-2016-00401 | 2014-08-13 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -19.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE; MEASUREMENT FOUND THE LENS WITHIN SPECIFICATIONS. PREVIOUS (B)(4) (NEW INCISION) WAS INCORRECT; ONLY SECONDARY SURGERY FOR EXPLANT WAS REPORTED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140257&pc=MTA | 2023826-2016-01700 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING A CATARACT AND RECEIVING PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND THE LENS SCRATCHED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS OVER THE AGE OF (B)(6). CORRECTED DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING CATARACT. SECONDARY SURGICAL INTERVENTION, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION, WAS PERFORMED ON PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140262&pc=MTA | 2023826-2016-01699 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT DEVELOPING A CATARACT AND RECEIVING PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS OVER THE AGE OF 45 YEARS OLD. CORRECTED DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING CATARACT. SECONDARY SURGICAL INTERVENTION, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION, WAS PERFORMED ON PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293149&pc=MTA | 2023826-2015-01656 | 2013-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -15.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2013. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAIR ON SURFACE BUT NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: DIOPTER: -15.50. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266255&pc=MTA | 2023826-2015-01603 | 2014-12-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM12.1, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. A ANTERIOR SUBCAPSULAR CATARACT OPACITY WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A VICMO LENS, SAME LENGTH AND SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. THE CATARACT IS BEING MONITORED. POST-OP, PATIENT'S UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293173&pc=MTA | 2023826-2015-01654 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM 115V4 -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2007. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Event Description: CATARACT SURGERY AND IOL IMPLANTATION WERE PERFORMED ON (B)(6) 2015. Manufacturer Narrative: (B)(4).CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). Manufacturer Narrative: DATE OF EVENT: PREVIOUS DATE INCORRECT. DESCRIBE EVENT OR PROBLEM: PREVIOUS DATE INCORRECT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304862&pc=MTA | 2023826-2015-01615 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED AND OBSERVED NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (TISSUE). A DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305142&pc=MTA | 2023826-2015-01614 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378594&pc=MTA | 2023826-2015-01668 | 2015-04-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CORRECTIONAL DATA: PATIENT CODE - (B)(4) PREVIOUS CODE NOT APPLICABLE, UNKNOWN CAUSE FOR LENS OPACITY. PATIENT CODE - (B)(4) (NEW INCISION) - CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. PATIENT CODE - (B)(4) (EXPLANTED) - SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. DEVICE CODE - (B)(4) (EXPLANTED) - SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5292426&pc=MTA | 2023826-2015-01655 | 2015-08-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -19.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2005. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DIOPTER: -19.00; DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549348&pc=MTA | 2023826-2016-00387 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -22.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2013. AN ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED AND ON (B)(6) 2015, THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -22.00. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919743&pc=MTA | 2023826-2016-01203 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE REPORTER INDICATED THERE WAS ADHESION AND DEGENERATION OF THE IRIS. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/400. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5931126&pc=MTA | 2023826-2016-01226 | 2016-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM ICM120V4 12.0 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT RECEIVED PHACO-EMULSIFICATION (CATARACT SURGERY), AND AN IOL WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934314&pc=MTA | 2023826-2016-01225 | 2016-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM ICM120V4 12.0 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT RECEIVED PHACO-EMULSIFICATION (CATARACT SURGERY), AND AN IOL WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMGE AND DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923611&pc=MTA | 2023826-2016-01208 | 2016-05-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (SPOTS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. EXPLANTED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918280&pc=MTA | 2023826-2016-01181 | 2016-06-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (HAIRS) AND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PMA 510(K): - THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972078&pc=MTA | 2023826-2016-01311 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: SERIAL # (B)(4), EXPIRATION DATE - 10/31/2018. DEVICE MANUFACTURE DATE - 11/07/2015. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CONCLUSION CODE, AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553040&pc=MTA | 2023826-2016-00403 | 2015-06-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919728&pc=MTA | 2023826-2016-01201 | 2015-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -03.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS DEVICE IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PATIENT CODE: (B)(4) - PREVIOUS CODE NOT APPLICABLE; NO INDUCED CATARACT REPORTED. DEVICE CODE: (B)(4) (EXPLANTED) - CODE NOT APPLICABLE; SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919733&pc=MTA | 2023826-2016-01202 | 2015-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -03.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE THAT IS NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553044&pc=MTA | 2023826-2016-00405 | 2015-11-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955097&pc=MTA | 2023826-2016-01269 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 03/14/2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/21. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917886&pc=MTA | 2023826-2016-01178 | 2016-02-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 . THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS OPACITY. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED ON THE PATIENT, AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4): ONLY SECONDARY INTERVENTION, LENS EXPLANT WAS REPORTED. PATIENT ALSO HAD A PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE THAT IS NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: EVALUATION FOUND LENS ARE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579726&pc=MTA | 2023826-2016-00480 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), OBSERVED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE CORRECT EVENT DATE REPORTED WAS (B)(6) 2016'. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5971116&pc=MTA | 2023826-2016-01228 | 2016-08-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS OPACITY. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. CORRECTED INFORMATION: (B)(4). THE REPORTER DID NOT STATE THE CATARACT WAS INDUCED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5931212&pc=MTA | 2023826-2016-01224 | 2015-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT IS NOT MANUFACTURED IN THE U.S. AND NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002463&pc=MTA | 2023826-2018-01537 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO A NUCLEAR SCLEROTIC CATARACT AND A CAPSULAR TEAR. THE REPORTER INDICATED DURING THE SURGERY, THE PATIENT'S BILATERAL SUSPENSORY LIGAMENT WAS LOOSE AND THE LEFT ANTERIOR SUSPENSORY LIGAMENT WAS BROKEN. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENT WAS NERVOUS). Manufacturer Narrative: CORRECTED DATA: (B))(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908388&pc=MTA | 2023826-2019-01481 | 2019-03-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: H6- DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908395&pc=MTA | 2023826-2019-01482 | 2019-03-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15009430&pc=MTA | 2023826-2022-02285 | 2022-06-15 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022; AS A REPLACEMENT LENS FOR A PREVIOUS REFRACTIVE SURPRISE. THE PATIENT DEVELOPED LENS OPACITY (ASC) WHICH WAS ASSESSED ON (B)(6) 2022; CORNEAL/STROMAL OEDEMA AND DESCEMET FOLDS AND PHOTOPHOBIA WAS ALSO REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACO-EMULSIFICATION WAS PERFORMED AND AN IOL FROM A DIFFERENT MANUFACTURER WAS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. POSTOPERATIVE COMMENTS NOTED "THE UCVA AND BCVA HAS NOT INCREASED YET BECAUSE THERE ARE STILL CORNEAL & STROMAL OEDEMA AND DESCEMET FOLDS. THE PATIENT IS UNDER CLOSE MONITORING". THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "DURING THE EXCHANGE THE CRYSTALLINE LENS WAS TOUCHED WITH THE ICL MANIPULATOR BY THE SURGEON". Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464587&pc=MTA | 2023826-2023-00588 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) IN 2022. THE LENS WAS REMOVED DUE TO POOR VISION, SINEQUIAS OF THE IRIS TO THE CORNEA AND CORNEAL EDEMA WITH VERY LOW ECC COUNTING. THE LENS WAS REPLACED WITH A SHORTER LENS. POST-OP, THERE IS AN ANTERIOR SUBCAPSULAR CATARACT AND OVAL PUPIL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - SUTURE WAS OBSERVED. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) IN 2022. THE LENS WAS REMOVED DUE TO ANTERIOR SUBCAPSULAR CATARACT, PUPIL OVALIZATION, POOR VISION, SINEQUIAS OF THE IRIS TO THE CORNEA AND CORNEAL EDEMA WITH VERY LOW ECC COUNTING. THE LENS WAS REPLACED WITH A SHORTER LENS (12.1MM, -07.00 DIOPTER, IMPLANT VERTICAL), THEN ONE WEEK AFTER EXCHANGE NORMAL. A SUTURE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - POOR VISION, SINEQUIAS OF IRIS TO CORNEA, LOW ECC COUNTING, OVAL PUPIL. CLAIM # (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148280&pc=MTA | 2023826-2021-02469 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Endophthalmitis; Blurred Vision | N/A | NA | 1 | Event Description: A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE SURGEON REPORTS PERMANENT LOSS OF BCVA AND ENDOPHTHALMITIS. DIAGNOSTICS WERE PERFORMED; NO ORGANISM WAS FOUND. A VITRECTOMY, INTRAVITREAL ANTIBIOTICS, AND A VITREOUS TAP WERE PERFORMED (B)(6) 2021. ON (B)(6) 2021 A VITRECTOMY WAS PERFORMED AND THE LENS WAS EXPLANTED, PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED AS THE PATIENT HAD DEVELOPED A CATARACT. PROBLEM IS RESOLVED. Manufacturer Narrative: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519015&pc=MTA | 2023826-2019-00607 | 2019-03-06 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Endophthalmitis; Vitrectomy; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -16.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ONE DAY POST-OP THE PATIENT HAD EXAGGERATED INFLAMMATION AND EXUDATES IN THE EYE. THE PATIENT WAS DIAGNOSED WITH ACUTE ENDOPHTHALMITIS AND IMMEDIATELY WAS INJECTED WITH INTRAOCULAR ANTIBIOTICS. VITREOUS BIOPSY SAMPLE REVEALED PSEUDOMONAS INFECTION. AUTOCLAVE INDICATORS PASSED. OT AND AIR CULTURES REPORTED NEGATIVE. INSTRUMENTS AND CONSUMABLES REPORTED NEGATIVE FOR ANY MICROBIOLOGICAL GROWTH. VITREO-RETINAL SURGERY, CATARACT EXTRACTION, AND ICL EXPLANT PERFORMED. BCVA=6/18. INFECTION IS CURRENTLY INACTIVE. Manufacturer Narrative: REPORTER INDICATED THAT VITRECTOMY SURGERY WAS PERFORMED AND THE PATIENT'S STATUS IS STABLE, WITH VISION AT 6/24P. CLAIM#: (B)(4). Manufacturer Narrative: METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11435324&pc=MTA | MW5099834 | 2019-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-05 | MTA | STAAR SURGICAL ICL IMPLANT | Insufficient Information | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | NA | NA | 1 | Event Description: I HAD ICL SURGERY ON (B)(6) 2009. IN 2011, I STARTED TO HAVE ISSUES WITH MY VISION. I HAD TO START WEARING EYE GLASSES TO DRIVE AND THEN ENDED UP WITH -1.25 VISION TO THIS DATE. IN (B)(6) 2019 I STARTED TO HAVE AN ITCHING/POKING FEELING IN MY RIGHT EYE. MY EYE CONTINUED TO WATER AND ITCH AS IF SOMETHING WAS POKING ME. I VISITED SEVERAL EYE DOCTORS, INCLUDING MY SURGEON WHO HAD DONE MY ICL SURGERY. IT FEELS/FELT AS IF SOMETHING WAS WRONG WITH MY ICL. THE PAIN IS BEHIND MY EYE. THE SURGEON STATED THAT I HAVE A SMALL CATARACT NOW ON MY EYE AND THAT IS CAUSING MY BLURRY VISION AND THE ISSUE. I AM HAVING BLURRY VISION, BUT I AM STILL WITNESSING THE WATERY/ITCHING SENSATION. THE SURGEON DR. LIU ADVISED I NEED TO GET MY ICL IMPLANT TAKEN OUT AND INPUT A CATARACT LENSES. I STATED THAT I DO NOT FEEL THIS IS RIGHT. SOMETHING IS NOT RIGHT WITH THE CURRENT ILC LENSES INSIDE MY EYE. I HAVE SINCE DEALT WITH THIS CONTINUED EFFECT FOR THE PAST TWO YEARS. EYE DOCTORS CANNOT FIGURE IT OUT. IT IS NOT MY CATARACT CAUSING THIS SENSATION. IT'S ALMOST LIKE MY ICL IS TRYING TO POKE OUT OF MY EYE OR COME OUT. I AM FINALLY GETTING A SECOND OPINION THIS NEXT WEEK WITH ANOTHER EYE SURGEON TO CHECK MY ICL. TO SEE IF THAT IS IN FACT THE ISSUE, AS MY RIGHT EYE IS STILL ITCHING, WATERY. I CAN'T PUT A CONTACT IN MY RIGHT EYE AS IT IS TOO BOTHERSOME. I WANT TO REPORT THIS AS I FEEL AFTER THIS MANY YEARS PERHAPS THE LENSES HAS GONE BAD OR IT NEEDS REPLACED AFTER A CERTAIN TIME. WHEN I DID ICL BACK IN 2009 IT WAS NEW AND JUST BECOMING A THING. PERHAPS THE ICL'S BACK THEN ARE NOT AS UP TO DATE AS WHAT THEY HAVE CREATED NOW. STAAR SURGICAL ICL IMPLANTS. IMPLANT ICL: MICL12.1 RIGHT EYE. DIOPTER. -10.0/ OPTIC DIAMETER 4.9-5.8MM OVERALL DIAMETER: 12.1M. SN: (B)(4) LEFT EYE: DIOPTER.-10.5/ OPTIC DIAMETER 4.9- 5.8MM OVERALL DIAMETER: 12.2M SN: (B)(4). FDA SAFETY REPORT ID # (B)(4). | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620596&pc=MTA | 2023826-2021-00861 | NA | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2009. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY CORTICAL, REFRACTIVE CHANGE OVER TIME, THE PATIENT HAD TO WEAR GLASSES TO DRIVE, A WATERY/ITCHING/POKING FEELING AND PAIN IN THE EYE. BLURRY VISION ALSO REPORTED. THE SURGEON STATES THE PATIENT'S ACD CHANGED FROM 3.93 IN 2009 TO 3.39 IN 2020. THE PATIENT'S AXIAL LENGTH INCREASED. "NO ONE HAS CONTROL OVER THAT!" Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651554&pc=MTA | 2023826-2021-00885 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2021 DUE TO LOW VAULT, LENS OPACITY (ASC), BLURRED VISION AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DERIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. AFTER LENS EXPLANT, PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. HEALTH IMPACT CODE 4625. ADDITIONAL SURGERY (PHACOEMULSIFICATION & IOL IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814249&pc=MTA | 2023826-2020-00475 | 2020-01-17 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Headache; Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION, LOSS OF BEST CORRECTED VISUAL ACUITY AND LENS OPACITY (POSTERIOR SUBCAPSULAR) WAS NOTED. THE PATIENT IS SENSITIVE TO BRIGHT LIGHT AND HEADACHE WHEN READING. THE LENS REMAINED IMPLANTED AND PATIENT WILL RETURN IN 3 MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED WITH AN INTRAOCULAR LENS ON (B)(6) 2020. CATARACT SURGERY WAS ALSO PERFORMED ON (B)(6) 2020, DURING LENS REPLACEMENT. THE REPORTER INDICATED THAT THE PATIENT WAS SEEN ON (B)(6) 2020, AND THAT SHE IS "STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." H6 - HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY, CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090066&pc=MTA | 2023826-2018-01697 | 2018-09-02 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Hemorrhage/Bleeding; Hyphema; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC). REPORTEDLY, DURING THE INITIAL PROCEDURE, HEMORRHAGE OF AN IRIS VESSEL OCCURED. TO MANAGE THE IRIS BLEED THE SURGEON INTRODUCED BUBBLES INTO THE EYE. REPORTEDLY, THE BUBLLES WERE STILL PRESENT 3 DAYS LATER AND THE ICL WAS IN CONTACT WITH THE ANTERIOR CAPSULE. THE SURGEON STATES THE CAUSE OF THE ICL TOUCHING THE ANTERIOR CAPSULE WAS DUE TO THE BUBBLES, NOT THE ICL DEVICE. THE CAUSE OF THE IRIS HEMORRHAGE AND LENS OPACITY (ASC) IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CORRECTED DATA: EVENT - "HEMORRHAGE OF AN IRIS VESSEL OCCURRED" SHOULD BE CORRECTED TO "HEMORRHAGE OF AN IRIS VESSEL OCCURED LEADING TO HYPHEMA" CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6305835&pc=MTA | 2023826-2017-00223 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND WAS PRESCRIBED ORAL MEDICATIONS OR EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO INSIDE OF EYE WAS INFLAMED. INFORMATION RECEIVED FROM THE FACILITY CONFIRMED THE PATIENT REPORT AND THE DATE OF THE CATARACT DIAGNOSIS WAS (B)(6) 2016. THEY INDICATED PATIENT HAS EPISCLERITIS EXTERNALLY WHICH REQUIRED PRED FORTE TO CLEAR. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107373&pc=MTA | 2023826-2019-01752 | 2019-05-30 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED AND THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED AND A MONOFOCAL LENS WAS IMPLANTED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1932: INFLAMMATION. DEVICE CODE 1494:OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10714478&pc=MTA | 2023826-2020-02494 | 2020-01-15 | Injury | STAAR SUGICAL COMPANY | 2020-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; Insufficient Information | NA | NA | 1 | Event Description: IT WAS REPORTED THAT A LENS IS BEING RETURNED. THE REASON FOR RETURN IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (PSC, OBSERVED: (B)(6) 2020) AND CATARACT SURGERY WAS PERFORMED WITHIN THE SAME SURGERY AS REMOVAL. THE REPORTER STATED "PATIENT WAS SEEN ON (B)(6) 2020, STABLE FOLLOWING SURGERY, HAS BEEN INFORMED ABOUT THE MILD POSTERIOR CAPSULE OPACIFICATION AND THE NEED FOR A YAG IF THIS WORSENS. SHE DEVELOPED A CATARACT FROM LENS IMPLANTATION. THE CATARACT HAS BEEN REMOVED AND AN INTRAOCULAR LENS IMPLANTED. PATIENT WAS SEEN ON (B)(6) 2020, STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." FOR CAUSE OF EVENT THE REPORTER STATED "ICL RELATED - INITIAL LENS 13.2 WAS OVERSIZED AND THE EDGES MAY HAVE INCISED THE CAPSULE. THE 12.6 LENS WAS USED TO REPLACE THIS AT THE SAME SITTING BUT SHE DEVELOPED A CATARACT WITHIN A FEW DAYS. THE 12.6 WAS ALSO TOO SHALLOW. IDEALLY SHE WOULD HAVE HAD A 12.9." H6- WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED FROM THE REPORTER THAT PREVIOUS INFORMATION PROVIDED WAS INTENDED FOR CLAIM# (B)(4), MEDWATCH REPORT# 2023826-2020-00475, THERE IS NO COMPLAINT FOR LENS SN# (B)(6). REPORTED INDICATED "LEFT EYE HAS ICL IN AND VISION WAS GOOD STRAIGHT AWAY POST OP AND IS STILL NOW, NO ISSUES". CLAIM#: (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516827&pc=MTA | 2023826-2021-03729 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1494: OFF-LABEL USE: ACD<3.0MM CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098622&pc=MTA | 2023826-2020-03254 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND PUNTIFORM OPACITY IN THE POSTERIOR CAPSULE. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: NON RECOMMENDED OVD AND EXCESSIVE USE OF OVD DUE TO A NON TOTAL DILATION DURING SURGERY. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. REPORTEDLY, THE IOP WAS MANAGED AFTER THE EVACUATION OF OVD. PUNTIFORM OPACITY HAS DECREASED AFTER 48 HOURS OF SURGERY. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PROBLEM IS NOW RESOLVED. (B)(4). Manufacturer Narrative: B5. ELEVATED IOP IS AN EXPECTED OR FORESEEABLE SIDE EFFECT. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669737&pc=MTA | 2023826-2020-02329 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE. THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - LOW VAULT, TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669743&pc=MTA | 2023826-2020-02328 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE . THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS REPORTED AS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. HOWEVER, MILD PUPIL CORECTOPIA IS REPORTED POST EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155887&pc=MTA | 2023826-2017-01982 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, ON (B)(6) 2015 IN THE PATIENT'S RIGHT EYE (OD). LENS OPACITY (ASCO) WAS OBSERVED ON (B)(6) 2017,ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE CHANGE OVERTIME DUE TO CATARACT WAS ALSO NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS REPORTED. Manufacturer Narrative: FINAL PATIENT STATUS IS REPORTED AS "GOOD RESULT, PATIENT IS OK" WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: REPORTER STATED THAT "THE LENS IS NOT AVAILABLE (DISCARDED)". DATE OF THIS REPORT: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM "19-JUN-2017" TO "19-JUN-2014". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772247&pc=MTA | 2023826-2019-01176 | 2019-06-08 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "DUE TO FLOPPY IRIS AND HIGH PRESSURE RESULTING IN ANGLE CLOSURE", THE LENS WAS REMOVED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 1426, SHOULD BE REMOVED FROM SUPPLEMENTAL #2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD HAVE BEEN ADDED IN SUPPLEMENTAL MEDWATCH REPORT. WORK ORDER SEARCH: "NO SIMILAR COMPLAINTS" SHOULD BE CORRECTED TO "THREE SIMILAR COMPLAINTS." RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION :THE REPORTER INDICATED THAT EXCESSIVE VAULT (INITIALLY 1.5, THEN CHANGED TO 5) WAS OBSERVED. PUPIL BLOCK WITH ELEVATED IOP (55 MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (55 MMHG) WERE OBSERVED ON (B)(6) 2019, LENS OPACITY (ASC) OBSERVED ON (B)(6) 2019. REPORTEDLY ENLARGEMENT OF PI AND ADDITION OF NEW PI WERE SURGICALLY PERFORMED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND "TOPOMAX RELATED SHIFT? SURGICALLY ENLARGED PI FAILED TO RESOLVE EDEMATUOUS CB WORSENED BY TOPOMAX?" PER REPORTER ON (B)(6) 2019, "THERE IS CATARACT FORMATION AND PATIENT IS SCHEDULED FOR CE WITH IOL IMPLANTATION." (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6859975&pc=MTA | 2023826-2017-01422 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LENS OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. THE LENS WAS EXCHANGED. A POSSIBLE IRIDOPLASTY IS PLANNED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. THE REPORTER ALSO STATED THAT THE SOFTWARE DOES NOT ALWAYS PREDICT CORRECT SIZE IMPLANT. (B)(4). CORRECTION DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LEN OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS AND A PIECE OF THE HAPTIC MISSING. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233814&pc=MTA | 2023826-2019-00072 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE INITIAL LENS WAS NOT EXCHANGED FOR ANOTHER LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233815&pc=MTA | 2023826-2019-00071 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. A SECONDARY YAG WAS PERFORMED FOR ENLARGEMENT OF THE PERIPHERAL IRIDECTOMIES. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. SEE MFR REPORT # 2023826-2019-00073 FOR INITIAL LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024717&pc=MTA | 2023826-2016-01448 | NA | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PRESSURE AND PAIN FOR PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRY VISION. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. POST OP THE PATIENT HAD TRACE CORTICAL OPACITY AND BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE VISUAL ACUITY WAS MORE STABLE BUT THE PUPIL WAS DILATED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291699&pc=MTA | 2023826-2017-00197 | NA | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291665&pc=MTA | 2023826-2017-00196 | NA | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Missing Value Reason | Cataract; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | unknown | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358993&pc=MTA | 2023826-2017-00252 | 2015-09-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED AND REPLACED. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THEY CONFIRMED THE PATIENTS REPORT. THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS EXPLANTED ON (B)(6) 2015, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. THE FACILITY REPORTED THE EVENT WAS RELATED TO THE DEVICE. Manufacturer Narrative: (B)(4). METHOD: LENS WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6408305&pc=MTA | 2023826-2017-00414 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2016 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE DATE OF EXPLANT WAS (B)(6) 2016, THE PATIENTS POST-OP VA WAS 20/20 AND PROGNOSIS WAS EXCELLENT. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601777&pc=MTA | 2023826-2015-00318 | NA | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED AND THE CATARACT REMOVED ON (B)(6) 2015. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - EVEN THOUGH THE POSITION OF ICL IMPLANTATION COULD CONTRIBUTE TO CATARACT, OTHER FACTORS COULD ALSO CAUSE THE DEVELOPMENT OF CATARACT SUCH AS AGING, TRAUMA (EYE RUBBING, EXTERNAL FORCES), DIABETES, SMOKING AND DRINKING, DRUG USE (STEROIDS) AND RADIATION EXPOSURE. ACCORDING TO THE LITERATURE, MYOPIA PATIENTS ARE AT EARLIER AGE OF DEVELOPING CATARACT THAN PEOPLE OF NORMAL REFRACTION. HIGH DEGREE MYOPIA PATIENT COULD BE 10 YEARS YOUNGER THAN NORMAL PEOPLE TO GET CATARACT. GIVEN THE AGE AND HIGH DEGREE OF MYOPIA (-15D) OF THIS PATIENT IT IS LIKELY THAT THE MEDICAL DEVICE IS NOT THE ROOT CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12215415&pc=MTA | 2023826-2021-02797 | 2021-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -16.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY ON (B)(6) 2021 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE INITIAL LENS WAS IMPLANTED (B)(6) 2016. ON (B)(6) 2021 LENS OPACITY PSC/NS WAS OBSERVED. THE LENS WAS REMOVED, PHACO-EMULSIFICATION PERFORMED AND AN IOL WAS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "POSTERIOR SUBCAPSULAR POLAR AGE-RELATED CATARACT." H6-HEALTH EFFICATED IMPACT CODE 4650 (PHACO-EMULSIFICATION, IOL IMPLANTATION). CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722429&pc=MTA | 2023826-2023-03743 | 2023-07-03 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM512.6, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS OPACITY(CATARACT) ASC OBSERVED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS HAPPY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5 - THE REPORTER ALSO INDICATED THAT PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS ALSO PERFORMED THAT RESOLVED THE PROBLEM. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505351&pc=MTA | 2023826-2020-02057 | 2020-08-10 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE ( LENS TOO SMALL) WAS REPORTED TO BE THE CAUSE OF THIS EVENT. THE REPORTER STATED " LENS IS TOO CLOSE TO NATURAL LENS NOT SITTING CORRECTLY." LENS EXCHANGE IS PLANNED BUT NOT YET TAKEN PLACE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED AT A LATER DATE THAT THE PATIENT DEVELOPED CATARACTS. LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505372&pc=MTA | 2023826-2020-02056 | 2020-08-10 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE ( LENS TOO SMALL) WAS REPORTED TO BE THE CAUSE OF THIS EVENT. THE REPORTER STATED " LENS IS TOO CLOSE TO NATURAL LENS NOT SITTING CORRECTLY." LENS EXCHANGE IS PLANNED BUT NOT YET TAKEN PLACE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT DEVELOPED CATARACTS. THE LENS STILL REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062446&pc=MTA | 2023826-2023-04853 | 2023-09-07 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY ASC THAT WAS OBSERVED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222347&pc=MTA | 2023826-2021-02864 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY (ASC, OBSERVED: (B)(6) 2021). "ANTERIOR SUBCAPSULAR CATARACT FORMATION TO EDGE OF VISUAL AXIS" WAS REPORTED. IT WAS REPORTED THAT THE DOCTOR NOTICED THE LENS OPACITY WAS WORSENING AND WANTED TO EXCHANGE ICL FOR A DIFFERENT SIZE, SO IT WOULDN'T RUB ON THE LENS. THE REPORTER STATED, "LENS OPACITY IS STILL PRESENT, WE HOPE IT STAYS STABLE AND WE DON'T NOTICE ANY IMMEDIATE CHANGES." CAUSE OF THIS EVENT WAS REPORTED TO BE THE DEVICE. CAUSE DETAILS "SMALL ICL." BCVA WAS REPORTED TO BE (B)(6) 2021. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222348&pc=MTA | 2023826-2021-02865 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM,MICL12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT AN ICL EXCHANGED IS SCHEDULED FOR THE PATIENT, CAUSE IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY (ASC, OBSERVED: (B)(6) 2021). "ANTERIOR SUBCAPSULAR CATARACT FORMATION TO EDGE OF VISUAL AXIS" WAS REPORTED. IT WAS REPORTED THAT THE DOCTOR NOTICED THE LENS OPACITY WAS WORSENING AND WANTED TO EXCHANGE ICL FOR A DIFFERENT SIZE, SO IT WOULDN'T RUB ON THE LENS. THE REPORTER STATED, "LENS OPACITY IS STILL PRESENT, WE HOPE IT STAYS STABLE AND WE DON'T NOTICE ANY IMMEDIATE CHANGES." CAUSE OF THIS EVENT WAS REPORTED TO BE THE DEVICE. CAUSE DETAILS "SMALL ICL." BCVA WAS REPORTED TO BE 20/70 ON (B)(6) 2021. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFO:UNK. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012227&pc=MTA | 2023826-2020-03022 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY AG | 2020-12-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2003. LENS OPACITY (ASC WAS OBSERVED ON 15/OCT/2020. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2003. LENS OPACITY (ASC) WAS OBSERVED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LENS OF A DIFFERENT MODEL WITH A LONGER LENGTH ON (B)(6) 2020 WHICH RESOLVED THE PROBLEM. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE" AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271733&pc=MTA | 2023826-2023-05506 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. A PLANNED LENS REMOVAL AND REPLACMENT WAS REPORTED, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354764&pc=MTA | 2023826-2019-00272 | 2013-07-10 | Injury | STAAR SURGICAL COMPANY AG | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL 12.1, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 DUE TO LENS OPACITY, CORTICAL, BEING OBSERVED (B)(6) 2018. IT WAS REPORTED THE LENS WAS REMOVED TO PREVENT THE PROGRESSION THE CORTICAL CATARACT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354778&pc=MTA | 2023826-2019-00273 | 2013-07-25 | Injury | STAAR SURGICAL COMPANY AG | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. IT WAS REPORTED THAT PATIENT IS DEVELOPING EARLY CORTICAL CATARACTS AND THE ICL WILL BE REMOVED AND EXCHANGED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. IT WAS REPORTED THAT PATIENT STARTED DEVELOPING EARLY CORTICAL CATARACTS. THE ICL WAS REMOVED ON (B)(6) 2019. THE ICL WAS EXCHANGED FOR A SAME MODEL LARGER LENGTH LENS WITH SAME DIOPTER. THE REPORTER WILL NOT KNOW IF THE ISSUE WAS RESOLVED UNTIL FURTHER FOLLOW UP. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153415&pc=MTA | 2023826-2016-01727 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AS THE SURGEON WAS IMPLANTING THE LENS, THE PATIENT TURNED AND THE TIP OF THE SFC-45 FP CARTRIDGE TORE THE ANTERIOR CAPSULE. A CATARACT DEVELOPED AND THE LENS WAS EXPLANTED ON (B)(6) 2016. CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: (B)(4). RESULTS CODES: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6560729&pc=MTA | 2023826-2017-00806 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE SURGEON NOTED A CATARACT. THE LENS WAS CUT TO REMOVE WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. A SUTURE WAS REQUIRED. THE PATIENT'S POST-OP UCVA WAS 20/30. THE REPORTER STATED THERE WAS NO PROBLEM WITH THE LENS VAULT AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7207877&pc=MTA | 2023826-2018-00107 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.1, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2009. ON (B)(6) 2017, THE LENS WAS REMOVED BECAUSE THE PATIENT DEVELOPED A CATARACT. CATARACT REMOVAL WAS PERFORMED AT THE SAME TIME OF ICL REMOVAL AND A MX60 ( BAUSCH + LOMB) INTRAOCULAR LENS WAS IMPLANTED. IT IS REPORTED THAT INCIDENT/INJURY WAS NO PRODUCT RELATED, THERE WAS NO PROBLEM WITH THE LENS. PATIENT IS DOING GOOD. Manufacturer Narrative: LENS WAS RETURNED IN A SPECIMEN CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: IN FOLLOW-UP #1, CORRECTED TO 10-JAN-2018. (B)(4) Manufacturer Narrative: CONCLUSION: (B)(4) "THE LENS WAS REMOVED BECAUSE THE PATIENT HAD A CATARACT. NO PROBLEMS WITH LENS." (B)(4) Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7288989&pc=MTA | 2023826-2018-00264 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-02-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON EXPLANTED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS FROM THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED DUE TO A CATARACT. CATARACT SURGERY WAS PERFORMED AT THE SAME TIME OF ICL REMOVAL AND A SYMFONY (AMO) INTRAOCULAR LENS WAS IMPLANTED. IT IS REPORTED THAT THE CAUSE OF EVENT WAS NOT PRODUCT RELATED, THERE WAS NO PROBLEM WITH THE LENS. PATIENT STATUS WAS REPORTED AS "GOOD". Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL 13.2; -11.5 DIOPTER; IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED FROM THE EYE ON (B)(6) 2018 DUE TO CATARACT. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8274513&pc=MTA | 2023826-2019-00158 | 2018-05-10 | Injury | STAAR SURGICAL COMPANY AG | 2019-01-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR/NUCLEAR SCLEROTIC/CORTICAL). PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8210581&pc=MTA | 2023826-2019-00012 | 2018-11-01 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 LENS OPACITY (ASC) WAS OBSERVED AND THE LENS WAS REMOVED ON 15 NOV 2018. IT WAS REPORTED THAT BECAUSE THE PATIENT IS RELATIVELY OLD AND OPACITY OF THE LENS IS MILD, THE PATIENT IS BEING CORRECTED WITH GLASSES. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED ON 15-FEB-2019: "THIS WAS THE SURGEON'S FIRST CASE OF ICL IMPLANTATION. DR. (B)(6) DECIDED TO EXPLANT BOTH LENSES DUE TO TWO REASONS: PATIENT'S OVER REFRACTION (OR) WAS +1.50, OU AND OBVIOUSLY THIS PRESBYOPIC PATIENT WAS NOT HAPPY AT ALL; DR. (B)(6) FOUND ASCO IN THE PERIPHERY IN BOTH LENSES. AFTER EXPLANTATIONS, PATIENT'S VISION IS CORRECTED BY SPECTACLE AND HER BCVA IS 0.8, OU. NO FURTHER SURGICAL OPTION IS PLANNED. DR. ADVISED FUTURE IOL IMPLANTATION IN CASE SHE DEVELOPS CATARACT." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00011 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00012 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8210647&pc=MTA | 2023826-2019-00011 | 2018-11-01 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 LENS OPACITY (ASC) WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. IT WAS REPORTED THAT BECAUSE THE PATIENT IS RELATIVELY OLD AND OPACITY OF THE LENS IS MILD, THE PATIENT IS BEING CORRECTED WITH GLASSES. Manufacturer Narrative: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. SEE MDR# 2023826-2019-00012 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. SEE MDR# 2023826-2019-00011 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED ON 15-FEB-2019: "THIS WAS THE SURGEON'S FIRST CASE OF ICL IMPLANTATION. DR. (B)(6) DECIDED TO EXPLANT BOTH LENSES DUE TO TWO REASONS: PATIENT'S OVER REFRACTION (OR) WAS +1.50, OU AND OBVIOUSLY THIS PRESBYOPIC PATIENT WAS NOT HAPPY AT ALL; DR. (B)(6) FOUND ASCO IN THE PERIPHERY IN BOTH LENSES. AFTER EXPLANTATIONS, PATIENT'S VISION IS CORRECTED BY SPECTACLE AND HER BCVA IS 0.8, OU. NO FURTHER SURGICAL OPTION IS PLANNED. DR. ADVISED FUTURE IOL IMPLANTATION IN CASE SHE DEVELOPS CATARACT." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795471&pc=MTA | 2023826-2019-01236 | 2019-01-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE RIGHT (OD) EYE ON (B)(6) 2019 DUE TO AN ANTERIOR SUBCAPSULAR CATARACT. THE SURGERY IS REPORTED TO HAVE BEEN SUCCESSFUL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN CONTACT LENS CONTAINER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RETURNED WITH RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 09/12/2019. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANTERIOR SUBCAPSULAR CATARACT IN THE LEFT EYE (OS). NO CATARACT WAS PRESENT IN THE RIGHT EYE, BUT THE PATIENT STILL UNDERWENT CATARACT SURGERY AND EXPLANT IN THE RIGHT EYE (OD). REPORTEDLY THE PATIENT HAS 'HEALED' AND 'HE IS FINE'. PATIENT CODE 1766 NO LONGER APPLICABLE. CATARACT NOT PRESENT IN RIGHT EYE (OD), PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, CATARACT SURGERY), LENS WORKS ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6 - RESULT CODE 114 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#2. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8528238&pc=MTA | 2023826-2019-00631 | 2019-02-11 | Injury | STAAR SURGICAL COMPANY | 2019-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON THE SAME DAY IN A DIFFERENT SURGERY DUE TO LENS OPACITY (ASC). THE REPORTER STATED " THE DOCTOR NOTED THE OPACITY AFTER IMPLANTATION. THE OPACITY IS MORE RESOLVED THAN BEFORE, IT WAS A LITTLE IMPROVED AFTER BEING TREATED WITH MEDICATION. IMMEDIATELY AFTER THE REMOVAL OF THE LENS BCVA WAS DECREASED BUT RECOVERED. AT THE LAST VISIT BCVA WAS 20/20 ((B)(6) 2019)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION: LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736590&pc=MTA | 2023826-2019-01119 | 2019-05-29 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED AND THE SURGEON PERFORMED PHACO-EMULSIFICATION AND AN IOL WAS IMPLANTED. THE SURGEON REPORTS LENS OPACITY-CORTICAL. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K #: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276016&pc=MTA | 2023826-2019-02092 | 2019-09-19 | Injury | STAAR SURGICAL COMPANY AG | 2019-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: LENS WAS EXPLANTED DUE TO CATARACT SURGERY. PATIENT CODE 3191: NO CODE AVAILABLE (CATARACT SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT THEY FIRST NOTED THE CATARACT ON (B)(6) 2019. WHEN THE CATARACT WAS REMOVED IT WAS REPLACED WITH AN IOL FROM A DIFFERENT MANUFACTURER. POST-OP VISIT WAS REPORTED AS THE PATIENT WAS DOING VERY WELL. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276024&pc=MTA | 2023826-2019-02093 | 2019-09-19 | Injury | STAAR SURGICAL COMPANY AG | 2019-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT RECEIVED CATARACT SURGERY. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (CATARACT SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THE CATARACT WAS NOTICED ON 19-SEPT-2019. WHEN THE CATARACT WAS REMOVED IT WAS REPLACED WITH AN IOL FROM ANOTHER MANUFACTURER. POST-OP REPORT STATED THE "PATIENT WAS DOING VERY WELL". CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441175&pc=MTA | 2023826-2019-02347 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2013. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. AN IOL FROM ANOTHER MANUFACTURER WAS IMPLANTED ON (B)(6) 2019. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140605&pc=MTA | 2023826-2020-01204 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY NS, DATE ASSESSED (B)(6) 2019. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "THE PROBLEM WAS RESOLVED." SHOULD BE ADDED. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(6). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703060&pc=MTA | 2023826-2020-00268 | NA | Injury | STAAR SURGICAL COMPANY AG | 2020-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 14.5 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CATARACT DEVELOPED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE INCISION WAS NOT ENLARGED TO REMOVE THE LENS AND SUTURES WERE NOT REQUIRED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, WRAPPED IN GAUZE. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS. BUT NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592014&pc=MTA | 2023826-2015-00279 | 2012-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 AND GLARES/HALOES WERE OBSERVED ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(4) 2014 AND WAS NOT REPLACED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20. SEE MFR REPORT # 2023826-2015-00278 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: PT : UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, DEVICE HISTORY AND MEDICAL REVIEWS. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE AND A HAPTIC WAS TORN. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH WERE REMEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. A DEVICE HISTORY REVIEW WAS PERFORMED AND CONCLUDED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTER EVENT. MEDICAL REVIEW - GLARES/HALOS ARE SUBJECTIVE PERCEPTION OF STARBURST OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING THE DAY. POTENTIAL CAUSES ARE: RESIDUE REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PREOPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ECT. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF A WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS HE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY AND MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600860&pc=MTA | 2023826-2015-00309 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (CATARACT). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION SURGERY AND AN IOL WAS IMPLANTED. POST-OP, THE PATIENT WAS PSEUPHAKIC AND THE UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592739&pc=MTA | 2023826-2015-00233 | 2015-01-03 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (NUCLEAR SCLEROSIS). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION SURGERY AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ICL RELATED OPACITIES ARE USUALLY ANTERIOR LENS LOCATED (ASCO = ANTERIOR SUB-CAPSULAR OPACITY). NUCLEAR SCLEROSIS IS KNOWN TO HAVE A STRONG ASSOCIATION WITH AXIAL MYOPIA. SEVERAL FACTORS MAY PLAY A ROLE IN CATARACT DEVELOPMENT (AGING, DEGREE OF MYOPIA, SYSTEMIC DISEASES, MEDICATIONS, SURGICAL MANIPULATION, INFLAMMATION, TRAUMA, ETC). A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING PROCESS OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4639048&pc=MTA | 2023826-2015-00387 | 2015-03-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON EXPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF A TRAUMATIC CATARACT . THE LENS WAS IMPLANTED ON (B)(6) 2011. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. A SUTURE WAS REQUIRED. THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND TWO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - EXTERNAL FACTORS SUCH AS TRAUMA COULD CAUSE CATARACT. IT IS LIKELY THAT THE MEDICAL DEVICE IS NOT THE CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4822542&pc=MTA | 2023826-2015-00645 | NA | Injury | STAAR SURGICAL COMPANY AG | 2015-06-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN MICL IMPLANTABLE COLLAMER LENS DUE TO THE PATIENT DEVELOPED A CATARACT. THE PATIENT HAD THE MICL IMPLANTED IN 2008. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015. THE CATARACT WAS REMOVED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER INDICATED THE INCIDENT/INJURY WAS NOT PRODUCT-RELATED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, ICL WAS EXPLANTED 7 YEARS POSTOPERATIVELY TO ADDRESS CATARACT DEVELOPMENT. ACCORDING TO CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY (CCQS), RECEIVED ON (B)(6) 2016, ICL AND CATARACT WERE REMOVED AT THE SAME SURGERY DATE AND IOL WAS SUCCESSFULLY IMPLANTED. THE VA WAS 20/20 POSTOPERATIVELY. THE SAME FORM STATED THAT CATARACT DEVELOPMENT WAS NOT PRODUCT RELATED. IT SHOULD BE NOTED THAT CATARACT HAD DEVELOPED BILATERALLY. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. DEVICE EVALUATED BY MANUFACTURER?: LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY FACTORS SUCH AS AGING, TRAUMA, DIABETES, RADIATION, SMOKING, ALCOHOL USE AND EVEN SOME MEDICATIONS (EX. STEROID) COULD CAUSE CATARACT. GIVEN THE LATE ONSET OF THE EVENT, THESE FACTORS COULD NOT BE RULED OUT. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358894&pc=MTA | 2023826-2017-00260 | 2013-04-16 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.50 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AND THE PATIENT REPORTS NO RESTRICTIONS WITH ASCO. Manufacturer Narrative: CONCLUSION CODE: PER DFU, PATIENT'S AGE IS OVER 45 YEARS OF AGE WHICH IS OFF LABEL USE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358899&pc=MTA | 2023826-2017-00261 | 2013-11-19 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICMO13.2, -8.00 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AS THE PATIENT IS STILL HAPPY AND FEELS NO LIMITATIONS FROM THE ASCO. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AS THE PATIENT IS STILL HAPPY AND FEELS NO LIMITATIONS FROM THE ASCO. THE SURGEON HAS NO PLANS TO EXPLANT OR EXCHANGE THE ICL. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/25. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OF AGE AND WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6303775&pc=MTA | 2023826-2017-00143 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY ASCO (ANTERIOR SUBCAPSULAR OPACITY) WHICH WAS OBSERVED ON (B)(6) 2016. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/30 WHICH IS THE SAME AS PREOPERATIVE DATA AS OF (B)(6) 2015. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE'S CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FIBERS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4). LENS NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557824&pc=MTA | 2023826-2018-00812 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS OPACITY ASC. REPORTEDLY, ON (B)(6) 2018 CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED WHICH RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY & IOL IMPLANTATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389922&pc=MTA | 2023826-2017-00343 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | NA | 1 | Event Description: ON (B)(6) 2017 THE REPORTER INDICATED AWARENESS "OF A CASE OF V4C LENS + LENS OPACITY." THE LENS WAS REMOVED AND PHACO IOL IMPLANTATION WAS PERFORMED. AT THE DATE OF MDR SUBMISSION, ATTEMPS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, DIOPTER -11.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LENS OPACITY. PHACO IOL IMPLANTATION WAS PERFORMED. Manufacturer Narrative: PATIENT INFORMATION: UNK. EVENT DATE: UNK. SUSPECT MEDICAL DEVICE: UNK. INITIAL REPORTER ADDRESS: UNK. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. MFG DATE: UNK. (B)(4). Manufacturer Narrative: (B)(4): LENS EXCHANGE; SHOULD STATE: LENS EXPLANT. CONCLUSION: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. DEVICE REMAINS IMPLANTED. | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740169&pc=MTA | 2023826-2017-01236 | NA | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTED STATED THAT THE LENS IS SCHEDULED TO BE REMOVED ON (B)(6) 2017 DUE TO THE PATIENT FORMING CATARACTS. THE REPORTER INDICTED THAT CATARACT SURGERY IS BEING PERFORMED AND A TECNIS IOL FROM A DIFFERENT MANUFACTURER IS BEING IMPLANTED. REPORTED STATED THAT IT IS UNKNOWN IF THE ADVERSE EVENT WAS RELATED TO THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL AND THERE WERE NO ISSUES WITH THE REMOVAL OF THE ICL LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740468&pc=MTA | 2023826-2017-01237 | NA | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2015. THE REPORTED STATED THAT THE LENS IS SCHEDULED TO BE REMOVED ON (B)(6) 2017 DUE TO THE PATIENT FORMING CATARACTS. THE REPORTER INDICTED THAT CATARACT SURGERY IS BEING PERFORMED AND A TECNIS IOL FROM ANOTHER MANUFACTURER IS BEING IMPLANTED. REPORTED STATED THAT IT IS UNKNOWN IF THE ADVERSE EVENT WAS RELATED TO THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL AND THERE WERE NO ISSUES WITH THE REMOVAL OF THE ICL LENS. (B)(4). Manufacturer Narrative: SCHEDULED TO BE EXPLANTED ON (B)(6) 2017. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879893&pc=MTA | 2023826-2017-01480 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -14.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASCO). REPORTEDLY, THE TOP OF THE LENS WAS NICKED BY THE SURGEON IN ERROR, WHICH LED TO "CAPSULE MEMBRANE DAMAGE." THE ICL WAS THEN REMOVED AND REPLACED WITH AN IOL. THE PATIENT HAS NO PRE-EXISTING CONDITION OF LENS OPACITY. ACCORDING TO THE SURGEON, THE LENS CAPSULE MEMBRANE DAMAGE IS THE MAIN CAUSE OF THE SITUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577503&pc=MTA | 2023826-2016-00438 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. A LENS OPACITY (CORTICAL) WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND A IOL WAS IMPLANTED. PATIENT'S LAST VISIT ON 02/22/2016, BCVA WAS 20/20. SEE MFR. #2023826-2016-00447 FOR PREVIOUS SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS OLD AND WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50. DEVICE NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289852&pc=MTA | 2023826-2017-00169 | 2016-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER STATED CAPSULAR TEAR AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY) WAS OBSERVED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) THERE WAS A CAPSULAR TEAR FOR THE ICL LENS THAT WAS EXPLANTED. THE SECONDARY SURGERY WHICH WAS A PHACO-EMULSIFICATION (CATERACT SURGERY) AND IOL IMPLANTATION HAPPENED DUE TO THE PATIENT HAVING CATARACT. (B)(4) LENS NOT RETURNED. | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6080846&pc=MTA | 2023826-2016-01598 | 2014-11-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2014. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. A 16.0 DIOPTER IS INCORRECT, SHOULD BE 16.5 DIOPTER. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404792&pc=MTA | 2023826-2017-00401 | 2016-07-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED DRY IN A SMALL VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT EXPERIENCED LENS OPACITY (ASCO). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289966&pc=MTA | 2023826-2017-00149 | 2016-08-06 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM (B)(4) -12.5 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT REPORTING LOW VAULT AND LATE LENS OPACITY (ASCO). ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXCHANGED AGAIN DUE TO LOW VAULT, REFERENCE MDR# 2023826-2017-00194. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/15. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214422&pc=MTA | 2023826-2016-01843 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS OPACITY, AND ANTERIOR SUBCAPSULAR CATARACT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS MEASUREMENTS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214458&pc=MTA | 2023826-2016-01844 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS OPACITY, AND ANTERIOR SUBCAPSULAR CATARACT. Manufacturer Narrative: ADDITIONAL EVENT DESCRIPTION - PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25 AS OF (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED BENT AND DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE LENS OPTIC WAS TORN AND THE HAPTIC WAS BROKEN. FOREIGN MATERIAL (FIBERS) WAS FOUND ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495937&pc=MTA | 2023826-2018-00713 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY-ASC. THE SURGEON THEN PERFORMED CATARACT SURGERY WITH A TRIFOCAL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND THE LENS IS MISSING A PIECE OF HAPTIC. A-9.5 DIOPTER. REPORT SOURCE - USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177027&pc=MTA | 2023826-2018-00008 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), OBSERVED ON (B)(6) 2017. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED 4 DAYS LATER. THE PATIENT'S CURRENT BCVA IS 20/40. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN OFF AND MISSING. THE LENS WIDTH WAS MEASURED AND WAS VERIFIED. WAS UNABLE TO VERIFY DIAMETER/LENGTH DUE TO MISSING HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339941&pc=MTA | 2023826-2018-00412 | 2017-12-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LENS OPACITY (CATARACT). CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE REPORTER INDICATED THE PATIENT IS ALLERGIC TO SULPHUR MEDICATIONS HENCE DIAMOX WAS NOT PRESCRIBED. THE CATARACT COULD HAVE POSSIBLY FORMED DUE TO FLUCTUATIONS IN INTRAOCULAR PRESSURE (IOP) AND AC DEPTH POST-OP. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311086&pc=MTA | 2023826-2020-01496 | NA | Injury | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATE THAT A 13.2MM, MICL13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULT AND ASC CATARACT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED, PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592589&pc=MTA | 2023826-2015-00235 | 2014-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS OPACITY (CORTICAL). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592661&pc=MTA | 2023826-2015-00277 | 2014-07-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE AND THERE WAS NO VISIBLE DAMAGE WAS OBSERVED. BOTH THE WIDTH AND LENGTH WERE REMEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASE WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655201&pc=MTA | 2023826-2016-00685 | 2014-09-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4, -20 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT AND AN ANTERIOR SUBCAPSULAR CATARACT WERE OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED, AND AN IOL WAS IMPLANTED ON (B)(6) 2015. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592687&pc=MTA | 2023826-2015-00281 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LASTEST UCVA WAS 20/30. SEE MFR REPORT #2023826-2015-00282 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592699&pc=MTA | 2023826-2015-00282 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LASTEST UCVA WAS 20/20. SEE MFR REPORT #2023826-2015-00281 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND A CLEAR SURGICAL RESIDUE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL CREASES THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154946&pc=MTA | 2023826-2015-01267 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. LENS OPACITY (ASCO) WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/26. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -16.5. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727980&pc=MTA | 2023826-2015-00472 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. LENS OPACITY (ANTERIOR) WAS NOTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER INDICATED THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO SIZING. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT LOW VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF LOW VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE REPORTER INDICATED THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO SIZING. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY THE INSUFFICIENT VAULTING WAS A CONSEQUENCE OF WRONG LENS USE. HOWEVER, NO DEFINITE ROOT CAUSE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161950&pc=MTA | 2023826-2015-01278 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/24. SEE MFR #2023826-2015-01277 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161958&pc=MTA | 2023826-2015-01277 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/24. SEE MFR #2023826-2015-01278 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: TWO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4) (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5095987&pc=MTA | 2023826-2015-01180 | 2015-08-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. LOW VAULT, REFRACTIVE CHANGE OVER TIME AND LENS OPACITY WAS OBSERVED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SURGERY IS PENDING TO EXPLANT AND EXCHANGE FOR LONGER LENS. UNCORRECTED VISUAL ACUITY WAS 20/10 ON (B)(6) 2015. SEE MFR #20238236-2015-01179 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). LENS EXPLANTED, BUT NOT RETURNED. Manufacturer Narrative: DIOPTER: -16.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5707624&pc=MTA | 2023826-2016-00696 | 2016-04-01 | Injury | STAAR SURGICAL COMPANY | 2016-06-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT HAD A MILD SUP. ANTERIOR CORTICAL OPACITY. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199510&pc=MTA | 2023826-2015-01455 | 2015-02-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER -12.50 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. LOW VAULT AND A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE ICL WAS EXPLANTED, CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015 - BCVA WAS 20/20 AND UCVA WAS 20/32. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. INITIAL MDR DATE 10/13/2015. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.50. (B)(4). CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592752&pc=MTA | 2023826-2015-00227 | 2013-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING, LENS ROTATION AND LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2013. THE PATIENT HAD CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4) DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR SULCUS OR CILIARY SULCUS CYST, ETC). ALTHOUGH SEVERAL FACTORS MAY PLAY A ROLE IN CATARACT DEVELOPMENT (AGE, DEGREE OF MYOPIA, SYSTEMIC DISEASES, MEDICATIONS, SURGICAL MANIPULATION, ETC.) WE CANNOT RULE OUT LOW VAULTING AS A CONTRIBUTING FACTOR. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7562910&pc=MTA | 2023826-2018-00836 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, CATARACT SURGERY AND IOL IMPLANT CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962469&pc=MTA | 2023826-2019-01587 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE DEVELOPMENT OF A CATARACT. THE CATARACT WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. B5 - THE LENS WAS EXPLANTED ON (B)(6) 2019. THE CAUSE OF THE EVENT WAS REPORTED AS THE ICL. PATIENT HAD A PRE-EXISTING CONDITION OF PROBABLY AMBLYOPIA. POSTOPERATIVE REPORT STATED PATIENT IS "DOING WELL" ON (B)(6) 2019 LAST VISIT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889749&pc=MTA | 2023826-2019-01444 | 2019-06-22 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND LENS REMAINS IMPLANTED. REPORTER INDICATED "THE VAULT WAS 271 MICROMETERS ONE WEEK AFTER SURGERY ((B)(6) 2019), BUT AFTER THAT THE VAULT DECREASED." SURGEON CONTINUES TO MONITOR AND IS "CONSIDERING EXPLANTATION OR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION THAT SHOULD HAVE BEEN REPORTED IN INITIAL MDR: THE SURGEON ALSO FOUND PUNCTATE CATARACT DURING A POST-OP VISIT 6 MONTHS AFTER IMPLANTATION ((B)(6) 2019). ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT. PATIENT CODE 1766 ADDED, PATIENT CODE 2692 NOT APPLICABLE. RESULT CODE 213 ADDED, RESULT CODE 114 NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889750&pc=MTA | 2023826-2019-01445 | 2019-06-22 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND LENS REMAINS IMPLANTED. REPORTER INDICATED "THE VAULT WAS 241 MICROMETERS ONE WEEK AFTER SURGERY ((B)(6) 2019), BUT AFTER THAT THE VAULT DECREASED." SURGEON CONTINUES TO MONITOR AND IS "CONSIDERING EXPLANTATION OR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION THAT SHOULD HAVE BEEN REPORTED IN INITIAL MDR: THE SURGEON ALSO FOUND PUNCTATE CATARACT DURING A POST-OP VISIT 6 MONTHS AFTER IMPLANTATION ((B)(6) 2019). ADVERSE EVENT OR PRODUCT PROBLEM CORRECTED TO ADVERSE EVENT. PATIENT CODE 1766 ADDED, PATIENT CODE 2692 NOT APPLICABLE. RESULT CODE 213 ADDED, RESULT CODE 114 NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577133&pc=MTA | 2023826-2017-00836 | 2017-04-18 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO DEVELOPMENT OF A CATARACT OR COULD BE A FILM BETWEEN NATURAL LENS AND ICL OR THE START OF CATARACTS. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2017 AND THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE REPORTER INDICATED THE EVENT WAS RELATED TO THE DEVICE. NO MEDICATIONS OR TREATMENTS WERE REQUIRED. IMPLANTED DATE: (B)(6) 2011. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426147&pc=MTA | 2023826-2017-00471 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED EYE IS WORSE FOR READING. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY REPORTED THEY CONFIRMED THE PATIENT REPORT. DATE OF LAST VISIT WAS (B)(6) 2016 AND REPORTED PATIENT HAD PRESBYOPIA, WITH A MILD CATARACT. NO MEDICATIONS OR TREATMENTS WERE REQUIRED AND THE EVENT WAS NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546330&pc=MTA | 2023826-2017-00772 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED REDUCED VISUAL ACUITY AND WAS WEARING GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THE PATIENT HAD DECREASED VISUAL ACUITY DUE TO A CATARACT HAD DEVELOPED. SURGERY IS PLANNED AT A FUTURE DATE FOR REMOVAL OF CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE FACILITY REPORTED THE PATIENT HAD LASIK ON (B)(6) 2011 AND WAS PRESCRIBED MEDICATION FOR NIGHT DRIVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323983&pc=MTA | 2023826-2019-00238 | NA | Injury | STAAR SURGICAL COMPANY AG | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2009. THE REPORTER INDICATED A CATARACT HAS DEVELOPED AND VISION IS SLOWLY CLOUDING. THE PLAN IS TO EXPLANT THE LENS BUT CURRENTLY THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: REPORT # MW5082245. AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. OCCUPATION: UNK. (B)(4). | Lens Extraction | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17872920&pc=MTA | 2023826-2023-04387 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS REMOVED AND REPLACED FOR A SMALLER LENGTH LENS DUE TO LENS OPACITY. THE PROBLEM WAS RESOLVED. REPORTEDLY, "QUIET ICL VAULT GOOD." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "1766" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3001" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS OPACITY (CATARACT). ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "QUIET. ICL VAULT GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873747&pc=MTA | 2023826-2018-01363 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01364 FOR FELLOW EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873772&pc=MTA | 2023826-2018-01364 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863580&pc=MTA | 2023826-2017-01430 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A CATARACT, AND REPLACED WITH AN INTRAOCULAR LENS (IOL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN SPECIMEN CUP WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS. LENS HAS AN "80" STAMPED /EMBEDDED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362600&pc=MTA | 2023826-2019-00300 | NA | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS STILL UNHAPPY WITH HIS VISION. THE DOCTOR FOUND THE START OF CATARACTS AT HIS LAST EXAM AND RECOMMENDED REMOVING THE LENS AT THE SAME TIME. PATIENT HAS SOUGHT A 2ND OPINION. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT - CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. (DEVICE MANUFACTURER DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362754&pc=MTA | 2023826-2019-00301 | NA | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT- CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE 3191: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9233429&pc=MTA | 2023826-2019-02023 | NA | Injury | STAAR SURGICAL COMPANY | 2019-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Adverse Event Without Identified Device or Use Problem | Cataract; Vitreous Floaters; Retinal Detachment; Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM FDA: # MW5089673. THE PATIENT REPORTED ON (B)(6)2019, A "SERIOUS INJURY" REPORT WITH OUTCOME OF "DISABILITY/PERMANENT DAMAGE". THE PATIENT STATED "GOT VISIAN ICL PLASTIC LENS IMPLANTS IN BOTH EYES. IN LEFT EYE OVER A YEAR LATER I HAD SYMPTOMS OF FOGGY VISION, NEW FLOATERS AND TODAY MY VISION IS ALMOST GONE. THERE'S A HUGE BLOB OF WATER BLOCKING MY SIGHT. WHEN I MOVE MY EYE I CAN SEE THE WATER MOVE." PER THE SAME REPORT CONCOMITANT MEDICAL PRODUCTS LISTED AS: "ALPHAGAN AND PRENATAL VITAMINS", WITH PATIENT PROBLEMS OF: "IMPAIRED VISION, VITREOUS FLOATERS" AND DEVICE PROBLEMS: "ADVERSE EVENTS WITHOUT IDENTIFIED DEVICE OR USE PROBLEM" WAS REPORTED. PER PATIENT EMAIL OF 30 SEP 2019 "VISIAN ICL WAS REMOVED ON THE LEFT EYE (B)(6) TO PREPARE FOR VITRECTOMY FOR RETINAL DETACHMENT REPAIR". INFORMATION RECEIVED FROM THE IMPLANTING SURGEON INDICATED THAT A 13.2MM, MICL13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017 AND REMOVED ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (NS), RETINAL DETACHMENT, BLURRED VISION, AND FLOATERS WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION/TREATMENT (PRK, ICL REMOVAL + CE/IOL, RETINAL REPAIR) WERE PERFORMED. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS BCVA WAS 20/80-, UCVA 20/200. CLEAR CORNEA, SUPERIOR PI, PC/IOL WITH MILD TO MODERATE PCO, S/P VITRECTOMY, 45 % GAS, AND RETINA ATTACHED WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). PER RETINAL DETACHMENT REPAIR SURGEON AND ICL REMOVAL AND CATARACT SURGERY SURGEON THIS INCIDENT DID NOT INVOLVE AN ADVERSE EVENT RELATED TO THE VISIAN ICL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRK TREATMENT WAS PERFORMED ON (B)(6) 2018. ICL REMOVAL AND CE/IOL PROCEDURES WERE COMPLETED DURING ONE SURGICAL EVENT BY A 2ND SURGEON. RD REPAIR WAS PERFORMED BY A 3RD SURGEON ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT # 2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved