## Vision Issue-1
Web Address...29 | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153422&pc=MTA | 2023826-2016-01745 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER VISION WAS SLIGHTLY BLURRY AT NIGHT AND IN DIM LIGHTING. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY CONFIRMED THE INFORMATION AND INDICATED THE EVENT WAS DUE TO THE PATIENT HAD LARGE PUPILS AND WAS NOT DEVICE RELATED. NO OTHER COMPLICATIONS WERE REPORTED. PATIENT'S POST-OP VISUAL ACUITY WAS 20/15 AND THE OUTCOME WAS EXCELLENT. Manufacturer Narrative: ON (B)(6) 2012. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153430&pc=MTA | 2023826-2016-01744 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HER VISION WAS SLIGHTLY BLURRY AT NIGHT AND IN DIM LIGHTING. THE LENS WAS IMPLANTED ON (B)(6) 2010 AND REMAINS IMPLANTED. THE FACILITY CONFIRMED THE INFORMATION AND INDICATED THE EVENT WAS DUE TO THE PATIENT HAD LARGE PUPILS AND WAS NOT DEVICE RELATED. NO OTHER COMPLICATIONS WERE REPORTED. PATIENT'S POST-OP VISUAL ACUITY WAS 20/20 AND THE OUTCOME WAS EXCELLENT. Manufacturer Narrative: MFR DATE- 05-25-2012. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887955&pc=MTA | 2023826-2019-01354 | NA | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY, "PATIENT WANTED BETTER NEAR VISION IN EYE AND SOME ASTIGMATISM WAS FOUND TO BE RELATED TO IT AS WELL." ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A TORIC ICL. THIS RESOLVED THE PROBLEM. PER REPORTER ON (B)(6) 2019, "PATIENT IS HAPPY WITH VISION AND SAYS CONTINUALLY IMPROVING." Manufacturer Narrative: RESULT CODE: 213 REPORTED IN INITIAL MDR IS CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS EXPLANT DATE OF (B)(6) 2019 IS NOT APPLICABLE. THE EXPLANT DATE WAS ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN CONTACT LENS CASE, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513718&pc=MTA | 2023826-2015-00143 | NA | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS). THE PATIENT REPORTED THE LENS WAS EITHER REMOVED, REPLACED OR REPOSITIONED? ALSO ASTIGMATLISM IS RETURNING AND THERE IS SORENESS AND BLURRINESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON INDICATED HE DID NOT AGREE WITH THE PATIENT REPORT AND AS OF THE DATE OF THE LAST POST-OP VISIT ON (B)(6) 2015, THE PATIENT HAS NOT HAD ANY ADDITIONAL PROCEDURES. THE SURGEON INDICATED THE ASTIGMATISM WAS PRE-EXISTING AND MINOR. THE PATIENT IS 20/20 OU WITHOUT COMPLICATION. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING THE DIRECTION FOR USE (DFU), THE MICL LENS IS USED TO ADDRESS MILD TO MODERATE MYOPIA. ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH A MICL LENS. MOST LIKELY THE CAUSE OF THE EVENT IS THE PATIENT'S PREEXISTED CONDITION: ASTIGMATISM. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE EVENT IS THE PATIENT'S PREEXISTED CONDITION: ASTIGMATISM. (B)(4). Manufacturer Narrative: (B)(4): EVALUATION(B)(6): METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213979&pc=MTA | 2023826-2018-00137 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6, -5.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE SURGEON NOTES LENS ROTATION, LOSS OF BCVA, AND BLURRED VISION. REPORTEDLY, LENS REMAINS IMPLANTED. SURGEON REPORTS THE "ICL ON DEAD VERTICAL POSITION." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205806&pc=MTA | 2023826-2018-00105 | 2015-11-30 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTICED BY THE SURGEON. THE PATIENT HAS UNSTABLE REFRACTION AND THE SURGEON IS CONTROLLING HIS REFRACTION PERIODICALLY. THE CONTACT REPORTED THAT WITHOUT A STABLE REFRACTION LENS EXCHANGE CANNOT TAKE PLACE. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6366306&pc=MTA | 2023826-2017-00276 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE LENS, DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. ADDITIONAL INFO: 03/15/2017. ADDITIONAL INFO: DEVICE RETURNED 04/05/2017; NOT YET EVALUATED. ADDITIONAL INFO: DEVICE RETURNED, EVALUATION ANTICIPATED; NOT YET PERFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) , AND LENS STUCK IN VIAL. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879884&pc=MTA | 2023826-2017-01477 | 2017-08-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON. ON 28-08-2017, DURING THE PATIENT'S LAST REPORTED VISIT, THE UCVA WAS MEASURED AT 20/20 AND THE RX +0.25 SPH. PER THE SURGEON NO LENS EXCHANGE IS NEEDED. THE CLAIM REPORTER STATED THAT THERE POSSIBLY WAS SOME VISCO LEFT IN THE EYE AT 1-DAY POST-OP, WHICH MAY HAVE CAUSED THE REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268465&pc=MTA | 2023826-2018-00179 | 2018-01-15 | Malfunction | STAAR SURGICAL COMPANY AG | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0 MM ICM130V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S (OS) ON (B)(6) 2001. ON (B)(6) 2018, REFRACTIVE CHANGE OVER TIME WAS NOTICED. THE SURGEON IS MAKING PLANS TO EXCHANGE THE LENS. THE REPORTER STATED THAT THE REASON FOR THE REFRACTIVE CHANGE OVER TIME WAS BECAUSE OF PENETRATING KERATOPLASTY (PKP) PERFORMED 18 YEARS AGO. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0MM ICM130V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2001. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484263&pc=MTA | 2023826-2018-00639 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660090&pc=MTA | 2023826-2018-01003 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS >2-LINE DROP OF BCVA AT THE 1-MONTH POSTOPERATIVE VISIT. UPON THE REPEAT EXAM OF THE BCDVA DURING A FOLLOW UP VISIT, IT WAS DETERMINED THAT, ALTHOUGH THERE WAS A 1-LINE IMPROVEMENT IN THE BEST CORRECTED VISION OF THE EYE, IT IS STILL A 2-LINE DROP FROM PREOPERATIVE VALUES. THE LENS REMAINS IMPLANTED AND INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IN THE SURGEON'S OPINION, THE CAUSE OF THIS EVENT IS THE DEVICE. THE STATUS OF THE PATIENT'S EYE IS "RESOLVING." (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT IS DETERMINED TO BE RESOLVED. SUBJECT COMPLETED THE STUDY ON (B)(6) 2018 WITH BCDVA OF 20/20. IT IS THE MEDICAL MONITOR'S OPINION THAT THE TRANSIENT DECREASE AND SUBSEQUENT IMPROVEMENT IN BCDVA REFLECT'S THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6839063&pc=MTA | 2023826-2017-01374 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED AS THE PATIENT IS NOW APPROXIMATELY +1.25D SEQ POST-OP. THE SURGEON CONFIRMED THAT THE ICL IS SITTING IN THE CORRECT POSITION AND NOT INVERTED. HE ALSO VERIFIED THAT THE CYCLO WAS NOT SIGNIFICANTLY DIFFERENT THAT THE MANIFEST REFRACTION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: METHOD: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6839072&pc=MTA | 2023826-2017-01375 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL12.1, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978049&pc=MTA | 2023826-2017-01625 | NA | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.00 INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, IMMEDIATELY AFTERWARD, THE PATIENT COMPLAINED OF A CIRCULAR "CLOUD" (BLURRED SPOT) IN HER FIELD OF VIEW JUST OFF-CENTER, ADJACENT TO HER CENTRAL VISION. THE SURGEON STATES THE "HOLE" OF THE LENS IS "VERY CLOSE TO CENTER." UPON EXAM, PATIENT'S CORNEA IS CLEAR WITH NO CATARACT FORMATION PRESENT. SURGEON STATES THAT "ON PAPER" THE PATIENT HAS A GREAT RESULT AND SUGGESTS POSSIBLE "NON-ORGANIC" EXAGGERATION OF SYMPTOMS. STAAR WAS NOTIFIED OF THIS CASE BY A SECOND SURGEON WHO PERFORMED LASIK SURGERY ON SAME PATIENT IN (B)(6) 2017. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542565&pc=MTA | 2023826-2015-00174 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME (OVERCORRECTION) AND BLURRED VISION. THE ICL WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT A DIFFERENT DIOPTER POWER (-15.5), WHICH RESOLVED THE PROBLEM. THE PATIENT IS HEALING WELL OU. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED BUT WITH THE RE-MEASUREMENT OF THE LENS, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180685&pc=MTA | 2023826-2018-00036 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE; INCIDENT DATE OF (B)(6) 2017. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180690&pc=MTA | 2023826-2018-00037 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE; INCIDENT DATE OF (B)(6) 2017. Manufacturer Narrative: CORRECTED INFORMATION: THE LENS REMAINS IMPLANTED. THE LENS IS PLANNED TO BE EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524328&pc=MTA | 2023826-2020-02041 | 2020-08-03 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.00, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WERE REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MANUFACTURED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524391&pc=MTA | 2023826-2020-02040 | 2020-08-03 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WERE REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030056&pc=MTA | 2023826-2022-04436 | 2021-04-08 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029739&pc=MTA | 2023826-2021-01917 | 2021-05-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VICM5_ 12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THIS EXCHANGED RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592022&pc=MTA | 2023826-2023-03621 | 2023-07-19 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS EXPLANTED. THE LENS WAS LATER REPLACED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS: "THE ARCH HEIGHT IS TOO LOW". CLAIM# (B)(4). Manufacturer Narrative: H6: PATIENT RELATED FACTOR: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784581&pc=MTA | 2023826-2022-04080 | 2018-03-27 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022, WITH A VICM5_12.6 LENS, 08.00 DIOPTER. THE PROBLEM WAS RESOLVED. THE INCISION IS HEALING WELL, ICL IS POSITIONED WELL WITH GOOD VAULT. VA IS 20/25. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823161&pc=MTA | 2023826-2021-01483 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751185&pc=MTA | 2023826-2021-01281 | 2018-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696925&pc=MTA | 2023826-2018-01042 | 2018-06-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS A 2-LINE DROP IN BCDVA COMPARED TO THE PREOPERATIVE EXAM. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IN THE SURGEON'S OPINION, THE CAUSE OF THIS IS EVENT IS THE DEVICE. THE PATIENT'S CONDITION IS "RESOLVING." (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT DETERMINED TO BE RESOLVED BY INVESTIGATOR. ON (B)(4) 2018. THE SUBJECT COMPLETED THE STUDY WITH BCDVA OF 20/20. THE MEDICAL MONITOR BELIEVES THE TRANSIENT DECREASE AND SUBSEQUENT IMPROVEMENT IN BCDVA REFLECTS THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7799794&pc=MTA | 2023826-2018-01251 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2018-08-20 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THERE WAS A >/= 2-LINE DROP IN BCDVA NOTED AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. IN THE SURGEON'S OPINION, THE CAUSE OF THIS EVENT IS THE DEVICE. AT THE PATIENT'S LAST VISIT THE SURGEON STATES THERE IS A SIGNIFICANT IMPROVEMENT IN BINOCULAR UNCORRECTED VISUAL ACUITY. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: AT 3 MONTHS POSTOPERATIVE VISIT THE PATIENT'S BCDVA WAS RECORDED AS 20/20. AT 6 MONTHS POSTOPERATIVE VISIT UCVA WAS RECORDED AS 20/40. THE EVENT STATUS IS NOTED AS RECOVERED/RESOLVED. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: AT THE 6-MONTH VISIT THE CDVA DROPPED AGAIN TO 20/25. AT AN UNSCHEDULED VISIT APPROXIMATELY 2 MONTHS AFTER STUDY EXIT, THE VA RETURNED TO 20/20 AND THE NEW EVENT WAS CLASSIFIED AS RESOLVED BY THE SURGEON. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE RECURRENT TRANSIENT DECREASE IN BCDVA AND SUBSEQUENT IMPROVEMENT REFLECT THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351020&pc=MTA | 2023826-2021-00491 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351101&pc=MTA | 2023826-2021-00492 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721899&pc=MTA | 2023826-2021-01018 | 2018-08-04 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A FOR A SAME LENGTH TORIC LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED/BROKEN AND THE HAPTIC BROKEN/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452978&pc=MTA | 2023826-2019-00489 | 2018-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721905&pc=MTA | 2023826-2021-01025 | 2018-10-14 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE CHANGE OVER TIME ON (B)(6) 2019. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351898&pc=MTA | 2023826-2021-00494 | 2019-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6 -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8740571&pc=MTA | 2023826-2019-01152 | 2019-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, - 08.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764776&pc=MTA | 2023826-2019-01147 | 2019-04-04 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VICM5_12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: FILLED IN KNOWN LOT NUMBERS. CONCLUSION CODE 4310: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322922&pc=MTA | 2023826-2021-00422 | 2019-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VIICMO12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION FOR WIDTH DUE TO BROKEN HAPTIC. DIMENSIONAL INSPECTION FOR OVERALL AND FUNCTIONAL VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846158&pc=MTA | 2023826-2019-01210 | 2019-06-04 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "PATIENT WANTS TO WAIT AND WEAR GLASSES AGAIN." Manufacturer Narrative: LENS RETURNED TO MANUFACTURER IN PREVIOUS MDR IS INCORECT. THE CORRECT DATE IS 09-AUG-2019. DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR IS INCORRECT. THE CORRECT DATE IS 02-OCT-2019. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS NOT REMOVED ON (B)(6) 2019, IT WAS REMOVED ON THE (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY- UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492339&pc=MTA | 2023826-2019-02405 | 2019-11-05 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED, " H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: IS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814218&pc=MTA | 2023826-2020-00408 | 2020-02-19 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS, SEE MFR REPORT # 2023826-2020-00407 FOR INITIAL LENS. THE REPORTER INDICATED POST-OP, THE PATIENT HAD LOSS OF VISUAL ACUITY. TO DATE, THE BCVA HAS NOT INCREASED BUT VAULT IS BETTER. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: CORRECTED DATA: D4: SERIAL # (B)(6). CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988384&pc=MTA | 2023826-2020-00955 | 2020-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS IMPLANTED AS AN EXCHANGE LENS AND IT WAS UNKNOWN IF THE PROBLEM WAS RESOLVED, IF THE PATIENT STILL EXPERIENCED A REFRACTIVE SURPRISE. SEE MFR REPORT # 2023826-2020-00954 FOR THE INITIAL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002861&pc=MTA | 2023826-2020-00977 | 2020-04-01 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -2.25 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: ON (B)(6) 2021; THE LENS WAS REMOVED AND A SAME MODEL; SAME LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. REPORTEDLY; "THE PATIENT DID NOT SHOW UP FOR THE FOLLOW-UP EXAMINATIONS. SHE IS CURRENTLY UNAVAILABLE." CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459229&pc=MTA | 2023826-2022-00385 | 2020-04-16 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: "ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION." H6 - HEALTH EFFECT - CLINICAL CODE: 2137 BLURRED VISION SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453113&pc=MTA | 2023826-2021-00603 | 2020-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO 12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH TORIC LENS ( -6.00/1.0/082 (SPHERE/CYLINDER/AXIS)) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. RESUT CODE 213 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577512&pc=MTA | 2023826-2020-02194 | 2020-05-30 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPLACED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE REPORTER STATES "POST-OPERATIVE ASTIGMATISM DUE TO MULTIPLE SURGERIES REMAINS." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE / DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH STAINS / FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098617&pc=MTA | 2023826-2020-03256 | 2020-06-24 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098631&pc=MTA | 2023826-2020-03255 | 2020-06-24 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171163&pc=MTA | 2023826-2021-02584 | 2020-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148282&pc=MTA | 2023826-2021-02429 | 2020-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B5: PREVIOUSLY STATED THE LENS WAS EXPLANTED AND THEN REPLACED ON (B)(6) 2020. CORRECTION: THE LENS WAS EXCHANGED ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620598&pc=MTA | 2023826-2021-00877 | 2020-07-24 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SHORTER LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620605&pc=MTA | 2023826-2021-00876 | 2020-07-24 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE REC¿D BY MFR: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171147&pc=MTA | 2023826-2021-02585 | 2020-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322376&pc=MTA | 2023826-2021-03118 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -11.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2015. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM IS RESOLVED. THE SURGEON REPORTS, "TRACE CELLS, GOOD VAULT AT 1-WEEK POSTOP" AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322379&pc=MTA | 2023826-2021-03117 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -11.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2015. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM IS RESOLVED. THE SURGEON REPORTS, "TRACE CELLS, GOOD VAULT AT 1-WEEK POSTOP" AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755799&pc=MTA | 2023826-2020-02463 | 2020-09-07 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED TO: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.. CLAIM# 717567. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938834&pc=MTA | 2023826-2020-02844 | 2020-09-28 | Injury | STAAR SUGICAL COMPANY | 2020-12-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788143&pc=MTA | 2023826-2021-01393 | 2020-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTICS BROKEN,BENT,DEFORMED AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788188&pc=MTA | 2023826-2021-01394 | 2020-10-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536792&pc=MTA | 2023826-2021-00734 | 2020-10-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937930&pc=MTA | 2023826-2020-02779 | 2020-10-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON APR/2015. REFRACTIVE SURPRISE WITH GLARE/HALOES WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WITH GLARE/HALOES WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2015; CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563239&pc=MTA | 2023826-2021-00762 | 2020-10-21 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH LENS WITH A -5.5 DIOPTER TO RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 -04.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE REPORTED. THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS/ HIGHER POWER DUE TO REFRACTIVE SURPRISE ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317799&pc=MTA | 2023826-2022-00205 | 2020-11-01 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11710516&pc=MTA | 2023826-2021-01107 | 2020-11-08 | Injury | STAAR SURGICAL COMPANY | 2021-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416066&pc=MTA | 2023826-2021-00564 | 2020-11-13 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN "UNEXPECTED REFRACTIVE OUTCOME" AND A 25 DEGREES ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED, THAT A 13.2MM, MICL 13.2, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED. A NEW LENS IMPLANT IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -4.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT EXPERIENCED REFRACTIVE SURPRISE AND SHADOWING OF LETTERS. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE AND REPORTEDLY, A NEW LENS IMPLANT IS PLANNED. D6B- EPLANT DATE IS CORRECTED TO UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416418&pc=MTA | 2023826-2021-00513 | 2020-11-29 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, HIGHER POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "THERE IS NO PROBLEM WITH THE LENS MODEL AND LENS SIZE. THE PROBLEM WAS SOLVED BY SLIGHTLY CHANGING THE POWER OF THE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544735&pc=MTA | 2023826-2021-03791 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -2.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135923&pc=MTA | 2023826-2021-00059 | 2020-12-11 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE SURPRISE WAS OBSERVED." SHOULD BE IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855179&pc=MTA | 2023826-2021-01509 | 2020-12-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO 12.1, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DATE RECD BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366536&pc=MTA | 2023826-2021-00479 | 2020-12-13 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) STATEMENT IN INITIAL MDR IS NOT APPLICABLE. SHOULD HAVE BEEN REPORTED IN PREVIOUS MDR FOR DIMENSIONAL/ FUNCTIONAL INVESTIGATION CONCLUSION. ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256327&pc=MTA | 2023826-2022-00151 | 2020-12-15 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178442&pc=MTA | 2023826-2021-02711 | 2020-12-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207945&pc=MTA | 2023826-2021-02747 | 2021-01-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED AND STRETCHED OUT WITH FIBER ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942684&pc=MTA | 2023826-2021-01770 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308762&pc=MTA | 2023826-2021-00308 | 2021-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SAME LENGTH, LESSER POWER LENS ON (B)(6) 2021 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY PRODUCT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343195&pc=MTA | 2023826-2021-00398 | 2021-01-20 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE CHANGE OVERTIME WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER LENS (-10.50 DIOPTER) DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942674&pc=MTA | 2023826-2021-01779 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400353&pc=MTA | 2023826-2021-00514 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON THE SAME DATE IN A SEPARATE SURGERY, A SAME SIZE DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751223&pc=MTA | 2023826-2021-01284 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207967&pc=MTA | 2023826-2021-02748 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620582&pc=MTA | 2023826-2021-00867 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378621&pc=MTA | 2023826-2021-03375 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207976&pc=MTA | 2023826-2021-02749 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683173&pc=MTA | 2023826-2021-01074 | 2021-03-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR. (B)(6)) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID ALREADY BEFORE TRANS-PRK ELSWHERE)". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR. BECHMANN,SMILEYES) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID ALREADY BEFORE TRANS-PRK ELSEWHERE)". CLAIM # (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION CORRECTED TO "THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND PROBLEM NOT RESOLVED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR.BECHMANN,SMILEYES) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID BEFORE TRANS-PRK ELSEWHERE.)". REFERENCE MFR FILE#(B)(4) OF THE REPLACEMENT LENS CLAIM." IN SUPPLEMENTAL MDR#1. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4) | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781991&pc=MTA | 2023826-2021-01385 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT , IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11782012&pc=MTA | 2023826-2021-01386 | 2021-03-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A MODEL AND DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PROODUCT , IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942685&pc=MTA | 2023826-2021-01769 | 2021-04-15 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK: RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457492&pc=MTA | 2023826-2021-03552 | 2021-04-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN WITH FIBERS ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SDU A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT REPORTED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12208015&pc=MTA | 2023826-2021-02759 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524257&pc=MTA | 2023826-2021-03720 | 2021-04-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A2: DOB: (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524286&pc=MTA | 2023826-2021-03719 | 2021-04-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A2: DOB (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072245&pc=MTA | 2023826-2021-02086 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072278&pc=MTA | 2023826-2021-02085 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544805&pc=MTA | 2023826-2021-03793 | 2021-05-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, LENGTH AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072479&pc=MTA | 2023826-2021-02033 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS LATER EXPLANTED DUE TO VISUAL ACUITY NOT BEING AS EXPECTED AFTER SURGERY. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE OUT OF SPECIFICATION. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4): NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986341&pc=MTA | 2023826-2021-01874 | 2021-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED NIGHT VISION DUE TO LARGE PUPIL SIZE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK, DATE OF EVENT-UNK, PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK, CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791894&pc=MTA | 2023826-2021-04319 | 2021-05-26 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A REPLACEMENT LENS OF DIFFERENT POWER WAS LATER IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS AS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS PERFORMED, AND THE DIAMETER VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION PERFORMED FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257193&pc=MTA | 2023826-2021-02977 | 2021-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524316&pc=MTA | 2023826-2021-03722 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170553&pc=MTA | 2023826-2021-02698 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT DIOPTER AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. TYPE OF INVESTIGATION 3331 DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222316&pc=MTA | 2023826-2021-02809 | 2021-06-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/ BROKEN/BENT AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872779&pc=MTA | 2023826-2021-04511 | 2021-06-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS "OTHER", FOR CAUSE DETAILS THE REPORTER STATED " WRONG REFRACTION BY OPTOMETRIST." IT WAS REPORTED THAT THERE WAS AN ERROR IN THE PRE-OPERATIVE REFRACTION BEING THE PATIENT POST-PKP. OPTOMETRIST REPEATED REFRACTION TWICE BEFORE ORDERING THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178441&pc=MTA | 2023826-2021-02712 | 2021-06-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12644985&pc=MTA | 2023826-2021-04019 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6_12.6,-8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. FOR PATIENTS CURRENT CONDITION "FAR SIGHT NOT GOOD" WAS REPORTED. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTED STATED " DROP OF VISION AFTER VICM6 SURGERY." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATE INFORMATION THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. ON (B)(6) 2021 A REPLACEMENT LENS OF DIFFERENT MODEL AND DIOPTER WAS IMPLANTED AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: SHOULD BE CORRECTED TO STATE: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378568&pc=MTA | 2023826-2021-03363 | 2021-06-17 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS-BOTH REFRACTIVE CHANGE OVER TIME AND REFRACTIVE SURPRISE ARE REPORTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY REPORTED REFRACTIVE CHANGE OVER TIME IN ERROR. THIS IS A CASE OF REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222324&pc=MTA | 2023826-2021-02814 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725063&pc=MTA | 2023826-2021-04166 | 2021-07-25 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082453&pc=MTA | 2023826-2021-04969 | 2021-08-01 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENGTH, DIFFERENT SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725238&pc=MTA | 2023826-2021-04178 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL WITH . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566510&pc=MTA | 2023826-2021-03842 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, INSIDE MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523480&pc=MTA | 2023826-2021-03690 | 2021-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE ERROR. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND OTHER: "CLINICAL REFRACTIONS" AND "OVER-MINUSED IN CLINIC WITH REFRACTION." REPORTEDLY THE DEVICE DID NOT FAIL TO PERFORM AS EXPECTED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725195&pc=MTA | 2023826-2021-04172 | 2021-09-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -14.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT LENGTH AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12909815&pc=MTA | 2023826-2021-04616 | 2021-09-21 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -3.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840513&pc=MTA | 2023826-2021-04393 | 2021-09-25 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION : THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATONS. CLAIM # (B)(4). Manufacturer Narrative: B5: LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246071&pc=MTA | 2023826-2022-01311 | 2021-09-26 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246554&pc=MTA | 2023826-2022-01312 | 2021-09-26 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNABLE TO IDENTIFY, WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. UNABLE TO IDENTIFY, WHICH LENS PERTAINS TO, WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14659120&pc=MTA | 2023826-2022-01745 | 2021-10-09 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH AN IDENTICAL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146229&pc=MTA | 2023826-2022-00053 | 2021-10-12 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER, REFRACTIVE ERROR".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER, REFRACTIVE ERROR". Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911555&pc=MTA | 2023826-2022-00961 | 2021-10-12 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -15.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182540&pc=MTA | 2023826-2022-00029 | 2021-10-25 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182541&pc=MTA | 2023826-2022-00030 | 2021-10-25 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE LENS WAS REPLACED WITH A TICL VERTICALLY PLACED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983927&pc=MTA | 2023826-2022-01048 | 2021-10-27 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATIENT IS WEARING GLASSES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS IMPLANTED ON (B)(6) 2021 AND NOT ON (B)(6) 2021 AS INITIALLY REPORTED. IT WAS REPORTED THAT THE PATIENT CONTINUES WITH UNEXPECTED RESIDUAL OF -1.25 IN HER LAST CONTROL. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973349&pc=MTA | 2023826-2021-04699 | 2021-11-08 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "ASTIGMATISM AND KERATOMETRY FROM THE START." THE REPORTER STATES THAT POSSIBLY THE SURGEON UNDERESTIMATED THE ASTIGMATISM-KERATOMETRY GIVES A CYLINDER OF 1.4. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459224&pc=MTA | 2023826-2022-00389 | 2021-11-19 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR REP# 2023826-2022-00392 FOR REPLACEMENT LENS. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317681&pc=MTA | 2023826-2022-00204 | 2021-11-24 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -14.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862585&pc=MTA | 2023826-2022-01899 | 2021-11-24 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173031&pc=MTA | 2023826-2022-00099 | 2021-12-06 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ALSO REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15738720&pc=MTA | 2023826-2022-03992 | 2021-12-13 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THE PATIENT HAD PERMANENT LOSS OF BCVA AND BLURRED VISION. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "RESIDUAL ASTIGMATISM REMAINS". THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT MODEL (TORIC) LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE COMMENTS INDICATED "GOOD APPEARANCE OF THE ANTERIOR CHAMBER MENISCUS NORMAL BUT NORMAL". Manufacturer Narrative: ADDITIONAL INFORMATION: D6B: EXPLANT DATE: (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862904&pc=MTA | 2023826-2022-01905 | 2021-12-30 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13981610&pc=MTA | 2023826-2022-01046 | 2022-01-05 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998315&pc=MTA | 2023826-2022-00978 | 2022-01-15 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE AND GHOSTING. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862403&pc=MTA | 2023826-2022-02091 | 2022-02-03 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504752&pc=MTA | 2023826-2022-01688 | 2022-02-09 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN MOISTURE WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795182&pc=MTA | 2023826-2022-00797 | 2022-02-16 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRY/FOGGY VISION WHEN THE LIGHT CONDITIONS GO FROM DARK TO LIGHT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944339&pc=MTA | 2023826-2022-01025 | 2022-02-24 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE WITH A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-02285 FOR REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944425&pc=MTA | 2023826-2022-01024 | 2022-02-24 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE WITH A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH AND MODEL BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL AND FUNCTIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: "THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND "BLOOD STAINS" ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS." CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365578&pc=MTA | 2023826-2022-01518 | 2022-02-25 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT -9.0 DIOPTER. THE EXCHANGE LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16720053&pc=MTA | 2023826-2023-01357 | 2022-03-18 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON 11-FEB-2023 THE LENS WAS EXCHANGED WITH THE SAME LENGTH LENS BUT TORIC MODEL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220682&pc=MTA | 2023826-2022-01098 | 2022-03-24 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM6_12.1 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE REPORTED WITH -1.50 SPHERE RESIDUAL ERROR. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220740&pc=MTA | 2023826-2022-01099 | 2022-03-24 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM6_12.1 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE REPORTED WITH -1.75 SPHERE RESIDUAL ERROR. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862698&pc=MTA | 2023826-2022-02025 | 2022-04-01 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801811&pc=MTA | 2023826-2022-03954 | 2022-04-04 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. Manufacturer Narrative: DEVICE EVALUATED BY MFR: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND EXCHANGED ON (B)(6) 2022. IT WAS REPORTED AS USER ERROR AND, "WRONG LENS ORDERED" WAS STATED. ALSO, STATUS OF THE EYE WAS REPORTED AS, "ISSUE RESOLVED." CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR THE SAME MODEL/ LENGTH LENS; MORE MYOPIC POWER AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS " ISSUE RESOLVED". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY "WRONG LENS ORDERED". DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14417748&pc=MTA | 2023826-2022-01591 | 2022-04-06 | Injury | STAAR SURGICAL COMPANY | 2022-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS PERMANENT LOSS OF BCVA AND IT WAS NOTED "PATIENT HAS COMPLAINTS OF POOR FAR AND NEAR VISION". LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS PERMANENT LOSS OF BCVA AND IT WAS NOTED "PATIENT HAS COMPLAINTS OF POOR FAR AND NEAR VISION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14455997&pc=MTA | 2023826-2022-01621 | 2022-04-11 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT HAS COMPLAINTS IN LEFT EYE (OS) OF POOR FAR AND NEAR VISION POST VIVA IMPLANTABLE COLLAMER LENS IMPLANTATION. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VICM6 12.6 -6.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PERMANENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED, PATIENT HAS NOT DECIDED FOR AN EXCHANGE. IN THE REPORTER OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM #(B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6: INVESTIGATION TYPE 4110 ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153652&pc=MTA | 2023826-2022-02582 | 2022-05-10 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN/BROKEN HAPTIC. BLOOD STAINS WERE OBSERVED ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE: SHOULD BE CORRECT TO 31-JUL-2024. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15061062&pc=MTA | 2023826-2022-02329 | 2022-05-12 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGE WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND IT WAS NOTED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423702&pc=MTA | 2023826-2023-00619 | 2022-05-18 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC MODEL AND SAME SIZE LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862210&pc=MTA | 2023826-2022-01907 | 2022-05-31 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE REPORTED. REPORTEDLY RESIDUAL ASTIGMATISM WAS CORRECTED BY INCISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697668&pc=MTA | 2023826-2022-03892 | 2022-06-10 | Injury | STAAR SURGICAL COMPANY | 2022-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO PRODUCT NON-CONFORMITY. REPORTEDLY, "FELLOW EYE LENS VAULTING WAS EXPLANTED DUE TO EXCESSIVE INFLAMMATION AND CELLS POST OP, PATIENT IS TRAUMATISED WITH THE WHOLE EXPERIENCE, PATIENT CURRENTLY ON GLAUCOMA MEDICATIONS G GANFORT ON AND G SIMBRINZA TDS ON RE." CAUSE OF THE EVENT IS UNKNOWN. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: MODEL NUMBER- SHOULD BE CORRECTED TO STATE VICM5_12.1. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298835&pc=MTA | 2023826-2022-02920 | 2022-06-11 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 03-31-2025. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17228874&pc=MTA | 2023826-2023-02567 | 2022-06-17 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE TREATMENT WAS PERFORMED ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15276405&pc=MTA | 2023826-2022-02828 | 2022-06-24 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: B5: CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE- 11/30/2024. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822417&pc=MTA | 2023826-2023-04280 | 2022-06-28 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (TORIC LENS) AND POWER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PER REPORTER THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED "PATIENT DIDN'T REACH EXPECTED VA" WAS REPORTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921202&pc=MTA | 2023826-2022-04343 | 2022-07-06 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18101043&pc=MTA | 2023826-2023-05054 | 2022-07-07 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, SLIGHTLY MORE HYPEROPIC THAN PLANNED, UDVA IS 1.0. PATIENT IS DISSATISFIED WITH NEAR VISION AND COMPUTER DISTANCE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15160364&pc=MTA | 2023826-2022-02614 | 2022-07-17 | Injury | STAAR SURGICAL COMPANY | 2022-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. IT WAS ADDITIONALLY NOTED AFTER INITIAL IMPLANTATION OF POSTOPERATIVE VISION; IT COULD NOT MEET THE DAILY WORK AND LIFE NEEDS; AND WAS DECIDED TO BE EXCHANGED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882900&pc=MTA | 2023826-2022-04253 | 2022-07-17 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -14.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712752&pc=MTA | 2023826-2022-03897 | 2022-07-24 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712753&pc=MTA | 2023826-2022-03898 | 2022-07-24 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -8.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728673&pc=MTA | 2023826-2023-01334 | 2022-08-07 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR DUE TO REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712749&pc=MTA | 2023826-2022-03896 | 2022-08-22 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO TRANSIENT LOSS OF BCVA. IN A SUBSEQUENT SURGERY LATER THAT DAY, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: CORRECTED TO - 07640145865114 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849481&pc=MTA | 2023826-2022-04262 | 2022-09-01 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0. DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH A DIFFERENT DIOPTER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. D6B: ON (B)(6) 2022. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937239&pc=MTA | 2023826-2022-04285 | 2022-09-07 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE WAS REPORTED AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT A DIFFERENT DIOPTER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3: DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR IS CORRECTED TO 07-NOV-2022. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577434&pc=MTA | 2023826-2022-03585 | 2022-09-12 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE DURING THE POSTOPERATIVE EXAMINATION WITH THE PATIENT. ON (B)(6) 2022 THE SURGEON EXCHANGED THE LENS WITH A SAME LENGTH/MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE SURGEON SUSPECTS THE LENS DIOPTER WAS INCORRECT WITH THE INITIAL IMPLANTED LENS AND WANTS INVESTIGATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301418&pc=MTA | 2023826-2023-00379 | 2022-09-13 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: LENS RETURNED 20-FEB-2023. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154822&pc=MTA | 2023826-2023-00129 | 2022-09-15 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODE AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9:(B)(6)2022. DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154821&pc=MTA | 2023826-2023-00131 | 2022-09-29 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16414710&pc=MTA | 2023826-2023-00520 | 2022-10-08 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REPLACED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "POSTOPERATIVELY, THE PATIENT HAS DIFFICULTY OPENING HER EYES EVEN AFTER MYDRIASIS, AND ALTHOUGH THE OD HAS NOT REPLACED LENSES, PATIENT'S VISUAL ACUITY HAS INCREASED FROM 0.8 TO 1.2 WHICH I WONDER MAY HAVE A PSYCHOLOGICAL COMPONENT." ALSO, "I CONSIDER THAT POOR CORNEAL CONDITION SEEMS TO HAVE HAD A SIGNIFICANT IMPACT ON LENS POWER." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16828266&pc=MTA | 2023826-2023-01532 | 2022-10-08 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17338886&pc=MTA | 2023826-2023-02982 | 2022-10-12 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -09.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THE PATIENT HAD VISUAL ACUITY DISSATISFACTION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344588&pc=MTA | 2023826-2023-00474 | 2022-11-09 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT - CLAIM# 727957 FOR FELLOW EYE (OD). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16356403&pc=MTA | 2023826-2023-00475 | 2022-11-09 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT CLAIM#: (B)(4) FOR FELLOW EYE (OS). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464988&pc=MTA | 2023826-2023-00719 | 2022-12-02 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16150999&pc=MTA | 2023826-2023-00139 | 2022-12-08 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119689&pc=MTA | 2023826-2023-02165 | 2022-12-25 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6 12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 27/NOV/2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. REPORTEDLY, LENS MAYBE EXCHANGED. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119700&pc=MTA | 2023826-2023-02166 | 2022-12-25 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. REPORTEDLY, LENS MAYBE EXCHANGED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250313&pc=MTA | 2023826-2023-00320 | 2023-01-04 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED "PATIENT HAPPY WITH DISTANCE VISION BUT DISSATISFIED WITH INTERMEDIATE AND NEAR VISION". LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250339&pc=MTA | 2023826-2023-00319 | 2023-01-04 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -10.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE; IT WAS ADDITIONALLY NOTED "PATIENT HAPPY WITH DISTANCE VISION BUT DISSATISFIED WITH INTERMEDIATE AND NEAR VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17121541&pc=MTA | 2023826-2023-02331 | 2023-01-10 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2023. THE PATIENT EXPERIENCED BLURRED VISION WITH VISION QUALITY ISSUES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281081&pc=MTA | 2023826-2023-00403 | 2023-01-15 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -5.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED; "DUE TO THE LOW DEGREE OF ASTIGMATISM OF THE PATIENT BEFORE SURGERY -- 0.75D, THE PATIENT DID NOT CHOOSE ASTIGMATISM CORRECTION IN THE SURGICAL DESIGN AFTER PREOPERATIVE COMMUNICATION WITH THE PATIENT, AND THE PATIENT SHOWED VISUAL SENSITIVITY AND DISCOMFORT AFTER SURGERY, SO THE PATIENT REPORTED POOR CORRECTED VISION WITHOUT ADDING ASTIGMATISM CORRECTION, AND HIS VISION WAS RESTORED TO THE BEST CORRECTED VISION AFTER REPLACING THE ASTIGMATISM CORRECTION CRYSTAL!". Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281082&pc=MTA | 2023826-2023-00395 | 2023-01-15 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED; "DUE TO THE LOW DEGREE OF ASTIGMATISM OF THE PATIENT BEFORE SURGERY -- 0.75D, THE PATIENT DID NOT CHOOSE ASTIGMATISM CORRECTION IN THE SURGICAL DESIGN AFTER PREOPERATIVE COMMUNICATION WITH THE PATIENT, AND THE PATIENT SHOWED VISUAL SENSITIVITY AND DISCOMFORT AFTER SURGERY, SO THE PATIENT REPORTED POOR CORRECTED VISION WITHOUT ADDING ASTIGMATISM CORRECTION, AND HIS VISION WAS RESTORED TO THE BEST CORRECTED VISION AFTER REPLACING THE ASTIGMATISM CORRECTION CRYSTAL!". Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678113&pc=MTA | 2023826-2023-01214 | 2023-01-24 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -2.75 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON 16 MAY 2023 FOR A SAME MODEL DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892521&pc=MTA | 2023826-2023-01748 | 2023-02-07 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A DEFORMED OPTIC AND THE HAPTICS BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131352&pc=MTA | 2023826-2023-02389 | 2023-02-09 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED REFRACTIVE TREATMENT (LASIK/PRK/ETC). THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "ALL IS FINE AND PATIENT IS HAPPY, LASER TOUCH UP TREATMENT ON (B)(6) 2023, HAS CORRECTED THE POST ICL" THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066600&pc=MTA | 2023826-2023-02225 | 2023-02-10 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -8.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH, BUT DIFFERENT MODEL/POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792960&pc=MTA | 2023826-2023-01479 | 2023-02-11 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BEFORE REPLACING THE FIRST LENS, THE DOCTOR PERFORMED LENS ALIGNMENT AND ROTATION IN THE EYE FOUND THAT THE ASTIGMATIC AXIS OF THE PATIENT'S SUBJECTIVE OPTOMETRY HAD CHANGED. THEREFORE, IT IS SUSPECTED THAT THE FIRST LENS IMPLANTED IS AN TICL." Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678778&pc=MTA | 2023826-2023-01178 | 2023-02-14 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6:EXPLANT DATE (B)(6) 2023. B5-ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: PT INFORMATION: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17696360&pc=MTA | 2023826-2023-03777 | 2023-02-14 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971404&pc=MTA | 2023826-2023-04572 | 2023-02-17 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SUPRISE. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17992929&pc=MTA | 2023826-2023-04742 | 2023-02-18 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748688&pc=MTA | 2023826-2023-01354 | 2023-03-01 | Injury | STAAR SURGICAL COMPANY | 2023-04-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16746899&pc=MTA | 2023826-2023-01350 | 2023-03-14 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS ON (B)(6) 2023 AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761001&pc=MTA | 2023826-2023-04001 | 2023-03-17 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16810401&pc=MTA | 2023826-2023-01482 | 2023-03-20 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS "GOOD". Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. UNABLE TO PERFORM LENS WIDTH DUE TO HAPTIC DAMAGE, LENSES HAVE BEEN SWITCHED IN YELLOW BAGS WHILE RETURNING. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNKNOWN. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16810453&pc=MTA | 2023826-2023-01481 | 2023-03-20 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS "GOOD." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. UNABLE TO PERFORM LENS WIDTH DUE TO HAPTIC DAMAGE, LENSES HAVE BEEN SWITCHED IN YELLOW BAGS WHILE RETURNING. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592020&pc=MTA | 2023826-2023-03620 | 2023-03-24 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF A -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A2- " ON (B)(6) 2023" SHOULD BE REMOVED AND " ON (B)(6) 1989" SHOULD BE ADDED. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954043&pc=MTA | 2023826-2023-04567 | 2023-03-24 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SUPRISE AND REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE TORIC MODEL AND STRONGER MYOPIC POWER, SAME LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. 6B: "ON (B)(6) 2023" SHOULD BE REMOVED AND "ON (B)(6) 2023" SHOULD BE ADDED. DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213970&pc=MTA | 2023826-2023-02554 | 2023-03-30 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17214401&pc=MTA | 2023826-2023-02555 | 2023-03-30 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991214&pc=MTA | 2023826-2023-01769 | 2023-04-03 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL . VISUAL INSPECTION FOUND THE OPTIC TORN AND SPLIT AND THE HAPTIC TORN AND BROKEN. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DIMENSIONAL INSPECTION COULD NOT BE PERFORMED DUE TO EXCESSIVE LENS DAMAGE. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410063&pc=MTA | 2023826-2023-03193 | 2023-04-04 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434430&pc=MTA | 2023826-2023-03192 | 2023-04-09 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477568&pc=MTA | 2023826-2023-03350 | 2023-04-10 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PER REPORTER PATIENT WAS DISSATISFIED WITH VA. ON (B)(6) 2023 LENS WAS EXCHANGED FOR TORIC LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840230&pc=MTA | 2023826-2023-04285 | 2023-04-12 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330141&pc=MTA | 2023826-2023-02928 | 2023-04-14 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT TORIC LENS/DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION, LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148654&pc=MTA | 2023826-2023-02406 | 2023-04-26 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -3.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6)2023 THE LENS WAS EXPLANTED AND LATER ON THIS DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17818024&pc=MTA | 2023826-2023-04288 | 2023-05-12 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND OPTIC SPLIT. UNABLE TO PERFORM DIMENSIONAL/FUNCTIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477662&pc=MTA | 2023826-2023-03352 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BROKEN/BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840052&pc=MTA | 2023826-2023-04284 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18803156&pc=MTA | 2023826-2024-00837 | 2023-05-26 | Injury | STAAR SURGICAL COMPANY | 2024-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2024, THE LENS WAS EXCHANGED FOR SAME SIZE, DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4: A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: 734061. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633295&pc=MTA | 2023826-2023-03533 | 2023-05-29 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - PRODUCT WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862008&pc=MTA | 2023826-2023-04063 | 2023-06-09 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2021. REFRACTIVE CHANGE OVERTIME AND LOW VAULT WERE OBSERVED(ASSESSED ON (B)(6) 2022) . LENS WAS EXCHANGE WITH A DIFFERENT POWER LENS ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455639&pc=MTA | 2023826-2023-03219 | 2023-06-15 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "REPLACE ASTIGMATIC CRYSTALS BASED ON POSTOPERATIVE OPTOMETRY RESULTS." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455641&pc=MTA | 2023826-2023-03220 | 2023-06-15 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "REPLACE ASTIGMATIC CRYSTALS BASED ON POSTOPERATIVE OPTOMETRY RESULTS." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358296&pc=MTA | 2023826-2023-02969 | 2023-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF A -16.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358386&pc=MTA | 2023826-2023-02968 | 2023-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF A -16 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874320&pc=MTA | 2023826-2023-04417 | 2023-07-04 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM612.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581258&pc=MTA | 2023826-2023-03589 | 2023-07-13 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO TRANSIENT LOSS OF BCVA . THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581262&pc=MTA | 2023826-2023-03588 | 2023-07-13 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO TRANSIENT LOSS OF BCVA . THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17545681&pc=MTA | 2023826-2023-03517 | 2023-07-16 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761235&pc=MTA | 2023826-2023-04075 | 2023-08-08 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395190&pc=MTA | 2023826-2023-05776 | 2023-08-13 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18370009&pc=MTA | 2023826-2023-05621 | 2023-08-21 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. "DOCTOR CALCULATED LENS WRONGLY." Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375462&pc=MTA | 2023826-2023-05765 | 2023-08-22 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -1.25 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(3) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(3) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560295&pc=MTA | 2023826-2024-00361 | 2023-08-22 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(6). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17694988&pc=MTA | 2023826-2023-03947 | 2023-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT HAS REFRACTIVE SURPRISE. THE SURGEON PERFORMED LRI SURGERY. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS REPORTED AS A USER ERROR WITH NO DEVICE CAUSALITY AND NOTED "SURGEON PLANNED ASTIGMATISM CORRECTION WITH LRI WHICH DOESN'T WORK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H1 - TYPE OF EVENT - SERIOUS CORRECTED TO MALFUNCTION. H4 - MANUFACTURE DATE: 31-MAR-2026 CORRECTED TO 05-APRIL-2023. Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17694989&pc=MTA | 2023826-2023-03948 | 2023-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 20-JUL-2023. IT WAS REPORTED THAT THE PATIENT HAS REFRACTIVE SURPRISE. THE SURGEON PERFORMED LRI SURGERY. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS REPORTED AS A USER ERROR WITH NO DEVICE CAUSALITY AND NOTED "SURGEON PLANNED ASTIGMATISM CORRECTION WITH LRI WHICH DOESN'T WORK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H1 - TYPE OF EVENT: SERIOUS CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271396&pc=MTA | 2023826-2023-05501 | 2023-08-25 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17891552&pc=MTA | 2023826-2023-04271 | 2023-08-27 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. EXCESSIVE VAULT WAS REPORTED AND BLURRED VISION. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18233713&pc=MTA | 2023826-2023-05366 | 2023-08-31 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18057401&pc=MTA | 2023826-2023-04943 | 2023-09-14 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2M VICMO13.2; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18066236&pc=MTA | 2023826-2023-04855 | 2023-09-15 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136064&pc=MTA | 2023826-2023-05082 | 2023-09-15 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136489&pc=MTA | 2023826-2023-05058 | 2023-09-16 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18478055&pc=MTA | 2023826-2024-00053 | 2023-09-28 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416892&pc=MTA | 2023826-2023-05827 | 2023-09-29 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062883&pc=MTA | 2023826-2023-04964 | 2023-10-04 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS USER ERROR. THE REPORTER STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS USER ERROR. THE REPORTER STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4) Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323057&pc=MTA | 2023826-2023-05653 | 2023-10-10 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED DUE TO REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2023 A REPLACEMENT LENS OF DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18478523&pc=MTA | 2023826-2024-00054 | 2023-10-17 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH TORIC LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271708&pc=MTA | 2023826-2023-05505 | 2023-10-19 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258189&pc=MTA | 2023826-2023-05312 | 2023-10-25 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258205&pc=MTA | 2023826-2023-05311 | 2023-10-25 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196960&pc=MTA | 2023826-2023-05303 | 2023-10-27 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 18- SEP-2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18321863&pc=MTA | 2023826-2023-05652 | 2023-11-02 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. CAUSE OF EVEN WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18522597&pc=MTA | 2023826-2024-00230 | 2023-11-27 | Injury | STAAR SURGICAL COMPANY | 2024-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856340&pc=MTA | 2023826-2020-02698 | NA | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS RESIDUAL REFRACTIVE ERROR WHICH HAD TO BE CORRECTED WITH LASIK. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308938&pc=MTA | 2023826-2021-00226 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ON (B)(6) 2020 REFRACTIVE TREATMENT (LASIK, PRK, ETC.) WAS PERFORMED. REPORTEDLY, "THE AMOUNT OF HIPER CORRECTION IS ONLY XPLAINED BY A MISTAKE IN THE LENS POWER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148271&pc=MTA | 2023826-2021-02474 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, THE LENS WAS "CHANGED" DUE TO REFRACTION ERROR. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657017&pc=MTA | 2023826-2022-00665 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM6_13.7; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE POST OP REPORT INDICATES "SHE SAYS THAT SEEING IS BETTER OVERALL. ESPECIALLY IN THE DISTANCE. HOWEVER; SHE IS NOW USING READING GLASSES WITH +2.00DPT. I GAVE HER CONTACT LENSES WITH +0.50 DPT FOR THE LEFT EYE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT - "LENS WILL REMAIN INSIDE. THE PATIENT DID NOT SEE BETTER WITH CONTACT LENSES; ONLY IN THE INTERMEDIATE AREA. SO WE'RE GOING TO TELL HER IT'S BEST TO LEAVE IT AS IS". CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657018&pc=MTA | 2023826-2022-00664 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE POST OP REPORT INDICATES "SHE SAYS THAT SEEING IS BETTER OVERALL. ESPECIALLY IN THE DISTANCE. HOWEVER; SHE IS NOW USING READING GLASSES WITH +2.00DPT. I GAVE HER CONTACT LENSES WITH +0.50 DPT FOR THE LEFT EYE". ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT "LENS WILL REMAIN INSIDE. THE PATIENT DID NOT SEE BETTER WITH CONTACT LENSES; ONLY IN THE INTERMEDIATE AREA. SO WE'RE GOING TO TELL HER IT'S BEST TO LEAVE IT AS IS". CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14283152&pc=MTA | 2023826-2022-01402 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2022; AS A REPLACEMENT LENS TO A PREVIOUS LOW VAULTING AND CATARACT ISSUE. IT WAS MENTIONED THE PROBLEM WAS RESOLVED; HOWEVER NOTED "AFTER LENS REPLACEMENT THERE IS A DECREASE IN VISUAL ACUITY; BUT THE PATIENT IS UNDER OBSERVATION. IT IS EXPECTED THAT VISION WILL RECOVER IN THE FUTURE". LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-01401 FOR INITIAL LENS CLAIM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: (B)(4): LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14749667&pc=MTA | 2023826-2022-01944 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THERE WAS REFRACTIVE SURPRISE REPORTED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14749668&pc=MTA | 2023826-2022-01943 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS REFRACTIVE SURPRISE REPORTED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282555&pc=MTA | 2023826-2022-02879 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SURPRISE. STATUS OF THE EYE IS REPORTED AS "RESOLVED." CAUSE OF THE EVENT IS USER ERROR. REPORTEDLY, "NEEDED A TORIC RATHER THAN A SPHERICAL LENS." Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: G3: SUPP1 DATE: 10/04/2022. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15709881&pc=MTA | 2023826-2022-03909 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15709888&pc=MTA | 2023826-2022-03910 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801769&pc=MTA | 2023826-2022-04131 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -7.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON 12-JUL-2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801803&pc=MTA | 2023826-2022-04130 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON 27-JUL-2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801953&pc=MTA | 2023826-2022-04132 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2024. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED ON (B)(6)2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15808829&pc=MTA | 2023826-2022-04009 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1MM VICM6 OF -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811139&pc=MTA | 2023826-2022-04010 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1MM VICM6 OF -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15831959&pc=MTA | 2023826-2022-04018 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED AND INCREASES OF CYL. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15834977&pc=MTA | 2023826-2022-03242 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6 MM VICM612.6, -15.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S RIGHT EYE (OD) (B)(6) 2021. ON (B)(6) 2021, LVC WAS PERFORMED DUE TO RESIDUAL REFRACTION AND POOR UDVA. Manufacturer Narrative: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15835000&pc=MTA | 2023826-2022-03243 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM VICM612.6, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S LEFT EYE (OS) (B)(6) 2021. ON (B)(6) 2021, LVC WAS PERFORMED DUE TO RESIDUAL REFRACTION AND POOR UDVA. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946952&pc=MTA | 2023826-2022-04449 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S EYE ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S EYE ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "LENS WAS IMPLANTED UPSIDE DOWN". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND TORN. SUBSEQUENT IMPLANTATION OF NEW LENS WAS INITIATED BUT REPORTEDLY, "LENS GOT UNSTERILE AS IT FELL DOWN ON FLOOR AND THE SURGERY HAS TO BE ABORTED". PROBLEM NOT RESOLVED. WORK ORDER SEARCH FOUND NO SIMILAR EVENTS. PATIENT INJURY NOT REPORTED. D6: EXPLANT DATE (B)(6) 2023. H6: 4627-LENS EXPLANTED. CLAIM#(B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: B5: REPORTEDLY, "PATIENT IS NOT COMING AND IS HAPPY WITH THE VISION BILATERALLY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE:03-MAY-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC BROKEN AND THE HAPTIC BROKEN. DUE TO SIGNIFICANT LENS DAMAGE DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE NOT POSSIBLE TO BE PERFORMED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16517890&pc=MTA | 2023826-2023-00844 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON AN UNKNOWN DATE THE LENS WAS EXPLANTED. D9: RETURN DATE: 30-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL L INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: LENS DIOPTER SHOULD BE CORRECTED TO: -7.0. ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. THE PROBLEM HAS NOT RESOLVED. ADDITIONAL INFORMATION, A RELATIVELY LARGE HYPEROPIC SHIFT WAS FOUND. REPORTEDLY, THE EXCHANGE IS GOING TO BE DONE IN (B)(6) BUT THE EXACT DATE HAS NOT BEEN DECIDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, "I CAN SEE MORE COMFORTABLY THAN IN THE PREVIOUS VISIT AND BEFORE THE EXCHANGE AND WILL WAIT AND SEE...ANOTHER EXCHANGE OR SOME INTERVENTION HAS NOT BEEN PLANNED". PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16517916&pc=MTA | 2023826-2023-00843 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 11-NOV-2022. THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. THE PROBLEM HAS NOT RESOLVED. REPORTEDLY, THE EXCHANGE IS GOING TO BE DONE IN (B)(6) BUT THE EXACT DATE HAS NOT BEEN DECIDED. CLAIM#(B)(4). Manufacturer Narrative: DIOPTER SHOULD BE CORRECTED TO STATE: -7.50. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON AN UNKNOWN DATE THE LENS WAS EXPLANTED. D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, "I CAN SEE MORE COMFORTABLY THAN IN THE PREVIOUS VISIT AND BEFORE THE EXCHANGE AND WILL WAIT AND SEE...ANOTHER EXCHANGE OR SOME INTERVENTION HAS NOT BEEN PLANNED". PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS" SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17607594&pc=MTA | 2023826-2023-03630 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023, DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "BLURRED VISION DUE TO RESIDUAL ASTIGMATISM. SUBJECTIVE REFRACTION AT FINAL VISIT. +0-1.00X130." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187772&pc=MTA | 2023826-2023-04974 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. RESIDUAL REFRACTIVE ERROR IS REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. - PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. - IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18421048&pc=MTA | 2023826-2023-05810 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -12.5 WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656495&pc=MTA | 2023826-2019-00942 | NA | Malfunction | STAAR SURGICAL COMPANY AG | 2019-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), BACK IN 2003/2004. THE PATIENT WAS EXPERIENCING REFRACTIVE CHANGE OVER TIME AND THE SURGEON WAS PLANNING TO EXPLANT AND EXCHANGE THE LENS. BUT THE SURGEON CHANGED HIS MIND AND DECIDED TO MONITOR THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656503&pc=MTA | 2023826-2019-00943 | NA | Malfunction | STAAR SURGICAL COMPANY AG | 2019-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), BACK IN 2003/2004. THE PATIENT WAS EXPERIENCING REFRACTIVE CHANGE OVER TIME AND THE SURGEON WAS PLANNING TO EXPLANT AND EXCHANGE THE LENS. BUT THE SURGEON CHANGED HIS MIND AND DECIDED TO MONITOR THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14713784&pc=MTA | 2023826-2022-01864 | 2022-05-16 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.25 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!". ADDITIONAL INFORMATION: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14714037&pc=MTA | 2023826-2022-01865 | 2022-05-16 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.25 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!". ADDITIONAL INFORMATION: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15557432&pc=MTA | 2023826-2022-03451 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OBSERVED REFRACTIVE CHANGE OVER TIME IN THE PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.5 DIOPTER, WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT ON (B)(6) 2018. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ASKING FOR EXCHANGE BECUASE OF MYOPICIZATION OF -1.00. SIZING IS OK." PLANNED EXCHANGED FOR (B)(6) 2023. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND WITH ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA SUPP1 G3: (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: B1: PRODUCT PROBLEM. B2: ENTRY SHOULD BE BLANK/VOID. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5 WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 07/NOV/2018. THE PATIENT REFRACTIVE CHANGE OVER TIME. ON 11/JAN/2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, BUT RESIDUE AND DEBRIS ON FOOTPLATE/HAPTIC OF LENS. CLAIM# (B)(4). Manufacturer Narrative: PT. INFO:UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941306&pc=MTA | 2023826-2019-01580 | 2019-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941316&pc=MTA | 2023826-2019-01579 | 2019-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932874&pc=MTA | 2023826-2019-01548 | 2019-07-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THREE MONTHS LATER, THE PATIENT IS NOT SATISFIED WITH HIS VISUAL ACUITY AND COMPLAINS OF BLURRED VISION. THE LENS HAS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAS DECIDED TO WAIT A COUPLE OF MONTHS AND THEN WILL HAVE HIS VISION CHECKED AGAIN. TO DATE THE PATIENT IS FINE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10744678&pc=MTA | 2023826-2020-02533 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE AND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856354&pc=MTA | 2023826-2020-02692 | 2020-09-23 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS, THE SURGEON REPORTING EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND PERMANENT LOSS OF BCVA. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "1 PER VAULT, CHANGE ICL WITH 12.6 LENGTH." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343188&pc=MTA | 2023826-2021-00400 | 2021-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. REPORTEDLY, "ADDITIAONAL REFRACTION INCLUDING CR WILL BE MEASURED. THEN, THE SURGEON WILL CONSIDER A REPLACEMENT SURGERY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343198&pc=MTA | 2023826-2021-00399 | 2021-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS WAS EXPLANTED ON (B)(6)2021 DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. REPORTEDLY, "ADDITIAONAL REFRACTION INCLUDING CR WILL BE MEASURED. THEN, THE SURGEON WILL CONSIDER A REPLACEMENT SURGERY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE.VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. THE WIDTH OF THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. UNABLE TO PERFORM DIAMETER VERIFICATION AS THE LENS WAS DAMAGED. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214206&pc=MTA | 2023826-2021-02660 | 2021-06-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM613.2, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 . THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, VAULT IS OK NOW AND PATIENT IS MORE HAPPY WITH DISTANCE AND NEAR VISION. PATIENT UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PATIENT IS SEEING MUCH BETTER. HAS TO GET USED SEEING IN DARK AND LIGHT CONDITIONS. REPORTEDLY, "IT SEEMS PROBLEM IS RESOLVED." CLAIM #(B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214232&pc=MTA | 2023826-2021-02659 | 2021-06-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM613.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, VAULT IS OK NOW AND PATIENT IS MORE HAPPY WITH DISTANCE AND NEAR VISION. PATIENT UNDER OBSERVATION. Manufacturer Narrative: PATIENT IS SEEING MUCH BETTER. HAS TO GET USED SEEING IN DARK AND LIGHT CONDITIONS. REPORTEDLY, "IT SEEMS PROBLEM IS RESOLVED." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220828&pc=MTA | 2023826-2022-01102 | 2022-02-01 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE REPORTED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS USER ERROR - ICL 13.7 INSTEAD OF 12.6. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN,RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS USER ERROR - ICL 13.7 INSTEAD OF 12.6. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: G3: DATE RECEIVED BY MANUFACTURER: 09/08/2022. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14649879&pc=MTA | 2023826-2022-01839 | 2022-02-13 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND DIFFERENT MODEL. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED "THE PATIENT'S RIGHT EYE HAD ASTIGMATISM OF -0.75D BEFORE OPERATION. AND ARCH HEIGHT IS LOW. THE ICL WITHOUT ASTIGMATISM WAS SELECTED DURING THE FIRST OPERATION. IT WAS FOUND THAT THE PATIENT HAD ASTIGMATISM OF -1.50D AFTER OPERATION. CONSIDERING THAT IT WAS CAUSED BY THE SURGICAL INCISION; THE ICL WITH ASTIGMATISM WAS SELECTED DURING THE SECOND OPERATION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS IN A RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED; HAPTIC DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(6). Manufacturer Narrative: CORRECTED DATA: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862479&pc=MTA | 2023826-2022-01906 | 2022-05-31 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE REPORTED. REPORTEDLY RESIDUAL ASTIGMATISM WAS CORRECTED BY INCISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, "THE MAIN COMPLAINT OF THIS CASE IS THE LOW VAULT. DUE TO LOW VAULT AFTER THE INITIAL SURGERY, THE SURGEON DECIDED TO EXCHANGE WITH A LARGER SIZE . HE ALSO CONSIDERED CHANGING THE DIOPTER AT THE SAME TIME.' THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1583 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: A2- CORRECTED TO : 25-JAN-1997 IN INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109368&pc=MTA | 2023826-2023-00085 | 2022-08-04 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109425&pc=MTA | 2023826-2023-00086 | 2022-08-04 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588754&pc=MTA | 2023826-2022-03592 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: UNK. DEVICE EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED AN 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE, IRIDOCORNEAL ANGLES ARE OPEN AND THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT WAS REPORTED AS "UNKNOWN". H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 15-FEB-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588810&pc=MTA | 2023826-2022-03593 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE, IRIDOCORNEAL ANGLES ARE OPEN AND THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT WAS REPORTED AS "UNKNOWN". H6: WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15810735&pc=MTA | 2023826-2022-04175 | 2022-10-10 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -03.50 DIOPTER IMPLANTALE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR# 2023826-2022-04176 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811248&pc=MTA | 2023826-2022-04177 | 2022-10-25 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT NOW HAS EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 DIFFERENT MODEL LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865570&pc=MTA | 2023826-2023-01628 | 2023-02-04 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT, CLAIM# (B)(4) FOR PATIENT'S FELLOW EYE (OS). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865571&pc=MTA | 2023826-2023-01627 | 2023-02-04 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT, CLAIM# (B)(4) FOR PATIENT'S FELLOW EYE (OD). CLAIM (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17193864&pc=MTA | 2023826-2023-02529 | 2023-05-17 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN CONTACT LENS CASE DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302269&pc=MTA | 2023826-2023-02759 | 2023-05-30 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A,12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND BIG PUPIL, PHOTOPHOBIA, AND BLURRED VISION WERE OBSERVED. REPORTEDLY, "NO CORRELATION BETWEEN WTW, SULCUS-SULCUS AND ICL SIZE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "PENDING EXCHANGE ICL." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344449&pc=MTA | 2023826-2023-02859 | 2023-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2023. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344451&pc=MTA | 2023826-2023-02860 | 2023-06-24 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2023. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17591904&pc=MTA | 2023826-2023-03631 | 2023-06-27 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. IT WAS INDICATED THAT THE PATIENT EXPERIENCED REFRACTIVE CHANGE WHICH WAS ASSESSED ON (B)(6) 2023, THE PATIENT ALSO HAD LOW VAULTING WITH NO ROTATION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT POWER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN WITH NO DEVICE CAUSALITY AND NOTED "PATIENT HAS NOTICED DECREASED VA OS OVER LAST 6 MONTHS SINCE (B)(6) 2023. Manufacturer Narrative: CORRECTED DATA: B5: IMPLANT DATE SHOULD BE CORRECTED TO: (B)(6) 2015. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017240&pc=MTA | 2023826-2023-04814 | 2023-07-31 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986858&pc=MTA | 2023826-2023-04599 | 2023-09-02 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH DIFFERENT POWER LENS (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11375703&pc=MTA | 2023826-2021-00501 | NA | Malfunction | STAAR SUGICAL COMPANY | 2021-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM, VICM5_12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN (B)(6) 2019, THE SURGEON NOTED REFRACTIVE SURPRISE OF -1.25D ALTHOUGH THEY EXPECTED 0D. IN ADDITION THE LENS WAS SMALL AND THEY NOTICED LOW VAULTING. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8261656&pc=MTA | 2023826-2019-00159 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: (OS) LEFT EYE: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON SEEN PATIENT WHO WAS SEEKING REFRACTIVE SURGERY. UPON EXAMINATION, THE PATIENTS UNCORRECTED DISTANCE VISUAL ACUITY WAS COUNTING FINGERS AT 2 METERS OF CORRECTED 20/22 IN EACH EYE THE PATIENT WAS MYOPIC AND PRESBYOPIA AND WAS THOROUGHLY EXPLAINED TO PATIENT BEFORE SURGERY. ON (B)(6) 2018, THE PATIENT HAD UNDERGONE REFRACTIVE SURGERY BY IMPLANTING A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO HER LEFT EYE (OS). UDVA ON THE FIRST POST-OPERATIVE FOLLOW UP WAS 20/28. ON PATIENTS LAST FOLLOW UP, PATIENT ENDED UP HYPEROPIC WITH UDVA OF 20/50. REFRACTION WAS +0.50 SPH 0.75 CYL X 150. THE VAULT WAS ALSO MEASURED TO BE A BIT HIGH WHICH WAS AROUND 800 MICRONS. THE SURGEON WOULD LIKE TO EXCHANGE THE LENS WITH A SMALLER SIZE (12.1MM) WITH LENS POWER OF -8.00 DIOPTER. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURER DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8261800&pc=MTA | 2023826-2019-00160 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-01-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: (OD) RIGHT EYE: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON SEEN PATIENT WHO WAS SEEKING REFRACTIVE SURGERY. UPON EXAMINATION, THE PATIENTS UNCORRECTED DISTANCE VISUAL ACUITY WAS COUNTING FINGERS AT 2 METERS OF CORRECTED 20/22 IN EACH EYE THE PATIENT WAS MYOPIC AND PRESBYOPIA AND WAS THOROUGHLY EXPLAINED TO PATIENT BEFORE SURGERY. ON (B)(6) 2018, THE PATIENT HAD UNDERGONE REFRACTIVE SURGERY BY IMPLANTING A 12.6MM, VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO HER RIGHT EYE (OD). UDVA ON THE FIRST POET-OPERATIVE FOLLOW UP WAS 20/40. ON PATIENTS LAST FOLLOW UP, PATIENT ENDED UP HYPEROPIC WITH UDVA OF 20/100. REFRACTIVE WAS +1.75 SPH -1.00 CYL X20. THE VAULT ALSO MEASURED TO BE A BIT HIGH WHICH WAS AROUND 800 MICRONS. THE SURGEON WOULD LIKE TO EXCHANGE THE LENS WITH A SMALLER SIZE (12.1MM) WITH LENS POWER OF -8.00 DIOPTER. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE CODE 1583: CODE SHOULD BE ADDED TO ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491987&pc=MTA | 2023826-2020-01957 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN UNEXPECTED REFRACTIVE RESULT WITH ANTERIOR VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, 14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN UNEXPECTED REFRACTIVE RESULT WITH ANTERIOR VAULT. LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931407&pc=MTA | 2023826-2020-02739 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT REFREACTIVE SURPRISE HAS BEEN OBSERVED IN PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - "EXCESSIVE VAULT" SHOULD BE CORRECTED TO "LOW VAULT" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119243&pc=MTA | 2023826-2021-02332 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS 0 VAULT AND SLIGHTLY BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: THE REASON FOR THE REPLACEMENT LENS IS "LOW ARCH AND HIGH". THE PATIENT WAS CONCERNED ABOUT THE RISK OF SURGERY. THE LAST REPLACEMENT LENS WAS AN ICL. THE ASTIGMATISM WAS CORRECTED BY LASER SURGERY. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457498&pc=MTA | 2023826-2021-03551 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S EYE. THE SURGEON REPORTS "300% VAULT" AND +2.00D HYPEROPIC OUTCOME. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT CONTINUES TO IMPROVE AND SO THE DATA IS SHIFTING. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE : UNK. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15964379&pc=MTA | 2023826-2022-04176 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -03.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022; AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT EXPERIENCED LOW VAULTING WITH RESIDUAL ASTIGMATISM. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: G2: "OTHER: ITALY" SHOULD BE CORRECTED TO "OTHER: FRANCE". CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739724&pc=MTA | 2023826-2023-01368 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THERE IS ZERO VAULTING WITH REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5_13.2 -6.50 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THERE IS ZERO VAULTING WITH REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924450&pc=MTA | 2023826-2023-01822 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.; -9.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH, DIFFERENT MODEL/POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 24-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484564&pc=MTA | 2023826-2023-03356 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PATIENTS VISION AFTER SURGERY WAS 20/50 +1 WITH VAULT AROUND 1000. AFTER 2 WEEKS WITH NO VISUAL IMPROVEMENT, THE LENS WAS EXCHANGED FOR A TORIC LENS OF SHORTER LENGTH AND DIFFERENT POWER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17894899&pc=MTA | 2023826-2023-04407 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT DIOPTER LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981074&pc=MTA | 2023826-2023-04657 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH REFRACTIVE CHANGE OVER TIME WAS REPORTED. ON (B)(6) 2023 THE LENS WAS EXCHANGE FOR A LONGER LENGTH AND THE PROBLEM WAS RESOLVED. D6B- 10-OCT-2023. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15890872&pc=MTA | 2023826-2022-04363 | 2019-12-11 | Injury | STAAR SURGICAL COMPANY | 2022-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412314&pc=MTA | 2023826-2019-00423 | 2018-08-09 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -01.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT, ACD<3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552452&pc=MTA | 2023826-2019-00649 | 2019-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRY VISION. THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: SERIAL NUMBER, EXP. DATE, UDI: UNK. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11352467&pc=MTA | 2023826-2021-00490 | 2019-09-02 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412440&pc=MTA | 2023826-2019-02275 | 2019-10-29 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES LOSS OF BCVA AND REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988414&pc=MTA | 2023826-2020-00949 | 2020-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED AS AN EXCHANGE LENS. POST-OP, THE NEAR VISION WAS SLIGHTLY IMPROVED BUT STILL NOT IDEAL. THE PATIENT STILL EXPERIENCED A REFRACTIVE SURPRISE AND WILL BE MONITORED. THE LENS REMAINED IMPLANTED. SEE MFR REPORT # 2023826-2020-00948 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155602&pc=MTA | 2023826-2021-02441 | 2020-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS SPHERE TO BE OUT OF SPECIFICATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15288032&pc=MTA | 2023826-2022-02900 | 2020-08-31 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. OFF LABE - ACD <3.0. (B)(4). Manufacturer Narrative: A2: CORRECTED TO - AGE AT TIME OF EVENT, DATE OF BIRTH: (B)(6) 1995 IN THE INITIAL MDR. B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." IN INITIAL MDR. D4: EXPIRATION DATE - 30-JUN-2022 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453120&pc=MTA | 2023826-2021-00602 | 2020-11-05 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH TORIC LENS (-10.50/1.0/171 (SPHERE/CYLINDER/AXIS)) DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697121&pc=MTA | 2023826-2021-01090 | 2020-11-05 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY AFTER EXCHANGE THE POSTOP VISUAL ACUITY WAS 1.2 AND NO GHOSTING WAS REPORTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11714206&pc=MTA | 2023826-2021-01112 | 2020-12-31 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222296&pc=MTA | 2023826-2021-02819 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -1.75 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. UNABLE TO PERFORM A DIMENSIONAL AND REFRACTIVE VERIFICATION SINCE LENS HAS SIGNIFICANT DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378560&pc=MTA | 2023826-2021-03376 | 2021-03-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072264&pc=MTA | 2023826-2021-02071 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON REPORTS, "THE PATIENT HAD POOR NEAR VISION AFTER THE FIRST LENS IMPLANTATION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599748&pc=MTA | 2023826-2021-03907 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A LIQUID VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN LENS. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED WITHIN THE SAME LOT. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS, INCLUDING THE SUSPECTED DEVICE,HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COALLMER LENS; -15.0 DIOPTER; INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599734&pc=MTA | 2023826-2021-03917 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363349&pc=MTA | 2023826-2021-03286 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT WITH FIBERS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524337&pc=MTA | 2023826-2021-03717 | 2021-07-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1,-7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874462&pc=MTA | 2023826-2021-04518 | 2021-08-30 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE RECEIVED BY MANUFACTURER: THIS PRODUCT IS NOT MARKETED IN THE US .(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13319427&pc=MTA | 2023826-2022-00244 | 2021-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -02.25 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRED VISON AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS, SEE MFR. REPORT # 2023826-2022-00190 FOR THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -02.25 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRED VISON AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS, SEE MFR. REPORT # 2023826-2022-00190 FOR THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- PER EMAIL COMMUNICATION ON (B)(6), IT WAS REPORTED THAT "THE PATIENT DIDN'T HAVE 'BLURRED VISION' AFTER THE REPLACEMENT LENS WAS IMPLANTED". THIS EVENT WAS REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12903199&pc=MTA | 2023826-2021-04615 | 2021-09-29 | Injury | STAAR SURGICAL COMPANY | 2021-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIEN'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13581963&pc=MTA | 2023826-2022-00591 | 2021-11-16 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -4.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639725&pc=MTA | 2023826-2022-01509 | 2022-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION COULD NOT BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15462009&pc=MTA | 2023826-2022-03136 | 2022-01-28 | Injury | STAAR SURGICAL COMPANY | 2022-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE PATIENT IS 43 YEARS OLD. CONSIDERING THE PREBYOPIA PROBLEM, IN ORDER TO DELAY THE TIME OF WEARING READING GLASSES, HE COMMUNICATED REPEATEDLY. FINALLY, HE DECIDED TO RETAIN 50 DEGREES OF MYOPIA DURING THE SURGICAL DESIGN. HOWEVER, THERE IS A BIG DIFFERENCE BETWEEN POSTOPERATIVE CORRECTION OF VISION AND ESTIMATED CORRECTION OF VISION, WHICH CANNOT MEET THE PATIENT'S LONG-TERM NEEDS. AFTER CONSULTATION WITH THE SURGEON, HE DECIDED TO REPLACE IT WITH A CRYSTAL WITH AN UNRETAINED DEGREE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK.ACD<3.0 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15751843&pc=MTA | 2023826-2022-04040 | 2022-06-25 | Injury | STAAR SURGICAL COMPANY | 2022-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR TORIC LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H10- CLAIM# (B)(4) SHOULD BE CORRECTED TO CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE: 1494 OFF-LABEL USE (ACD<3.0MM). DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN/BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640034&pc=MTA | 2023826-2022-03702 | 2022-09-13 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6, -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 18-MAY-2022. IT WAS REPORTED THAT LIMBAL RELAXING INCISIONS (LRIS) WERE MADE TO REDUCE ASTIGMATISM BUT IT DID NOT REDUCE THE EXPECTED AMOUNT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC LENS TO ADDRESS THE ASTIGMATISM (HIGH CYLINDER POWER) AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: DEVICE CODE 1494- OFF LABEL USE- PATIENT ACD<3.00MM. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640127&pc=MTA | 2023826-2022-03703 | 2022-09-13 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6, -8.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT LIMBAL RELAXING INCISIONS (LRIS) WERE MADE TO REDUCE ASTIGMATISM BUT IT DID NOT REDUCE THE EXPECTED AMOUNT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC LENS TO ADDRESS THE ASTIGMATISM (HIGH CYLINDER POWER) AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: DEVICE CODE 1494- OFF LABEL USE- PATIENT ACD<3.00MM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993591&pc=MTA | 2023826-2022-04600 | 2022-10-05 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. THE REPORTER STATED " AFTER THE OPERATION, HER BINOCULAR VISION WAS 1.2, BUT SHE FELT UNCOMFORTABLE NEAR VISION, SO SHE ASKED HER DOMINANT EYE TO BE LOWERED BY 50 DEGREES. AT THE REPEATED REQUEST OF THE PATIENT, THE RIGHT EYE INTRAOCULAR LENS REPLACEMENT WAS PERFORMED ON THE MORNING OF NOVEMBER 5, 2022. THE POSTOPERATIVE VISUAL ACUITY WAS 1.0. THE PATIENT SAID THAT HE WAS VERY SATISFIED WITH THE CURRENT STATE, AND HE FELT COMFORTABLE LOOKING NEAR AND FAR." Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916214&pc=MTA | 2023826-2023-01799 | 2022-11-19 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2022. THE SURGEON REPORTS THERE WAS REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "PATIENT HAS SEEN NO PROBLEMS AT CLOSE RANGE WITH SATISFACTORY RESULTS". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075853&pc=MTA | 2023826-2022-04748 | 2022-11-25 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM12.0 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 15-FEB-2013. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. H4: DEVICE MANUFACTURE DATE: 23-OCT-2012. H6: 1494- OFF-LABEL: AGE<21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16542287&pc=MTA | 2023826-2023-00794 | 2022-11-29 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE PATIENT'S AGE WAS OLD AND HAD DIFFICULTY SEEING CLOSE AFTER SURGERY. THE DIOPTER OF THE LENS WAS ADJUSTED." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-EHTNICITY :UNK. OFF LABEL - AGE <45. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347483&pc=MTA | 2023826-2023-02986 | 2023-03-06 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347676&pc=MTA | 2023826-2023-02987 | 2023-03-06 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17075871&pc=MTA | 2023826-2023-02258 | 2023-04-14 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT HAD 1D ASTIGMATISM BEFORE IMPLANTATION, BUT DUE TO THE PRICE A VICMO MODEL WAS REQUIRED. AFTER IMPLANTATION THE POSTOPERATIVE VISUAL QUALITY WAS LOWER THAN EXPLANTED. ON (B)(6)2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: MODEL# SHOULD BE CORRECTED TO VICMO12.1 CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18212758&pc=MTA | 2023826-2023-05372 | 2023-06-06 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH AND DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THERE WAS NO DEVICE CAUSALITY. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549360&pc=MTA | 2023826-2023-03448 | 2023-07-11 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437836&pc=MTA | 2023826-2024-00030 | 2023-07-11 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - AGE<21. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778122&pc=MTA | 2023826-2023-04109 | 2023-08-26 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO POOR VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395378&pc=MTA | 2023826-2023-05777 | 2023-09-16 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416672&pc=MTA | 2023826-2023-05825 | 2023-10-31 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416770&pc=MTA | 2023826-2023-05826 | 2023-10-31 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525783&pc=MTA | 2023826-2024-00175 | 2023-11-22 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS REPLACED WITH A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18496445&pc=MTA | 2023826-2024-00117 | 2023-11-27 | Injury | STAAR SURGICAL COMPANY | 2024-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17393777&pc=MTA | 2023826-2023-03125 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER AS A REPLACEMENT LENS INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PERILIMABAL ON (B)(6) 2023, DUE TO BLURRED VISION AND "OD HORIZONTAL JUXTAPOSED METAMORPHIC EYE." CAUSE OF THE EVENT IS UNKNOWN AND THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "MIOCHOL WAS ON BACK ORDER MIOSTATE USED OU FOR INITIAL SURGERY. PT HAS RETINAL EVAL PENDING ALTHOUGH MOCULAR OCT SHOWS NO ABNORMALITIES WHEN COMPARED TO PREOP MOC OCT. APPT (B)(6) 2023 - RETINA." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PERILIMABAL. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869317&pc=MTA | 2023826-2020-00627 | 2019-12-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Improper or Incorrect Procedure or Method | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; AN ACD VALUE OF 3.2MM WAS USED INSTEAD OF THE ACTUAL ACD VALUE OF 2.53MM IN ORDER TO CALCULATE THE LENS. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721909&pc=MTA | 2023826-2021-01028 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11773680&pc=MTA | 2023826-2021-01383 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -3.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH TORIC LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS WAS FOUND ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683469&pc=MTA | 2023826-2021-01068 | 2020-10-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM, CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190874&pc=MTA | 2023826-2022-01297 | 2020-10-22 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, BLURRED VISION, AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190879&pc=MTA | 2023826-2022-01296 | 2020-10-22 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009098&pc=MTA | 2023826-2021-01886 | 2021-03-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -03.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT THE LENS WAS REMOVED ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459227&pc=MTA | 2023826-2022-00384 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#(B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12921591&pc=MTA | 2023826-2021-04632 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075679&pc=MTA | 2023826-2022-04744 | 2022-01-26 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -8.0 DIOPTER, WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592023&pc=MTA | 2023826-2023-03623 | 2022-05-04 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, THE PROBLEM IS NOT RESOLVED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075680&pc=MTA | 2023826-2022-04745 | 2022-05-20 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS WITH A -8.0 DIOPTER, WAS IMPLANTED AS A REPLACEMENT INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: B3 - DATE OF EVENT: 20-MAY-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO "SULCUS TOO SMALL FOR 12.1." AND BLURRY VISION. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2022 SHOULD BE ADDED TO THE INTIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15810623&pc=MTA | 2023826-2022-03960 | 2022-09-18 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL, LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: (B)(4). OFF LABEL ACD<3.0 <45 YEARS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 20-DEC-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787278&pc=MTA | 2023826-2023-01468 | 2023-03-24 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER EMAIL REPLY, "PATIENT WAS SEEN (B)(6) 2023, EVO ICL NICELY CENTERED. DRYNESS FROM SURGERY HAS RESOLVED AND VISION CONTINUING TO IMPROVED." CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2023. REPORTEDLY, "EYE VERY DRY. STARTED TEARS." DRYNESS FROM SURGERY HAS RESOLVED AND VISION CONTINUIING TO IMPROVED. CLAIM# (B)(4). Manufacturer Narrative: H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569168&pc=MTA | 2023826-2023-03513 | 2023-06-22 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL USE ACD<3.0. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999784&pc=MTA | 2023826-2023-04729 | 2023-09-04 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17845437&pc=MTA | 2023826-2023-04238 | 2022-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15611640&pc=MTA | 2023826-2022-03644 | 2022-04-08 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -5.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "PATIENT FELT ABNORMAL NEAR VISION AFTER SURGERY AND STRONGLY REQUESTED TO REPLACE THE LENS". Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598566&pc=MTA | 2023826-2022-01744 | NA | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -1.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE CHANGE OVERTIME AND ASTIGMATISM POST-SURGERY WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT: UNK. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178423&pc=MTA | 2023826-2021-02627 | 2021-02-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Unintended Movement | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING BEEN IMPLANTED UPSIDE DOWN AND PATIENT COMPLAINED OF BLURRED VISON. VISUAL ACUITY WAS POOR. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865569&pc=MTA | 2023826-2023-01626 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM612.6 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND A TORIC LENS WAS IMPLANTED DUE TO REFRACTIVE SUPRISE, REFRACTIVE CHANGE OVER TIME, GLARES/HALOS, AND BLURRED VISION. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. REPORTEDLY, "EVERYTHING IS GOOD, NOTHING TO ADD." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#(B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135976&pc=MTA | 2023826-2021-00042 | 2020-11-05 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT MODEL LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187759&pc=MTA | 2023826-2023-05236 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. PATIENT IS HYPEROPIC +25 +.50 X086 = 20/15. UNCORRECTED VA IS 20/25. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. - PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. - IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523962&pc=MTA | 2023826-2021-03725 | 2021-06-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/501K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560960&pc=MTA | 2023826-2024-00365 | 2023-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA (BLURRY VISION WITH SHADOW) AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10956213&pc=MTA | 2023826-2020-02850 | 2020-10-22 | Injury | STAAR SURGICAL COMPANY | 2020-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VICMO 13.2, -17.50 DIOPTER,IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT REPORTED THAT THE PROBLEM RESOLVED, PATIENT IS OKAY. THE REPORTER INDICATED " THERE WAS AN IRIS PROLAPS DURING EXPLANTATION OF THE ICL. THE SURGEON DECIDED TO WAIT WITH NEW IMPLANTATION OF ICL AND THE PATIENT WEARS CL." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372191&pc=MTA | 2023826-2021-00442 | 2021-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, ELEVATED IOP (26MMHG). CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON(B)(6) 2021, "LENS EXCHANGE IS PLANNED". Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT IMPACT CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939017&pc=MTA | 2023826-2021-01668 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -12.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION, BLURRED VISION AND HYPHEMA, INFLAMMATION. THIS RESOLVED THE PROBLEM. REPORTEDLY, AC BLEEDING WAS REMOVED BY IRRIGATION. INFLAMMATION WAS TREATED BY MEDICATION. A NEW LENS IMPLANT IS PLANNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944355&pc=MTA | 2023826-2022-01026 | 2022-01-18 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICM5_13.2; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AS WELL AS HYPEROPIA. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE IT WAS NOTED A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT POWER WAS IMPLANTED. THIS DID NOT RESOLVE THE PROBLEM. SEE MFR# 2023826-2022-01027 FOR REPLACEMENT LENS. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICAITONS. CLAIM#:(B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17802053&pc=MTA | 2023826-2023-04080 | 2023-04-11 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED UNREACTIVE (FIXED) PUPIL, BLURRED VISION AND IRIS ATROPHY. POST-OP, THE PUPIL IS STILL LARGE. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK . H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL, IRIS ATROPHY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: POST-OP THE PUPIL WAS STILL LARGE AND POORLY CONSTRICTED. AT THE PATIENTS LAST VISIT ON (B)(6) 2023, THE DOCTOR INFORMED PATIENT TO REDUCE VIGAMOX + PREDFORTE ON TILL FINISHED. USE BACK PILOCARPINE TDS ON LE , OPTIVE FUSION WHEN NECESSARY. PATIENT WILL BE REVIEWED IN (B)(6) 2023. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14376862&pc=MTA | 2023826-2022-01317 | NA | Injury | STAAR SURGICAL COMPANY | 2022-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND DIZZINESS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS; HOWEVER, FUNCTIONAL INSPECTION FOUND THE LENS SPHERE TO BE OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15332386&pc=MTA | 2023826-2022-03005 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM6_12.1; -04.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022 AND (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). EXPIRATION DATE: 09-30-2023. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15332387&pc=MTA | 2023826-2022-03006 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022 AND (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442847&pc=MTA | 2023826-2022-03082 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6 OF -06.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2021 AND A REVISION WAS ALSO PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442849&pc=MTA | 2023826-2022-03081 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1 MM VICM6 OF -3.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480051&pc=MTA | 2023826-2022-03260 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480541&pc=MTA | 2023826-2022-03261 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587053&pc=MTA | 2023826-2021-03865 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED CQ RECEIVED ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH TORIC LENS (13.2MM VTICMO13.2, DIOPTER -15.5/+1.50/074) ON (B)(6) 2021. PROBLEM RESOLVED. H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16427707&pc=MTA | 2023826-2023-00686 | 2022-11-07 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16427709&pc=MTA | 2023826-2023-00687 | 2022-11-07 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15396076&pc=MTA | 2023826-2022-03059 | 2022-06-09 | Injury | STAAR SURGICAL COMPANY | 2022-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD DIFFICULTY LOOKING CLOSE AND EXPERIENCED BLURRED VISION. THE PATIENT STRONGLY REQUESTED TO REMOVE THE LENS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE: 2993 SHOULD BE CORRECTED TO 1401. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457491&pc=MTA | 2023826-2021-03554 | NA | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LOW VAULT, A DECREASE IN BCVA, AND A DECREASE IN THE NUMBER OF INDOTHELIAL CELLS. REPORTEDLY, THE SURGEON DOES NOT PLAN TO SEE THE PATIENT AGAIN FOR ONE YEAR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. MODEL#: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. HEALTH EFFECT CLINICAL CODE: (B)(4) (ENDOTHELIAL CELL LOSS). DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17194710&pc=MTA | 2023826-2023-02531 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM), KERATOCONUS. CLAIM# (B)(4). Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND LOSS OF VISUAL ACUITY. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT'S POST-OP REFRACTION ON (B)(6) 2023: PLANO -3.752 X 025 W/ VAULT OF 1070"". H6-HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED AND "4614" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND LOSS OF VISUAL ACUITY. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT'S POST-OP REFRACTION ON (B)(6) 2023: PLANO -3.752 X 025 W/ VAULT OF 1070". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. H6 - MEDICAL DEVICE PROBLEM CODE: "4614" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH THE SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS OTHER. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. D6B.- "(B)(6) 2023" SHOULD BE CORRECTED TO "(B)(6) 2023" . B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS OTHER. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359537&pc=MTA | 2023826-2021-03269 | 2021-01-24 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS BEING IMPLANTED UPSIDE DOWN. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THERE WAS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON DECIDED TO ADD NEW PIS DURING ICL IMPLANTATION. THE SURGEON PLANS TO EXPLANT/EXCHANGE THE LENS BUT CURRENTLY THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND RE-IMPLANTED CORRECTLY. THE PROBLEM WAS RESOLVED AND THE PATIENT IS PERFECT, VISION IS GOOD. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15513701&pc=MTA | 2023826-2022-03377 | 2022-06-09 | Injury | STAAR SURGICAL COMPANY | 2022-09-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAD BLURRED VISION AND NEAR VISION WAS DIFFICULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM AND IT WAS NOTED "THE PATIENT HAD DIFFICULTY LOOKING CLOSE AND STRONGLY REQUESTED TO TAKE IT OUT". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372192&pc=MTA | 2023826-2021-00443 | 2021-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "LENS EXCHANGE IS PLANNED." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, ELEVATED IOP (26MMHG) BUT OPEN ANGLE. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "LENS EXCHANGE IS PLANNED." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15055918&pc=MTA | 2023826-2022-02273 | NA | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO_13.2; -16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND BLURRED VISION WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT. B2:REQUIRED INTERVENTION. B5: ADDITIONAL INFORMATION: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO_13.2; -16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND BLURRED VISION WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. H6: HEALTH EFFECT CLINICAL CODE: 2137-BLURRED VISION. 4581- SIGNIFICANT REDUCTION IRIDOCORNEAL ANGLES. H6: HEALTH EFFECT IMPACT CODE: 4629-DEVICE REPLACEMENT. H6: MEDICAL DEVICE PROBLEM CODE: 1401-MISFOCUSING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591731&pc=MTA | 2023826-2022-03483 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WILL BE REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME, KERATOCONUS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE VISION WAS WORSENING AND CORNEA TOPOGRAPHY SHOWED KERATOCONUS DEVELOPING OU VERSUS 6 MONTHS PRIOR. THE LENS WAS REPLACED WITH A DIFFERENT MODEL (VICM5) BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE VA IS GOOD AND PATIENT IS DOING WELL, HAS SOME NIGHT TIME GLARE CAUSED BY LIGHTS. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: PATIENT SEX, WEIGHT, ETHNICITY AND RACE: UNK. DATE OF EVENT, IMPLANT DATE: UNK. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591791&pc=MTA | 2023826-2022-03484 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WILL BE REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME, KERATOCONUS. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE VISION WAS WORSENING AND CORNEA TOPOGRAPHY SHOWED KERATOCONUS DEVELOPING OU VERSUS 6 MONTHS PRIOR. THE LENS WAS REPLACED WITH A DIFFERENT MODEL (VICM5) BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE VA IS GOOD AND PATIENT IS DOING WELL, HAS SOME NIGHT TIME GLARE CAUSED BY LIGHTS. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8197958&pc=MTA | 2023826-2018-01925 | 2018-09-05 | Injury | STAAR SURGICAL COMPANY | 2018-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Capsular Bag Tear; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR/ DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 IMPLANTABLE COLLAMRE LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THIS COMPLAINT LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 M WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER INDICATED THAT THE INITIAL LENS WAS 12.6MM BUT PROVIDED THE SERIAL NUMBER OF A 13.2MM LENS. NO ADDITIONAL INFORMATION HAS BEEN FORTHCOMING. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190415&pc=MTA | 2023826-2022-01286 | 2022-03-14 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VIVA IMPLANTABLE COLLAMER LENS WAS REMOVED FROM THE PATIENTS LEFT EYE (OS) AND REPLACED WITH AN EVO IMPLANTABLE COLLAMER LENS ON (B)(6) 2022 DUE TO UNCOMFORTABLE VISION "NAUSEA IF PATIENT LOOKS VERY CLOSE (NEAR)." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION A 13.2MM, VICM6_13.2, -5.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON 14-MAR-2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (TORIC IMPLANTABLE LENS) AND POWER LENS DUE TO PERMANENT LOSS OF BCVA AND THE PROBLEM RESOLVED. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT-RELATED FACTOR, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965166&pc=MTA | 2023826-2020-00897 | 2019-04-02 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD DISCOMFORT DUE TO ASTIGMATISM AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL ICL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965173&pc=MTA | 2023826-2020-00898 | 2019-04-02 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD DISCOMFORT DUE TO ASTIGMATISM AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL ICL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696913&pc=MTA | 2023826-2018-01041 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT, AND BLURRED VISION. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE TORN HAPTIC PIECES ARE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696917&pc=MTA | 2023826-2018-01043 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT, AND BLURRED VISION. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE.(B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755755&pc=MTA | 2023826-2021-04210 | 2021-07-04 | Injury | STAAR SURGICAL COMPANY | 2021-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Eye Pain | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT HAS A VAGUE OF THE LEFT EYE BRINGING A VERY SERIOUS IMPACT ON THEIR LIVES TO MAKE IT VERY PAINFUL AND ANXIOUS, COMMUNICATED, AND DECIDED TO REMOVE CRYSTALS." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL INSPECTION DUE TO SIGNIFICANT DAMAGE OF THE LENS OPTIC PART. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555190&pc=MTA | 2023826-2015-00193 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO A REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT DIFFERENT DIOPTER POWER (-18.0) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NEW INCISION. REFRACTIVE SURPRISE. GLARE. LENS EXPLANTED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYES MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587339&pc=MTA | 2023826-2023-00950 | 2022-11-24 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT HAD SOME HALOES/BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED (HALOES/BLURRED VISION HAD RESOLVED). THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567853&pc=MTA | 2023826-2022-03574 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. PATIENT EXPERIENCED HALOS AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057498&pc=MTA | 2023826-2022-04706 | 2022-08-03 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -12.0 WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078375&pc=MTA | 2023826-2022-04754 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT REPORTED HAS DEVELOPED POSITIVE DYSPHOTOPSIA AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT EXPERIENCED "POSITIVE DYSPHOTOPSIA" (GLARES/ HALOES). CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - POSITIVE DYSPHOTOPSIA. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078608&pc=MTA | 2023826-2022-04753 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT REPORTED HAS DEVELOPED POSITIVE DYSPHOTOPSIA AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT EXPERIENCED "POSITIVE DYSPHOTOPSIA" (GLARES/ HALOES). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030859&pc=MTA | 2023826-2022-04696 | 2022-07-27 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -1.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. STATUS OF THE EYE IS "QUIET.". Manufacturer Narrative: DEVICE EVALUATED BY MFR: PATIENT RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.75 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES; AND GLARES/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "QUIET". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR; AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030862&pc=MTA | 2023826-2022-04697 | 2022-07-27 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -2.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. STATUS OF THE EYE IS "QUIET.". Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES; AND GLARES/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "QUIET". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR; AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012303&pc=MTA | 2023826-2015-01021 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012304&pc=MTA | 2023826-2015-01022 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154940&pc=MTA | 2023826-2015-01345 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154947&pc=MTA | 2023826-2015-01346 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589778&pc=MTA | 2023826-2017-00880 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS EXPLANT RESOLVED THE PROBLEM. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. CORRECTED DATA: CONCLUSION CODE: AS PER DFU, VICMOS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. DEVICE CODE: (B)(4) PREVIOUS CODE IS NOT APPLICABLE, BLURRED VISION AND GLARE/HALOS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589780&pc=MTA | 2023826-2017-00879 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS EXPLANT RESOLVED THE PROBLEM. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. CORRECTED DATA: DEVICE CODE: (B)(4) PREVIOUS CODE IS NOT APPLICABLE, BLURRED VISION AND GLARE/HALOES WERE REPORTED. CONCLUSION CODE: AS PER DFU, VICMOS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041103&pc=MTA | 2023826-2016-01094 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY 20/25. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: (B)(4) SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041106&pc=MTA | 2023826-2016-01093 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. FOREIGN MATERIAL (DRIED, DISCOLOURED MATERIAL) WAS FOUND ON THE LENS SURFACE. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5797034&pc=MTA | 2023826-2016-00924 | 2016-03-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC ICL AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614664&pc=MTA | 2023826-2015-00346 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); BLURRED VISION; HALO; (GLARE); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614665&pc=MTA | 2023826-2015-00347 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664100&pc=MTA | 2023826-2016-00661 | 2016-04-07 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED AND IT IS PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THERE WAS FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE DIOPTER MEASUREMENT IS WITHIN SPECIFICATION. CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT'S AGE OVER 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216360&pc=MTA | 2023826-2023-02162 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -06.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON EXCHANGED THE LENS FOR A SAME LENGTH BUT DIFFERENT POWER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H10 - DISREGARD SUPPLEMENTAL #1 MDR AS IT WAS SUBMITTED IN ERROR WITH THE WRONG CLAIM DETAIL CORRECTIONS. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627897&pc=MTA | 2023826-2022-01722 | 2022-02-15 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL/ CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+5.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED RESIDUAL ASTIGMATISM. REPORTEDLY "THE SURGEON IS PLANNING TO EXCHANGE THE LENS WITH SPHERICAL ICL AND PERFORM THE LRI FOR ASTIGMATISM". IT WAS LATER REPORTED "AS I CHECKED WITH THE SURGEON, THE PATIENT STILL HAS THE RESIDUAL ASTIGMATISM (-1.0 CYL). HOWEVER, THE PATIENT DIDN'T FEEL UNCOMFORTABLE AND HASN'T COME BACK TO THE CLINIC FOR FOLLOW-UP. THE SURGEON SAID THAT IT DOESN'T SEEM TO HAVE FURTHER TREATMENT." THE LENS REMAINS IMPLANTED. CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778983&pc=MTA | 2023826-2021-01321 | 2021-02-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -1.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PRESCRIBED DISTANCE GLASSES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. D6A: IMPLANT DATE CORRECTED TO (B)(6) 2021. H6-HEALTH IMPACT CODE: 4642-DOCTOR PRESCRIBED DISTANT GLASSES. INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. SERIAL # - UNK. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778984&pc=MTA | 2023826-2021-01320 | 2021-02-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -1.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PRESCRIBED DISTANCE GLASSES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. H6-HEALTH IMPACT CODE: 4642-DOCTOR PRESCRIBED DISTANT GLASSES. INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. D6A: IMPLANT DATE CORRECTED TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER PLANNED EXCHANGE IN THE FUTURE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15230266&pc=MTA | 2023826-2022-02685 | 2021-10-14 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -3.00 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODE SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15554010&pc=MTA | 2023826-2022-03517 | 2021-10-25 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS (B)(6), PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." Manufacturer Narrative: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136767&pc=MTA | 2023826-2022-02525 | 2022-03-01 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME/ MODEL AND SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136769&pc=MTA | 2023826-2022-02526 | 2022-05-30 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS OF THE SAME SIZE DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352414&pc=MTA | 2023826-2022-03015 | 2022-06-17 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH A DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4: DATE OF THIS REPORT-04-AUG-2022 FOR SUPPLEMENTAL #1. G3: DATE RECEIVED BY MANUFACTURER: 19-OCT-2022 FOR SUPPLEMENTAL #1. CLAIM# (B)(4). . Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 31-JAN-2025. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15674658&pc=MTA | 2023826-2022-03576 | 2022-09-05 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -17.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE AND DIFFERENT DIOPTER WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. " THE DEFECT WAS MAINLY DUE TO THE WRONG REFRACTION SUPPLIED FOR CALCULATION." Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16603910&pc=MTA | 2023826-2023-00955 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | UNK | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM6_12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON(B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTICS BENT, WITH BLOOD STAINS ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTION (FUNCTIONAL) INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669961&pc=MTA | 2023826-2020-02332 | 2020-09-03 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION : LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS DUE TO PATIENT REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. HEALTH EFFECT CODE IN INITAL MDR 4582 CORRECTED TO 2137. MEDICAL DEVICE PROBLEM CODE IN INITIAL MDR 1401 CORRECTED TO 2993. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048073&pc=MTA | 2023826-2023-04718 | 2023-10-11 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE WRONG POST-OPERATIVE CORRECTION FOR CALCUATION WAS USE BY SURGEON BY ERROR. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. REPORTEDLY "HE USED THE WRONG PARAMETER FOR CALCULATION." ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS: "SURGEON USED THE WRONG POSTOPERATIVE CORRECTION FOR CALCULATIONS". D6B. - (B)(6) 2023 SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" AMD "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983359&pc=MTA | 2023826-2020-00933 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT IS HAPPY. H6- HEALTH EFFECT IMPACT CODE 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. UNABLE TO ENTER CODE IN H6 FIELD. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# : (B)(4). Manufacturer Narrative: CORRRECTED DATA: H6: CLINICAL CODE 4582 IS NOT APPLICABLE IN PREVIOUS MDR. H6: RESULT CODE 114 IS NOT APPLICABLE IN PREVIOUS MDR. H6: CONCLUSION CODE 4310 SHOULD BE IN SUPPLEMENTAL MDR #1. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983362&pc=MTA | 2023826-2020-00932 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT IS HAPPY. H6- HEALTH EFFECT IMPACT CODE 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. UNABLE TO ENTER CODE IN H6 FIELD. CLAIM# (B)(4). Manufacturer Narrative: CORRRECTED DATA: H6-CLINICAL CODE 4582 IS NOT APPLICABLE IN PREVIOUS MDR. H6-RESULT CODE 114 IS NOT APPLICABLE IN PREVIOUS MDR. H6- CONCLUSION CODE 4310 SHOULD BE IN SUPPLEMENTAL MDR #1. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686315&pc=MTA | 2023826-2022-03836 | 2022-09-23 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A ICM120V4 IMPLANTABLE COLLAMER LENS OF A -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2022, LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686372&pc=MTA | 2023826-2022-03837 | 2022-09-23 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A ICM120V4 IMPLANTABLE COLLAMER LENS OF A -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. ON(B)(6)2022 , LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT/ ETHINICTY/ RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14059294&pc=MTA | 2023826-2022-01159 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298674&pc=MTA | 2023826-2022-02908 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT AND REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: 11/30/2015. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE CHANGE OVER TIME" SHOULD BE REMOVED FROM THE ORIGINAL SUBMISSION FOR CORRECTION. H10: "ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHININ THE SAME LOT" SHOULD BE REMOVED AND CORRECTED TO SAY "NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT". Manufacturer Narrative: CORRECTED DATA: H6 HEALTH EFFECT - CLINICAL CODE-2137 SHOULD BE OMITTED AND CODE 4582 SHOULD BE ADDED FOR CORRECTION. DEVICE CODE-1401 SHOULD BE OMITTED FOR CORRECTION. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298880&pc=MTA | 2023826-2022-02909 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT AND REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 HEALTH EFFECT - CLINICAL CODE-2137 SHOULD BE OMITTED AND CODE 4582 SHOULD BE ADDED FOR CORRECTION. DEVICE CODE: 1401 SHOULD BE OMITTED FOR CORRECTION. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE CHANGE OVER TIME" SHOULD BE REMOVED FROM THE ORIGINAL SUBMISSION FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 09/28/2012. THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691191&pc=MTA | 2023826-2023-01211 | 2023-01-14 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO_12.6 IMPLANTABLE COLLAMER LENS -15.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2032 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2023-01210 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: PATIENT IDENTIFIERS: (B)(6). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION: LENS DISLOCATION/ SUBLUXATION AND REFRACTIVE SURPRISE. REPORTEDLY "INTRAOCULAR LENS DISPLACEMENT, SEE THE EDGE OF INTRAOCULAR LENS". ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, TWO SIZES LONGER LENGTH, 1.0D STRONGER POWER AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE PATIENT IS IN GOOD CONDITION". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. H6- TYPE OF INVESTIGATION CODE: 4110 - LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"2993, 1401, 2923" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12806826&pc=MTA | 2023826-2021-04330 | 2021-09-28 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15821272&pc=MTA | 2023826-2022-03964 | 2022-08-08 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS WITH A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "DUE TO THE PATIENT'S DIFFICULTY IN VISUAL PROXIMITY, THE ICL DEGREE WAS ADJUSTED TO -8.5." Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119844&pc=MTA | 2023826-2021-00022 | 2020-12-05 | Injury | STAAR SURGICAL COMPANY AG | 2021-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2008. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINED IMPLANTED. THE PLAN IS TO EXPLANT AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING, REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS SATISFIED WITH VISUAL OUTCOME. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR (PROGRESSIVE MYOPIA). H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CORRECTED DATA: A2: DATE OF BIRTH (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119845&pc=MTA | 2023826-2021-00021 | 2020-12-05 | Injury | STAAR SURGICAL COMPANY AG | 2021-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2008. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINED IMPLANTED. THE PLAN IS TO EXPLANT AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING, REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS SATISFIED WITH VISUAL OUTCOME. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR (PROGRESSIVE MYOPIA). H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CORRECTED DATA: A2: DATE OF BIRTH - (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15673922&pc=MTA | 2023826-2022-03635 | 2022-03-18 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. OFF LABEL USE: ACD <3.0. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE." DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH "RESIDUE"/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796051&pc=MTA | 2023826-2017-01254 | 2017-04-21 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25 AND THE PATIENT IS HAPPY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928358&pc=MTA | 2023826-2018-01437 | 2018-07-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT REPORTED HER NEAR VISION WAS POOR. THE PATIENT COULDN'T SEE CLEARLY FROM CLOSE (40CM), SO SHE REQUESTED TO REMOVE THE ICL LENS AND DIDN'T WANT ANOTHER LENS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254835&pc=MTA | 2023826-2019-00097 | NA | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND DIOPTER LENS DUE REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949670&pc=MTA | 2023826-2015-00894 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT RECOVERED WELL AND IS SATISFIED WITH THE RESULT. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. CLAIM # (B)(4), LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244431&pc=MTA | 2023826-2015-01536 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL WITH DIFFERENT DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO DAMAGE; LENS MEASUREMENT WAS WITHIN SPECIFICATION. RESULT (LENS MEASURED WITHIN SPECIFICATION). CONCLUSION (LENS MEASURED WITHIN SPECIFICATION). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.00. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7161172&pc=MTA | 2023826-2018-00004 | 2017-11-23 | Injury | STAAR SURGICAL COMPANY | 2018-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND WHITE MATERIAL ON THE LENS. ON (B)(6) 2017, THE PATIENT'S BCVA WAS 20/13. Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: THE DEVICE WAS RETURNED TO THE MANUFACTURER OF 15-DEC-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH SMALL AMOUNTS OF LIQUID PRESENT, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND SMALL PARTICULATES ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6465367&pc=MTA | 2023826-2017-00586 | 2015-08-13 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2015, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN. ONE HAPTIC WAS TORN AND BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205788&pc=MTA | 2023826-2018-00104 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS PLEASED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197059&pc=MTA | 2023826-2016-01802 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LENS THE SAME SIZE, MODEL, BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993869&pc=MTA | 2023826-2017-01661 | 2016-11-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: REPORTEDLY, THE SURGEON IMPLANTED A 13.2MM VICMO13.2 DIOPTER -15.50 DIOPTER INTO THE PATIENT'S EYE ON (B)(6) 2016. THE SURGEON NOTES A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, PLANS TO BE EXCHANGED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7200518&pc=MTA | 2023826-2018-00100 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -01.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON NOTED A REFRACTIVE SURPRISE AND PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -1.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED DUE TO REFRACTIVE SURPRISE ON (B)(6) 2018. THE PROBLEM WAS NOT RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801614&pc=MTA | 2023826-2017-01300 | 2017-01-18 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6)2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, SO THE LENS WAS EXPLANTED ON (B)(6)2017 AND WAS REPLACED WITH THE SAME MODEL LENS BUT WITH A -8.5 DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION- DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918306&pc=MTA | 2023826-2017-01535 | 2017-01-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 DIOPTER -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE SURGEON EXPLANTED THE LENS DUE TO REFRACTIVE SURPRISE. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ON (B)(6) 2017, THE CLAIMANT DISCLOSED THE SERIAL NUMBER FOR THE ASSOCIATED LENS. ON THIS DATE, IT WAS DETERMINED THAT THIS CASE IS A DUPLICATE OF 2023826-2017-01630 (CLAIM# (B)(4)). PLEASE DISREGARD THE INFORMATION IN THE CURRENT CLAIM AND INSTEAD REFERENCE 2023826-2017-01630. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE INFORMATION: UNK, AS SERIAL NUMBER IS UNK. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS EXPLANT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398191&pc=MTA | 2023826-2018-01499 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED. THE PROBLEM WAS RESOLVED AND "THE PATIENT IS REALLY HAPPY". (B)(4) Manufacturer Narrative: DEVICE EVALUATIONS - LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND TWO FEET WERE BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS FOR A 13.7MM LENS. FUNCTIONAL INSPECTION FOUND THE POWER OF THE RETURNED LENS TO BE -8.6 D INSTEAD OF THE LABELLED POWER OF -120 D. WORK ORDER SEARCH - A WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATIONS - THE CORRECT LABELLED POWER SHOULD BE -12.0 D NOT -120 D. (B)(4). Manufacturer Narrative: CORRECTIVE DATA: EXPLANT DATE: CHANGED FROM UNKNOWN TO(B)(6) 2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT IS IS AMBLYOPIC IN THE FELLOW EYE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6648555&pc=MTA | 2023826-2017-00993 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT DIOPTER LENS, DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: STATED LENS HAD NOT YET BEEN EVALUATED. UPDATED: LENS HAS BEEN EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (SAME PATIENT, FELLOW EYE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6651664&pc=MTA | 2023826-2017-00994 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT DIOPTER LENS, DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (SAME PATIENT, FELLOW EYE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6388842&pc=MTA | 2023826-2017-00317 | 2017-02-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THAT THE FIRST LENS WAS ORDERED BY ERROR 13.2 -2.25 INSTEAD OF 13.2 -18.0. ON (B)(6) 2017, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, THE LENS WAS EXCHANGED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6402737&pc=MTA | 2023826-2017-00376 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, WAS EXPLANTED AND REPLACED WITH A LESSER POWER ICL DUE TO HYPEROPIC RESULT . THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE UCVA AFTER ICL EXCHANGE WAS 20/20-2. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6717761&pc=MTA | 2023826-2017-01197 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR SIZE LENS OF A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/13 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE HAPTIC SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721360&pc=MTA | 2023826-2017-01191 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND THE LENS WAS EXCHANGED FOR A SIMILAR SIZE LENS OF A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017 THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE LENS WAS EXCHANGED FOR A SAME SIZE LENS OF A DIFFERENT DIOPTER. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND THE LENS WAS RETURNED WITH SOME MOISTURE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: PATIENT AGE >(B)(6). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6808731&pc=MTA | 2023826-2017-01352 | 2017-04-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT, REFRACTIVE SURPRISE, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRY VISION WAS NOTICED. THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN THE OPERATIVE EYE WITH ANY CATARACT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595444&pc=MTA | 2023826-2017-00893 | 2017-04-26 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -6.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CONCLUSION: NO FAILURE DETECTED; DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670093&pc=MTA | 2023826-2017-01029 | 2017-05-05 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIFFERENT POWER, THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670109&pc=MTA | 2023826-2017-01030 | 2017-05-05 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIFFERENT POWER LENS, THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183462&pc=MTA | 2023826-2018-00048 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON 26 APR 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS THEN EXCHANGED FOR A DIFFERENT LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF CLEAR AND WHITE RESIDUE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE PERFORMED. CORRECTED DATA: DATE OF THIS REPORT: IT IS CORRECTED FROM "18-AUG-2017" TO " 15-DEC-2017". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A NEW LENS WAS IMPLANTED LATER AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213916&pc=MTA | 2023826-2018-00123 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME AND THE PATIENT WANTED THE LENS REMOVED DUE TO DECREASED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WAS NO PATIENT INJURY WHEN LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND SCRATCHES ON THE OPTIC. THERE WAS PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213918&pc=MTA | 2023826-2018-00122 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME AND THE PATIENT WANTED THE LENS REMOVED DUE TO DECREASED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WAS NO PATIENT INJURY WHEN LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653553&pc=MTA | 2023826-2017-01005 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS FOUND RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725514&pc=MTA | 2023826-2017-01204 | 2017-06-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZED, SIMILAR DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUS: THE LENS WAS EXCHANGED FOR SAME SIZE, SIMILAR DIOPTER. _CORRECTED: THE LENS WAS EXCHANGED FOR A SAME SIZE , LOWER DIOPTER. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED FOUND THAT THE PRODUCT IS WITHIN THE VERIFICATION. (B)(4). Manufacturer Narrative: PREVIOUSLY: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED FOUND THAT THE PRODUCT IS WITHIN THE VERIFICATION. _CORRECTED: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED AND FOUND THAT THE PRODUCT IS WITHIN THE SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993839&pc=MTA | 2023826-2017-01659 | 2017-06-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 DIOPTER -11.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS ON (B)(6) 2017, THE SURGEON REPORTING A REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725922&pc=MTA | 2023826-2017-01216 | 2017-06-22 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -7.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER AND TYPE LENS DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN LNES CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874037&pc=MTA | 2023826-2017-01445 | 2017-06-30 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND FAR SIGHT TENDENCY. THE LENS WAS EXCHANGED FOR ANOTHER LENS, SAME MODEL BUT DIFFERENT DIOPTER, AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S HYPEROPIA IMPROVED, BUT NOT COMPLETELY. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN, AND FOREIGN MATERIALS ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF 45 YEARS OLD AND IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908143&pc=MTA | 2023826-2017-01500 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 16 NOV 2017 AND EXCHANGED WITH A SAME MODEL BUT LOWER DIOPTER LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND THAT LENS HAD A VERY YELLOW VISUAL ASPECT. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS MEASURED TO BE -9.31D AT THE TIME OF RELEASE. WORK ORDER SEARCH FOUND NO OTHER SIMILAR DEFECT WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. CLAIM# (B)(4). Manufacturer Narrative: IMPLANTED MAY 2017. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021358&pc=MTA | 2023826-2017-01670 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOSS OF BCVA. BCVA WENT FROM 20/25 TO 20/35. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BCVA ON (B)(6) 2017 WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146627&pc=MTA | 2023826-2017-01944 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS +3.0D FROM THE TARGET. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE PERFORMED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183428&pc=MTA | 2023826-2018-00050 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND IT RESOLVED THE PROBLEM. THERE IS NO PLANS IMPLANT ANOTHER LENS AS THE PATIENT WAS DISSATISFIED WITH THE BLURRED VISION. ON (B)(6) 2017, THE PATIENT'S BCVA IS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE ITS HAPTIC AND PRESENCE OF CLEAR SURGICAL RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006958&pc=MTA | 2023826-2017-01687 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 A REFRACTIVE SURPRISE WAS REPORTED BY THE SURGEON. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON. THE LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7140439&pc=MTA | 2023826-2017-01950 | 2017-10-16 | Malfunction | STAAR SURGICAL COMPANY AG | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2005. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME. IT WAS STATED THAT THE LENS REMAINS IMPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS TO RESOLVE THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7140505&pc=MTA | 2023826-2017-01951 | 2017-10-16 | Malfunction | STAAR SURGICAL COMPANY AG | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2005. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME. IT WAS STATED THAT THE LENS REMAINS IMPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS TO RESOLVE THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213955&pc=MTA | 2023826-2018-00120 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213961&pc=MTA | 2023826-2018-00121 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISUAL DAMAGE AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE . WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343058&pc=MTA | 2023826-2018-00427 | 2017-12-27 | Injury | STAAR SURGICAL COMPANY | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, -2.75 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. IT WAS REPORTED THAT THIS EYE HAD A LVC (LASER VISION CORRECTION) 15 YEARS AGO. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR 12.6 VTICM5_12.6, -2.00/+1.00/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. DOCTOR AND PATIENT ARE HAPPY AND SATISFIED WITH THE RESULTS. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660091&pc=MTA | 2023826-2018-00997 | 2018-03-21 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499550&pc=MTA | 2023826-2018-00709 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT LOWER POWER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499555&pc=MTA | 2023826-2018-00710 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -10.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571599&pc=MTA | 2023826-2018-00845 | 2018-05-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER LENS, THE SURGEON NOTING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6755542&pc=MTA | 2023826-2017-01248 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE AND CLOSE TO LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS EXCHANGED FOR A VICM5_13.2, -6.50 DIOPTER, AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE HAPTICS WERE SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711472&pc=MTA | 2023826-2017-01181 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -15.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON ASSUMED THAT THE LENS IS UPSIDE DOWN, AND PLANNED TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2017 LENS WAS EXCHANGED FOR SIMILAR SIZE AND DIOPTER LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244802&pc=MTA | 2023826-2015-01534 | 2015-01-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. REFRACTIVE SURPRISE WAS NOTED IN (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL BUT DIFFERENT DIOPTER SIZE. THE RESOLVED THE PROBLEM. , POST-OP, UCVA WAS 20/20 SEE MFR. #2023826-2015-01535 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900669&pc=MTA | 2023826-2016-01136 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694868&pc=MTA | 2023826-2016-00800 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION CONCLUDES THE LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: PATIENT CODE - (B)(4) (EXPLANTED) - SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. PATIENT CODE - (B)(4) (NEW INCISION) - CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. DEVICE CODE - (B)(4) (EXPLANTED) - SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613052&pc=MTA | 2023826-2016-00537 | 2016-05-17 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTED A REFRACTIVE SURPRISE AND DECIDED TO EXCHANGE THE LENS WITH ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/20. THE DEVICE WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. (B)(4). INITIAL REPORT IS INCORRECT; REPORT SOURCE WAS A DISTRIBUTOR, NOT A USER FACILITY (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. (B)(4) REFRACTIVE SURPRISE. LENS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074347&pc=MTA | 2023826-2016-00910 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER ON FOLLOW-UP REPORT #2 IS INCORRECT; CORRECT DATE IS (B)(6) 2017. LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639951&pc=MTA | 2023826-2018-00967 | 2018-01-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON REPORTS MILD REFRACTIVE ERROR (MYOPIA, ASTIGMATISM) AND WHITE FOREIGN MATTER ADHERED TO THE SURFACE OF THE LENS, WHICH HAS INCREASED FROM THE FIRST EXAM TO THE FOLLOW-UP EXAM. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS UNDERGOING STEROID TREATMENT AND IS UNDER FOLLOW-UP. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. CORRECTED DATA: DATE OF EVENT: IT IS CORRECTED TO (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: EXACT DATE OF EVENT IS UNKNOWN; (B)(6) 2018 IS REPORTED. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639952&pc=MTA | 2023826-2018-00966 | 2018-01-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON REPORTS MILD REFRACTIVE ERROR (MYOPIA, ASTIGMATISM) AND WHITE FOREIGN MATTER ADHERED TO THE SURFACE OF THE LENS, WHICH HAS INCREASED FROM THE FIRST EXAM TO THE FOLLOW-UP EXAM. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS UNDERGOING STEROID TREATMENT AND IS UNDER FOLLOW-UP. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: EXACT DATE OF EVENT IS UNKNOWN; (B)(6) 2018 IS REPORTED. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213549&pc=MTA | 2023826-2016-01833 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | PHAKIC INTRAOCULAR LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: CONCLUSION: PER DFU FOR THIS LENS MODEL, VICMOS ARE CONTRAINDICATED FOR PATIENTS ABOVE THE AGE OF 45 YEARS OLD. THIS PATIENT WAS ABOVE 45 AT THE AGE OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, PRE-EXISTING DIAGNOSIS OF KERATOCONUS IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. DATE RECEIVED BY MANUFACTURER ON INITIAL REPORT IS 11/30/2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6210714&pc=MTA | 2023826-2016-01830 | 2016-09-13 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT.Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE DATE FOR THE INITIAL MDR IS (B)(6) 2016. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.50/+3.0/16 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED." (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE: PREVIOUS DATE INCORRECT. (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227597&pc=MTA | 2023826-2017-00015 | 2016-10-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE . AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR PRODUCT EVALUATION. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXCHANGED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DATE RECEIVED BY MANUFACTURER FOR INITIAL REPORT WAS 12/09/2016. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED, BUT NOT SOLD IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179489&pc=MTA | 2023826-2016-01765 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AT THE TIME OF REPORTING IT IS UNKNOWN IF THE LENS WILL BE EXCHANGED. Manufacturer Narrative: LENS DIOPTER DESCRIPTION IS -9.50/+2.0/90 (SPHERE/CYLINDER/AXIS). DATE RECEIVED BY THE MANUFACTURER FOR THE INITIAL MDR FOR THIS CLAIM IS 11/16/2016. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) TO OTHER SERIOUS (IMPORTANT MEDICAL EVENTS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT AVAILABLE TO RETURN. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978337&pc=MTA | 2023826-2017-01630 | 2017-01-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE/REFRACTIVE CHANGE OVERTIME. THE POSTOPERATIVE REFRACTION WAS NOT PREDICTED (HYPEROPIA-MANIFEST REFRACTION, CYCLOPLEGIC REFRACTION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE CORRECTED DATE IS 07-SEPT-2017. THE COMPLAINT WAS FIRST RECEIVED ON CLAIM # (B)(4), MFR REPORT #: 2023826-2017-01535. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438806&pc=MTA | 2023826-2018-00564 | 2010-11-24 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED-ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM IS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438807&pc=MTA | 2023826-2018-00566 | 2010-11-24 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. CORRECTED INFORMATION-ON 10-JUL-2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED.DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS.(B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7599401&pc=MTA | 2023826-2018-00908 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LOW VAULT,REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED AND THE SURGEON REPORTED THAT HE PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. THIS EXCHANGE RESOLVED THE PROBLEM AND "PATIENT IS HAPPY". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOR DIOPTER FOUND LENS WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7601418&pc=MTA | 2023826-2018-00907 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LOW VAULT,REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED AND THE SURGEON REPORTED THAT HE PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. 'PATIENT IS HAPPY', PROBLEM RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6199353&pc=MTA | 2023826-2016-01798 | 2016-10-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT REPORTED A LOW VAULT AND REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON LENS SURFACE. (B)(4). - WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373632&pc=MTA | 2023826-2017-00315 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (CORRECTED DATA): LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017. (PREVIOUS DATA): STATED LENS WAS EXCHANGED FOR A SHORTER LENS. (B)(4). (CORRECTED DATA): LENS EXCHANGED FOR A LONGER LENS. (PREVIOUS DATA)-LENS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SCRATCH ON THE LENS OPTIC. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) AND THE LENS HAVING A SCRATCH ON OPTIC. (B)(4). Manufacturer Narrative: DEVICE EXPLANTED IN 2017. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6388872&pc=MTA | 2023826-2017-00330 | 2017-02-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE), AND THE LENS MISSING A PIECE OF THE HAPTIC. INCORRECT DIOPTER SIZE -17.0, SHOULD BE -15.5. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812399&pc=MTA | 2023826-2017-01361 | 2017-04-05 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND CLEAR AND RED RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE (PRE-EXISTING CONDITION OF CATARACT IN THE OPERATIVE EYE OR NON-TRAUMATIC CATARACT IN THE FELLOW EYE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565214&pc=MTA | 2023826-2017-00790 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY NEAR VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SMALL CONTAINER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801617&pc=MTA | 2023826-2017-01279 | 2017-06-11 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: BCVA WAS 20/25 AT LAST EXAM (B)(6) 2017). IT IS UNKNOWN WHY BCVA DECREASED COMPARED TO BASELINE (20/20). (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908164&pc=MTA | 2023826-2017-01489 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT/DEFORMED AND THE HAPTIC BENT/DEFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890299&pc=MTA | 2023826-2017-01470 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -3.75 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON IT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS FOUND TO BE 20/24. THIS INFORMATION SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050172&pc=MTA | 2023826-2017-01749 | 2017-10-23 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS A REFRACTIVE SURPRISE. THE DOCTOR PLANS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AND DOCTOR NOTED THAT EVERYTHING IS FINE NOW. (B)(4). REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152456&pc=MTA | 2023826-2017-01961 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE SURGEON ALSO NOTED THE FOLLOWING: LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS NOTED AS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152530&pc=MTA | 2023826-2017-01960 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE SURGEON ALSO NOTED THE FOLLOWING: LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS NOTED AS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213972&pc=MTA | 2023826-2018-00139 | 2017-12-05 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, DIOPTER -13.0 INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO BLURRED VISION AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL, AND IN AN ENVELOPE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, MISSING A PIECE OF HAPTIC, AND WHITE RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7308043&pc=MTA | 2023826-2018-00308 | 2018-01-15 | Injury | STAAR SURGICAL COMPANY | 2018-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP STATUS OF THE EYE APPEARS NORMAL, GOOD OUTCOME AND PATIENT IS PLEASED. (B)(4). Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LEN- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597242&pc=MTA | 2023826-2018-00900 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PROGRESSIVE MYOPIA WAS REPORTED AS A PRE-EXISTING CONDITIONS. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULTING, REFRACTIVE SURPRISE, AND BLURRED VISION. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE MISSING) AND DEBRIS ON THE LENS SURFACE. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597627&pc=MTA | 2023826-2018-00899 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PROGRESSIVE MYOPIA WAS REPORTED AS A PRE-EXISTING CONDITIONS. Manufacturer Narrative: REPORT SOURCE SHOULD NOT BE "UNKNOWN", REPORT SOURCE IS IN THE UNITED STATES. REPORT SOURCE: HEALTH PROFESSIONAL AND USER FACILITY. SHOULD HAVE BEEN RECORDED IN INITIAL MDR. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002182&pc=MTA | 2023826-2018-01540 | 2018-10-04 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS "EXCHANGED" FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988487&pc=MTA | 2023826-2020-00984 | 2020-03-20 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 20 MAR 2013. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A VICMO13.2, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODE/LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. D7-EXPLANT DATE: (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094948&pc=MTA | 2023826-2016-01635 | NA | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON EXPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, INTERMITTENT ANGLE CLOSURE, BLURRY VISION AND DILATED EYE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PATIENT IS DOING WELL. THE REPORTER STATED THE LENS WAS IMPLANTED BY ANOTHER SURGEON IN (B)(6). THE CUSTOMER IN (B)(6) STATED THE LENS WAS IMPLANTED ON (B)(6) 2016. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244089&pc=MTA | 2023826-2017-00054 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO MYOPIC REFRACTION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.1MM) BUT DIFFERENT DIOPTER POWER (-8.0). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL WITH LIQUID. A VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SMALL TEAR AND A SCRATCH ON THE HAPTIC. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244134&pc=MTA | 2023826-2017-00055 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO MYOPIC REFRACTION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.1MM) BUT DIFFERENT DIOPTER POWER (-7.5). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL WITH LIQUID. A VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR ON THE HAPTIC AND A SMALL PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993874&pc=MTA | 2023826-2017-01660 | NA | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. DATE IS UNKNOWN. LENS WAS IMPLANTED IN (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542595&pc=MTA | 2023826-2015-00175 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME (OVERCORRECTION) AND BLURRED VISION. THE ICL WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT A DIFFERENT DIOPTER POWER (-15.0), WHICH RESOLVED THE PROBLEM. THE PATIENT IS HEALING WELL OU. Manufacturer Narrative: (B)(4). NEW INCISION. REFRACTIVE CHANGE OVERTIME/OVERCORRECTION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED BUT WITH THE RE-MEASUREMENT OF THE LENS, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161953&pc=MTA | 2023826-2015-01270 | 2015-04-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LENS ROTATION WITH LOW VAULT WAS OBSERVED ON (B)(6) 2015. PATIENT HAD BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370894&pc=MTA | 2023826-2016-00057 | 2015-12-08 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO MECHANICAL COMPLICATION OF ICL INITIAL ENCOUNTER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED MILD BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND STABLE. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. RETURNED LENS WAS FOUND WITHIN DIMENSIONAL SPECIFICATIONS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE EXCESSIVE VAULT OF THE IMPLANTED LENS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH MICL (12.6 MM) FOUR DAYS AFTER IMPLANTATION TO ADDRESS EXCESSIVE VAULTING AND SUBSEQUENT NARROWING OF THE ANGLE WITHOUT IOP ELEVATION. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2016, THE CAUSE OF THE EVENT WAS OVERSIZED LENS (NOT PRODUCT RELATED) AND VAULTING WAS ASSESSED ON THE FIRST POSTOP DAY. THE SAME REPORT STATED THAT PATIENT WAS STABLE FOLLOWING ICL EXCHANGE AND BCVA WAS 20/20. GIVEN THE INFORMATION THAT AFTER IMPLANTATION OF A SHORTER LENGTH ICL THE VAULTING RESOLVED, IT IS VERY LIKELY THAT PREOPERATIVE BIOCALCULATION ERROR HAS CAUSED THE SELECTION OF WRONG ICL LENGTH (OVERSIZED). CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, IT IS VERY LIKELY THAT PREOPERATIVE BIOCALCULATION ERROR HAS CAUSED THE SELECTION OF WRONG ICL LENGTH (OVERSIZED). (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923619&pc=MTA | 2023826-2016-01221 | 2016-05-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 1999. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.2MM) AND THE PROBLEM WAS NOT RESOLVED. THE VAULT IS LOW AND THE PLAN IS TO EXCHANGE FOR A 13.7MM. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 1999. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.2) OF A DIFFERENT MODEL. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. NOT APPLICABLE. THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. IMPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. MFR DATE: UNK. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826240&pc=MTA | 2023826-2016-00255 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, LENS DISLOCATION/SUBLUXATION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.6. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BENT AND FIBERS ON LENS SURFACE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594279&pc=MTA | 2023826-2018-00891 | 2013-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -2.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS LOW VAULT, REFRACTIVE CHANGE OVER TIME AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING, REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172583&pc=MTA | 2023826-2016-01747 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 20106 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT, DRIED DISCOLORED MATERIAL) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE: AT THE TIME OF IMPLANTATION, THE PATIENT HAD CATARACT IN THE OPERATIVE EYE OR NON-TRAUMATIC CATARACT IN THE FELLOW EYE, PREVIOUS CORNEAL OR REFRACTIVE SURGERY, LOW ENDOTHELIAL CELL COUNT, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PREVIOUS HISTORY OF IRITIS, SYNECHIA, PIGMENT DISPERSION, PSEUDOEXFOLIATION, ALL OF WHICH ARE CONTRAINDICATED AS PER THE DFU OF THIS LENS MODEL. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268468&pc=MTA | 2023826-2018-00181 | 2016-11-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS MAKING PLANS TO REPLACE THE LENS. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXPLANT. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A KERATOCONUS EYE. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6376873&pc=MTA | 2023826-2017-00305 | 2016-12-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, DIOPTER -5.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCES LOW VAULT, REFRACTIVE SURPRISE, AND RESIDUAL MYOPIA. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7118994&pc=MTA | 2023826-2017-01865 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO DECREASE VISUAL ACUITY. AT THE TIME OF SUBMISSION, ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED TAPED TO AN EMPTY CARTRIDGE CASE. VISUAL INSPECTION FOUND DEBRIS ON LENS AND THE HAPTIC MISSING A PIECE. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION CODE: - PER DFU FOR THIS LENS MODEL, IT IS RECOMMENDED ONLY TO USE MICROSTAAR INJECTORS, MODEL MSI-TF OR MSI-PF WITH SFC-45 OR SFC-45 FP CARTRIDGE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7119003&pc=MTA | 2023826-2017-01866 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO DECREASE VISUAL ACUITY. AT THE TIME OF SUBMISSION, ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: DESCRIBE EVENT OR PROBLEM - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED TAPED TO AN EMPTY CARTRIDGE CASE. VISUAL INSPECTION FOUND DEBRIS AND CLEAR RESIDUE ON LENS AND NO VISIBLE DAMAGE TO LENS. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, IT IS RECOMMENDED ONLY TO USE MICROSTAAR INJECTORS, MODEL MSI-TF OR MSI-PF WITH SFC-45 OR SFC-45 FP CARTRIDGE. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281454&pc=MTA | 2023826-2018-00259 | 2017-10-11 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REPORTEDLY, THE LENS WAS LATER EXPLANTED DUE TO THE SIZE BEING "TOO SMALL" AND REFRACTIVE SURPRISE. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7200513&pc=MTA | 2023826-2018-00099 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2006. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS A REFRACTIVE SURPRISE. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED DECREASED VISUAL ACUITY. LASEK WAS PERFORMED ON THE PATIENT ON (B)(6) 2007 FOR CORRECTING RESIDUAL REFRACTION AFTER ICL IMPLANTATION. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED AND THE PATIENT WAS SATISFIED WITH HIS VISUAL ACUITY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488976&pc=MTA | 2023826-2018-00678 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR ANOTHER LENS. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS HAPPY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PIECE MISSING, FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7527890&pc=MTA | 2023826-2018-00736 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE TORN AND THE LENS MISSING THE TORN PIECE OF HAPTIC. THERE WAS PRESENCE OF RESIDUE ON THE PRODUCT. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, IMPLANTABLE COLLAMER LENS OF, -18.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS OF DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975336&pc=MTA | 2023826-2016-01313 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, REFRACTIVE SURPRISE AND RETINAL REPAIR. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT (VAULT VALUE: 1360 UM) AND REFRACTIVE SURPRISE. RETINAL REPAIR WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER SAME MODEL LENS WITH DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING FINE WITH VAULT VALUE 500 UM AND AV 9/10; THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) IS 0.9 BETTER THAN THE PREOPERATIVE BEST CORRECTED VISUAL ACUITY (BCVA) = 0.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826245&pc=MTA | 2023826-2016-00256 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, LENS DISLOCATION/SUBLUXATION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.6. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION CONCLUDED LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191011&pc=MTA | 2023826-2016-01785 | 2016-01-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2016. AFTER IMPLANTATION, REFRACTIVE SURPRISE AND A LOSS OF BCVA WAS OBSERVED (20/40 PRE-OPERATIVE TO 20/70 POST-OPERATIVE). THE PATIENT COMPLAINED OF BLURRED VISION IMMEDIATELY AFTER SURGERY. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: DATE OF BIRTH IS UNKNOWN; BIRTH YEAR IS (B)(6). (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900450&pc=MTA | 2023826-2017-01487 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO SURGEON NOTED REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH (13.2) BUT DIFFERENT DIOPTER (-12.0), AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6962643&pc=MTA | 2023826-2017-01592 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (VTICMO12.6) AND DIOPTER (-5.0/+1.5/098) LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS ON LENS). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451075&pc=MTA | 2023826-2020-01854 | 2016-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, -03.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451111&pc=MTA | 2023826-2020-01853 | 2016-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, -04.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8055981&pc=MTA | 2023826-2018-01622 | 2017-09-29 | Injury | STAAR SURGICAL COMPANY | 2018-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773081&pc=MTA | 2023826-2018-01192 | 2017-11-24 | Injury | STAAR SURGICAL COMPANY AG | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE PATIENT DECIDED TO PROCEED WITH LASEK, WHICH WAS PERFORMED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773096&pc=MTA | 2023826-2018-01191 | 2017-11-24 | Injury | STAAR SURGICAL COMPANY AG | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM, ICM120V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE PATIENT DECIDED TO PROCEED WITH LASEK, WHICH WAS PERFORMED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT SHOULD BE CORRECTED TO 16-JUL-2018 FROM 17-JUL-2018 IN ORIGINAL MDR. DATE RECEIVED SHOULD BE CORRECTED TO 16-JUL-2018 FROM 17-JUL-2018 IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135976&pc=MTA | 2023826-2020-01174 | 2018-01-01 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672610&pc=MTA | 2023826-2019-01040 | 2018-03-05 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH BLACK STAINS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538838&pc=MTA | 2023826-2020-00001 | 2018-04-09 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE OVERTIME ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME MODEL, SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- DATE OF EVENT IN INITIAL MDR IS CORRECTED TO (B)(6) 2018. B5-THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME MODEL, SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721256&pc=MTA | 2023826-2018-01075 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -0.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS A >/= 2-LINE DROP IN BCDVA AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE DECREASE IN BCDVA REFLECTS THE SUBJECT'S CHALLENGE WITH NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PREVIOUS STATEMENT: THE LENS REMAINS IMPLANTED. CORRECTED AND ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2018. THE VISION LOSS THAT WAS INITIALLY REPORTED HAS BEEN RESOLVED. THE CAUSE IS LISTED AS DEVICE. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721259&pc=MTA | 2023826-2018-01073 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -0.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS A >/= 2-LINE DROP IN BCDVA AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE DECREASE IN BCDVA REFLECTS THE SUBJECT'S CHALLENGE WITH NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED AND ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2018. THE CAUSE IS LISTED AS DEVICE. THE VISION LOSS INITIALLY REPORTED HAS BEEN RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CURVED. (B)(4). Manufacturer Narrative: UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432773&pc=MTA | 2023826-2020-01790 | 2018-07-19 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHAGED FOR A TORIC LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746178&pc=MTA | 2023826-2019-01140 | 2018-07-21 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR DIFFERENT MODEL AND DIOPTER LENS DUE TO OBSERVATION OF REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272659&pc=MTA | 2023826-2019-00145 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: G4- "COMBINATION PRODUCT" SHOUDL BE REMOVED FROM PREVIOUS MDR. H6- INVESTIGATION CONCLUSION CODE: 4310 SHOULD BE CORRECTED TO 4315 IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021132&pc=MTA | 2023826-2018-01579 | 2018-08-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS NOTED AS UNKNOWN. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS.THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149315&pc=MTA | 2023826-2018-01806 | 2018-11-02 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS.THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT AND CLEAR SURGICAL RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451066&pc=MTA | 2023826-2020-01852 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451118&pc=MTA | 2023826-2020-01851 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2015. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141191&pc=MTA | 2023826-2020-01229 | 2018-12-11 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT.. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8333964&pc=MTA | 2023826-2019-00237 | 2019-01-18 | Injury | STAAR SURGICAL COMPANY | 2019-02-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBER ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404547&pc=MTA | 2023826-2019-00357 | 2019-01-31 | Injury | STAAR SURGICAL COMPANY AG | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2019. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE LENS WAS DISCARDED. (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140603&pc=MTA | 2023826-2020-01211 | 2019-02-01 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440522&pc=MTA | 2023826-2019-02332 | 2019-03-04 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS EXCHANGED FOR A SAME SIZE BUT DIFFERENT MODEL (TORIC) LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440662&pc=MTA | 2023826-2019-02333 | 2019-03-04 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -4.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS EXCHANGED FOR A SAME SIZE BUT DIFFERENT MODEL (TORIC) LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574170&pc=MTA | 2023826-2019-00636 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAIN IMPLANTED. REPORTER INDICATED EXCHANGE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795449&pc=MTA | 2023826-2019-01240 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836679&pc=MTA | 2023826-2019-01348 | 2019-05-24 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1,-11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836776&pc=MTA | 2023826-2019-01349 | 2019-05-24 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICM5_12.1,-11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836981&pc=MTA | 2023826-2019-01350 | 2019-06-08 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICM5_12.6,-4.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CONCLUSION CODE: 4310 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. RESULT CODE: 114 SHOULD BE REMOVED FROM SUPPLEMENTAL #1 MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY WITH SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC BENT, AND FIBERS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: "DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." SHOULD HAVE ALSO BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9339100&pc=MTA | 2023826-2019-02181 | 2019-06-19 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE(OD) ON (B)(6)2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUAITION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9010196&pc=MTA | 2023826-2019-01654 | 2019-06-26 | Injury | STAAR SURGICAL COMPANY | 2019-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076363&pc=MTA | 2023826-2019-01717 | 2019-07-04 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.75 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852361&pc=MTA | 2023826-2019-01384 | 2019-07-09 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE MANUFACTURE DATE: CORRECTED TO 25-APR-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - ON (B)(6) 2019 IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED AND PATIENT HAS ISSUES WITH HER BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILLIAR BODY. SEE MFR# 2023826-2019-02272 FOR EXCHANGED REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852363&pc=MTA | 2023826-2019-01383 | 2019-07-10 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - ON (B)(6) 2019 IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED AND PATIENT HAS ISSUES WITH HER BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILLIAR BODY. REPORTER PROVIDED PATIENT STATUS, AND STATED "WITH WEARING CONTACT LENSES THERE ARE NO ISSUES, OS IS HER DOMINANT EYE, MAYBE THE HIGHER RX ON OS CAUSE THE PROBLEMS. THE PATIENT ALSO TAKES PSYCHOTROPIC DRUGS. SHE WILL COME BACK FOR EXAM IN 1 MONTH". SEE MFR# 2023826-2019-02271 FOR EXCHANGED REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588684&pc=MTA | 2023826-2020-00072 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. G4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703107&pc=MTA | 2023826-2020-00263 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350976&pc=MTA | 2023826-2019-02184 | 2019-08-23 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTICS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WIDTH TO BE WITHIN SPECIFICATION. HOWEVER, OVERALL LENGTH VERIFICATION WAS NOT ABLE TO BE PERFORMED DUE TO THE RETURNED CONDITION OF THE LENS. FUNCTIONAL VERIFICATION WAS FOUND TO BE OUT OF SPECIFICATIONS DUE TO THE RETURNED CONDITION OF THE LENS (FOOTPLATE BROKEN/MISSING), LENS NOT SITING PROPERLY ON LENS VERIFICATION HOLDER. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - DATE OF EVENT: (B)(6) 2019. H10 - ADDITIONAL MANUFACTURER NARRATIVE: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299869&pc=MTA | 2023826-2019-02082 | 2019-09-03 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299940&pc=MTA | 2023826-2019-02083 | 2019-09-03 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND A HAPTIC TEAR. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372949&pc=MTA | 2023826-2019-02244 | 2019-09-04 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND BLACKISH RESIDUE AND THE HAPTIC SCRATCHED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "PROBLEM WAS RESOLVED." SHOULD BE REMOVED AND REPLACED WITH "PROBLEM WAS NOT RESOLVED." IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9513144&pc=MTA | 2023826-2019-02481 | 2019-09-09 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND STAINS ON THE HAPTIC/OPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9506333&pc=MTA | 2023826-2019-02479 | 2019-10-16 | Injury | STAAR SURGICAL COMPANY | 2019-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -7.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION FINDINGS 213 SHOULD BE ADDEX TO THE PREVIOUS MDR. H6- INVESTIGATION CONCLUSION 4310 SHOUDL BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403872&pc=MTA | 2023826-2020-01738 | 2019-10-16 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.75 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE PROBLEM WAS RESOLVED. CORRECTED STATEMENT: PROBLEM WAS NOT RESOLVED. REFERENCE FILE# (B)(4) FOR REPLACEMENT LENS. H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS FIBER ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WAS DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10- REFERENCE MFR# 2023826-2020-01755 (CLAIM# (B)(4)). CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857171&pc=MTA | 2023826-2020-00580 | 2019-11-06 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE LENS MAY BE EXCHANGED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5 13.2 OF -8.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A TORIC LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE: (B)(6)2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857186&pc=MTA | 2023826-2020-00581 | 2019-11-06 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE LENS MAY BE EXCHANGED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5 13.2 OF -8.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A TORIC LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE: (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824585&pc=MTA | 2023826-2020-00491 | 2019-11-19 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE TORIC LENS ON (B)(6) 2019 DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION CODE 10 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913389&pc=MTA | 2023826-2020-00740 | 2019-11-27 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE DIMENSIONAL VERIFICATION FOUND THE LENS IN SPECIFICATION BUT REFRACTIVE (FUNCTIONAL) VERIFICATION FAILED. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516586&pc=MTA | 2023826-2019-02472 | 2019-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND PUPIL HYPER REACTION WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "DUE TO THE PUPIL HYPER REACTION, PATIENT WAS COMPLAINING ABOUT QUALITY OF VISION." LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE 2137 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988391&pc=MTA | 2023826-2020-00954 | 2020-03-14 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND IT WAS UNKNOWN IF THE PROBLEM WAS RESOLVED, SEE MFR REPORT # 2023826-2020-00955 FOR THE EXCHANGED LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335265&pc=MTA | 2023826-2020-01562 | 2020-04-01 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED BY THE REPORTER THE LENS WAS REMOVED AND EXCHANGED ON (B)(6)2020. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025652&pc=MTA | 2023826-2020-01005 | 2020-04-02 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "INDUCED ASTIGMATISM." THE SURGEON BELIEVES THE ASTIGMATISM WAS CAUSED NOT BY THE LENS BUT THE "LOAD" ON THE WOUND AND CORNEA DURING SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH FOREIGN BLOOD STAINS/FIBERS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION FINDINGS CODE 213 SHOULD BE INCLUDED IN SUPPLEMENTAL MDR#1. H6 - INVESTIGATION CONCLUSION CODE 4310 SHOULD BE INCLUDED IN SUPPLEMENTAL MDR#1. CLAIM# : (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476385&pc=MTA | 2023826-2020-01929 | 2020-04-19 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H3-DEVICE EVALUATION: FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135977&pc=MTA | 2023826-2020-01175 | 2020-04-21 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075976&pc=MTA | 2023826-2020-01088 | 2020-04-22 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. THE PROBLEM IS RESOLVED. H6-PATIENT CODE 3191 (LENS EXCHANGE, SECONDARY SURGERY). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - DATE OF LENS EXCHANGE: (B)(6) 2020, SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308537&pc=MTA | 2023826-2020-01498 | 2020-06-01 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2020 A DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10404437&pc=MTA | 2023826-2020-01755 | NA | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -02.25 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928374&pc=MTA | 2023826-2018-01397 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT REPORTED HER NEAR VISION WAS POOR. THE PATIENT COULDN'T SEE CLEARLY FROM CLOSE (40CM), SO SHE REQUESTED TO REMOVE THE ICL LENS AND DIDN'T WANT ANOTHER LENS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; LENS WAS CAUGHT IN VIAL CAP. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8198020&pc=MTA | 2023826-2018-01926 | 2018-09-27 | Injury | STAAR SURGICAL COMPANY | 2018-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A VAULT VALUE OF 300, LENS TEAR BEFORE/DURING LOADING, LENS DISLOCATION/SUBLUXATION AND BLURRED VISION WAS REPORTED. REPOSITIONING WAS REPORTED AS A SECONDARY INTERVENTION. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTION: EVENT - THE REPORTER INDICATED THAT A VICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT THE LENS WAS FOUND TORN/BROKEN BEFORE OPENING THE VIAL. LENS DISLOCATION/ SUBLUXATION ALONG WITH BLURRED VISION AND LENS REPOSITIONING WERE ALSO REPORTED. THE VAULT VALUE WAS REPORTED AS 300 M. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. DEVICE CODE 2993: PREVIOUS CODE NOT APPLICABLE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050057&pc=MTA | 2023826-2017-01758 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REPORTEDLY, ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO LOSS OF BCVA. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND MOISTURE ON THE LENS BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048973&pc=MTA | 2023826-2020-01038 | 2020-03-25 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; CALCULATIONS WERE PERFORMED INCORRECTLY BY THE USER. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048977&pc=MTA | 2023826-2020-01037 | 2020-03-25 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; CALCULATIONS WERE PERFORMED INCORRECTLY BY THE USER. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227424&pc=MTA | 2023826-2019-00053 | 2009-03-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. "PATIENT IS HAPPY AND EXCHANGE WENT WELL!" PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (LENS EXCHANGE, SECONDARY SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107521&pc=MTA | 2023826-2018-01749 | 2016-02-27 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN ICM120V4, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. "DUE TO OBSERVATION OF LOW VAULT" SHOULD BE CORRECTED TO "DUE TO OBSERVATION OF LOW VAULT AND REFRACTIVE CHANGE OVERTIME." CONCOMITANT MEDICAL PRODUCT: CARTRIDGE MODEL AND LOT NOT APPLICABLE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8108029&pc=MTA | 2023826-2018-01748 | 2016-02-27 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN ICM120V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. "DUE TO OBSERVATION OF LOW VAULT" SHOULD BE CORRECTED TO "DUE TO OBSERVATION OF LOW VAULT AND REFRACTIVE CHANGE OVERTIME." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107279&pc=MTA | 2023826-2018-01671 | 2018-02-09 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A LONGER LENS DUE TO BLURRED VISION AND LOW VAULT. THE REPORTER STATES, "EXCHANGE DIDN'T RESOLVE LOW VAULT, PATIENT COMPLAINT ABOUT BLURRED VISION BUT UCVA ACHIEVED PRE-OP BCVA." REFERENCE MFR REPORT# 2023826-2018-01665 FOR CASE CONTINUATION. Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624221&pc=MTA | 2023826-2020-00147 | 2018-05-02 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE AND DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824085&pc=MTA | 2023826-2018-01269 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS NOTED AS "MISMATCH OF LENS SIZE." ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER, DIFFERENT POWER AND MODEL LENS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164913&pc=MTA | 2023826-2018-01819 | 2018-10-14 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS "RPORTED" AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). "WORD" ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241995&pc=MTA | 2023826-2019-00087 | 2018-11-01 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE STATUS OF THE EYE IS NOTED AS "GOOD". CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843596&pc=MTA | 2023826-2019-01339 | 2019-06-17 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED WITH A LONGER, DIFFERENT MODEL LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SPHERICAL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY AND THE POST OP UCVA IS 20/20. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D7-EXPLANT DATE : (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SPHERICAL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY AND THE POST OP UCVA IS 20/20. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D7- EXPLANT DATE (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894316&pc=MTA | 2023826-2019-01430 | 2019-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL WITHIN A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894317&pc=MTA | 2023826-2019-01429 | 2019-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON B)(6) 2019. ON B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL WITHIN A ZIPLOC BAG. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979191&pc=MTA | 2023826-2020-00927 | 2019-06-22 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, SURGEON "USED LRI TECHNIQUE MAKING THE MAIN INCISION SUPERIORLY TO REDUCE WTR CYLINDER AND WAS ABLE TO REDUCE BUT NOT FULLY DUE TO HIGH AMOUNT OF CYLINDER POWER. SURGEON FELT CURRENT VAULTING WAS TOO LOW(230R/262L) AND NEEDS TO CHANGE THE LENS WITH DIFFERENT POWER AND LENGTH" WHILE "RE-ENTERING MAIN INCISION TO DECREASE ASTIGMATISM FURTHER AND INCREASE -0.50 MORE MYOPIC LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN SUPPLEMENTAL MDR #1 B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMAPAIRMENT IN SUPPLEMENTAL MDR #1 H6 - MEDICAL DEVICE PROBLEM CODE 1583 SHOULD BE INCLUDED IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979197&pc=MTA | 2023826-2020-00926 | 2019-06-22 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, SURGEON "USED LRI TECHNIQUE MAKING THE MAIN INCISION SUPERIORLY TO REDUCE WTR CYLINDER AND WAS ABLE TO REDUCE BUT NOT FULLY DUE TO HIGH AMOUNT OF CYLINDER POWER. SURGEON FELT CURRENT VAULTING WAS TOO LOW (230R/262L) AND NEEDS TO CHANGE THE LENS WITH DIFFERENT POWER AND LENGTH" WHILE "RE-ENTERING MAIN INCISION TO DECREASE ASTIGMATISM FURTHER AND INCREASE -0.50 MORE MYOPIC LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN SUPPLEMENTAL MDR #1. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMAPAIRMENT IN SUPPLEMENTAL MDR #1. H6 - MEDICAL DEVICE PROBLEM CODE 1583 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960165&pc=MTA | 2023826-2020-00866 | 2019-07-24 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416976&pc=MTA | 2023826-2019-02201 | 2019-08-04 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782973&pc=MTA | 2023826-2020-00407 | 2020-02-05 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/40 (PRE-OP BCVA WAS 20/20) SO CLAIM # (B)(4) (MFR REPORT # 2023826-2020-00408) WAS OPENED. TO DATE, THE BCVA HAS NOT INCREASED BUT VAULT IS BETTER. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT SHOULD HAVE BEEN SPECIFIED THAT THE LENS WAS EXPLANTED AND REPLACED IN A SEPERATE SURGERY ON THE SAME DATE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508629&pc=MTA | 2023826-2018-00723 | 2010-06-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2010. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS REFRACTIVE CHANGE OVERTIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.7) BUT DIFFERENT MODEL AND DIOPTER - VTICM5_13.7, -6.00/+1.0/X086 (SPHERE/CYLINDER/AXIS). THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT. MALFUNCTION TO SERIOUS INJURY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887952&pc=MTA | 2023826-2019-01353 | 2019-04-22 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. BLURRED VISION, INTERMITTENT ANGLE CLOSURE WITHOUT ELEVATED IOP, AND EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER, "DOCTOR THOUGHT THERE WAS NO RELATED FACTOR WITH PRODUCT." Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND DEBRIS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8693088&pc=MTA | 2023826-2019-01082 | 2019-04-09 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR THE SAME MODEL AND DIOPTER LENS DUE TO BLURRED VISION AND ACCRETION ON THE LENS, LENS OPACITY. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "THERE WERE WHITE MATERIALS ADHERED ON THE LENS AND THE LEVEL DID NOT GET WELL BUT WORSE, THEY DECIDED TO REMOVE THE LENS". Manufacturer Narrative: H6 - DEVICE CODE 1426 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - FOREIGN BODY ANALYSIS: LIGHT-COLORED DEBRIS ON THE LENS WAS ANALYZED USING FOURIER TRANSFORM INFRARED SPECTROSCOPY (FTIR) AND WAS FOUND TO BE A MIXTURE OF PROTEIN AND CARBOHYDRATE, CONSISTENT WITH A BIOLOGICAL MATERIAL. H6 - RESULT CODE 213 CORRECTED TO RESULT CODE 114 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, NO WHITE PARTICULATE. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965322&pc=MTA | 2023826-2019-01611 | 2019-06-04 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; No Apparent Adverse Event | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. RINDERON EYE DROPS INITIALLY. CURRENT TREATMENT FLUOROMETHOLONE EYE DROPS (PATIENT IS AN ATHLETE, CANNOT USE STEROIDS FOR LONG). THE LENS REMAINS IMPLANTED. PER THE REPORTER ON PATIENT'S (B)(6) 2019 VISIT, "BLURRED VISION HAS BEEN GENERATED FROM THREE WEEKS BEFORE THIS DAY." Manufacturer Narrative: RESULT CODE: 213 SHOULD BE CORRECTED TO 114, AS IT IS NOT APPLICABLE. DEVICE CODE: 1401 SHOULD HAVE BEEN INCLUDED. PATIENT CODE 3191: NO CODE AVAILABLE (PATIENT WAS TREATED WITH MEDICATIONS, RINDERON AND FLUOROMETHLONE) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. ADVERSE EVENT SHOULD HAVE BEEN SELECTED IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN SELECTED IN INITIAL MEDWATCH REPORT. AT A FOLLOW UP VISIT IT WAS REPORTED THAT THE "WHITE ACCRETIONS OF THE PATIENT'S GOT SMALLER WITHOUT SPECIFIC MEDICAL TREATMENT. DOCTOR THINKS THE RISK OF RECURRENCE WILL BE LOW. REPORTED LENS OPACITY MEANS THAT ICL LENS OPACITY WAS DUE TO WHITE ACCRETIONS AND NOT PATIENT'S CATARACT." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136933&pc=MTA | 2023826-2019-01806 | 2018-10-01 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138555&pc=MTA | 2023826-2020-01118 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - PATIENT CODE: 2137 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574235&pc=MTA | 2023826-2019-00674 | 2019-02-28 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM512.1, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE CODE: 1494 - OFF-LABEL USE SHOULD BE REMOVED FROM PREVIOUS MEDWATCH REPORT, AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: AT THE LATEST DISCUSSION WITH THE SURGEON, THE SURGEON BELIEVES THAT IT IS POSSIBLE HE MISTAKINGLY IMPLANTED THE LENS MEANT FOR THE PATIENT'S OTHER EYE. THE SURGEON ISN'T CONSIDERING EXCHANGING THE LENS. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE SPHERE TO BE OUT OF SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552454&pc=MTA | 2023826-2019-00645 | 2019-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE, BLURRY VISION, NARROW ANGLES, AND "1+ PIGMENTED CELL". THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: SERIAL NUMBER, EXP. DATE, UDI: UNK. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8617956&pc=MTA | 2023826-2019-00819 | 2019-03-28 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED ALONG WITH THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION OF THE LENS SPHERE WAS FOUND TO BE 0.45D OUT OF SPECIFICATION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMAITON: THE REPORTER INDICATED THAT THAT A 13.7MM, VICM5_13.7, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR, PATIENT HAD SOME PREBYOPIA DUE TO AGE. ALTHOUGH THE PATIENT HAD GOOD VISION AFTER INITIAL LENS IMPLANT, THE PATIENT WAS DISSATISFIED WITH NEAR VISION. THE INITIAL LENS WAS EXCHANGED BECAUSE THE PATIENT WANTED BETTER NEAR VISION. THERE ARE CURRENTLY NO ADVERSE EVENTS IN THE PATIENT AND NO COMPLAINTS FROM THE PATIENT. CORRECTED DATA DATE OF BIRTH: CORRECTED FROM (B)(6) 2019 TO (B)(6) 1968. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: PREVIOUS MDR SHOULD INCLUDE "USE/HANDLING AND ADDITIONAL PROCESSING IMPACTS LENS DIMENSIONS/FUNCTIONS; RETURNED LENS MAY NOT MEET ORIGINAL SPECIFICATIONS." DATE RECEIVED BY MANUFACTURER: DATE IN INITIAL MDR SHOULD BE CORRECTED TO 18-APR-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618086&pc=MTA | 2023826-2019-00820 | 2019-03-28 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THAT THAT A 13.7MM, VICM5_13.7, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR, PATIENT HAD SOME PREBYOPIA DUE TO AGE. ALTHOUGH THE PATIENT HAD GOOD VISION AFTER INITIAL LENS IMPLANT, THE PATIENT WAS DISSATISFIED WITH NEAR VISION. THE INITIAL LENS WAS EXCHANGED BECAUSE THE PATIENT WANTED BETTER NEAR VISION. THERE ARE CURRENTLY NO ADVERSE EVENTS IN THE PATIENT AND NO COMPLAINTS FROM THE PATIENT. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTICS WERE BROKEN, BENT AND DEFORMED. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). PATIENT DATE OF BIRTH: CORRECTED FROM (B)(6) 2019 TO (B)(6) 1968. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF REPORT: "17-APR-2019" SHOULD BE CORRECTED TO "18-APR-2019" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: "17-APR-2019" SHOULD BE CORRECTED TO "18-APR-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9513282&pc=MTA | 2023826-2019-02482 | 2019-09-09 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND STAINS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. D4- EXPIRATION DATE: 31-DEC-2021. H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). B5 - THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043596&pc=MTA | 2023826-2020-01031 | 2019-10-21 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -03.50 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND LENS WAS EXCHANGED ON (B)(6) 2019 WITH A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED MOISTURIZED IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9379814&pc=MTA | 2023826-2019-02245 | 2019-10-24 | Injury | STAAR SURGICAL COMPANY | 2019-11-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS/ CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF-LABEL USE; ACD < 3.00MM SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412469&pc=MTA | 2023826-2019-02274 | 2019-10-29 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -16.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES LOSS OF BCVA AND REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A VTICMO12.6, -17.50/0.5/088 (SPHERE/CYLINDER/AXIS), LENS ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE) CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988419&pc=MTA | 2023826-2020-00948 | 2019-11-18 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS BUT THE NEAR VISION WAS SLIGHTLY IMPROVED BUT STILL NOT IDEAL. PATIENT WILL BE MONITORED. SEE MFR REPORT # 2023826-2020-00949 FOR EXCHANGED LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398560&pc=MTA | 2023826-2020-01750 | 2020-06-19 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D11- NOT APPLICABLE IN THE INITAL MDR. ADDITIONAL INFORMATION: H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAMEE LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814242&pc=MTA | 2023826-2020-00473 | 2019-12-07 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8850039&pc=MTA | 2023826-2019-01369 | 2019-06-13 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, MICL 12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE REPORTER STATED, "BLURRY AND DOUBLE VISION POST OP, NO PATIENT INJURY. PROBLEM RESOLVED AFTER EXPLANTING THE LENS". PATIENT IS STABLE, VISION WAS BETTER AND CLOUDINESS WAS GONE. NO PLANS TO IMPLANT ANOTHER LENS, PATIENT DID NOT WANT THE LENS REPLACED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE; ACD < 3.00MM. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440448&pc=MTA | 2023826-2020-01894 | 2020-01-02 | Injury | STAAR SUGICAL COMPANY | 2020-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -6.5 DIOPTER, INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT MODEL LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717957&pc=MTA | 2023826-2020-00334 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE ISSUE. Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE: 3001 SHOULD BE CORRECTED TO 1401. CLAIM#:(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628975&pc=MTA | 2023826-2019-00891 | 2019-02-18 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 17 FEB 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 2682 PATIENT-DEVICE INCOMPATIBILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639513&pc=MTA | 2023826-2020-00188 | 2019-12-19 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - DEVICE CODE: 2682 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416217&pc=MTA | 2023826-2020-01720 | 2020-07-09 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (-06.50) LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066004&pc=MTA | 2023826-2018-01665 | 2018-02-20 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO BLURRED VISION AND "ICL TILTED." THE REPORTER STATES, "AFTER FIRST EXCHANGE, ICL LENS TILTED, SURGEON DIDN'T WANT TO EXCHANGE ANOTHER ICL. PATIENT IS OK." REFERENCE MFR REPORT# 2023826-2018-01671 FOR INFORMATION REGARDING FIRST LENS IMPLANT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: METHOD CODE 4110: LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS IN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729355&pc=MTA | 2023826-2019-01113 | 2019-05-30 | Injury | STAAR SURGICAL COMPANY | 2019-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENTS NEXT VISIT IS IN (B)(6), NO TREATMENT HAS BEEN PERFORMED AND THE PATIENTS CURRENT CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, DOCTOR ADMITTED THAT THE LENS WAS UPSIDE DOWN. HE REMOVED THE LENS AND RE-IMPLANTED THE LENS. IT WAS REPORTED THAT THE PATIENT IS DOING VERY WELL NOW. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REMOVED AND RE-IMPLANTED LENS. CLAIM#: (B)(4). Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN INDICATED IN SUPPLEMENTAL MDR#1 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729476&pc=MTA | 2023826-2019-01114 | 2019-05-30 | Injury | STAAR SURGICAL COMPANY | 2019-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENTS NEXT VISIT IS IN (B)(6), NO TREATMENT HAS BEEN PERFORMED AND THE PATIENTS CURRENT CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, DOCTOR ADMITTED THAT THE LENS WAS UPSIDE DOWN. HE REMOVED THE LENS AND RE-IMPLANTED THE LENS. IT WAS REPORTED THAT THE PATIENT IS DOING VERY WELL NOW. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REMOVED AND RE-IMPLANTED LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8051091&pc=MTA | 2023826-2018-01593 | 2018-06-28 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. REFRACTIVE SUPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. REMOVAL AND EXCHANGED IS CURRENTLY BEING PLANED TO RESOLVE THE ISSUE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THE ORIGINAL MDR STATED "LENS REMAINS IMPLANTED" WHICH SHOULD BE UPDATED TO "LENS REMOVED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT SHOULD HAVE BEEN ADDED IN SUPPLEMENTAL MEDWATCH REPORT #2. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #2. UNK SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #2. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: 12-SEP-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111670&pc=MTA | 2023826-2016-01655 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION TO DATE OF REPORT: 09/30/2016. ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4): SECONDARY SURGERY, LENS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, SAME DIOPTER AND SIZE. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169824&pc=MTA | 2023826-2018-00053 | 2017-12-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Consequences Or Impact To Patient | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS IMPLANTED AND REFRACTIVE SURPRISE WAS OBSERVED AFTER SURGERY. THE REPORTED STATED THAT TIME CORRECTED THE PROBLEM AND THAT THE LENS REMAINS IMPLANTED. LENS REMOVAL WAS NOT NECESSARY, PATIENT IS DOING GREAT AND IS VERY HAPPY. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365365&pc=MTA | 2023826-2017-00279 | 2017-01-13 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION-LENS RETURNED DRY IN SMALL CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DRY RESIDUE AND DEBRIS ON LENS AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965298&pc=MTA | 2023826-2019-01562 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; No Apparent Adverse Event; Appropriate Term/Code Not Available | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS, BLURRED VISION, AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: 1426 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2137. DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1401. DEVICE CODE: 3191 NO CODE AVAILABLE (OPACITY ON THE LENS) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: 2019. IMPLANTED: 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764121&pc=MTA | 2023826-2019-01156 | 2019-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REPORTEDLY "POST-OP REFRACTION, SHOWED HYPEROPIC SHIFT (+1.0D)" IN PATIENT EYE. PER REPORTER ON (B)(6) 2019, "PATIENT'S UCVA WAS 20/13 AND POST-OP REFREATION SHOWED +1.0/-0.25/109 ON THE LAST VISIT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2137 IN INITIAL MEDWATCH REPORT. DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1401 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039570&pc=MTA | 2023826-2023-01884 | 2023-04-06 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME(TASS) AND BLURRED VISION WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITION OF ORAL STEROIDS AND OINTMENTS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WASREPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6: HEALTH IMPACT CODE: "4625 - ADDITIONAL SURGERY: ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163013&pc=MTA | 2023826-2020-01257 | NA | Malfunction | STAAR SURGICAL COMPANY AG | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -09.0, IMPLANTABLE COLLAMER LENS INTO PATIENT'S LEFT EYE (OS) IN 2009. REPORTEDLY, LENS "HAS ROTATED." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - SHOULD BE CORRECTED TO "H6 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT", IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME ("VISION STARTED CHANGING"). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163014&pc=MTA | 2023826-2020-01256 | NA | Malfunction | STAAR SURGICAL COMPANY AG | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -09.0, IMPLANTABLE COLLAMER LENS INTO PATIENT'S RIGHT EYE (OD) IN 2009. REPORTEDLY, LENS "HAS ROTATED." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THE PATIENT HAD EXPERIENCED REFRACTIVE CHANGE OVER TIME (PATIENT'S VISION STARTED TO CHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - SHOULD BE CORRECTED TO "H6 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT", IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: IMPLANTED DATE: 2009. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595254&pc=MTA | 2023826-2023-00997 | 2023-01-09 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A DIOPTER OF -9.0 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL DUE TO REFRACTIVE SURPRISE AND BLURRED VISON. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "SURGICAL INCISION INDUCED MORE ASTIMATISM THAN EXPECTED," AND "REFRACTIVE MISS - PATIENT HAD 1.D OF ASTIGMASTISM AFTER SPHERICAL ICL PLACEMENT, DESPITE HAVING ONLY 0.5D PRE-OP SO I EXCHANGED THE SPHERICAL ICL FOR A TORIC." THE STATUS OF THE EYE IS REPORTED AS, "HEATHY." Manufacturer Narrative: WEIGHT-RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- RETURNED TO MANUFACTURER: "13-MAR-2023" SHOULD BE CORRECTED TO "10-MAR-2023." B5- "CAUSE IS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED." DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00D INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY "1D ASTIGMATISM AFTER SPHERICAL ICL PLACEMENT, DESPITE HAVING ONLY 0.50D PRE-OP". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE DIFFERENT MODEL/ SAME LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS NOTED AS "HEALTHY". THE CAUSE FOR THIS EVENT WAS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED.- "SURGICAL INCISION INDUCED MORE ASTIGMATISM THAN EXPECTED". H6- MEDICAL DEVICE PROBLEM CODE: "2017" SHOULD BE REMOVED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987063&pc=MTA | 2023826-2018-01508 | 2018-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED LENS MICL12.1, -9.5 DIOPTER, IMPLANTED IN PATIENTS RIGHT EYE ON (B)(6) 2018 IMPROVED VAULT BUT VAULT CONTINUES TO BE HIGH. AT 3 MONTH EXAM, PATIENT WAS 20/20 UDVA, WITH A +0.25 SPH (OD), AND A 2+ VAULT. PATIENT COMPLAIN IS WITH INTERMEDIATE AND NEAR VISION AFTER LONG DAYS. PATIENT HAS CHOSEN NOT TO EXPLANT LENS. ADDITIONAL INFORMATION, PILO 1% AND LIMIFY HAVE BENN OF MINIMAL HELP. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7995562&pc=MTA | 2023826-2018-01555 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.1, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018 AS A REPLACEMENT LENS. AT 3 MONTH POST-OP EXAM, THE PATIENT WAS 20/20 UDVA, WITH A PI +0.25X075 (OS), AND A 2+ VAULT. THE PUPIL WAS MINIMALLY REACTIVE AT 6-7MM. PATIENT'S COMPLAIN IS WITH INTERMEDIATE AND NEAR VISION AFTER LONG DAYS. PATIENT HAS CHOSEN NOT TO EXPLANT THE LENS. PILO 1% AND LUMIFY OF MINIMAL HELP. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794738&pc=MTA | 2023826-2022-01891 | 2021-02-20 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY GLARES/HALOES AND BLURRED VISION OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE. STATUS OF EYE IS WITHIN NORMAL LIMITS. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880650&pc=MTA | 2023826-2021-04519 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL13.7, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT GLARE/HALOES BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ABERATIONS IMPROVED BUT STILL PRESENT. PER UPDATED INFORMATION THE REPORTER REPORTED THAT THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS RESIDUAL REFRACTIVE ERROR AND GLARE. LASIK FOR RESIDUAL ASTIGMATISM IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880665&pc=MTA | 2023826-2021-04520 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL13.7, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT GLARE/ HALOES BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ ABERATIONS IMPROVED BUT STILL PRESENT. PER UPDATED INFORMATION THE REPORTER REPORTED THAT THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS RESIDUAL REFRACTIVE ERROR AND GLARE. LASIK FOR RESIDUAL ASTIGMATISM IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6329973&pc=MTA | 2023826-2017-00236 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-02-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.50 INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2016 (EXACT DATE UNKNOWN). AFTER IMPLANTATION, THE PATIENT COMPLAINED OF SEEING GLARES, HALOES, AND BLURRED VISION. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AS OF DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN PLANNED TO EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE IMPLANT DATE REPORTED WAS (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED DUE TO GLARE/HALOES, WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. CORRECTIONAL DATA: THE CORRECT IMPLANT DATE REPORTED WAS (B)(6) 2017. (B)(4). Manufacturer Narrative: EXACT IMPLANT DATE NOT KNOWN; IMPLANT IN (B)(6) 2016. THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153640&pc=MTA | 2023826-2023-00127 | 2022-12-03 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D9: RETURN DATE (B)(6) 2023. G2: FOREIGN; OTHER: (B)(6). H3: DEVICE EVALUATED: NO. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153746&pc=MTA | 2023826-2023-00128 | 2022-12-03 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: D9: RETURN DATE 06-FEB-2023. G2: FOREIGN OTHER: MALAYSIA H3: DEVICE EVALUATED: NO. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483348&pc=MTA | 2023826-2020-01951 | 2018-12-01 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2016. LENS OPACITY (ASC) WAS OBSERVED ON 15/JAN/2019. GLARE/HALOES AND BLURRED VISION WAS ALSO OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HALOS SINCE SURGERY, AT NIGHT. ALPHAGAN HELPED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: B3-DATE OF EVENT IN INITIAL MDR SHOULD BE REPORTED AS (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483350&pc=MTA | 2023826-2020-01952 | 2018-12-01 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2016. LENS OPACITY (ASC) WAS OBSERVED ON (B)(6) 2019. GLARE/HALOES AND BLURRED VISION WAS ALSO OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HALOS SINCE SURGERY, AT NIGHT. ALPHAGAN HELPED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: B3-DATE OF EVENT IN INITIAL MDR SHOULD BE REPORTED AS (B)(6) 2018. CLAIM# : (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930997&pc=MTA | 2023826-2020-02815 | 2020-09-24 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS WAS STUCK IN CARTRIDGE/INJECTION SYSTEM AND TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE PATIENT COMPLAINED OF GLARE/HALOS AND BLURRED VISION. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE PATIENT IS STABLE WITH NO DIPLOPIA. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052011&pc=MTA | 2023826-2023-02086 | 2023-03-13 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052113&pc=MTA | 2023826-2023-02087 | 2023-03-13 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364576&pc=MTA | 2023826-2019-02196 | 2019-10-18 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT COMPLAINED ABOUT GLARE, BLURRED VISION, AND THE IMPLANTED LENS. BUT HE STILL COMPLAINED ABOUT VISUAL IMPAIRMENT AFTER REMOVING THE LENS. PER THE REPORTER ON (B)(6) 2019, "CURRENTLY THE PATIENT IS UNDER CONTINUOUS OBSERVATION AND NO TREATMENT." Manufacturer Narrative: B5: IT WAS LATER MENTIONED THAT THE PATIENT'S VISUAL IMPAIRMENT HAS INCLUDED TO HAVE FLOATERS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED BY THE PATIENT THAT THEY HAD TRIED OUT RGP, SCLERALS LENSES BUT THIS STILL DID NOT RESOLVE THE PROBLEM AND VISION IS STILL MUCH WORSE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16375402&pc=MTA | 2023826-2023-00543 | 2022-11-26 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT TOLERATE THE HALO. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991203&pc=MTA | 2023826-2023-01764 | 2023-03-28 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -3.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT PREFERRED LENS EXPLANT AND USE OF GLASSES. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991204&pc=MTA | 2023826-2023-01765 | 2023-03-28 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION ON (B)(6)2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT PREFERRED LENS EXPLANT AND USE OF GLASSES. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11661950&pc=MTA | 2023826-2021-01047 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES, BLURRED VISION, CIRCLES IN VISION AND RAINBOW HALOES. LENS REMAINS IMPLANTED. THE REPORTER STATED " I BELIEVE HE IS GOING TO WAIT A COUPLE OF MONTHS AND SEE IF HE CAN ADAPT TO THE LENSES." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION " THE PATIENT IS NOTICING AN IMPROVEMENT IN THEIR SYMPTOMS AND IS HOLDING OFF ON MAKING A DECISION UNTIL WINTER." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11662621&pc=MTA | 2023826-2021-01048 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE REPORTER STATED " HE IS COMPLAINING OF HALOS/GLARES AND THEN A REFRACTIVE SURPRISE." IT WAS LATER REPORTED THAT THE PATIENT IS ALSO EXPERIENCING BLURRED VISION, CIRCLES IN VISION AND RAINBOW HALOES. LENS REMAINS IMPLANTED. THE REPORTER STATED " I BELIEVE HE IS GOING TO WAIT A COUPLE OF MONTHS AND SEE IF HE CAN ADAPT TO THE LENSES." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION " THE PATIENT IS NOTICING AN IMPROVEMENT IN THEIR SYMPTOMS AND IS HOLDING OFF ON MAKING A DECISION UNTIL WINTER." CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279827&pc=MTA | 2023826-2020-01502 | 2020-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE PATIENT COMPLAINED SHE COULD NOT SEE WELL, WAS SEEING HALOS, LINES AND VIBRATION OF THE LENS. HER NEAR VISION WAS BLURRED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 STATEMENT - "PATIENT IS DEPRESSED" SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H1: TYPE OF REPORTABLE EVENT IN INITIAL MDR SHOULD BE CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391264&pc=MTA | 2023826-2023-03041 | 2022-12-21 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON 08/SEP/2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND BLURRED VISION. PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391269&pc=MTA | 2023826-2023-03040 | 2022-12-21 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, 16.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND BLURRED VISION. PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412661&pc=MTA | 2023826-2019-00418 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO 13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AND THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. PATIENT CURRENT CONDITION IS GOOD. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414779&pc=MTA | 2023826-2019-00419 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATD THAT THE SURGEON IMPLANTED A 13.2MM, VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO GLARE/HALOES AND BLURRED VISION. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER SHOULD BE 14-FEB-2019 IN INITIAL MDR (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473809&pc=MTA | 2023826-2024-00133 | 2023-12-12 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: EXCESSIVE VAULT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES, PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16329543&pc=MTA | 2023826-2023-00464 | 2022-11-16 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND DISCOMFORT DUE TO OVER CORRECTION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- OVER CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191998&pc=MTA | 2023826-2016-01811 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND GLARE/HALOES. THE LENS REMAINS IMPLANTED IN THE PATIENTS EYE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). METHOD: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438817&pc=MTA | 2023826-2018-00563 | 2013-01-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS LOW VAULT, REFRACTIVE CHANGE OVER TIME, GLARE/HALOS, AND BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684682&pc=MTA | 2023826-2017-01095 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, GLARE/HALOES AND PUPIL MODERATELY DILATED. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH VICMO12.6 AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). SECONDARY SURGERY INTERVENTION, EXCHANGE, PIGMENT DISPERSION, SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES,PUPIL MODERATELY DILATED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7074520&pc=MTA | 2023826-2017-01826 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT, REFRACTIVE CHANGE OVERTIME, READING PROBLEMS, BLURRED VISION, AND GLARE/HALOS WAS NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7075520&pc=MTA | 2023826-2017-01827 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT, REFRACTIVE CHANGE OVERTIME, READING PROBLEMS, BLURRED VISION, AND GLARE/HALOS WAS NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540586&pc=MTA | 2023826-2017-00742 | 2016-11-29 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT STATED THAT HE EXPERIENCED GLARE/HALOES AT NIGHT. THE LENS REMAINS IMPLANTED, THE PATIENT FEELS THE SIDE EFFECT ARE NOT SIGNIFICANT. Manufacturer Narrative: PREVIOUS: (B)(6) 2017, CORRECTED: (B)(6) 2016. ADDITIONAL INFORMATION: PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, AND IRIS ATROPHY. PREVIOUS: (B)(4; CORRECTED: (B)(4). (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT STATED THAT HE EXPERIENCED GLARE/HALOES AT NIGHT, EXCESSIVE VAULTING, LOSS OF BCVA AND IRIS ATROPHY. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906082&pc=MTA | 2023826-2018-01439 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. GLARE/HALOES AND BLURRED VISION WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7918895&pc=MTA | 2023826-2018-01454 | 2018-06-03 | Injury | STAAR SURGICAL COMPANY | 2018-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND BLURRED VISION. ON (B)(6) 2018 A SHORTER LENGTH LENS WAS IMPLANTED. THIS REPLACEMENT LENS RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE MAIN REASON OF REMOVAL WAS THE HIGH VAULT THAT RESULTED IN A DILATED UNREACTIVE PUPIL. PATIENT COMPLAINED MAINLY FROM HALOES AND DISRUPTED IMAGES. Manufacturer Narrative: (B)(6). (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "30-AUG-2018" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269224&pc=MTA | 2023826-2019-00146 | 2018-11-27 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT STILL COMPLAINS OF GLARE/HALOS AND BLURRED VISION. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. POST-OP PATIENT CONDITION - DRY CORNEAL EPITHELIUM. SEE MFR REPORT # 2023826-2019-00147 FOR EXCHANGED LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634216&pc=MTA | 2023826-2020-00202 | 2018-11-29 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. PRK, REFRACTIVE TREATMENT WAS ALSO REPORTED AS A SECONDARY INTERVENTION. IT WAS REPORTED THAT THE EXCHANGE LENS RESOLVED THE PROBLEM. AT PATIENTS POST OP FOLLOW UP VISIT " SOME GLARE AT NIGHT " AND " S/P PRK DOING BETTER, MUCH BETTER VAULT" WAS NOTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS ICL (HIGH VAULT CAUSING SECONDARY ABERRATIONS, REFRACTORY TO CONTACT LENS CORRECTION). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN SPECIMEN CONTAINER WITH RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634632&pc=MTA | 2023826-2020-00203 | 2018-11-29 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. PRK, REFRACTIVE TREATMENT WAS ALSO REPORTED AS A SECONDARY INTERVENTION. IT WAS REPORTED THAT THE EXCHANGE LENS RESOLVED THE PROBLEM. AT PATIENTS POST OP FOLLOW UP VISIT " GLARE AT NIGHT "AND " S/P PRK DOING BETTER, MUCH BETTER ICL VAULT" WAS NOTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS ICL (HIGH VAULT CAUSING SECONDARY ABERRATIONS, REFRACTORY TO CONTACT LENS CORRECTION). Manufacturer Narrative: CORRECTED DATA: D6: SHOULD BE "(B)(6) 2018" IN THE INITIAL MDR. ADDITIONAL INFORMATION: H6- WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN SMALL CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7911897&pc=MTA | 2023826-2018-01448 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER EXPLANTING THE LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7912042&pc=MTA | 2023826-2018-01449 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATINETS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER EXPLANTING THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190948&pc=MTA | 2023826-2020-01349 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: (B)(6) 2017. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189334&pc=MTA | 2023826-2020-01281 | 2017-12-12 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906085&pc=MTA | 2023826-2018-01440 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. GLARE/HALOES AND BLURRED VISION WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A MICRO-CENTRIFGUE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269212&pc=MTA | 2023826-2019-00147 | 2018-12-21 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT COMPLAINED OF GLARE/HALOS AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THIS IS THE EXCHANGED LENS FOR MFR REPORT # 2023826-2019-00146. Manufacturer Narrative: CHANGED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". CLAIM#: (B)(4). Manufacturer Narrative: DATE OF REPORT, DATE RECEIVED BY MANUFACTURER CORRECTED TO 24-DEC-2018 FOR INITIAL MDR. CONCOMITANT MEDICAL PRODUCTS: PREVIOUS CARTRIDGE MODEL LOT NUMBER AND FOAM TIP PLUNGER MODEL LOT NUMBER NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2019 THE LENS WAS EXPLANTED AND POSTOPERATIVE THE EYE IS "QUIET" AND THE "PATIENT WAS FEELING FINE". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364572&pc=MTA | 2023826-2019-02197 | 2019-04-17 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARE AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT COMPLAINED ABOUT GLARE, BLURRED VISION, AND THE IMPLANTED LENS. BUT HE STILL COMPLAINED ABOUT VISUAL IMPAIRMENT AFTER REMOVING THE LENS. PER THE REPORTER ON (B)(6) 2019, "CURRENTLY THE PATIENT IS UNDER CONTINUOUS OBSERVATION AND NO TREATMENT." Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 3191 SHOULD BE IN THE INITIAL MDR. H6- PATIENT CODE 2227 SHOULD BE ADDED IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS LATER INDICATED BY THE PATIENT THAT THEY HAD TRIED RGP/SCLERALS LENSES BUT THIS DID NOT RESOLVE THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516581&pc=MTA | 2023826-2019-02471 | 2019-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL HYPER REACTION, GLARE/HALOES, AND BLURRED VISION WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "DUE TO THE PUPIL HYPER REACTION. PATIENT WAS COMPLAINING ABOUT QUALITY OF VISION." LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE 3191 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR H6 - PATIENT CODE 3191 - 'SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE' CORRECTED TO 'SECONDARY SURGICAL INTERVENTION, LENS EXPLANT' IN INITIAL MDR H6 - PATIENT CODE 3191 - NO CODE AVAILABLE, PUPIL HYPER REACTION SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6396523&pc=MTA | 2023826-2017-00367 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS HAS YET TO BE EXPLANTED AS OF (B)(6) 2017 AND THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. Manufacturer Narrative: THE PATIENT EXPERIENCED GLARE AND HALOS DUE TO CENTRAL AQUAPORT HOLE. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS. THE PROBLEM IS RESOLVED, SYMPTOMS DISAPPEARED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213739&pc=MTA | 2023826-2016-01837 | 2014-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE LENS WAS EXCHANGED DUE TO LOW VAULTING AND 'REFRACTIVE CHANGE OVER TIME'. NO REPORT OF REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727937&pc=MTA | 2023826-2015-00470 | NA | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT WAS HAVING VISUAL ACUITY PROBLEMS. THE PATIENT COMPLAINED OF GHOSTING/BLURRING VISION. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE SURGEON IS GOING TO WAIT UNTIL THE EYE HEALS BEFORE IMPLANTING ANOTHER LENS. SUTURES WERE NOT REQUIRED AND THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: PT WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED BY FACILITY. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - MANY FACTORS COULD CAUSE VISUAL ACUITY CHANGE SUCH AS UNDER CORRECTED REFRACTION AND PATHOPHYSIOLOGICAL CHANGE OF VISUAL SYSTEM. GIVEN THE REPORTED INFORMATION, THE CAUSE OF THE EVENT IS UNKNOWN. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY, ICL (12.6 MM;-8.5 D) WAS EXPLANTED SIX WEEKS POSTOPERATIVELY TO ADDRESS PATIENT VISUAL DISTURBANCES ("GHOSTING"/"BLURRY VISION"). ACCORDING TO THE REPORT IMPLANTATION OF ANOTHER LENS WAS NOT SCHEDULED YET. THE PATIENT WAS DOING WELL. THE SAME REPORT STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. VISUAL DISTURBANCES HAVE BEEN IDENTIFIED AS POTENTIAL COMPLICATIONS AFTER ICL IMPLANTATION. THE OPTICAL DIAMETER OF THE ICL RANGES FROM 4.9 MM TO 5.8 MM AND THE PRECAUTIONS SECTION OF THE DFU INSTRUCTS PHYSICIANS THAT "PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE". IT FURTHER PRECAUTIONS THAT "THE EFFECT OF PUPIL SIZE ON VISUAL SYMPTOMS IS NOT KNOWN". THE DFU INDICATES THE "SMALLER THE OPTIC DIAMETER, THE GREATER THE INCIDENCE OF SUBJECTIVE PATIENT SYMPTOMS". (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727929&pc=MTA | 2023826-2015-00469 | NA | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT WAS HAVING VISUAL ACUITY PROBLEMS. THE PATIENT COMPLAINED OF GHOSTING/BLURRING VISION. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE SURGEON IS GOING TO WAIT UNTIL THE EYE HEALS BEFORE IMPLANTING ANOTHER LENS. SUTURES WERE NOT REQUIRED AND THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED BY FACILITY. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY FACTORS COULD CAUSE VISUAL ACUITY CHANGE SUCH AS UNDER CORRECTED REFRACTION AND PATHOPHYSIOLOGICAL CHANGE OF VISUAL SYSTEM. GIVEN THE REPORTED INFORMATION, THE CAUSE OF THE EVENT IS UNKNOWN. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6146503&pc=MTA | 2023826-2016-01717 | 2015-02-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -4.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE CUSTOMER EXPERIENCED REFRACTIVE SURPRISE AND LOSS OF BEST CORRECTED VISUAL ACUITY (PRE-OP 20/20 / POST-OP 20/25). CUSTOMER COMPLAINED FROM FIRST POST-OP DAY, BUT A FEW MONTHS LATER BECAME PREGNANT. THE LENS REMAINS IMPLANTED. PATIENT TO SEE SURGEON FOR CORRECTION AFTER DELIVERY OF BABY. Manufacturer Narrative: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5240530&pc=MTA | 2023826-2015-01521 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | NA | NA | 1 | Event Description: THE PATIENT REPORTED THROUGH SOCIAL MEDIA HAD A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IMPLANTED IN THEIR LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT REPORTED FOUR DAYS AFTER SURGERY, WAS UNABLE TO READ ANYTHING/BLURRY. TO DATE THE PATIENT'S VISION HAS NOT IMPROVED. THE PATIENT WAS PRESCRIBED EYE DROPS (NEVANAC AND GENTEAL GEL) FOR A MONTH. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572294&pc=MTA | 2023826-2016-00409 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE CHANGE OVER TIME. THE REPORTER INDICATED THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919703&pc=MTA | 2023826-2016-01192 | NA | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS TORE DURING DELIVERY INTO THE EYE AND SUBSEQUENTLY THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934940&pc=MTA | 2023826-2016-01143 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934943&pc=MTA | 2023826-2016-01144 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/12.5. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645517&pc=MTA | 2023826-2017-00956 | 2017-03-22 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE REPORTER INDICATED LENS OPACITY (ANTERIOR) WAS FIRST OBSERVED ON (B)(6) 2017 AND THE LENS WAS EXPLANTED DUE TO LOW VAULTING ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT HAD ADDITIONAL SYMPTOMS - MILD BLURRING. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652095&pc=MTA | 2023826-2016-00641 | 2013-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW/NO VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LOW VAULT OF THE LENS WAS FIRST OBSERVED ON (B)(6) 2013. LENS OPACITY WAS OBSERVED (1 TR ASC) AND PATIENT EXPERIENCED BLURRY VISION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2, WITH NO DISCOMFORT OR COMPLAINTS. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328083&pc=MTA | 2023826-2015-01750 | NA | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HIS VISION HAS CHANGED TO -0.5. THE FACILITY REPORTED THE PATIENT HAS AN EARLY CATARACT, DUE TO THE AGING PROCESS AND NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH RESULTS: PER MEDICAL REVIEW: ON (B)(6) 2015 THE PATIENT COMPLETED 36 MONTH POST-TREATMENT PATIENT FOLLOW-UP FORM AND REPORTED DECREASE OF VISUAL ACUITY FOLLOWING ICL IMPLANTATION. ACCORDING TO DCR FORM, RECEIVED ON 11/19/15, THE CAUSE OF THE EVENT WAS A CATARACT DEVELOPMENT DUE TO AGING PROCESS AND NOT RELATED TO THE DEVICE. ICL REMAINED IMPLANTED AND UCVA WAS 20/25 WITH GOOD PROGNOSIS. SERIOUSNESS IS BASED ON NO REPORTED LOSS OF BCVA AND NO REPORT OF ANY SECONDARY INTERVENTION PLANNED OR SCHEDULED. IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE", THEREFORE THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS IMPLANTED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760961&pc=MTA | 2023826-2015-00534 | 2014-02-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device remains implanted | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED PATIENT HAD A MICL 13.6 MM IMPLANTABLE COLLAMER LENS IMPLANTED IN THE PATIENT'S LEFT EYE ON (B)(6) 2012. AS PART OF THE POST MARKET STUDY, THE PATIENT REPORTED EXPERIENCING HAVING LOST ACUITY IN HIS NEAR VISION AFTER THE SURGERY AND HAVE DIFFICULTY FOCUSING WITHOUT THE USE OF CONTACTS OR GLASSES TO BALANCE HIS EYESIGHT. ALSO BEING TOLD BY HIS OPTOMETRIST THAT SHE CAN SEE THE BEGINNINGS OF CATARACTS. THE PHYSICIAN'S OFFICE WAS CONTACTED FOR FURTHER INFORMATION. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2014. THE PATIENT IS NOW HYPEROPIC. STATED ADVERSE EVENT IS RELATED TO THE DEVICE. PROGNOSIS IS GOOD. VISUAL ACUITY POST-TREATMENT BCDVA 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). ICL REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536601&pc=MTA | 2023826-2016-00303 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER RIGHT EYE (OD) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY CONFIRMED THE PATIENT REPORT AND INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015. THE TYPE OF CATARACT WAS CORTICAL PERIPHERAL. THE LENS HAD A LOW VAULT AND THIS PROBABLY WAS THE CAUSE OF THE SUBCAPSULAR HAZE. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328002&pc=MTA | 2023826-2015-01751 | 2014-04-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR AND POST SUBCAPSULAR). THE LENS WAS EXPLANTED ON (B)(6) 2014, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S VA WAS 20/25 AND THE PROGNOSIS IS GOOD. Manufacturer Narrative: PER MEDICAL REVIEW: REPORTEDLY, ICL WAS EXPLANTED FROM A (B)(6) - YEAR OLD PATIENT, TWO YEARS POSTOPERATIVELY, TO ADDRESS CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST (DCR), RECEIVED ON 11/19/15, CATARACT WAS EXTRACTED AND IOL WAS IMPLANTED AT THE SAME SURGERY DATE. THE SAME REPORTED ATTRIBUTED CATARACT (ANTERIOR SUBCAPSULAR COMBINED WITH POSTERIOR SUBCAPSULAR) TO THE DEVICE. THE PROGNOSIS AFTER SECONDARY SURGICAL INTERVENTION WAS GOOD (VA 20/25). A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084245&pc=MTA | 2023826-2015-01194 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE IS STILL SEDIMENT (VISCOELASTIC) ON THE CRYSTALLINE LENS CENTER. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION: IMPLANTED DATE - (B)(6) 2015. ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION ((B)(4)) AND OCUCOAT ((B)(4)) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536606&pc=MTA | 2023826-2016-00302 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT REPORTED THE EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON REPORTED THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROGNOSIS IS EXCELLENT. THE VA POST-OP WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536972&pc=MTA | 2023826-2016-00300 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS REMOVED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2015. THE REPORTER INDICATED THE CATARACT WAS A CORTICAL CENTRAL AND THE EVENT WAS NOT DEVICE RELATED. AN IOL WAS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327964&pc=MTA | 2023826-2015-01752 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE LENS WAS EXPLANTED ON (B)(6) 2015, THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: PER MEDICAL REVIEW: REPORTEDLY ICL WAS EXPLANTED 3 YEARS POSTOPERATIVELY TO ADDRESS LOW VAULT AND A CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST, DATED ON (B)(6) 2015, THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR AND WAS ATTRIBUTED TO THE DEVICE. FOLLOWING ICL REMOVAL, CATARACT WAS EXTRACTED, AND IOL WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE PATIENT PROGNOSIS WAS GOOD (VA:20/20). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILIAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5139024&pc=MTA | 2023826-2015-01342 | NA | Injury | STAAR SURGICAL COMPANY AG | 2015-10-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT, THE DEVELOPMENT OF A CATARACT AND DECREASED VISUAL ACUITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER MEDICAL REVIEW - CENTRAL AND/OR PERIPHERAL CONTACT BETWEEN THE ICL AND THE CRYSTALLINE LENS IS ASSOCIATED WITH ANTERIOR SUBCAPSULAR CATARACT FORMATION. HOWEVER, THIS REPORT PROVIDES NO INFORMATION ON THE DEGREE OF VAULTING OR THE POSITION OF THE LENS OPACITIES. ADDITIONALLY, THERE IS NO INFORMATION ABOUT THE CRYSTALLINE LENS PRIOR TO THE INITIAL SURGERY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284955&pc=MTA | 2023826-2017-00114 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(6) 1966. NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284956&pc=MTA | 2023826-2017-00115 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(4). DATE OF BIRTH - (B)(6) 1966. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327956&pc=MTA | 2023826-2015-01753 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Retinal Tear; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE PATIENT HAD A SMALL RETINAL TEAR AND IT WAS REPAIRED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, (B)(6) YEAR OLD PATIENT RECEIVED ICL IN 2012 AND WAS COMPLAINING ABOUT DECREASE IN VA DUE TO CATARACT DEVELOPMENT BY COMPLETING 36 MONTH FOLLOW-UP FORM. THE PATIENT ALSO REPORTED SMALL RETINAL TEAR THAT WAS REPAIRED. DATA CLARIFICATION REQUEST FORM WAS RECEIVED ON 11/19/2015, STATING THAT ANTERIOR SUBCAPSULAR CATARACT WAS CAUSED BY LOW VAULT OF ICL DECREASING THE VA FROM 20/15 TO 20/20. ICL REMAINED IMPLANTED AND NO SECONDARY OR SURGICALLY INTERVENTION WAS PERFORMED. HOWEVER, THE SMALL RETINAL TEAR, THAT WAS TREATED AND RESOLVED BY CRYOTHERAPY, WAS NOT CAUSED BY DEVICE BUT BY PATIENT RELATED FACTOR (VITREOUS DEGENERATION). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) YEARS OLD AND PERF DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE "THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4) LENS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547267&pc=MTA | 2023826-2023-03522 | 2023-03-22 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158233&pc=MTA | 2023826-2020-01235 | 2020-03-23 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS EYE ON (B)(6) 2020. THE PATIENT COMPLAINED OF BLURRED VISION. EXPLANTATION OF THE LENS AND CLEAR LENS SURGERY IS PLANNED FOR (B)(6) 2020. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4 - CORRECTED TO (B)(6) 2020. B5 - "IT WAS REPORTED THAT CLEAR LENS SURGERY AND IMPLANTATION OF NON-STAAR LENS WAS PERFORMED, WHICH RESOLVED THE ISSUE." SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CORRECTED DATA: H6 - DEVICE CODE:1401 SHOULD HAVE BEEN INCLUDED. B3 - CORRECT DATE SHOULD BE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECT DATE SHOULD BE REVERTED BACK TO (B)(6) 2020. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED AS IT IS N/A. H6 - PATIENT CODE: 2137 SHOULD BE REMOVED AS IT IS N/A. B5 - "THE PATIENT COMPLAINED OF BLURRED VISION", SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS N/A. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914025&pc=MTA | 2023826-2021-04571 | NA | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN HIS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14252557&pc=MTA | 2023826-2022-01313 | 2022-02-15 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022 WITH PAINLESS PROGRESSIVE VISION LOSS IN EYE FOR HALF A YEAR. LATER, ON (B)(6) 2022, THE PATIENT RECEIVED ICL IOL EXTRACTION FOR PHACOEMULSIFICATION EXTRACTION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14253045&pc=MTA | 2023826-2022-01314 | 2022-02-15 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT HAD PAINLESS PROGRESSIVE VISION LOSS, AND WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022. ON (B)(6) 2022, THE PATIENT UNDERWENT ICL IOL EXTRACTION, PHACOEMULSIFICATION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13374207&pc=MTA | 2023826-2022-00351 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). IT WAS INDICATED THAT THE LENS WAS TOO FLAT AND RESULTED IN MINIMAL CLEARANCE OVER THE NATURAL LENS, AND THERE WAS A SLIGHT CATARACT WITH A "SOMEWHAT" HYPEROPIC OUTCOME. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PATIENT IS DOING QUITE WELL. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12904968&pc=MTA | 2023826-2021-04570 | NA | Injury | STAAR SURGICAL COMPANY | 2021-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A (B)(4) IMPLANTABLE (B)(6) LENS, -12.00 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE PATIENT SEES HALOES IN HIS RIGHT EYE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, EXPLANT DATE: UNKNOWN, INITIAL REPORTER OCCUPATION: UNKNOWN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191994&pc=MTA | 2023826-2016-01814 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT COMPLAINED OF BLURRY VISION AT NIGHT. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING FINE. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN PIECES AND UNABLE TO EVALUATE. THE LENS WAS RETURNED DRY. METHOD CODE(S): LENS WORK ORDER SEARCH. RESULTS CODE(S): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4959171&pc=MTA | 2023826-2015-00960 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY | 2015-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY) AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED BY PHACOEMULSIFICATION AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140284&pc=MTA | 2023826-2016-01711 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. AT 3 MONTHS POST-OP, PATIENT WAS 20/25 UNCORRECTED AND THE PRESSURES WERE NORMAL. THE REPORTER INDICATED THE BEGINNING OF AN ANTERIOR SUBCAPSULAR CATARACT WAS NOTED AT THAT VISIT. THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140273&pc=MTA | 2023826-2016-01712 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINS IMPLANTED BUT MAY BE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY (ANTERIOR) WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. THE LENS DIOPTER WAS -13.0. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978092&pc=MTA | 2023826-2017-01620 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -3.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS OPACITY-ASCO. SURGEON ALSO PERFORMED PHACO-EMULSIFICATION AND IOL IMPLANTATION. SURGEON STATES THAT THE PATIENT "DEVELOPED ASCO AFTER 5 MONTHS SURGERY." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/25 AT THE LAST EXAM WHICH WAS A DECREASE COMPARED TO BASELINE BCVA WHICH WAS 20/22. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021462&pc=MTA | 2023826-2017-01676 | 2017-09-06 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LOSS OF BCVA, FROM 20/40 TO 20/100, AND NS AND CORTICAL LENS OPACITY WAS NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. ON (B)(6) 2017, THE PATIENT'S POST-OP UCVA IS 20/30 AND THE STATUS OF THE STATUS IS PERFECT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH ANY CATARACT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438479&pc=MTA | 2023826-2018-00544 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT THE PROBLEMS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM 12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS OF SIMILAR MODEL ON (B)(6) 2018 AND THE "VAULT WAS NORMAL AROUND 400UM APROX.". THE FINAL BCVA VALUE FOR PATIENT IS 0.8. EXPLANTED DATE: IT IS CORRECTED TO (B)(6) 2018. (B)(4). DEVICE MANUFACTURE DATE: IT IS CORRECTED TO 08-MAR-2011. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190944&pc=MTA | 2023826-2020-01388 | 2019-05-12 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: A CASE REPORT WAS RECEIVED AND INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -22.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND THE PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT FEMTOSECOND LASER ASSISTED CATARACT SURGERY (FLACS) COMBINED WITH ICL EXTRACTION AND AN IOL WAS IMPLANTED. VISUAL ACUITY WAS REDUCED FOR THE LAST 2 YEARS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252472&pc=MTA | 2023826-2015-01553 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS HAD A LOW VAULT, REFRACTIVE SURPRISE AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE: (B)(4) PREVIOUS CODE INCORRECT. NO INDUCED CATARACT WAS REPORTED. DEVICE CODE: (B)(4) (IMPLANTED)-NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. ADDITIONAL DATA: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTION TORN, HAPTIC BROKEN, AND HAIR ON LENS SURFACE. LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS, DATE UNKNOWN. PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/25. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074711&pc=MTA | 2023826-2016-01569 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR AND POSTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2016. THE PATIENT EXPERIENCED BLURRED VISION. THE PATIENT WAS SCHEDULED FOR YAG CAPSULOTOMY. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WERE FIBERS ON THE LENS. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT AND THE LENS WAS FOUND TO BE IN SPECIFICATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6742631&pc=MTA | 2023826-2017-01242 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE SURGEON ALSO NOTED THAT THE PATIENT'S VISUAL ACUITY WAS 20/60 ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) AND EXPLANTED ON (B)(6) 2016. THE CATARACT WAS FIRST OBSERVED ON (B)(6) 2016 AND WAS LOCATED IN THE NUCLEAR PART OF THE EYE. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. (B)(4). DATE OF EVENT - (B)(6) 2016. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870372&pc=MTA | 2023826-2018-01382 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC); PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. ACCORDING TO THE SURGEON, THE EXACT CAUSE OF THE CATARACT IS UNKNOWN BUT MAY HAVE OCCURRED DUE TO HIGH MYOPIA. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFGUE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL DATA RECEIVED ON 15-FEB-2019: " THE PATIENT ALREADY HAD IOL IN THE LEFT EYE DUE TO CATARACT FORMATION. PATIENT CAME BACK TO THE CLINIC LAST SUMMER COMPLAINING BLURRY VISION AND DOCOTR FOUND NSC, NOT ASCO. HE ADVISED TO REMOVE ICL AND IMPLANTED IOL, (B)(6) 2018" (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438480&pc=MTA | 2023826-2018-00543 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT PROBLEMS. Manufacturer Narrative: THE BEST CORRECTED VISUAL ACUITY (BCVA) FOR THE PATIENT IS CURRENTLY 1.0. (B)(4) Manufacturer Narrative: "SURGERY IS PLANNED TO CORRECT PROBLEMS." IN INITIAL MDR SHOULD BE UPDATED TO "THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. NO INFORMATION REGARDING THE DETAILS OF THE SECONDARY SURGERY HAS BEEN PROVIDED DESPITE THE REQUESTS FOR DETAILS." PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938159&pc=MTA | 2023826-2020-00798 | NA | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE VAULT IS MINIMAL. LASIK ENHANCEMENT WAS PERFORMED 9 MONTHS LATER. THE PATIENT HAS NOW DEVELOPED AN ANTERIOR CORTICAL CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER ON (B)(6) 2019. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D6: IMPLANTED DATE - (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611974&pc=MTA | 2023826-2019-00792 | 2019-02-23 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -1.75 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS FOLLOWING ICL IMPLANTATION. PHACOEMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION WAS PERFORMED ON THE RIGHT EYE 3 WEEKS AFTER EXPLANTATION. THE PATIENT PASSED AWAY ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722066&pc=MTA | 2023826-2022-01932 | 2022-05-09 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT HAS LENS OPACITY (CATARACT ASC). IT WAS ALSO NOTED "THE INVESTIGATOR DISCUSSED THE DEVELOPMENT OF ANTERIOR SUBCAPSULAR OPACITIES WITH THE PATIENT AND THE DECISION WAS MADE TO MONITOR THE VISION FOR NOW". THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4614" AND "2199" SHOULD BE REMOVED AND CODE "4627" AND "4625" SHOULD BE ADDED. H6-HEALTH EFFECT- CLINICAL CODE: "2140" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S LEFT EYE (OS) ON 03/JUNE/2020 AS A BACKUP LENS. ON 09/MAY/2022 LENS OPACITY CATARACT (ASC) WAS OBSERVED. THE PATIENT ALSO EXPERIENCED BLURRY/HAZY VISION. ON 29/AUG/2023 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "SUBJ INITIALLY REPORTED BLURRY & HAZY VA OS AT V7 (09MAY22). PI NOTED AN ASC OS. PI & SUBJ DECIDED TO MONITOR. AT V8 (28JUL23), SUBJ REPORTED VA HAD WORSENED, ESP IN BRIGHT LIGHTS. PI & SUBJ DECIDED TO PROCEED W/ CATARACT SX OS.SAE EVENT SUMMARY 2 : SUBJ HAD UNEVENTFUL ICL EXPLANT & SIMULTANEOUS CATARACT EXTRACTION SURGERY ON (B)(6) 2023. UCVA OS IMPROVED TO 20/12.5 AT 1 DAY PO. SITE WILL CONT TO FOLLOW SUBJ TO ENSURE HEALING AND STABILITY". THE CAUSE OF THE EVENT WAS THE DEVICE. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069167&pc=MTA | 2023826-2017-01792 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069168&pc=MTA | 2023826-2017-01791 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877264&pc=MTA | 2023826-2017-01463 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM12.0 IMPLANTABLE COLLAMER LENS, -21.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2016 THE LENS WAS EXPLANTED. THE SURGEON NOTED LOW VAULT, LENS ROTATION, LENS OPACITY ASCO, AND "ARTIPHAKIA" AS THE STATUS OF THE EYE. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION IS ALSO REPORTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT'S DATE OF BIRTH IS CORRECTED FROM "(B)(6) 1986" TO "(B)(6) 1983". BLURRED VISION AND PRESENCE OF GLARES/ HALOES EXPERIENCED BY THE PATIENT, SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.9MM)]. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816274&pc=MTA | 2023826-2016-00974 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, LENS OPACITY (CATARACT), AND LOSS OF VISION. THE PATIENT EXPERIENCED AN ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, BLURRED VISION, IRIS ATROPHY, AND CORNEAL EDEMA. THE ANTERIOR CHAMBER IRRIGATION OF REMAINING VISCO FLUIDS WAS PERFORMED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/200. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND SPOT ON LENS SURFACE. DIMENSIONAL INSPECTION CONLUDED THAT THE LENS IS IN SPECIFICATION. (B)(4). Manufacturer Narrative: NO REPORTED LOSS OF VISION. DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661678&pc=MTA | 2023826-2016-00727 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908388&pc=MTA | 2023826-2019-01481 | 2019-03-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: H6- DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908395&pc=MTA | 2023826-2019-01482 | 2019-03-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148280&pc=MTA | 2023826-2021-02469 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Endophthalmitis; Blurred Vision | N/A | NA | 1 | Event Description: A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE SURGEON REPORTS PERMANENT LOSS OF BCVA AND ENDOPHTHALMITIS. DIAGNOSTICS WERE PERFORMED; NO ORGANISM WAS FOUND. A VITRECTOMY, INTRAVITREAL ANTIBIOTICS, AND A VITREOUS TAP WERE PERFORMED (B)(6) 2021. ON (B)(6) 2021 A VITRECTOMY WAS PERFORMED AND THE LENS WAS EXPLANTED, PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED AS THE PATIENT HAD DEVELOPED A CATARACT. PROBLEM IS RESOLVED. Manufacturer Narrative: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11435324&pc=MTA | MW5099834 | 2019-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-05 | MTA | STAAR SURGICAL ICL IMPLANT | Insufficient Information | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | NA | NA | 1 | Event Description: I HAD ICL SURGERY ON (B)(6) 2009. IN 2011, I STARTED TO HAVE ISSUES WITH MY VISION. I HAD TO START WEARING EYE GLASSES TO DRIVE AND THEN ENDED UP WITH -1.25 VISION TO THIS DATE. IN (B)(6) 2019 I STARTED TO HAVE AN ITCHING/POKING FEELING IN MY RIGHT EYE. MY EYE CONTINUED TO WATER AND ITCH AS IF SOMETHING WAS POKING ME. I VISITED SEVERAL EYE DOCTORS, INCLUDING MY SURGEON WHO HAD DONE MY ICL SURGERY. IT FEELS/FELT AS IF SOMETHING WAS WRONG WITH MY ICL. THE PAIN IS BEHIND MY EYE. THE SURGEON STATED THAT I HAVE A SMALL CATARACT NOW ON MY EYE AND THAT IS CAUSING MY BLURRY VISION AND THE ISSUE. I AM HAVING BLURRY VISION, BUT I AM STILL WITNESSING THE WATERY/ITCHING SENSATION. THE SURGEON DR. LIU ADVISED I NEED TO GET MY ICL IMPLANT TAKEN OUT AND INPUT A CATARACT LENSES. I STATED THAT I DO NOT FEEL THIS IS RIGHT. SOMETHING IS NOT RIGHT WITH THE CURRENT ILC LENSES INSIDE MY EYE. I HAVE SINCE DEALT WITH THIS CONTINUED EFFECT FOR THE PAST TWO YEARS. EYE DOCTORS CANNOT FIGURE IT OUT. IT IS NOT MY CATARACT CAUSING THIS SENSATION. IT'S ALMOST LIKE MY ICL IS TRYING TO POKE OUT OF MY EYE OR COME OUT. I AM FINALLY GETTING A SECOND OPINION THIS NEXT WEEK WITH ANOTHER EYE SURGEON TO CHECK MY ICL. TO SEE IF THAT IS IN FACT THE ISSUE, AS MY RIGHT EYE IS STILL ITCHING, WATERY. I CAN'T PUT A CONTACT IN MY RIGHT EYE AS IT IS TOO BOTHERSOME. I WANT TO REPORT THIS AS I FEEL AFTER THIS MANY YEARS PERHAPS THE LENSES HAS GONE BAD OR IT NEEDS REPLACED AFTER A CERTAIN TIME. WHEN I DID ICL BACK IN 2009 IT WAS NEW AND JUST BECOMING A THING. PERHAPS THE ICL'S BACK THEN ARE NOT AS UP TO DATE AS WHAT THEY HAVE CREATED NOW. STAAR SURGICAL ICL IMPLANTS. IMPLANT ICL: MICL12.1 RIGHT EYE. DIOPTER. -10.0/ OPTIC DIAMETER 4.9-5.8MM OVERALL DIAMETER: 12.1M. SN: (B)(4) LEFT EYE: DIOPTER.-10.5/ OPTIC DIAMETER 4.9- 5.8MM OVERALL DIAMETER: 12.2M SN: (B)(4). FDA SAFETY REPORT ID # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620596&pc=MTA | 2023826-2021-00861 | NA | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2009. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY CORTICAL, REFRACTIVE CHANGE OVER TIME, THE PATIENT HAD TO WEAR GLASSES TO DRIVE, A WATERY/ITCHING/POKING FEELING AND PAIN IN THE EYE. BLURRY VISION ALSO REPORTED. THE SURGEON STATES THE PATIENT'S ACD CHANGED FROM 3.93 IN 2009 TO 3.39 IN 2020. THE PATIENT'S AXIAL LENGTH INCREASED. "NO ONE HAS CONTROL OVER THAT!" Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651554&pc=MTA | 2023826-2021-00885 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2021 DUE TO LOW VAULT, LENS OPACITY (ASC), BLURRED VISION AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DERIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. AFTER LENS EXPLANT, PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. HEALTH IMPACT CODE 4625. ADDITIONAL SURGERY (PHACOEMULSIFICATION & IOL IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814249&pc=MTA | 2023826-2020-00475 | 2020-01-17 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Headache; Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION, LOSS OF BEST CORRECTED VISUAL ACUITY AND LENS OPACITY (POSTERIOR SUBCAPSULAR) WAS NOTED. THE PATIENT IS SENSITIVE TO BRIGHT LIGHT AND HEADACHE WHEN READING. THE LENS REMAINED IMPLANTED AND PATIENT WILL RETURN IN 3 MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED WITH AN INTRAOCULAR LENS ON (B)(6) 2020. CATARACT SURGERY WAS ALSO PERFORMED ON (B)(6) 2020, DURING LENS REPLACEMENT. THE REPORTER INDICATED THAT THE PATIENT WAS SEEN ON (B)(6) 2020, AND THAT SHE IS "STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." H6 - HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY, CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6305835&pc=MTA | 2023826-2017-00223 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND WAS PRESCRIBED ORAL MEDICATIONS OR EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO INSIDE OF EYE WAS INFLAMED. INFORMATION RECEIVED FROM THE FACILITY CONFIRMED THE PATIENT REPORT AND THE DATE OF THE CATARACT DIAGNOSIS WAS (B)(6) 2016. THEY INDICATED PATIENT HAS EPISCLERITIS EXTERNALLY WHICH REQUIRED PRED FORTE TO CLEAR. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107373&pc=MTA | 2023826-2019-01752 | 2019-05-30 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED AND THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED AND A MONOFOCAL LENS WAS IMPLANTED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1932: INFLAMMATION. DEVICE CODE 1494:OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098622&pc=MTA | 2023826-2020-03254 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND PUNTIFORM OPACITY IN THE POSTERIOR CAPSULE. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: NON RECOMMENDED OVD AND EXCESSIVE USE OF OVD DUE TO A NON TOTAL DILATION DURING SURGERY. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. REPORTEDLY, THE IOP WAS MANAGED AFTER THE EVACUATION OF OVD. PUNTIFORM OPACITY HAS DECREASED AFTER 48 HOURS OF SURGERY. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PROBLEM IS NOW RESOLVED. (B)(4). Manufacturer Narrative: B5. ELEVATED IOP IS AN EXPECTED OR FORESEEABLE SIDE EFFECT. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669737&pc=MTA | 2023826-2020-02329 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE. THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - LOW VAULT, TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669743&pc=MTA | 2023826-2020-02328 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE . THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS REPORTED AS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. HOWEVER, MILD PUPIL CORECTOPIA IS REPORTED POST EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233814&pc=MTA | 2023826-2019-00072 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE INITIAL LENS WAS NOT EXCHANGED FOR ANOTHER LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233815&pc=MTA | 2023826-2019-00071 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. A SECONDARY YAG WAS PERFORMED FOR ENLARGEMENT OF THE PERIPHERAL IRIDECTOMIES. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. SEE MFR REPORT # 2023826-2019-00073 FOR INITIAL LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024717&pc=MTA | 2023826-2016-01448 | NA | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PRESSURE AND PAIN FOR PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRY VISION. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. POST OP THE PATIENT HAD TRACE CORTICAL OPACITY AND BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE VISUAL ACUITY WAS MORE STABLE BUT THE PUPIL WAS DILATED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291699&pc=MTA | 2023826-2017-00197 | NA | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291665&pc=MTA | 2023826-2017-00196 | NA | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Missing Value Reason | Cataract; Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358993&pc=MTA | 2023826-2017-00252 | 2015-09-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED AND REPLACED. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THEY CONFIRMED THE PATIENTS REPORT. THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS EXPLANTED ON (B)(6) 2015, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. THE FACILITY REPORTED THE EVENT WAS RELATED TO THE DEVICE. Manufacturer Narrative: (B)(4). METHOD: LENS WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6408305&pc=MTA | 2023826-2017-00414 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2016 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE DATE OF EXPLANT WAS (B)(6) 2016, THE PATIENTS POST-OP VA WAS 20/20 AND PROGNOSIS WAS EXCELLENT. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601777&pc=MTA | 2023826-2015-00318 | NA | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED AND THE CATARACT REMOVED ON (B)(6) 2015. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - EVEN THOUGH THE POSITION OF ICL IMPLANTATION COULD CONTRIBUTE TO CATARACT, OTHER FACTORS COULD ALSO CAUSE THE DEVELOPMENT OF CATARACT SUCH AS AGING, TRAUMA (EYE RUBBING, EXTERNAL FORCES), DIABETES, SMOKING AND DRINKING, DRUG USE (STEROIDS) AND RADIATION EXPOSURE. ACCORDING TO THE LITERATURE, MYOPIA PATIENTS ARE AT EARLIER AGE OF DEVELOPING CATARACT THAN PEOPLE OF NORMAL REFRACTION. HIGH DEGREE MYOPIA PATIENT COULD BE 10 YEARS YOUNGER THAN NORMAL PEOPLE TO GET CATARACT. GIVEN THE AGE AND HIGH DEGREE OF MYOPIA (-15D) OF THIS PATIENT IT IS LIKELY THAT THE MEDICAL DEVICE IS NOT THE ROOT CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577133&pc=MTA | 2023826-2017-00836 | 2017-04-18 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO DEVELOPMENT OF A CATARACT OR COULD BE A FILM BETWEEN NATURAL LENS AND ICL OR THE START OF CATARACTS. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2017 AND THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE REPORTER INDICATED THE EVENT WAS RELATED TO THE DEVICE. NO MEDICATIONS OR TREATMENTS WERE REQUIRED. IMPLANTED DATE: (B)(6) 2011. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426147&pc=MTA | 2023826-2017-00471 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED EYE IS WORSE FOR READING. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY REPORTED THEY CONFIRMED THE PATIENT REPORT. DATE OF LAST VISIT WAS (B)(6) 2016 AND REPORTED PATIENT HAD PRESBYOPIA, WITH A MILD CATARACT. NO MEDICATIONS OR TREATMENTS WERE REQUIRED AND THE EVENT WAS NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546330&pc=MTA | 2023826-2017-00772 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED REDUCED VISUAL ACUITY AND WAS WEARING GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THE PATIENT HAD DECREASED VISUAL ACUITY DUE TO A CATARACT HAD DEVELOPED. SURGERY IS PLANNED AT A FUTURE DATE FOR REMOVAL OF CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE FACILITY REPORTED THE PATIENT HAD LASIK ON (B)(6) 2011 AND WAS PRESCRIBED MEDICATION FOR NIGHT DRIVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323983&pc=MTA | 2023826-2019-00238 | NA | Injury | STAAR SURGICAL COMPANY AG | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2009. THE REPORTER INDICATED A CATARACT HAS DEVELOPED AND VISION IS SLOWLY CLOUDING. THE PLAN IS TO EXPLANT THE LENS BUT CURRENTLY THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: REPORT # MW5082245. AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. OCCUPATION: UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17872920&pc=MTA | 2023826-2023-04387 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS REMOVED AND REPLACED FOR A SMALLER LENGTH LENS DUE TO LENS OPACITY. THE PROBLEM WAS RESOLVED. REPORTEDLY, "QUIET ICL VAULT GOOD." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "1766" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3001" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS OPACITY (CATARACT). ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "QUIET. ICL VAULT GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873747&pc=MTA | 2023826-2018-01363 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01364 FOR FELLOW EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873772&pc=MTA | 2023826-2018-01364 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863580&pc=MTA | 2023826-2017-01430 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A CATARACT, AND REPLACED WITH AN INTRAOCULAR LENS (IOL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN SPECIMEN CUP WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS. LENS HAS AN "80" STAMPED /EMBEDDED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362600&pc=MTA | 2023826-2019-00300 | NA | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS STILL UNHAPPY WITH HIS VISION. THE DOCTOR FOUND THE START OF CATARACTS AT HIS LAST EXAM AND RECOMMENDED REMOVING THE LENS AT THE SAME TIME. PATIENT HAS SOUGHT A 2ND OPINION. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT - CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. (DEVICE MANUFACTURER DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362754&pc=MTA | 2023826-2019-00301 | NA | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT- CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE 3191: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9233429&pc=MTA | 2023826-2019-02023 | NA | Injury | STAAR SURGICAL COMPANY | 2019-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Adverse Event Without Identified Device or Use Problem | Cataract; Vitreous Floaters; Retinal Detachment; Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM FDA: # MW5089673. THE PATIENT REPORTED ON (B)(6)2019, A "SERIOUS INJURY" REPORT WITH OUTCOME OF "DISABILITY/PERMANENT DAMAGE". THE PATIENT STATED "GOT VISIAN ICL PLASTIC LENS IMPLANTS IN BOTH EYES. IN LEFT EYE OVER A YEAR LATER I HAD SYMPTOMS OF FOGGY VISION, NEW FLOATERS AND TODAY MY VISION IS ALMOST GONE. THERE'S A HUGE BLOB OF WATER BLOCKING MY SIGHT. WHEN I MOVE MY EYE I CAN SEE THE WATER MOVE." PER THE SAME REPORT CONCOMITANT MEDICAL PRODUCTS LISTED AS: "ALPHAGAN AND PRENATAL VITAMINS", WITH PATIENT PROBLEMS OF: "IMPAIRED VISION, VITREOUS FLOATERS" AND DEVICE PROBLEMS: "ADVERSE EVENTS WITHOUT IDENTIFIED DEVICE OR USE PROBLEM" WAS REPORTED. PER PATIENT EMAIL OF 30 SEP 2019 "VISIAN ICL WAS REMOVED ON THE LEFT EYE (B)(6) TO PREPARE FOR VITRECTOMY FOR RETINAL DETACHMENT REPAIR". INFORMATION RECEIVED FROM THE IMPLANTING SURGEON INDICATED THAT A 13.2MM, MICL13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017 AND REMOVED ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (NS), RETINAL DETACHMENT, BLURRED VISION, AND FLOATERS WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION/TREATMENT (PRK, ICL REMOVAL + CE/IOL, RETINAL REPAIR) WERE PERFORMED. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS BCVA WAS 20/80-, UCVA 20/200. CLEAR CORNEA, SUPERIOR PI, PC/IOL WITH MILD TO MODERATE PCO, S/P VITRECTOMY, 45 % GAS, AND RETINA ATTACHED WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). PER RETINAL DETACHMENT REPAIR SURGEON AND ICL REMOVAL AND CATARACT SURGERY SURGEON THIS INCIDENT DID NOT INVOLVE AN ADVERSE EVENT RELATED TO THE VISIAN ICL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRK TREATMENT WAS PERFORMED ON (B)(6) 2018. ICL REMOVAL AND CE/IOL PROCEDURES WERE COMPLETED DURING ONE SURGICAL EVENT BY A 2ND SURGEON. RD REPAIR WAS PERFORMED BY A 3RD SURGEON ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT # 2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159605&pc=MTA | 2023826-2018-01833 | 2018-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT; CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, INFLAMMATION, AND SHALLOWING OF ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159607&pc=MTA | 2023826-2018-01834 | 2018-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM, AND MILD MYDRIASIS, LIKELY CONTRIBUTING TO GLARE COMPLAINT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM AND MILD MYDRIASIS LIKELY CONTRIBUTING TO THE GLARE COMPLAINT. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, NARROWING OF ANGLES, SHALLOWING OF ACD, AC WASHOUT, ICL REPOSITIONING, YAG PI, IOP LOWERING DROPS AND PILLS, EXPLANT, MYDRIASIS (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878560&pc=MTA | 2023826-2022-00910 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICM5_13.2; -10.50 DIOPTER; INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED NON-CONFORMITY - UNRESPONSIVE HYPETONUS DESPITE SURGICAL PROCEDURE AND COMPLICATIONS. THE PATIENT EXPERIENCED LOSS OF BCVA; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; CORNEAL DECOMPENSATION; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THE POSTOPERATIVE NOTES INDICATED MILD TRANSPARENT CRYSTALLINE; CORNEAL EDEMA "ABOUT PRESENT PUPIL REACTIVE WAITING TO PERFORM VISUAL FIELD". IT WAS ALSO NOTED THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540579&pc=MTA | 2023826-2017-00740 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE DR. SAID THE PROBLEM IS SLOWLY GETTING BETTER. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS EXCHANGED RESOLVED THE PROBLEM FOR THE PATIENT. LENS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, DEBRIS AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942421&pc=MTA | 2023826-2021-04638 | 2021-11-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA AND THE PATIENT WAS UNHAPPY WITH THE RESULT (VA). REPORTEDLY THIS WAS DUE TO POSTOP CORNEAL EDEMA. THE PROBLEM HAS BEEN RESOLVED. POSTOP UCVA IS AS GOOD AS PREOP BCVA AND CORNEAL EDEMA IS GONE FROM THE EYE. "THE PATIENT HAS A SPECIAL ATTITUDE AND VERY HIGH EXPECTATIONS. SO SHE IS STILL NOT 100% HAPPY WITH THE COURSE OF THE SURGERY AND THE OUTCOME." Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042387&pc=MTA | 2023826-2018-01606 | 2018-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRA-OPERATIVELY, THE LENS WAS EXPLANTED. THE SURGEON REPORTS BLURRED VISION AND CORNEAL EDEMA. ADDITIONALLY STATED, THE "LENS WAS INJECTED UPSIDE DOWN WHEN THE DOCTOR IMPLANTED. THE DOCTOR FOUND CORNEAL EDEMA WHEN WAS REMOVED THE LENS." THE REPORTER INDICATED THAT THE LENS WAS INJECTED INTO THE PATIENT'S EYE IN SUCH A WAY THAT THE IRIS-FACING LENS SURFACE WAS INSERTED FACING THE CAPSULE. ALSO INDICATED IS, "THREE MINUTES HAS ELAPSED FROM THE LOADING OF THE LENS TO THE INJECTION. DURING THIS TIME THE CARTRIDGE WAS KEEPING UNDERWATER. THE INCISION SIZE WAS NARROW." THE SURGEON INDICATES THAT THE CAUSE OF THE EVENT IS NOT KNOWN. THE CORNEAL EDEMA IS RESOLVING. Manufacturer Narrative: "THE DOCTOR HAS IMPLANTED A NEW LENS ON (B)(6) 2018 BECAUSE THE PATIENT'S CORNEAL CONDITION (BLURRED VISION) HAS BEEN IMPROVED". THE PROBLEM WAS RESOLVED AFTER THE NEW LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS FOREIGN MATERIAL (FIBER) ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17393103&pc=MTA | 2023826-2023-03145 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER TORE DURING INJECTION, INTRAOPERATIVELY THE LENS WAS REMOVED WITH PATIENT INJURY OF POSTOP CORNEAL EDEMA. SURGEON ALSO NOTED PATIENT HAD BLURRED VISION. A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED ON THIS SAME DATE. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND THERE IS DEVICE CAUSALITY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749903&pc=MTA | 2023826-2020-02525 | 2020-10-01 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRIS, IRIS PROLAPSE, EDEMA, AND IRIS HERNIA. REPORTEDLY EXCESSIVE MANIPULATION OCCURRED DURING IMPLANT SURGERY WHICH PROBABLY LED TO THE SYMPTOMS. 1-WEEK POSTOP THE UCVA IS 20/20 AND THE SYMPTOMS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652094&pc=MTA | 2023826-2016-00549 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Corneal Edema; Dry Eye(s); Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO PROBLEM WITH THE CORNEA, SUCH AS CORNEAL EDEMA (SWELLING). THE PATIENT HAD EPITHELIAL BASEMENT CELL DYSTROPHY, WHERE SCRAPING OF THE CORNEA WAS NEEDED. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY CONFIRMED THE PATIENT REPORT. THE DATE OF THE LAST VISIT WAS (B)(6) 2015. THE FACILITY INDICATED THE PATIENTS CONDITION WAS PRE-EXISTING AND THE PATIENTS VISION WILL FLUCTUATE/ONGOING. PATIENTS BCVA WAS 20/20. THE PATIENT ALSO HAS DRY EYE, WHICH ALSO IS PRE-EXISTING. THE PATIENTS PROGNOSIS IS GOOD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677856&pc=MTA | 2023826-2021-04046 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Corneal Edema; Intraocular Pressure Increased; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED DURING THE SURGERY AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS SOME CORNEAL EDEMA DUE TO THE INCREASED IOP AND BLURRED EYES. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711487&pc=MTA | 2023826-2017-01177 | 2017-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, DIOPTER -18.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. FOLLOW UP ON (B)(6) 2017, THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, BLURRED AND DIMNESS OF VISION, ELEVATED IOP, IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, CORNEA OEDEMATUS. NEXT DAY, THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS REPOSITIONING, AND IRIDECTOMY WAS PERFORMED. FOLLOW UP ON (B)(6) 2017, THE PATIENT'S RIGHT EYE CORNEA GOT CLEARER WITH FIXED DILATED ATONIC PUPIL, PIGMENTATION IN ICL LENS, AND ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE IOP WAS CONTROLLED, AND ANGLES OPEN, PUPIL SEMI DILATED 4MM TO 5MM. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISISION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, ANGLE CLOSURE, DIMNESS OF VISION, SECONDARY SURGERY, YAG, LENS REPOSITIONING, IRIDECTOMY, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS EXPLANT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8903520&pc=MTA | 2023826-2019-01478 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTERD INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION WAS OBSERVED. ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. THE REPORTER STATED THAT THE PATIENT IS HAPPY WITH THE RESULTS AND DONT WANT TO CHANGE THE LENS . IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: PHOTOPHOBIA, PATIENT CODE 1791: CORNEAL OEDEMA, METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11043753&pc=MTA | 2023826-2020-03159 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Optical Nerve Damage; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND HIGH IOP (INTRAOCULAR PRESSURE). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5:THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (47MMHG), PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PIS) WAS PERFORMED. THE REPORTER STATED "THE PATIENT WAS CYCLOPLEGED WITH ATROPINE FOR AN EXTENDED PERIOD OF TIME UNTIL THE IOP COULD BE STABILIZED AND THE CORNEAL EDEMA IMPROVE WITH STEROIDS, ORAL AND TOPICAL IOP LOWERING MEDICATIONS. THIS TOOK SEVERAL WEEKS TO RESOLVE". IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/20--, VAULT:1.5CT, IOP:12MMHG). FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "HE SUBSEQUENTLY DEVELOPED UNEXPLAINED SYMPTOMATIC RIGHT SIDED OPTIC NERVE EDEMA AND WAS SENT TO THE ER FOR FULL NEUROLOGICAL AND OPHTHALMOLOGICAL EVALUATION. MRI WAS NEGATIVE EXCEPT FOR ENHANCEMENT OF THE RIGHT OPTIC NERVE POTENTIALLY CONSISTENT WITH OPTIC NEURITIS. THE HEAD NEUROLOGIST DID NOT FEEL THAT IT WAS CLASSIC NEURITIS BUT HE WAS TREATED FOR IT REGARDLESS". ADDITIONAL INFORMATION REPORTED "HE WAS GIVEN IV STEROIDS IN THE HOSPITAL AND SENT HOME WITH A FEW DAYS OF HIGH DOES ORAL STEROIDS." NO OTHER SYMPTOMS HAVE DEVELOPED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL CAUSING ANGLE CLOSURE AND PUPILLARY BLOCK. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: CORRECTED DATA: B5 - SHOULD HAVE BEEN SPECIFIED THAT A YAG PI WAS PERFORMED. H6 - 4581: ANGLE CLOSURE, PIGMENT DISPERSION, ENLARGMENT OF PI (YAG), SHOULD HAVE BEEN SPECIFIED. CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DOB: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # 431266 | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097942&pc=MTA | 2023826-2017-01842 | 2017-10-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, ANGLE CLOSURE, NARROWING OF THE ANGLE, SHALLOWING OF THE ACD, BLURRY VISION, PAIN AND CORNEAL EDEMA. PARACENTESIS WAS DONE TO LOWER THE PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/25. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366387&pc=MTA | 2023826-2018-00456 | NA | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. THE LENS WAS FOUND WITH NO VISIBLE DAMAGE. DIMENSIONAL EVALUATION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366474&pc=MTA | 2023826-2018-00455 | NA | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439246&pc=MTA | 2023826-2017-00486 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT EXPERIENCED IRIS HERNIATION, AND REDUCTION SURGERY. ON THE NEXT DAY, THE PATIENT EXPERIENCED OCULAR PAIN AND BLURRY VISION, WITH SHALLOW ANTERIOR CHAMBER, MODERATE DIFFUSE CORNEAL EDEMA AND OCULAR HYPERTENSION. THE CENTRAL PORT OF THE LENS WAS BLOCKED BY IRIS PIGMENT. IRIDECTOMY SURGERY WAS PERFORMED, AND THE PROBLEM WAS RESOLVED. ENLARGEMENT OF PI OR ADDITION OF NEW PI SURGERY WAS ALSO PERFORMED. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, WILL NOT BE RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872434&pc=MTA | 2023826-2017-01443 | 2017-08-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, ELEVATED IOP, NARROWING OF THE ANGLE, BLURRY VISION, PAIN, PUPILLARY BLOCK, AND CORNEAL EDEMA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND VIGAMOX AND PREDNISOLONE WERE USED AS STANDARD POST OPERATIVE PROTOCOL. THE PROBLEMS WERE RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658366&pc=MTA | 2023826-2017-00957 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND CORNEAL EDEMA. THE REPORTER INDICATED MULTIPLE PARACENTISIS PERFORMED AND THE PERIPHERAL IRIDECTOMIES WERE REDONE. THE PATIENT WAS PRESCRIBED MEDICATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENTS POST-OP CONDITION IS - TRACE K EDEMA, 1+ CELLS WITH SHALLOW ANTERIOR CHAMBER, DILATED PUPIL WITH CORNEAL IRIS TOUCH PERIPHERALLY, AND BEST-CORRECTED VISUAL ACUITY 20/200. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE PROBLEM WAS NOT RESOLVED AFTER EXPLANTATION OF THE LENS. THE PATIENT CONTINUED TO HAVE INTRAOCULAR PRESSURE (IOP) INCREASE POST EXPLANT OF LENS. AT ONE WEEK POST-OP, THE PATIENT UNDERWENT AHMED VALVE SHUNT INSERTION FOR IOP, WITH GLAUCOMA SPECIALIST. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11234650&pc=MTA | 2023826-2021-00225 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (43MMHG), GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PI) WAS PERFORMED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/25-, VAULT:1.0 CT, IOP:14MMHG). THE REPORTER STATED "HE DEVELOPED A STEROID RESPONSE WITH CORNEAL EDEMA (B)(6) 2020 AND WAS SWITCHED TO LOTEMAX FROM PREDNISOLONE ACETATE 1%. HAS SINCE BEEN TAPERED OFF ALL STEROIDS AND IOP LOWERING MEDICATIONS SUCCESSFULLY." IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE, SHOULD HAVE BEEN SPECIFIED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675039&pc=MTA | 2023826-2017-01045 | 2017-05-16 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -6.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS, AND SWOLLEN CORNEA. SECONDARY SURGICAL INTERVENTIONS INCLUDE LENS REPOSITIONING AND AN ADDITIONAL SUTURE. THE SURGEON STATED THAT, DUE TO CORNEAL SWELLING, IT WAS TOO SOON TO DETERMINE WHETHER THE PROBLEM WAS RESOLVED. THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: IT IS TOO SOON TO DETERMINE WHETHER THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. LENS HAS NOT BEEN RETURNED FOR EVALUATION. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. THE LENS HAS BEEN RETURNED FOR EVALUATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842746&pc=MTA | 2023826-2020-00518 | 2020-02-18 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOS. THE PROBLEM WAS RESOLVED HOWEVER, PATIENT'S CONDITION AT FINAL VISIT IS MARKED AS "MILD CORNEAL EDEMA." THE CAUSE OF THE EVENT IS REPORTED AS "ANATOMICAL." REFERENCE MFR FILE# (B)(4) FOR 2ND IMPLANT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1791 SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1, AS IT IS REPORTED ON MFR FILE# (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171627&pc=MTA | 2023826-2015-01391 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Dry Eye(s); Blurred Vision; Halo; No Code Available | P030016 | NA | 1 | Event Description: A PATIENT REPORTED THROUGH (B)(6) HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED AND IS EXPERIENCING DRY, ITCHY EYE. VISION IS BLURRY AND NOT AS SHARP. ALSO GLARE AND HALOS AT NIGHT. THE PATIENT HAS REPORTED VISION HAS IMPROVED OVER TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.2 MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT HAS DRY EYES AND HAS BEEN PRESCRIBED DROPS. AT THE POST-OP VISIT ON (B)(6) 2016, THE PATIENT STILL REPORTED HALOS/GHOSTING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW - THE PATIENT REPORTED THAT HE WAS EXPERIENCING SUBJECTIVE VISUAL DISTURBANCES ("HALOS/GLARE''), ESPECIALLY DURING NIGHT TIME, FOLLOWING ICL IMPLANTATION ONE WEEK BEFORE. NO REPORT OF SECONDARY SURGICALLY OR MEDICALLY INTERVENTION PERFORMED. LENS REMAINS IMPLANTED. ACCORDING TO REPORT, FROM THE FACILITY, DATED ON (B)(6) 2016 ICL WAS IMPLANTED BILATERALLY AT THE SAME SURGERY DATE AND THE EXACT CAUSE OF THE EVENT WAS NOT PROVIDED. THE PATIENT HAD ASTIGMATISM AS WELL AS DRY EYE SYMPTOMS SO HE WAS ADVISED TO CONSIDER PRK FOR ASTIGMATISM AND TO USE EYE DROPS MORE THAN TWICE A DAY. VISUAL DISTURBANCES HAVE BEEN IDENTIFIED AS POTENTIAL COMPLICATIONS AFTER ICL IMPLANTATION. THE OPTICAL DIAMETER OF THE ICL RANGES FROM 4.9 MM TO 5.8 MM AND THE PRECAUTIONS SECTION OF THE DFU INSTRUCTS PHYSICIANS THAT "PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE". IT FURTHER PRECAUTIONS THAT "THE EFFECT OF PUPIL SIZE ON VISUAL SYMPTOMS IS NOT KNOWN". THE DFU INDICATES THE "SMALLER THE OPTIC DIAMETER, THE GREATER THE INCIDENCE OF SUBJECTIVE PATIENT SYMPTOMS". REASSESSMENT WILL BE PERFORMED WHEN (IF) ADDITIONAL INFORMATION IS RECEIVED FROM THE IMPLANTING SURGEON. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774155&pc=MTA | 2023826-2019-01205 | 2015-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Dry Eye(s); Headache; Pain; Red Eye(s); Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S LEFTT EYE (OS) IN (B)(6) OF 2015. HYPERMETROPIA OF +1, ASTIGMATISM OF "AROUND" 0.5, PAIN, AND RED EYE IS REPORTED. LASIK SURGERY WAS PERFORMED (B)(6) 2017 WHICH RESOLVED THE PROBLEM. HOWEVER, RED EYE STILL OCCURS APPROXIMATELY EVERY 2 MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: "HEADACHE, DRY EYE AND NAUSEA IS REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. PATIENT CODE 3191: NAUSEA. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284950&pc=MTA | 2023826-2017-00182 | 2012-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED GLARE AROUND LIGHT AT NIGHT. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY REPORTED DATE OF THE LAST VISIT WAS (B)(6) 2013. THE FACILITY INDICATED THE PATIENT REPORTED DRY EYE ON (B)(6) 2012. THE CONDITION WAS RESOLVED BY (B)(6) 2012 WITH ARTIFICIAL TEARS. NO OTHER CONDITIONS WERE NOTED AND PATIENTS VISUAL ACUITY POST-OP WAS 20/25. THE FACILITY INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. Manufacturer Narrative: PATIENT AGE: (B)(6) 1982. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284953&pc=MTA | 2023826-2017-00181 | 2012-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED NIGHTTIME GLARE AROUND LIGHTS. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY REPORTED DATE OF THE LAST VISIT WAS (B)(6) 2013. THE FACILITY INDICATED THE PATIENT REPORTED DRY EYE ON (B)(6) 2012. THE CONDITION WAS RESOLVED BY (B)(6) 2012 WITH ARTIFICIAL TEARS. NO OTHER CONDITIONS WERE NOTED AND PATIENTS VISUAL ACUITY POST-OP WAS 20/25. THE FACILITY INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. Manufacturer Narrative: (B)(6). WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853774&pc=MTA | 2023826-2017-01410 | NA | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HER EYESIGHT WAS NOT AS GOOD. THE PATIENT WAS USING EYE DROPS DUE TO EYES WERE DRY. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE: OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853794&pc=MTA | 2023826-2017-01411 | NA | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER EYES HAD CHANGED. THE PATIENT WAS USING EYE DROPS DUE TO EYES WERE DRY. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE: - OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652097&pc=MTA | 2023826-2016-00547 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); Visual Impairment | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED NOT BEING ABLE TO SEE TO DRIVE AT NIGHT AND HAS HAD TO GET GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THEY COULD NOT CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(4) 2011 AND VISUAL ACUITY WAS 20/20-2. THE REPORTER INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THERE WERE NO OTHER OCULAR COMPLICATIONS AND NO MEDICATIONS OR TREATMENTS WERE REQUIRED. THE PATIENT HAD DRY EYES AND THE VISION WOULD FLUX DAILY DUE TO DRYNESS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652100&pc=MTA | 2023826-2016-00548 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); Visual Impairment | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED NOT BEING ABLE TO SEE TO DRIVE AT NIGHT AND HAS HAD TO GET GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THEY COULD NOT CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2011 AND VISUAL ACUITY WAS 20/20-2. THE REPORTER INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THERE WERE NO OTHER OCULAR COMPLICATIONS AND NO MEDICATIONS OR TREATMENTS WERE REQUIRED. THE PATIENT HAD DRY EYES AND THE VISION WOULD FLUX DAILY DUE TO DRYNESS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO. LENS REMAINS IMPLANTED. (B)(4). LENS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633180&pc=MTA | 2023826-2019-00880 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INCIDENT DATE THE SAME DAY, THE SURGEON REPORTS ENDOPHTHALMITIS. EYE DROPS AND "INTERNAL REMEDY" PERFORMED. PATIENT GIVE STEROID INSTALLATION EVERY HOUR, 6 TABLETS OF PREDNISONE FOR 5 DAYS, AND DECADRON SUBCONJUNCTIVAL INJECTION. THE INFLAMMATION WAS THEN RELIEVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL SYMPTOM OF BLURRED VISION IS REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633248&pc=MTA | 2023826-2019-00879 | 2019-02-22 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INCIDENT DATE THE SAME DAY, THE SURGEON REPORTS ENDOPHTHALMITIS AND BLURRED VISION. EYE DROPS AND "INTERNAL REMEDY" PERFORMED. PATIENT GIVE STEROID INSTALLATION EVERY HOUR, 6 TABLETS OF PREDNISONE FOR 5 DAYS. THE INFLAMMATION WAS THEN RELIEVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180427&pc=MTA | 2023826-2019-01900 | 2019-03-18 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION ONE WEEK AFTER SURGERY AND THE SURGEON SUSPECTED POSTOPERATIVE ENDOPHTHALMITIS. THE SURGEON PERFORMED AN ANTERIOR CHAMBER WASHOUT AND INFUSED THE ANTERIOR CHAMBER WITH BSS CONTAINING VEGAMOX DILUTED 20-FOLD. THE INFLAMMATION SUBSIDED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167176&pc=MTA | 2023826-2015-01272 | 2014-12-17 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Endophthalmitis; Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER -10.00 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED AN INFLAMED EYE AND BLURRED VISION. SUSPECTED ENDOPHTALMITIS WAS DIAGNOSIS AND THE LENS WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.00, (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948891&pc=MTA | 2023826-2023-04438 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Endophthalmitis; Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1. -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON 23/SEP/2023. THE LENS WAS EXPLANTED DUE TO EDOPHTHALMITIS, BLURRED VISION, AND TOXIC ANTERIOR SEGMENT SYNDROM(TASS) ON 28/SEP/2023. PATIENT WAS GIVEN INTRAVITREAL ANTIBIOTICS AND VITREOUS TAP WAS STERILE. CAUSE OF EVENT REPORTED AS UNKNOWN. REPORTEDLY, INFLAMMATION REDUCED SIGNIFICANTLY POST EXPLANT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853704&pc=MTA | 2023826-2018-01352 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Red Eye(s); Blurred Vision; Vitritis; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018 AND OBSERVED TO HAVE INTRAOCULAR INFLAMMATION AND VITRITIS 1 DAY POST-OP. UPON EXAMINATION, THE SURGEON " FOUND REACTION AND MEMBRANES IN AC" AND "WAS OBLIGATED TO KEEP PATIENT INSIDE THE HOSPITAL". THE PATIENT "UNDERGOES INTRA-VITREAL ANTIBIOTIC INJECTION" AND " THE DOCTOR IS WORRIED ABOUT ENDOPHTHALMITIS". PER (B)(6) 2018 EMAIL FROM REPORTER, THE SURGEON HAS EXPLANTED THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). "NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 11 DEC 2018. THE CULTURE IS (B)(6). PATIENT IS STABLE NOW AND THE PROBLEM IS RESOLVED AFTER INTRAVITREAL INJECTION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT- ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6)2018: "FEMALE PATIENT SEEKING REFRACTIVE SURGERY RT UNEVENTFUL ICL IMPLANTATION WAS DONE ROUTINE POST OPERATIVE TTT VIGAMOX ED 1×5 ECONOPRED ED 1×5 2ND DAY POST OPERATIVE THE PATIENT PRESENTED WITH DIMINUTION OF VISION, PAIN AND RED EYE. BY EXAM CILIARY INJECTION, HYPOPYON IN AC, HIGH IOP 32, PUPILLARY MEMBRANE. FUNDUS NOT SEEN. ADMISSION OF THE PATIENT U/S WAS FREE. TREAT AS TASS FREQUENT TOPICAL STEROIDS, HOURLY TOPICAL ANTIBIOTICS, CYCLOPLIC, TOPICAL ANTIGLAUCOMA. BY THE END OF THE DAY NO IMPROVEMENT EXCEPT THE IOP DECREASED U/S SHOWED VITRITIS INTRAVITREAL VANCO AND CEFTAZIDIME 1MG WAS GIVEN 2ND DAY THE HYPOPYON INCREASED AND NO IMPROVMENT AC WASH WAS DONE BUT THE MEMBRANES WERE BEHIND ICL SO I DECIDED TO EXPLANT THE ICL AND PEELING THE MEMBRANES IN FRONT OF THE CRYSTALLINE LENS BUT THE YELLOW REFLEX WAS STILL PRESENT A VITRORETINAL COLLEAGUE STARTED TO PUT PPV SET EXAMINED THE POSTERIOR SEGMENT. WE FOUND DENSE INTRAVITREAL MEMBRANES; VITREOUS SAMPLE WITH THE AQUEOUS SAMPLE WAS TAKEN AND SENT TO MICROBIOLOGISTS. PPV WAS DONE AND INTRAVITREAL VANCO AND CEFTAZIDIME WAS GIVEN AT THE END OF THE SURGERY. THE PATIENT IMPROVED 2ND POST OPERATIVE AFTER 1 WEEK SHE REGAINED HER PREOP BCVA." EVENT- ADDITOINAL INFORMATION RECEIVED ON 30 NOV 2018 FROM REPORTER: A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6)2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SUSPECTED ENDOPHTALMITIS AND INFLAMMATION. MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WAS REPORTED. PATIENTS CONDITION AT LAST VISIT DATE: (B)(6)2018, RECOVERED POST OPERATIVE BCVA. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION IN COMPLAINT FORM: " IT IS INDICATED THAT ITS A CASE OF TASS WITH HYPOPYON IN AC AND PUPILLARY MEMBRANES, INTRAVITREAL MEMBRANES WAS ALSO FOUND AND CASE IMPROVED AFTER PPV WITH INTRAVITREAL VANCOMYCIN AND CEFTAZIDIME INJECTION. PATIENT CODE:(B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8724361&pc=MTA | 2023826-2019-01109 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Endophthalmitis; Inflammation; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATES THAT THE SURGEON IMPLANTED A 13.2MM, VICM5-13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON FOUND HIGH VAULT AND INFLAMMATION IN THE EYE, THE LENS WAS EXPLANTED, AND IRRIGATION OF THE ANTERIOR CHAMBER WAS PERFORMED. THE EXPLANTED LENS WAS SENT TO AN EXTERNAL TESTING LABORATORY FOR INSPECTION BECAUSE OF THE INFLAMMATION THAT THE EYE DEVELOPED, THE DOCTOR CONFIRMED THAT THERE WAS NO PROBLEM WITH THE LENS. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. PATIENTS CURRENT CONDITION IS GOOD. Manufacturer Narrative: LENS WAS RETURNED IN A STUART TRANSPORT SWAB WITH A GEL LIKE SUBSTANCE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826223&pc=MTA | 2023826-2016-00987 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Endophthalmitis; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE PATIENT EXPERIENCED AN ELEVATED IOP, SUSPECTED ENDOPHTHALMITIS, PIGMENT DISPERSION, AND BLURRED VISION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT HAD LOW CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY AND AN ANTERIOR CHAMBER DEPTH LOWER THAN 3.00MM. CORRECTED DATA: THERE WAS NO ENDOPHTHALMITIS OR ELEVATED INTRAOCULAR PRESSURE RISE OBSERVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278780&pc=MTA | 2023826-2019-02089 | 2019-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Intraocular Pressure Increased; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019. ELEVATED IOP (29.5 MMHG) WAS ASSESSED ON (B)(6) 2019. TREATMENT OF STEROID, PREDONIN, DIAMOX, EYE DROPS, AND INTERNAL MEDICINE. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, 'AFTER 3 DAYS FROM SURGERY, PATIENT CLAIMED CLOUDY VISION. FLARE VALUE HIGH. REVOFLO, AND SANBETA PRESCRIBED. AFTER 4 DAYS FROM THE SURGERY, CONFIRMED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) SYMPTOMS AND IOP WAS HIGH (29.5). ENFORCED STEROID AND STARTED INTERNAL USE OF PREDONIN, AND DIAMOX. AS OF (B)(6), THE INFLAMMATION HAS GOTTEN BETTER." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019." SHOULD BE REMOVED FROM THE INITIAL MDR. H6- PATIENT CODE 1835 IS NOT APPLICABLE IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- IT WAS REPORTED THAT AFTER 12-15 DAYS FROM SURGERY FIBRIN WAS RECOVERED, DEPOSIT WAS OBSERVED ON THE SURFACE OF ICL LENS, HYPEREMIA WAS RECOVERED AND AN IOP 14 MMHG WAS OBSERVED. AFTER A MONTH FROM SURGERY FIBRIN LIKE TURBIDITY APPEARED AGAIN AND BETAMETHASONE AND STEROID APPLIED ON PATIENT. AFTER 1.5 MONTH FROM SURGERY FIBRIN LIKE TURBIDITY GOT BETTER. THE PATIENT WAS CHANGED TO FLUOROMETHOLONE 0.1% X 4 TIMES PER DAY FROM BETAMETHASONE X4 TIMES PER DAY. AFTER 3 MONTH FROM SURGERY SLIGHT PIGMENT OBSERVED ON ICL BUT LESS INFLAMMATION. THE PATIENT WAS PRESCRIBED A DECREASED DOSE FREQUENCY TO FLUOROMETHOLONE 0.1% X 2 TIMES PER DAY. AFTER 4 MONTH FROM SURGERY THE PIGMENT OBSERVED ON ICL DECREASED AND THER WAS NO INFLAMMATION. THE PATIENT'S DOSE FREQUENCY WAS DECREASED TO FLUOROMETHOLONE 0.1% X 1 TIME PER DAY. THE PATIENT WENT THROUGH 4 MONTHS OF MEDICAL TREATMENT AND THEN FULLY RECOVERED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8315399&pc=MTA | 2023826-2019-00232 | 2018-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Iritis; Retinal Detachment; Uveitis; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOSS OF BCVA (POST-OP BCVA HM+), BLURRED VISION, UVEITIS/IRITIS AND TASS WAS REPORTED. AQUEOUS TEST REPORT WAS NEGATIVE FOR ENDOPHTHALMITIS. AC WASH, "INTENSIVE MEDICAL TREATMENT" (ANTIBIOTICS AND STEROIDS), AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WERE PERFORMED. IT WAS REPORTED THAT WITH INTENSIVE STEROIDS, THERE IS VERY GRADUAL DECREASE IN VITREOUS INFLAMMATORY EXUDATES. HOWEVER, THERE IS NOT MUCH OF VISION IMPROVEMENT AS OF NOW. B SCAN SHOWS RESOLVING VITREOUS EXUDATES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2019; OUR INITIAL IMPRESSION WAS TASS. BUT BY THE NEXT COUPLE OF DAYS,WE REALIZED IT WAS ENDOPHTHALMITIS OF THE LEFT EYE. THE PATIENT WAS TREATED AGGRESSIVELY IN THE EYE. THE INFECTION CAME UNDER CONTROL BUT UNFORTUNATELY,BY THE TIME IT WAS CONTROLLED, SHE DEVELOPED ENDOPHTHALMITIS DUE TO CONTRACTING VITREOUS. RETINAL DETACHMENT SURGERY WAS DONE. HER VISION NOW IS HAND MOVEMENTS. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2019; PATIENT HAS BEEN REFERRED TO A POSTERIOR SEGMENT SURGEON AT SANKARA EYE HOSPITAL, BANGALORE. RETINAL DETACHMENT SURGERY WAS PERFORMED. ORIGINAL ICL HAD BEEN REMOVED. PRESENT PATIENT VISION IS HAND MOVEMENTS ONLY. FURTHER INFORMATION WAS REQUESTED BUT HAS NOT YET BEEN PROVIDED. PATIENT CODE ADDITION INFORMATION - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANT, RETINAL DETACHMENT SURGERY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS, AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725776&pc=MTA | 2023826-2023-01250 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Endophthalmitis; Red Eye(s); Uveitis; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO ENDOPHTHALMITIS WAS OBSERVED. PATIENT HAD RED EYE AND BLURRED VISION. DIAGNOSTICS WERE PERFORMED FOR UVEITIS AND INFECTION. TREATMENT WAS PERFORMED - A/C IRRIGATION, PREDNILONE EYE DROPS (6.5ML) AND VIGAMOX 0.5% (5MIL). THE PATIENT WAS TREATED AND THE EYE HAS RECOVERED. THE LENS WAS SENT TO A LAB FOR BACTERIAL CULTURE TEST AND AND NO PRESENT OF BACTERIA WAS FOUND ON THE LENS. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - PREDNILINE EYE DROPS (6.5ML); VIGAMOX 0.5% (5ML). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - A/C IRRIGATION. (B)(4). H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6562081&pc=MTA | 2023826-2017-00789 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Endophthalmitis; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -8.00 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT IS EXPERIENCING PHOTOPHOBIA AND "SEEING THROUGH A DIRTY LENS." UPON EXAMINATION, THE SURGEON NOTES WHITISH DEPOSITS ON THE ANTERIOR AND POSTERIOR FACE OF THE LENS. THERE IS SUSPECTED ENDOPHTALMITIS. CULTURES ARE PENDING LAB ANALYSIS RESULTS. AS OF THE DATE OF MDR SUBMISSION, LENS REMAINS IMPLANTED. Manufacturer Narrative: SAMPLES WERE TAKEN FROM THE EYE. AN ANTIBIOTIC INJECTION WAS ADMINISTERED. CULTURES AND PCR TEST RESULTS WERE NEGATIVE. THE INFILTRATES HAVE DISAPPEARED, THOUGH IT IS POSSIBLE THAT SOME RESIDUAL REMAINS, ACCORDING TO CLAIM REPORTER. THE PATIENT IS SATISFIED AND THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12104313&pc=MTA | 2023826-2021-02224 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Injury; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT , TRANSIENT LOSS OF BCVA, AND BLEEDING ON THE BACK OF THE ICL WAS OBSERVED. MEDICAL INTERVENTION WAS PRESCRIBED (LINDERON 0.1%, EYE DROPS 7 TO 8 TIMES, ADNA TABLETS 30 MG TAKEN 3 TIMES). IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR ADDITIONAL INFORMATION THE REPORTER STATED " ACCORDING TO PATIENT COMMENTS, AFTER WAKING UP ON APR 19, HER VISION APPEARED TO BE VEILED, BUT THE SYMPTOMS DISAPPEARED TWO HOURS LATER. AFTER THAT, HER EYESIGHT BECAME HARD TO SEE FROM THE EVENING". IT WAS REPORTED THAT THERE ARE NO PLANS TO REPLACE THE LENSES BECAUSE THERE IS NO PROBLEM WITH THE EYESIGHT. CAUSE OF EVENT WAS UNKNOWN " DOCTORS SUSPECT THAT THE PATIENT SUFFERED A BRUISE WHILE SLEEPING." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983342&pc=MTA | 2023826-2022-02208 | 2022-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Fatigue; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE, BLURRED VISION AND VISUAL FATIGUE. IT NOTES "PATIENT PRESBYOPIA +1.25". ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14046187&pc=MTA | 2023826-2022-01093 | 2022-03-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Inflammation; Intraocular Pressure Increased; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND AC INFLAMMATION CELLS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- CORRECTED TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLNATABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), AC INFLAMMATION CELLS, PAIN, HEADACHES, NAUSEA, NUMBNESS OF LEG, GENERALIZED BODY WEAKNESS, INFLAMMATION, PIGMENT CELLS, MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; NUMBNESS OF LEG. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MED WATCH REPORT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590903&pc=MTA | 2023826-2020-02255 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPLANT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590933&pc=MTA | 2023826-2020-02256 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPALNT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE ON THE LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048692&pc=MTA | 2023826-2022-04678 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Glaucoma; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT REPORTED EYE PROBLEMS AND POSSIBLE LOSS OF VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY, "AS SOON AS I WOKE UP FROM ANESTHESIA I HAD UNBEARABLE PAIN IN THE LEFT EYE WITH VOMITING AND NAUSEA. IT TURNED OUT THAT I HAD AN ATTACK OF OCULAR HYPERTONIA. THE OPHTHALMOLOGIST PUT ME ON A MANNITOL INFUSION AND GAVE ME DIAMOX TABLETS WITH DROPS TO LOWER THE OCULAR TENSION WITHOUT RESULTS." ON (B)(6)2022, AN OPTICAL COHERENCE TOMOGRAPHY FOUND, "IRRIDOCORNEAN ANGLE IS CLOSED (2 DEGREES). THE LENS WAS EXPLANTED AND THE REPORTER STATED, "I HAD CORNEAL EDEMA, A LOSS OF ENDOTHELIAL CELLS (I ONLY HAVE 1000 AT THE MOMENT) AN OPACITY OF THE LENS (GLAUCOMAFLECKEN BECAUSE OF THE OCULAR HYPERTENSION CRISIS) PUPILLARY MYDRIASIS AND GLAUCOMA WITH SIGNIFICANT LOSS OF OPTIC NERVE FIBERS IN THE EYE." THE RIGHT LENS REMAINS IMPLANTED AND IS CAUSING CHRONIC PAIN AND DOUBLE VISION, AND WITH THE LEFT EYE DAMAGE THE REPORTER STATED, "SUFFERING ANISOMETROPIA DUE TO THE DIFFERENCE OF VISION." REPORTEDLY, "WE HAVE TWO HYPOTHESES TO KNOW WHAT HAPPENED TO MY LEFT EYE EITHER THE IMPLANT WAS BIGGER THAN MY EYE IN THIS CASE THE OPHTHALMOLOGIST HAD TO CHANGE THE IMPLANT IMMEDIATELY OR IT'S DUE TO OBSTRUCTION OF THE CENTRAL AQUAPORT WITH VISCOELASTIC AND INFLAMMATORY DEBRIS." H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT - 727444 FOR FELLOW RIGHT EYE (OD). CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. MANUFACTURE DATE: UNK. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179666&pc=MTA | 2023826-2016-01773 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED A HALO FOGGY RING AROUND LIGHTS WHEN IT IS DARK, CAUSING ISSUES WHEN DRIVING. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. THE PATIENT HAS NOT SEEN A DOCTOR IN OVER THREE YEARS AND NO MORE INFORMATION WAS AVAILABLE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179667&pc=MTA | 2023826-2016-01774 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED A HALO FOGGY RING AROUND LIGHTS WHEN IT IS DARK, CAUSING ISSUES WHEN DRIVING. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. THE PATIENT HAS NOT SEEN A DOCTOR IN OVER THREE YEARS AND NO MORE INFORMATION WAS AVAILABLE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375453&pc=MTA | 2023826-2023-05764 | 2023-11-13 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT HAS LOW VAULTING WITH ROTATION, THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE, AND THE PATIENT HAS GLARE/HALOES. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347362&pc=MTA | 2023826-2023-03025 | 2023-01-08 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAD GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY INDICATED THERE WILL BE NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273280&pc=MTA | 2023826-2023-02720 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -18.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THE NEXT DAY AFTER IMPLANT THE PATIENT SAW A SECTOR OF HALOS, HIS INITIAL LENS IMPLANT OF DOUBLE VISION DISAPPEARED, AND HE FELT MUCH BETTER. THE REPORTER INDICATED HIS HALO SYMPTOM ALMOST DISAPPEARED AND IT'S UNLIKELY AN EXCHANGE WILL OCCUR. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241368&pc=MTA | 2023826-2021-02899 | 2021-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT COMPLAINED OF MILD HALOS AND THE LENS REMAINED IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02043. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538491&pc=MTA | 2023826-2022-01436 | 2022-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL 13.2 IMPLANTABLE COLLAMER LENS OF -14.5 DIOPTER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS. "PATIENT IS BOTHERED BY NIGHTTIME HALOES." THE LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-01426 FOR THE INITIAL IMPLANT LENS. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584545&pc=MTA | 2023826-2022-03620 | 2022-09-13 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "DYSHOTOPSIA FOR AQUAPORT". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748577&pc=MTA | 2023826-2023-01351 | 2022-12-07 | Injury | STAAR SURGICAL COMPANY | 2023-04-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -3.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO HALO AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. FOR CAUSE DETAILS THE REPORTER INDICATED "THE PATIENT WAS NOT COMFORTABLE WITH HIS VISION". Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253088&pc=MTA | 2023826-2023-02690 | 2023-03-30 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO DISSATISFACTION AND GLARE/HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466238&pc=MTA | 2023826-2022-01633 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS; -14.5 DIOPTER INTO THE PATIENT'S EYE. 6 WEEKS POST OP REPORTS A HALO EFFECT IN ENVIRONMENTS THAT ARE NOT BRIGHTLY ILLUMINATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16080994&pc=MTA | 2023826-2022-04797 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | NA | 1 | Event Description: THE REPORTER (CONSUMER) INDICATED THAT AN EVO ICL WAS IMPLANTED INTO THE PATIENTS EYE AND "HALOS ARE SO BAD AT NIGHT."NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND MEDICATION WAS PRESCRIBED. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." IT WAS ALSO REPORTED "THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDIITONAL DATA: B5- "THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." THE LENS REMAINS IMPLANTED." H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND MEDICATION WAS PRESCRIBED. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." IT WAS ALSO REPORTED "THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". IT WAS LATER REPORTED " I HAD AN APPOINTMENT TODAY AND THEY SAID THE LENS' LOOK GOOD AND I COULD GET ANOTHER OPINION FROM ONE OF THE SURGEONS BUT THEY SUGGESTED THAT I KEEP THE LENS' AS IS AND USE GLAUCOMA DROPS TO SHRINK MY PUPIL TO HELP GET RID OF THE HALOS." THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152675&pc=MTA | 2023826-2023-00105 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO HALOS. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOES. REPORTEDLY "NO REPLACEMENT LENS WAS IMPLANTED". CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152716&pc=MTA | 2023826-2023-00104 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO HALOS. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.00 INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOES. REPORTEDLY "NO REPLACEMENT LENS WAS IMPLANTED". CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748170&pc=MTA | 2023826-2023-01388 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER (CONSUMER) INDICATED THAT A 13.2MM VICM5_13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 29-JUN-2022. REPORTER (CONSUMER) EXPERIENCED HALOS, HE STATES " ITS REALLY BAD AT NIGHT". LENS REMAINS IMPLANTED. REPORTER STATED " RIGHT NOW THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: IT WAS LATER REPORTED THAT THE PATIENT HAD AN APPOINTMENT ON (B)(6) 2023. THE REPORTER STATED " THEY SAID THE LENS' LOOK GOOD AND I COULD GET ANOTHER OPINION FROM ONE OF THE SURGEONS BUT THEY SUGGESTED THAT I KEEP THE LENS' AS IS AND USE GLAUCOMA DROPS TO SHRINK MY PUPIL TO HELP GET RID OF THE HALOS". LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802022&pc=MTA | 2023826-2023-01515 | 2023-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE VAULT IS STABLE BUT PATIENT IS FEELING MYOPIC. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- PHOTOPHOBIA. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.0D INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT WAS FEELING MYOPIC, THE PATIENT ALSO EXPERIENCED GLARES/HALOES AND PHOTOPHOBIA. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - MYOPIC. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802019&pc=MTA | 2023826-2023-01489 | 2023-03-27 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THAT THE PATIENT HAD A SLIGHT INCREASE IN VAULT OVER A FEW DAYS AND HYPEROPIA. THE LENS WAS REPOSITIONED IN THE VERTICAL POSITION; FROM 30 DEGREES TO 90 AND VAULT IS AT 480UM. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. H6: HEALTH EFFECT- CLINICAL CODE: 4581- PHOTOPHOBIA. B5: THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER +4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, HYPEROPIA, GLARES/HALOES, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND PHOTOPHOBIA. ON (B)(6) 2023 THE LENS WAS REPOSITIONED VERTICALLY. REPORTEDLY "ICL STILL TILTED. STILL VERY MUCH BOTHERED BY PHOTOPHOBIA". THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA, SLIGHTLY TILTED ICL". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, HYPEROPIA, GLARES/HALOES, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND PHOTOPHOBIA. ON (B)(6) 2023 THE LENS WAS REPOSITIONED VERTICALLY. REPORTEDLY "ICL STILL TILTED. STILL VERY MUCH BOTHERED BY PHOTOPHOBIA". THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA, SLIGHTLY TILTED ICL". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770483&pc=MTA | 2023826-2023-04112 | 2023-08-04 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LENS ROTATION, HALOS, STARBURST, DIPLOPIA, THE LENS WAS REPOSITIONED ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED AND IT WAS NOTED "VAULT IS HIGHER BUT OK, LENS IS CENTERED". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR WITH NO DEVICE CAUSALITY AND NOTED "DESCENTRATION CAUSED SHORT LENGTH". Manufacturer Narrative: CORRECTED INFORMATION: B5 - DECENTRATION SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DECENTRATION CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - STARBURST, DIPLOPIA, CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579377&pc=MTA | 2023826-2016-00469 | NA | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; Discomfort; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT HAVING DISCOMFORT WITH GLARES AND HALOS. THE REPORTER STATED THE VAULT AND OUTCOME WAS GOOD BUT THE PATIENT REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579747&pc=MTA | 2023826-2016-00470 | NA | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; Discomfort; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT HAVING DISCOMFORT WITH GLARES AND HALOS. THE REPORTER STATED THE VAULT AND OUTCOME WAS GOOD BUT THE PATIENT REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16690916&pc=MTA | 2023826-2023-01210 | 2022-11-21 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO_12.1 IMPLANTABLE COLLAMER LENS -15.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAUL WITH ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTICS TORN/BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2227 SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3001" AND "2923" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831218&pc=MTA | 2023826-2023-04210 | 2024-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS EYE ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO OVER CORRECTION. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734732&pc=MTA | 2023826-2017-01234 | NA | Injury | STAAR SURGICAL COMPANY | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO THE PATIENT WAS UNHAPPY WITH THE LENS DUE TO PERSISTENT GLARE AND HALOS. THE LENS WAS EXPLANTED WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAS RETURNED TO WEARING CONTACT LENSES AND IS DOING WELL. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6556157&pc=MTA | 2023826-2017-00720 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HALOS. THERE WAS NO REPORTED PATIENT INJURY OR SURGICAL INTERVENTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THEY IMPLANTED BILATERAL LENSES, WITH THE EXPECTATION OF THE LENSES BEING UNDERCORRECTED AND THEY PLANNED TO DO LASIK TO CORRECT THE VISION. IT WAS A 2 STEP PROCESS, BUT DUE TO A LANGUAGE BARRIER, THE PATIENT WAS UNCOOPERATIVE AND COMPLAINED OF HALOS. HE WANTED THE LENSES REMOVED. THE LENSES WERE EXPLANTED, WITH NO PATIENT INJURY AND THE PATIENT IS DOING WELL, VISION IS RECOVERING. THE LENSES WERE NOT EXCHANGED FOR OTHER LENSES. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6556158&pc=MTA | 2023826-2017-00719 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HALOS. THERE WAS NO REPORTED PATIENT INJURY OR SURGICAL INTERVENTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THEY IMPLANTED BILATERAL LENSES, WITH THE EXPECTATION OF THE LENSES BEING UNDERCORRECTED AND THEY PLANNED TO DO LASIK TO CORRECT THE VISION. IT WAS A 2 STEP PROCESS, BUT DUE TO A LANGUAGE BARRIER, THE PATIENT WAS UNCOOPERATIVE AND COMPLAINED OF HALOS. HE WANTED THE LENSES REMOVED. THE LENSES WERE EXPLANTED, WITH NO PATIENT INJURY AND THE PATIENT IS DOING WELL, VISION IS RECOVERING. THE LENSES WERE NOT EXCHANGED FOR OTHER LENSES. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304844&pc=MTA | 2023826-2015-01611 | 2012-02-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4702503&pc=MTA | 2023826-2015-00447 | 2014-09-09 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT REPORTED GLARE AND HALOES. AN ICL EXCHANGE IS PLANNED FOR (B)(6) 2015. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW-REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536611&pc=MTA | 2023826-2016-00357 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARES AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536840&pc=MTA | 2023826-2016-00358 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARES AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. EXPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # 708676 | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5009187&pc=MTA | 2023826-2015-01053 | 2014-12-30 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE PATIENT EXPERIENCED GLARES/HALOS. THE LENS WAS REMOVED AND REPLACED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -14.50. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7248713&pc=MTA | 2023826-2018-00206 | NA | Injury | STAAR SURGICAL COMPANY | 2018-02-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THAT HALOES AND INCREASE IN FLOATERS WERE NOTICED. ATTEMPT TO OBTAIN ADDITION INFORMATION WAS UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THAT HALOS AND INCREASE IN FLOATERS WERE NOTICED. THE PATIENT ALSO USED "EYE DROPS TO CONSTRICT THE PUPIL" FOR IMPROVING VISION. (B)(4) Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378433&pc=MTA | 2023826-2018-00469 | NA | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) APPROXIMATELY 2 YEARS AGO. THE PATIENT COMPLAINED OF HALOS AND STATED SHE SEES THE AQUAPORT. SHE THEN "DISAPPEARED" FROM HIS CONSULTATION. SHE RETURNED STATING SHE CAN NO LONGER TOLERATE THE HALOS, ESPECIALLY AT NIGHT, AND ON (B)(6) 2018 THE LENS WAS EXPLANTED. REPORTEDLY, THE PROBLEM IS RESOLVED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, ABLE TO SEE AQUAPORT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6286637&pc=MTA | 2023826-2017-00151 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A BENT CONDITION IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. AS OF (B)(6) 2017, THE PATIENT IS WELL ON TOWARDS RECOVERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870494&pc=MTA | 2023826-2016-01060 | 2015-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED GLARE AND HALOS AND THE PLAN IS TO EXPLANT THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY IS 20/20 AND UNCORRECTED VISUAL ACUITY IS 20/25. THE PATIENT PLANS TO EXCHANGE THE LENS AT A LATER DATE. DEVICE EVALUATED BY MANUFACTURER? LENS EXPLANTED BUT NOT RETURNED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORT SOURCE WAS INDICATED TO BE DISTRIBUTOR, NOT USER FACILITY. (B)(4). CLAIM #: (B)(4). LENS EXPLANTED BUT NOT RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923615&pc=MTA | 2023826-2016-01217 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE AND HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. PERFORMED VISUAL INSPECTION AND FOUND NO VISIBLE DAMAGE ON LENS AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926685&pc=MTA | 2023826-2016-00681 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND THAT THERE WERE SPOTS ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.83MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091629&pc=MTA | 2023826-2016-01605 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: REPORT SOURCE: USER FACILITY WAS INCORRECT; REPLACED BY DISTRIBUTOR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). METHOD: WORK ORDER SEARCH PERFORMED AND NO SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933380&pc=MTA | 2023826-2015-00904 | 2014-04-04 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - GLARES AND HALOS MAY BE NOTED BY PATIENTS EVEN IF THEY NEVER HAD ANY OCULAR SURGERY. THIS IS COMMONLY DUE TO ABERRATIONS OR IRREGULARITIES IN THE CORNEA. IT MAY BE NOTED MORE AT LOW LIGHT CONDITIONS WHEN THE PUPIL IS DILATED. IN ICL SURGERY, THE CENTRAL AND MID PERIPHERAL CORNEA IS NOT TOUCHED, THEREFORE PATIENTS WHO HAVE GLARES AND HALOS BEFORE THE SURGERY, WILL COMMONLY RETAIN THESE SYMPTOMS BUT LITTLE OR NO INCREASE IN SEVERITY SHOULD BE EXPERIENCED SINCE THE CENTRAL AND MID PERIPHERAL REGION REMAINS UNABLATED AND ONLY A CORNEAL INCISION WITH A MAXIMUM LENGTH OF 3.2MM AT THE TEMPORAL PERIPHERAL REGION IS MADE. GLARE AND HALOS MAY ALSO BE PERCEIVED MORE DUE TO THE INCREASED CLARITY OF VISION AFTER ICL SURGERY OR IF THE PATIENT HAS RETAINED ASTIGMATISM. ANOTHER FACTOR COULD ALSO BE THE EFFECTIVE OPTICAL CORNEAL ZONES (EOCZ) OF THE ICLS, WHICH HAVE A MAXIMUM DIAMETER OF 6.17 TO 7.30 MM DEPENDING ON THE ICL POWERS. THIS IS A DESIGN LIMITATION WHICH MAY CREATE SIMILAR SYMPTOMS OF GLARE AND HALOS DUE TO THE INTERFACE OF THE OPTICAL ZONE AND THE PATIENT'S OPTICAL FIELD OF VISION, WHICH MAY BE NOTED WHEN THE PUPIL SIZE IS GREATER THAN THE EOCZ. THE 0.36MM HOLE AT THE CENTER OF THE ICL HAS BEEN COMPARED TO THE ICL WITHOUT A HOLE WHICH SHOWED THAT MODULATION-TRANSFER FUNCTIONS (MTF) ARE SMALL AND CLINICALLY NEGLIGIBLE. GLARES AND HALOS ARE COMMONLY TEMPORARY, LASTING 1-6 MONTHS, BUT MAY ON VERY RARE OCCASIONS BECOME PERMANENT. THIS CONDITION IS CONSIDERED TRIVIAL AND TEMPORARY AND SHOULD NOT BE MDR REPORTABLE UNLESS MEDICAL OR SURGICAL INTERVENTION WAS DONE TO ELIMINATE SEVERE SYMPTOMS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701680&pc=MTA | 2023826-2017-01155 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7, DIOPTER -5.5 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, HALOES, AND WIDE PUPIL. THE SURGEON COMMENTED THAT THE PATIENT'S PUPIL > 8 MM ON (B)(6) 2017, AND PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP INDICATED NORMAL PUPIL SIZE, BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613548&pc=MTA | 2023826-2015-00328 | 2014-02-16 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARES/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); HALO; (GLARE); SIZE INCORRECT FOR PATIENT; (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS : A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676444&pc=MTA | 2023826-2016-00658 | 2016-04-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND IS PLANNED TO EXCHANGE THE LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREATIVE PUPIL (FIXED), AND GLARE/HALOES. THE PATIENT ALSO HAD YAG AND IS ON PILOCARPINE TREATMENT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/40. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371940&pc=MTA | 2023826-2016-00065 | NA | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9751864&pc=MTA | 2023826-2020-00384 | 2020-01-10 | Injury | STAAR SURGICAL COMPANY | 2020-02-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6)2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. REPORTER STATED "PATIENT HAD MASSIVE HALOS. CANCELLED SURGERY FROM 2ND EYE. COMES TO RLE WITHIN NEXT YEARS." THE REPORTER STATES THAT THE CAUSE OF THIS EVENT WAS DUE TO LIGHT SENSATION FROM AQUAPORT. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794397&pc=MTA | 2023826-2020-00433 | 2019-10-23 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO A PATIENT RELATED FACTOR, THE REPORTER STATED " PT. HAS A HUGE PUPIL (8.5MM) AND COULD NOT LIVE WITH THE HALOS. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2227 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794396&pc=MTA | 2023826-2020-00432 | 2019-10-28 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO A PATIENT RELATED FACTOR, THE REPORTER STATED " PT. HAS A HUGE PUPIL (8.5MM) AND COULD NOT LIVE WITH THE HALOS. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2227 SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8997601&pc=MTA | 2023826-2019-01704 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo; No Known Impact Or Consequence To Patient | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS DECENTERED, WITH LENS ROTATION OF 160 DEGREES. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON DOES NOT PLAN TO EXPLANT/EXCHANGE THE LENS AT THIS TIME, THE HALOS ARE IMPROVING AND THE SURGEON WANTS TO WAIT. DATE OF REPORT: 20-MAR-2019. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9886971&pc=MTA | 2023826-2020-00697 | 2020-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICM_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS AND IT IS REPORTED THAT "PATIENT STILL HAVE HALOES." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. REFERENCE MFR# 2023826-2019-02021 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6560720&pc=MTA | 2023826-2017-00791 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS OFF-TARGET REFRACTION (REFRACTIVE SURPRISE) AND LOW VAULTING WITH THE LENS. THE PATIENT WAS EXPERIENCING HEADACHES. THE LENS REMAINS IMPLANTED BUT WILL BE EXPLANTED AT A LATER DATE. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11510275&pc=MTA | 2023826-2021-00681 | 2021-01-11 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED FOR CAUSE OF EVENT, THE REPORTER STATED "PATIENT EXPERIENCED HEADACHE." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440020&pc=MTA | 2023826-2022-01618 | 2022-04-18 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, FIXED PUPIL AND IRIS ATROPHY WAS OBSERVED. ON (B)(6) 2022 PI YAG WAS PERFORMED. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATED, "GOOD SURGICAL OUTCOME, HAPPY VISION, BUT HAS HAD EPISODES OF HEADACHE AND BLURRED VISION ABOUT 2-3 WEEKS POST-OP. IOP ON EXAM WAS NORMAL, VAULT NORMAL." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PERSISTENT, MILD DILATED, FIXED PUPIL WITH DIFFUSE TID OD ONLY. ACUITY GOOD, BUT FLUCTUATING." FOR ADDITIONAL INFORMATION THE REPORTER STATED "THIS IS THE THIRD CASE OF FIXED PUPIL WITH DIFFUSE TIDS. ALL WITHIN LAST MONTH, ALL RIGHT EYES, NO COMMON THEMES. HAPPY WITH VISION, SO NOT LIKELY TO REMOVE LENS, BUT CONCERNED ABOUT LONG TERM ISSUES (IOP,GLARE,NIGHT VISION). AFTER DISCUSSION WITH COLLEAGUES, FEEL THIS MAY BE UNUSUAL CASE OF MOXIFLOXACIN TOXICITY." Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 - SHOULD BE CORRECTED TO IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587683&pc=MTA | 2023826-2017-00874 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND SEVERE HEADACHE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: LENS HAS NOT BEEN RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE: ADDITIONAL CODE ADDED. RESULT CODE: UPDATED BASED ON ADDITIONAL NFORMATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052958&pc=MTA | 2023826-2021-02041 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES OR VA LOSS. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052969&pc=MTA | 2023826-2021-02042 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES. Manufacturer Narrative: ENLARGE PIS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030225&pc=MTA | 2023826-2022-01092 | 2022-03-12 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, AC INFLAMMATION CELLS AND IRREGULAR PUPIL. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022. AT 3 DAY POST-OP, FIBRIN CONTRACTING AS PATIENT DEVELOPED SEVERE INFLAMMATION POST EXPLANT. EARLY CATARACT FORMATION NOTED. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: B5- PATIENT ALSO EXPERIENCED HEADACHES; NAUSEA; WEAKNESS; MEDICATION REQUIRED. SHOULD HAVE BEEN INCLUDED IN INITIAL MED WATCH REPORT. H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; WEAKNESS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15964027&pc=MTA | 2023826-2022-04484 | 2022-11-08 | Injury | STAAR SURGICAL COMPANY | 2022-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965745&pc=MTA | 2023826-2022-04483 | 2022-11-08 | Injury | STAAR SURGICAL COMPANY | 2022-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -8.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7404820&pc=MTA | 2023826-2018-00508 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE SURGEON INDICATED THAT A 13.2MM MICL13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. PATIENT REPORTED TO HAVE BLURRY VISION AND HEADACHE. ON (B)(6) 2018 THE SURGEON OBSERVED EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ACD (ANTERIOR CHAMBER DEPTH). PATIENT WAS PUT ON GLAUCOMA MEDICATION AND AN ANTERIOR CHAMBER WASH OUT WAS PERFORMED. THE LENS WAS EXPLANTED ON (B)(6) 2018. GLAUCOMA DROPS WERE GIVEN AND IOP (INTRAOCULAR PRESSURE) RETURNED TO NORMAL. SURGEON REPORTED THAT THE CAUSE OF EVENT IS UNKNOWN- COMBINATION OF TRAUMA (PATIENT ACCIDENTLY HIT HER EYE WITH HAND ON (B)(6) 2018), LENS LENGTH, SMALL 2ND LPI (LASER PERIPHERAL IRIDOTOMY). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE: SECONDARY SURGICAL INTERVENTION, EXPLANT. NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781720&pc=MTA | 2023826-2020-00391 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO HEADACHES, PATIENT REPORTS HAVING SLIGHT TWITCH IN THIS EYE NO PAIN." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6324404&pc=MTA | 2023826-2017-00233 | 2016-12-28 | Injury | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pain; Pupillary Block; Blurred Vision; Vomiting; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXTREME VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: THE REPORTER INDICATED THE PATIENT CALLED POST-SURGERY WITH EXTREME PAIN, HEADACHE, NAUSEA AND VOMITING. THE PATIENTS PRESSURE WAS 50. THE WOUND WAS BURPED BUT WOULD NOT GO DOWN BELOW 37. A SECONDARY YAG PI WAS PERFORMED BUT WITH NO PRESSURE RELIEF. THERE WAS NARROWING OF THE ANGLE AND MECHANICAL ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016. THE PATIENTS CURRENT VA IS 20/20 AND IS DOING WELL. THE REPORTER STATED THE EVENT WAS PRODUCT-RELATED. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS CLEAR RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065520&pc=MTA | 2023826-2017-01781 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK) AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH (ACD). A REVISION OF THE PERIPHERAL IRIDECTOMIES WAS PERFORMED. THE PATIENT EXPERIENCED BLURRY VISUAL ACUITY, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 21. Manufacturer Narrative: (B)(4) . DATE OF EVENT: IT IS CORRECTED FROM UNK TO (B)(6) 2017. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065648&pc=MTA | 2023826-2017-01782 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK). THE PATIENT EXPERIENCED BLURRY VISION, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 18. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND PIECES OF LENS HAPTIC TORN OFF AND MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138562&pc=MTA | 2023826-2020-01184 | 2019-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (29MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138567&pc=MTA | 2023826-2020-01185 | 2019-11-05 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (27MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO ADVERSE EVENT B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT B5 - PHOTOPHOBIA WAS ALSO REPORTED. SECONDARY YAG WAS PERFORMED. H6 - HEALTH EFFECT CLINICAL CODE 1880 H6 - HEALTH EFFECT IMPACT CODE 4644 H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (SECONDARY YAG) ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582873&pc=MTA | 2023826-2022-00595 | 2022-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -12.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL# (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582877&pc=MTA | 2023826-2022-00594 | 2022-01-12 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL # (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767216&pc=MTA | 2023826-2015-00531 | 2015-03-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. AT THE SAME DAY POST-OPERATIVE EXAMINATION, THE PATIENT HAD 20/60 UNCORRECTED VISUAL ACUITY, INTRAOCULAR PRESSURE BY APPLANATION WAS 41MM HG AND LENS HAD A 200% VAULT. THE PATIENT COMPLAINED OF SLIGHT HEADACHE AND SHE WAS EXPERIENCING "STARBURSTS SIMILAR TO MIGRAINE AURA". LPI ENHANCEMENT WAS NECESSARY AND "BURPING" OF THE WOUND WAS PERFORMED. THE INTRAOCULAR PRESSURE WAS LOWERED TO SATISFACTORY LEVELS AND MEDICATION WAS PRESCRIBED. AT ONE DAY POST-OPERATIVE VISIT, THE SURGEON DECIDED THE LENS WAS TOO LARGE AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT SAME DAY POST-OPERATIVE VISIT, THE INTRAOCULAR PRESSURE WAS 44MM HG, UNCORRECTED VISUAL ACUITY WAS 20/50, WITH 1+ CORNEAL EDEMA AND 75% VAULTING OF LENS. "BURPING" OF THE WOUND LOWERED THE INTRAOCULAR PRESSURE AND PATIENT WAS PRESCRIBED MEDICATION. AT POST-OPERATIVE VISIT ON (B)(6) 2015, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, INTRAOCULAR PRESSURE WAS 18MM HG WITH CLEAR CORNEA AND LENS HAD 75% VAULT. Manufacturer Narrative: (B)(4). INCISION OPENED. SECONDARY SURGERY. NEW IRIDECTOMY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, MOST LIKELY THE EXCESSIVE VAULTING CAUSED SHALLOW ANTERIOR CHAMBER, ANGLE NARROWING OR CLOSURE, AND/OR PUPILLARY BLOCK WHICH WERE ASSOCIATED WITH INCREASED IOP. THE SECOND IOP INCREASE AFTER THE LENS EXCHANGE COULD BE CAUSED BY SURGICAL RELATED ANTERIOR CHAMBER ANGLE DISTORTION, PRIOR SUSPECT GLAUCOMA, PRIOR OPEN ANGLE GLAUCOMA OR VISCOELASTIC SUBSTANCE USED IN THE SURGERY. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DARK SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614654&pc=MTA | 2023826-2015-00344 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614667&pc=MTA | 2023826-2015-00345 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335509&pc=MTA | 2023826-2020-01564 | 2020-07-02 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Nausea; Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF ICA AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED EYE PAIN, BLURRING OF VISION AND NAUSEA. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT DOES NOT FEEL HEADACHE ANYMORE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: UNK. EXPIRATION DATE UNK. NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671658&pc=MTA | 2023826-2019-01019 | 2015-11-12 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Headache; Nausea; Pain; Red Eye(s); Blurred Vision; Burning Sensation; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2015. THE SURGEON REPORTS PIGMENT DISPERSION, RED EYE, AND PERMANENT EYE PAIN SINCE IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE SURGEON THINKS THE FOOTPLATES OF THE ICL SEEM TO BE LOW, PLACED ON/IN BETWEEN ZONULAS OR ON/IN THE CILIARY BODY-WHICH COULD CAUSE THE PAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT CODE1880: HEADACHE. CLAIM#: (B)(4). Manufacturer Narrative: PER EMAIL ON 2019/07/25. OD IMPLANT SURGERY WAS ON "(B)(6)" OR (B)(6) 2015. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: DATE IS APPROXIMATE. INITIAL OPERATION OCCURRED EITHER (B)(6) 2015. AFTER IMPLANT PATIENT REPORTS HYPERMETROPIA OF +1 AND ASTIGMATISM OF 0.75. GLASSES WERE WORN AND PARTIALLY IMPROVED THE SITUATION. REPORTEDLY LASIK WAS PERFORMED IN (B)(6) 2017 WHICH IMPROVED THE PROBLEM. HOWEVER, IN (B)(6) 2018 THE RIGHT EYE BECAME WORSE AND THE PATIENT STATES PAIN IS CONSTANT ESPECIALLY WHEN READING. THERE IS ALSO A SLIGHT BURNING BEHIND THE EYE AND A FEELING OF NAUSEA ON THE RIGHT SIDE, EVEN AT NIGHT. IMPLANT DATE: DATE IS APPROXIMATE. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LASIK), LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788401&pc=MTA | 2023826-2016-00898 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Pain; Blurred Vision; Vomiting; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS BEING TOO LARGE AND HYPOTONY (LOW IOP). THE PATIENT EXPERIENCED BLURRED VISUAL ACUITY, EYE PAIN, HEADACHE, AND VOMITING. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN A SMALL CLEAR CONTAINER. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986867&pc=MTA | 2023826-2023-04600 | 2023-01-13 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694800&pc=MTA | 2023826-2017-01121 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A LENS HAPTIC SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119129&pc=MTA | 2023826-2015-01262 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Hyphema; Hypopyon; Inflammation; Irritation; Pain; Visual Impairment | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER STATED AT ONE DAY POST-OP ((B)(6) 2015), THE PATIENT'S EYE HAD HYPHEMA AND HYPOPHON. MEDICATIONS WERE ADMINISTERED FOR INFLAMMATION BUT THE REPORTER STATED IT LOOKED MORE LIKE AN IRRITATED IRIS. THE PATIENT WAS IN PAIN AND CULTURES WERE TAKEN. THE LENS VAULT IS GOOD. THE PATIENT RETURNED FOR A FOLLOW-UP VISIT ON (B)(6) 2015 AND THE VA WAS 20/20, THERE IS STILL SOME CELL IN THE EYE AND THE VISION IS NOT AS CLEAR. THE REPORTER STATED THE TEST ON THE CULTURE CAME BACK AS TOXIC ANTERIOR SEGMENT SYNDROME. THE REPORTER STATED THE PATIENT'S EYE HAS SETTLED DOWN AND IS RECOVERING. THE PATIENT'S VA WAS 20/20. Manufacturer Narrative: (B)(4). RESULTS: PER MEDICAL REVIEW - ENDOPHTHALMITIS IS A PAINFUL INFECTION OF OCULAR TISSUES THAT MAY OCCUR POSTOPERATIVELY. UNLESS IT IS VIRAL IN ORIGIN, A CULTURE WILL USUALLY BE POSITIVE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) IS A PAINLESS INFLAMMATORY POSTOPERATIVE CONDITION THAT HAS BEEN ASSOCIATED WITH DRUGS, DEVICES, AND SURGICAL ISSUES. CULTURE IS USUALLY NEGATIVE. IT IS NOT CLEAR WHICH OF THESE CONDITIONS DEVELOPED IN THE PATIENT'S EYE, AND THE CAUSE OF EITHER CONDITION IN THIS CASE IS UNKNOWN. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): LENS WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH, FOAM TIP PLUNGER LOT NUMBER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A FOAM TIP PLUNGER LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS, INJECTOR AND CARTRIDGE THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919174&pc=MTA | 2023826-2015-00841 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Inflammation; Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.0MM ICM120 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. ON (B)(6) 2015, THE SURGEON NOTED LOSS OF BCVA, INFLAMMATION AND DEPOSITS ON THE LENS. THE LENS WAS REMOVED AND REPLACED BUT THIS DID NOT RESOLVE THE ISSUE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: P030016 TO N/A - THIS PRODUCT IS NOT MARKETED IN THE U.S. MFR DATE - 03/01/2013 TO 12/16/2008, PREVIOUS DATE IS INCORRECT. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC WAS TORN AND THE HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322791&pc=MTA | 2023826-2019-00188 | 2019-01-03 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE INFLAMMATION WAS TREATED BUT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED.. THE PATIENT IS IMPROVING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, 11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY AND UNREACTIVE (FIXED) PUPIL. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS ON (B)(6) 2019. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THERE WAS SIGNIFICANT RESOLUTION UPON REMOVAL AND THE PATIENT IS IMPROVING. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLOUDY RESIDUE ON THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. PATIENT CODE: 3191 - UNREACTIVE (FIXED) PUPIL. DATE OF BIRTH: (B)(6) 1990. EXPLANTED DATE: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322830&pc=MTA | 2023826-2019-00187 | 2019-01-03 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THE PATIENT IS IMPROVING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY AND UNREACTIVE (FIXED) PUPIL. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS ON 01/02/2019. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THERE WAS SIGNIFICANT RESOLUTION UPON REMOVAL AND THE PATIENT IS IMPROVING. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLOUDY RESIDUE ON THE LENS. CORRECTED DATA: DATE OF BIRTH - (B)(6) 1990 EXPLANTED DATE - (B)(6) 2019 CLAIM # (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) NCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180447&pc=MTA | 2023826-2019-01903 | 2019-06-06 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT HAD BLURRED VISION AND TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED IN THE PATIENT'S EYE. THE PATIENT WAS GIVEN STEROIDS AND THE SYMPTOMS GRADUALLY IMPROVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 2993 CORRECTED TO DEVICE CODE 1401 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER APPROXIMATELY 3 MONTHS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873864&pc=MTA | 2023826-2016-01066 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, GLARE, HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED IOP AND VERY SHALLOW AC. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED AND THERE WAS ENLARGEMENT OF INCISION. THE PATIENTS CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON 06/20/2016 DUE TO EXCESSIVE VAULTING, GLARE/HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, AND VERY SHALLOW ANTERIOR CHAMBER. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THERE WAS ENLARGEMENT OF INCISION. THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS AND A PARS PLANA VITRECTOMY TO DRAIN THE ACQUEOUS MISDIRECTION. THE PATIENT'S CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. USER FACILITY IS THE INCORRECT REPORT SOURCE, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873698&pc=MTA | 2023826-2016-01062 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE AND INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, GLARE/HALOS AND BLURRED VISION. THE INCISION WAS ENLARGED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE PATIENT'S EYE IS IN GOOD CONDITION AFTER EXPLANTATION AND THE INFLAMMATION WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152546&pc=MTA | 2023826-2017-01967 | 2017-08-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Inflammation; Iritis; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED, THE SURGEON REPORTING INFLAMMATION AND IRITIS. THE CLAIMANT STATES THAT THE PATIENT HAD AN INFLAMMATORY RESPONSE AFTER THE SURGERY. THE SURGEON TREATED THE IRITIS BY PRESCRIBING ANTIBIOTICS AND ANTI-INFLAMMATORY MEDICINE. REPORTEDLY, THE SURGEON LEFT THE LENS IMPLANTED SINCE VA WAS 20/20. UPON CHECK-UP THE "VA WAS FAILING" AND THE LENS WAS EXPLANTED. PROBLEM RESOLVED WITH LENS EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PMA 510(K): THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528000&pc=MTA | 2023826-2019-02499 | 2019-10-21 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED HAD AN MICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN HER LEFT EYE (OS) ON (B)(6) 2019 WITH GOOD RESULTS. THE PATIENT REPORTED ON (B)(6) 2019, AFTER A 3 HOUR FLIGHT, THERE WAS A RED RING AROUND HER LEFT EYE AND HER PUPILS WERE DIFFERENT SIZES. TWO DAYS LATER, SHE DEVELOPED PAIN AND LIGHT SENSITIVITY IN THE LEFT EYE. THE PATIENT WENT TO HER DOCTOR AND THEY SAID THE PATIENT HAD REBOUND INFLAMMATION AND WAS PUT BACK ON STEROIDS. THE PATIENT HAS REPORTED CLOUDINESS AND HALF HER VISION WENT GRAY AND SHE COULD NOT SEE. TO THIS DAY, THE PATIENT HAS 2 DIFFERENT SIZE PUPILS AND VISION IN LEFT EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED AN 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED "GRAY VISION TRANSIENT". THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983737&pc=MTA | 2023826-2022-01034 | 2022-03-07 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Inflammation; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), BLEPHARITIS, IRIS BOMBE WITH NO APPARENT TOUCH POINT, "VH 1," AND NIGHTTIME HALOS IN PERIPHERY WHEN DRIVING AND ANY CEILING LIGHTS. THE SURGEON STATES, "VERY NARROW ANGLE...ANGLE SHUT OFF." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO BIG." Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428219&pc=MTA | 2023826-2017-00467 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Infection; Visual Impairment; Vitrectomy; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BEGINNING ON (B)(6) 2017 THE PATIENT BEGAN EXPERIENCING THE FOLLOWING: LOW VAULT, BCVA OF "ONLY HAND MOTIONS," DECREASED ACCOMMODATION, AND DIFFICULTY IN NEAR VISION. THE PATIENT UNDERWENT A VITRECTOMY, THEN WAS HOSPITALIZED FOR APPROXIMATELY 1 WEEK. PATIENT HAD AN INFECTION-SUSPECTED ENDOPHTHALMITIS BUT TESTS CAME OUT NEGATIVE. TESTED POSITIVE HOWEVER, FOR "FUNGAL-LIKE" ORGANISMS AND GRAM POSITIVE COCCI. PATIENT WAS TREATED WITH IV ANTIBIOTICS AND DROPS, PPV. LENS WAS EXPLANTED (B)(6) 2017. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4). Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT AND CLEAR RESIDUE AND FIBERS ON LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795690&pc=MTA | 2023826-2017-01321 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT HAS BEEN TREATED BY MEDICATIONS: COSOPT, SPERSADEX, AND DIAMOX. MULTIPLE ATTEMPS TO GET UPDATED INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE PATIENT'S ELEVATED IOP WAS TREATED WITH MEDICATIONS: COSOPT, SPERSADEX AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795695&pc=MTA | 2023826-2017-01322 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT IOP IS 26.7MMHG, AND HAS BEEN TREATED BY TRUSOPT, AND TOBRADEX MEDICATIONS. THE SURGEON ALSO NOTED THAT THE PATIENT IOP IN TWO WEEKS POST-OP IS 20.6 MMHG. MULTIPLE ATTEMPS TO GET ADDITIONAL INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE SURGEON NOTED THAT THE PATIENT'S IOP IN TWO WEEKS POST-OP IS 20.6MM HG AFTER THE PATIENT WAS TREATED WITH MEDICATIONS: TRUSOPT AND TOBRADEX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12761919&pc=MTA | 2023826-2021-04212 | 2021-06-24 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE SURGEON ENLARGED THE PIS BY YAG DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (28MMHG) AND SLIGHT VA BLURRINESS PER PATIENT. THE IOP RESOLVED AND WAS 25MMHG SAME DAY. BY (B)(6) 2021 THE IOP WAS 14MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02386. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK EXPLANT DATE: NA (B)(4). CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343177&pc=MTA | 2023826-2023-02983 | 2023-03-14 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343354&pc=MTA | 2023826-2023-02984 | 2023-03-14 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT - CLINICAL CODE: "1937" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135946&pc=MTA | 2023826-2021-00043 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP IS 16/16MMHG AND VAULT IS STILL HIGH AT 1243. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: AS OF (B)(6) 2021, THE VAULT IS (B)(4). TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATIENT IS OK" POST EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135969&pc=MTA | 2023826-2021-00044 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP 16/16MMHG AND VAULT IS STILL HIGH AT 1184. CLAIM# (B)(4). Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF 03-FEB-2021 THE VAULT IS 1093. TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATENT IS OK." POST EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195867&pc=MTA | 2023826-2022-02598 | 2022-07-08 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. MEDICATION WAS ADMINISTERED. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THERE WAS HIGH PRESSURE, ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATROPHY AND SHALLOWING OF AC. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATOPHY, SHALLOWING OF A/C. H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL AND ORAL ANTI-OCULAR HYPERTENSIVE MEDICATION. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750641&pc=MTA | 2023826-2023-01227 | 2023-02-19 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (PD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(60MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON 19/FEB/2023), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10544158&pc=MTA | 2023826-2020-02118 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THAT THE LENS IS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED INTO THE OD EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTATION. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY AND POSTERIOR SYNECHIAELYSIS ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, ENCOURAGED TO USE +1.25 READERS TO HELP WITH COMPUTER AND PHONE. CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR, "EYE COULDN'T HANDLE DEVICE." H6: LINICAL CODE 4581 (ANGLE CLOSURE). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551350&pc=MTA | 2023826-2020-02119 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED IN THE OS EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTED. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, "GAVE SCL'S +0.75 BASE CURVE 8.4 AND WILL SCHEDULE LASIK IN THE NEAR FUTURE." H6-CLINICAL CODE 4581 (ANGLE CLOSURE) CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDU ON THE LENS AND THAT THE LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309075&pc=MTA | 2023826-2023-00445 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ELEVATED INTRAOCULAR PRESSURE; GLARE/HALOES; BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED; IT WAS ADDITIONALLY INDICATED AN INCISION ENLARGEMENT AND ADDITION SUTURE WAS PERFORMED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHALLOW VAULT DUE TO LENGTH". Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14057555&pc=MTA | 2023826-2022-01148 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. ON (B)(6)2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTER THAT INTRAOCULAR PRESSURE WAS TAPER DOWN TO 17MMHG. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- RAINBOW LIGHT. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE; BLURRED VISION AND HALOS/RAINBOW LIGHT. ON (B)(6) 2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT MEDICATION WAS USED TO LOWER IOP. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089564&pc=MTA | 2023826-2015-01252 | 2015-08-19 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT REPORTED THE SURGERY WENT WELL BUT AT ONE DAY POST-OP, HE COULDN'T SEE OUT OF HIS EYE, HE HAD TROUBLE FOCUSING AND IT WAS LIKE LOOKING THROUGH A WHITE SHEET OF PAPER. HE WENT IN FOR A FOLLOW-UP VISIT THAT SAME DAY AND IT WAS NOTED THE PERIPHERAL IRIDOTOMIES WERE NOT OPEN AND THE PRESSURE WAS UP. THE PERIPHERAL IRIDOTOMIES WERE OPENED AND THE PATIENT WAS SENT HOME. ON (B)(6) 2015 THE PATIENT HAD THE SAME PROBLEM AND HIS VISION WAS BLURRY. HE WENT BACK TO HIS DOCTOR AND THE DOCTOR DECIDED TO EXPLANT THE LENS THAT SAME DAY. HE WAS STILL HAVING PROBLEMS WITH HIS EYE AND HIS DOCTOR GAVE HIM DROPS. THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO PERSISTENT ANGLE CLOSURE. Manufacturer Narrative: PER MEDICAL REVIEW - ANGLE CLOSURE FOLLOWING PIOL IMPLANTATION IS A RARE COMPLICATION. STAAR RECOMMENDS PREVENTION BY PERFORMING TWO LASER PERIPHERAL IRIDOTOMIES BEFORE SURGERY, AND IN THIS CASE, THE PIS HAD BEEN PERFORMED BUT THEY WERE NOT OPEN. EVEN AFTER THE PIS WERE OPENED UP WITH YAG LASER, IOP ROSE AGAIN. ONE POSSIBLE MECHANISM IS INITIAL PUPIL BLOCK DUE TO PUPIL-PIOL CONTACT THAT REMAINED BECAUSE SYNECHIAE DEVELOPED. ANOTHER POSSIBLE MECHANISM IS CILIOLENTICULAR BLOCK, BUT THERE ARE NO UBM OR OCT REPORTS TO DETERMINE THE CAUSE OF PERSISTENT ANGLE CLOSURE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE INCORRECT DIMENSIONS OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED INCORRECT SIZING OF THE LENS RESULTING IN ELEVATED IOP AND PERSISTENT ANGLE CLOSURE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6715656&pc=MTA | 2023826-2017-01142 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND ELEVATED IOP. AN AHMED VALVE WAS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: PATIENT IS FINE. BCVA IS 10/10 AND IOP IS CONTROLLED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467360&pc=MTA | 2023826-2018-00604 | 2018-03-26 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. THE LENS REMAINS IMPLANTED. REPORTEDLY, TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. NO UPDATED INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. THE REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED "AROUND (B)(6)" FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: B5 - UPDATED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. REPOSITIONING WAS PERFORMED ON (B)(6) 2018 BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED (B)(6) 2018 FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." D6B - (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661714&pc=MTA | 2023826-2017-01012 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -12.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. SECONDARY INTERVENTION INCLUDES ENLARGEMENT OF INCISION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802614&pc=MTA | 2023826-2017-01350 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, THE PATIENT EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ANGLE CLOSURE WITH ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRED VISION. AT A LATER DATE THE SURGEON EXPLANTED THE LENS, CUT IT IN ORDER TO SHORTEN IT, THEN RE-IMPLANTED THE LENS BACK INTO THE EYE. IT WAS REPORTED THAT THE VAULT AND EYE ARE GOOD NOW. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: PMA 50(K): - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL, UNAPPROVED, OR CONTRAINDICATED USE (KERATOCONUS EYE). (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7441301&pc=MTA | 2023826-2018-00565 | 2017-10-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, REFRACTIVE SURPRISE, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID; IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, SIGNIFICANT REDUCTION OF ICA). (B)(4) OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.81MM)]. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113305&pc=MTA | 2023826-2017-01883 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO VAULTING AND INCREASED INTRAOCULAR PRESSURE. A SHORTER LENGTH LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED , A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM -9.5 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, ELEVATED IOP, NARROWING OF ANGLE, SHALLOWING OF ACD. THE SURGEON ALSO PERFORMED A SECONDARY YAG ON (B)(6) 2017. THE PATIENT HAD BLURRY VISION DURING THE TIME OF POST-OP. IMPLANTED DATE IS CORRECTED TO (B)(6) 2017. (B)(4). CORRECTED DATA: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND REPLACED WITH SHORTER LENS. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267534&pc=MTA | 2023826-2018-00209 | 2017-12-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTION: LENS WAS IMPLANTED (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS HAPTIC MISSING A PIECE AND RESIDUE ON LENS. CONCLUSION: OFF-LABEL USE [ANTERIOR CHAMBER DEPTH < 3.00MM (2.84MM)]. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320724&pc=MTA | 2023826-2018-00319 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), BLURRED VISION, AND IRIS ATROPHY. SURGEON ALSO STATES HE PERFORMED IRRIGATION/ASPIRATION, PHACO-EMULSIFICATION, AND IOL IMPLANTATION. PATIENT IS NOW PSEUDOPHAKIC. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND REPLACED WITH A NON-STAAR LENS. AFTER SURGERY DAY 1 THERE WAS NO ANGLE CLOSURE (ONLY ELEVATED IOP), "BUT THEN THE IRIS WAS GETTING CLOSER TO THE CORNEA." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7504918&pc=MTA | 2023826-2018-00693 | 2018-04-19 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTIONS: ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION AND ELEVATED IOP. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF A HAPTIC MISSING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5037267&pc=MTA | 2023826-2015-01139 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-08-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND INCREASED ELEVATED INTRAOCULAR PRESSURE (UP TO 56MMHG). MEDICATIONS WERE ADMINISTERED BUT DID NOT HELP AND A YAG OF THE PERIPHERAL IRIDOTOMIES WAS PERFORMED BUT THE SURGEON DECIDED TO EXPLANT THE LENS. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT HAD ADDITIONAL SYMPTOMS OF BLURRY VISION AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS A MISCALCULATION/RECOMMENDED LENS SIZE. Manufacturer Narrative: THE PATIENT HAD SHALLOW ANTERIOR CHAMBER WITH NARROWED ANGLES OS, AND INCREASED IOP RESULTING FROM AN OVER-SIZED LENS. THE SURGEON DETERMINED SIZING USING THE FDA APPROVED LABELING. THE RISE IN IOP RESOLVED WITH LENS EXCHANGE, THEREFORE THE CHANGES IN ANTERIOR CHAMBER CONFIGURATION AND INCREASED IOP WERE RELATED TO THE PRESENCE OF THE OUTSIZE LENS IN THE EYE. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5731066&pc=MTA | 2023826-2016-00819 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), BLURRED VISION, AND PERMANENTLY DILATED PUPIL. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. (B)(4). SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5681611&pc=MTA | 2023826-2016-00759 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER . THE PATIENT EXPERIENCED BLURRY VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CONDITION IS GOOD AND BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923610&pc=MTA | 2023826-2016-01211 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE SURPRISE. THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: -DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN AND BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417003&pc=MTA | 2023826-2019-02307 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION AND CONSTANT VERTICAL BLACK LINE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK ONE WAS SEEN." PER REPORTER ON (B)(6) 2019, DOCTOR HAS REQUESTED A REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. BLURRED VISION, ELEVATED IOP, AND CONSTANT VERTICAL BLACK LINE WERE OBSERVED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK LINE WAS SEEN." THE ELEVATED IOP 'WAS BOUGHT UNDER CONTROL WITH DROPS' AND 'PBI WAS ALSO DONE.' THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT THE LINE REMAINED AND BECAME HORIZONTAL. THE LENS REMAINS IMPLANTED WITH PLANNED EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATIENT CODE 3191 - SECONDARY INTERVENTION: MEDICAL INTERVENTION (DROPS), LENS REPOSITIONING, PBI. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242379&pc=MTA | 2023826-2019-00102 | 2018-12-07 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER WAS TOO LARGE. Manufacturer Narrative: ADDITONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242395&pc=MTA | 2023826-2019-00101 | 2018-12-07 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -15.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER, WAS TOO BIG. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. PATIENT CODE: "2137" NOT INCLUDED IN THE INITIAL MDR. DEVICE CODE: "1401" NOT INCLUDED IN THE INITIAL MDR. ADDITIONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269193&pc=MTA | 2023826-2019-00142 | 2019-01-11 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, LOSS OF BEST CORRRECTED VISUAL ACUITY (BCVA), ELEVATED INTRAOCULAR PRESSURE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO RECOMMENDED LENS WAS TOO LARGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002877&pc=MTA | 2023826-2020-00983 | 2019-08-01 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE IN A SECOND SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774163&pc=MTA | 2023826-2019-01188 | 2019-06-10 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT, LOSS OF BCVA, PUPIL BLOCK, ELEVATED IOP, "FLAT CAMERA," AND ANGLE CLOSURE. ELEVATED IOP WAS TREATED WITH "MEDICAL DROPS" AND PAIN PILLS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794960&pc=MTA | 2023826-2020-00430 | 2019-12-26 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. H6- DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998035&pc=MTA | 2023826-2022-00825 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXESSIVE VAULT, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, UNREACTIVE (FIXED) PUPIL OF 5MM, AND SHALLOW ANTERIOR CHAMBER WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416096&pc=MTA | 2023826-2021-00507 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." D9- CHANGED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440682&pc=MTA | 2023826-2017-00502 | 2016-12-15 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF ICA, REFRACTIVE SURPRISE, LOSS OF BCVA, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), GLARE/HALOES, BLURRED VISION, AND IRIS ATROPHY. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SAME DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: METHOD, RESULTS CODES: UPDATED: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802597&pc=MTA | 2023826-2017-01349 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER IMPLANTATION, PATIENT EXPERIENCED EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP, GLARE/HALOS, AND BLURRED VISION. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546105&pc=MTA | 2023826-2017-00768 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -10.00 DIOPTER. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9746415&pc=MTA | 2023826-2020-00379 | 2019-11-25 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, ELEVATED IOP, GLARE/HALOES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. RETURNED PRODUCT EVALUATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288508&pc=MTA | 2023826-2023-05482 | 2023-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM513.7 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PUPIL BLOCK WITH ELEVATED IOP AND GLARES/HALOES ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4) | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10816077&pc=MTA | 2023826-2020-02662 | 2020-04-07 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON 06-APR-2020. AT ONE DAY POST-OP, THE VAULT AND VISION WAS IMPROVED BUT WITH MYDRIASIS AND HALOS. THE PATIENTS INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-00945. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE PATIENT IS DOING WELL WITH MILD RESIDUAL ASTIGMATISM, THAT WE HOPE TO ADDRESS WITH POSSIBLY PRK. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515451&pc=MTA | 2023826-2020-02022 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem; No Apparent Adverse Event | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. PRESSURE SPIKES WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONA INFORMATION: B5- THE REPORTER INDICATED THAT A 13.7MM MICL 13.7 OF -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS REPORTED. PATIENT STATED "SHE NOTICED HALOES AROUND ANY LIGHTS. THE LENS REMAINS IMPLANTED; HOWEVER THE SURGEON PERFORMED ENLARGEMENT OF PI; AND IT IS REPORTED THAT THE SURGEON HAS GONE BACK AND PERFORMED A SECOND IRIDOTOMY. PATIENT STATES NO PAIN OR VISION CHANGES. THE LEFT EYE PRESSURE IS OK. D6A: IMPLANT DATE: (B)(6) 2020. H1-SERIOUS . H6- CLINICAL CODE:4581-SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. HEALTH EFFECT:4625- ENLARGEMENT OF PI; IRIDOTOMY AND SECOND IRIDOTOMY. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 IN PREVIOUS MDR THE LENS DIOPTER REPORTED -7.0 IS CORRECTED TO -11.5. CLAIM# (B)(4). Manufacturer Narrative: NO APPARENT HARM OCCURRED IN RELATION TO THE ADVERSE EVENT. UNABLE TO ENTER CODE. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11795155&pc=MTA | 2023826-2021-01396 | NA | Injury | STAAR SURGICAL COMPANY | 2021-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VICMO12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA BUT IT DID NOT RESOLVE THE PROBLEM. FOR STATUS OF THE EYE " PRESENCE OF HALOS; UNREACTIVE PUPIL" WAS REPORTED. " DOCTOR BELIEVES THAT HIGH IOP, MAY BE DUE TO CORTICOSTEROIDS OR MANIPULATION" WAS REPORTED FOR ADDITIONAL INFORMATION. ON (B)(6) 2021 THE REPORTER REPORTED "CURRENT CONDITION: PATIENT COMPLAINING OF GLARE; LITTLE REACTIVE PUPILS AND PIO 18. THE USE OF CORTICOSTEROIDS IS PART OF THE POSTOPERATIVE TREATMENT. A PROGRESSIVE REDUCTION WAS MADE AND STOPPED YESTERDAY, (B)(6) 2021. CURRENTLY USING ANTI-GLAUCOMA DORZOLAMIDE EYE DROPS AND BRIMONIDINE. THE DOCTOR WILL TRY, NEXT WEEK, TO ADAPT THE CONTACT LENSES TO REDUCE THE RESIDUAL DEGREE AND CHECK IF THE VISUAL COMPLAINT IMPROVES." ON (B)(6) 2021 THE REPORTER REPORTED " PATIENT IS A LITTLE MORE SATISFIED. LAST REVIEW ON (B)(6) 2021: DOCTOR HAD TO USE PILOCARPINE, TO CLOSE THE PUPIL A LITTLE, WHICH WAS 6MM. 0.73 VAULT. THE PATIENT REMAINS UNDER OBSERVATION AND EVALUATION WEEKLY". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT "AFTER A WHILE THE GLARE DISAPPEARED." CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955095&pc=MTA | 2023826-2016-01266 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS BROKEN. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955096&pc=MTA | 2023826-2016-01267 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT WAS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028643&pc=MTA | 2023826-2016-01460 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. FOREIGN MATERIAL - HAIR ON THE LENS SURFACE PRESENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028650&pc=MTA | 2023826-2016-01459 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND FOREIGN MATERIAL (SPOT) ON SURFACE, BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6200082&pc=MTA | 2023826-2016-01824 | 2016-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES, AND ANGLE CLOSURE WITH ELEVATED IOP. AS OF THE DATE OF MDR FILING, THE LENS HAS NOT YET BEEN EXPLANTED AND REMAINS IMPLANTED IN THE PATIENT'S EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL EVENT DESCRIPTION: LENS DIOPTER: -17.0/+1.5/70 (SPHERE/CYLINDER/AXIS). THE SURGEON ADMINISTERED IOP LOWERING MEDICATION. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988375&pc=MTA | 2023826-2020-00945 | 2020-03-26 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON HAD PERFORMED A LIMBAL RELAXING INCISION BUT IT WAS LEAKING SO THE SURGEON SUTURED THE INCISION CLOSED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS IMPROVED VISION AND VAULT BUT WITH MYDRIASIS AT DAY ONE AND PATIENT COMPLAINED OF HALOS. INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A ZIP LOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE PREVIOUS MDR. H6- PATIENT CODE 1937 SHOULD BE REMOVED FROM INITIAL MDR. H6- PATIENT CODE 3191- (MYDRIASIS) SHOULD BE REMOVED FROM INITIAL MDR. H6- DEVICE CODE 3001 SHOULD BE REMOVED FROM INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572295&pc=MTA | 2023826-2016-00437 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS DISLOCATED/SUBLUXED, WITH THE PATIENT EXPERIENCING ELEVATED INTRAOCULAR PRESSURE AND GLARES/HALOS. THE PATIENT COULD SEE A CRESCENT SHAPE ON THE RIGHT SIDE OF THEIR VISUAL FIELD, WHICH, WITH REFLECTION OF LIGHT, WAS MAKING IT IMPOSSIBLE FOR THEM TO DRIVE AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631046&pc=MTA | 2023826-2018-00931 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Iritis; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFTT EYE (OS) ON (B)(6) 2018. ON (B)(6) /2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND IRIS ATROPHY. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT, ADDITION OF NEW PI AND REPOSITIONING) WERE PERFORMED. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP ((B)(6) 2018) THE PATIENT WAS ON FOUR DIFFERENT GLAUCOMA DROPS, IOP 25MMHG. ADDITIONAL INFORMATION REPORTED, PI PATENT, PATIENT REPORTS DECREASED GLARE. EXPLANTED ICL AFTER ROTATION, IRIS CELL AND FLARE INCREASED THE IOP. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A THERMA SCIENTIFIC CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: ADDITION OF NEW PI WAS NOT REPORTED. . Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT HAD IRITIS DUE TO THE RUBBING OF THE LENS. THE LENS WAS ROTATED AND THE IRITIS STILL CONTINUED. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS THE INFLAMMATION. THE PROBLEM OF THE CELL AND FLARE WAS RESOLVED AFTER THE REMOVAL OF THE ICL, THERE WERE NO CELL AND FLARE POST REMOVAL. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5458648&pc=MTA | 2023826-2016-00197 | NA | Injury | STAAR SURGICAL COMPANY | 2016-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Loss of Vision; Discomfort; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. AT THE SAME DAY OF SURGERY AND ONE DAY POST-OP THE PATIENT WAS DOING WELL AND VA WAS 20/15. THE PATIENT TRAVELLED OUT OF STATE THAT SAME WEEK OF SURGERY AND REPORTED A SLIGHT DECLINE IN VISION AND DISCOMFORT TO THE EYE. THE PATIENT WAS PRESCRIBED A STEROID AND PRESSURE DROP. THE PATIENT WAS ADVISED TO SEE A LOCAL DOCTOR AND PRESENTED WITH ANGLE CLOSURE. THE PATIENT RETURNED TO THE SURGEON THE FOLLOWING WEEK AND PRESENTED WITH A SHALLOW CHAMBER. THE SURGEON HAD DIFFICULTY OPENING AN ADDITIONAL PI, TEMPORIZING HER PRESSURE BUT THE PATIENT WANTED THE ICL REMOVED. THE LENS WAS EXPLANTED AND THE PATIENT HAS RECOVERED HER VISION WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INFORMATION RECEIVED - THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430437&pc=MTA | 2023826-2017-00461 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED WITH STICKY RESIDUE ON THE LENS AND UNABLE TO EVALUATE DUE TO RESIDUE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6468284&pc=MTA | 2023826-2017-00598 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (THE EVENTS WERE NOT REPORTED ON THE INITIAL MDR). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND ANGLE CLOSURE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574198&pc=MTA | 2023826-2017-00833 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE; ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN OFF AND MISSING. A PIECE OF THE OTHER HAPTIC WAS TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574253&pc=MTA | 2023826-2017-00834 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661682&pc=MTA | 2023826-2016-00729 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THERE WAS EVIDENCE OF DRY RESIDUE. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661683&pc=MTA | 2023826-2016-00730 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (OPERATIONAL PROBLEM) : THE PRODUCT FOR THIS EVENT WAS RETURNED BUT WAS UNABLE TO EVALUATE DUE TO THE LENS WAS STUCK TO THE LID OF THE CONTAINER THE LENS WAS RETURNED IN. THE LENS WAS RETURNED DRY. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762431&pc=MTA | 2023826-2017-01260 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0, INTO THE PATIENT'S EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND ANGLE CLOSURE. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL SYMPTOMS INCLUDE: IRIDECTOMY, BLURRED VISION, PAIN, PUPIL BLOCK. LENS EXCHANGE WAS PERFORMED ON THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM WAS RESOLVED WITH LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOM: IRIDECTOMY. _CORRECTED: IRIDECTOMY IS NOT A SYMPTOM, IT WAS PERFORMED PRIOR ICL IMPLANTATION SURGERY. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762437&pc=MTA | 2023826-2017-01262 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO HIGH VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. ADDITIONAL SYMPTOMS: NARROWING AND CLOSURE OF ANGLES, PAIN, BLURRED VISION, PUPILLARY BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND PI WAS REVISED PRIOR TO EXPLANTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE LEFT EYE (OS). (B)(4). Manufacturer Narrative: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762432&pc=MTA | 2023826-2017-01261 | 2017-07-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO "A LOT" OF VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED INFORMATION: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS IN MDR FOLLOW-UP#1: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. CORRECTED: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, PAIN, BLURRED VISION, AND PUPIL BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND THE SURGEON ALSO PERFORMED AC DECOMPRESSION AND YAG REVISION PRIOR TO EXPLANTATION. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743979&pc=MTA | 2023826-2015-00474 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE - 50MM HG (PUPILLARY BLOCK), PAIN AND BLURRY VISION. ANOTHER PERIPHERAL IRIDECTOMY WAS PERFORMED. A SHORTER LENS WAS IMPLANTED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25, THE CHAMBER WAS SHALLOW AND THE EYE IRRITATED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, THE EVENT WAS CAUSED BY EXCESSIVE VAULTING OF THE ICL LENS. THE HIGH IOP WAS THE CONSEQUENCE OF PUPILLARY BLOCK CAUSED BY EXCESSIVE VAULTING OF THE LENS. THE LENS EXPLANTATION, EXCHANGE AND ADDITIONAL IRIDOCTOMY WERE NECESSARY TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7119246&pc=MTA | 2023826-2017-01886 | 2017-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; Retinal Detachment; Vitreous Hemorrhage; Halo; Vitrectomy; No Code Available | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2017 AN ICL WAS REMOVED DUE TO RETINAL DETACHMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BROUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. VITRECTOMY WAS ALSO PERFORMED. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS (AND NATURAL LENS) WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. LASER RETINOPEXY WAS PERFORMED FOLLOWED BY PNEUMATIC CRYOPEXY ONLY TO HAVE THE CURTAIN IN PATIENT'S VISION RETURN. VITRECTOMY PERFORMED LED TO VITREOUS HEMORRHAGE. A SECOND RETINAL TEAR WAS NOTICED LEADING TO A SECOND VITRECTOMY, SCLERAL BUCKLE AND "LENSECTOMY" OR EXPLANT OF THE MICL AND NATURAL LENS.THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE, PAIN AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BOUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672736&pc=MTA | 2023826-2023-03646 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS 1+ PIGMENT ANTERIOR CHAMBER, THE IRIS WAS RESOLVING. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17673554&pc=MTA | 2023826-2023-03647 | 2023-05-19 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS 1+ PIGMENT ANTERIOR CHAMBER, THE IRIS WAS RESOLVING. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558798&pc=MTA | 2023826-2021-00703 | 2021-03-08 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS THE LENS "VAULTED TOO MUCH" AND HIGH PRESSURE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B1-CORRECTED B5-ADDITIONAL INFORMATION: ON (B)(6) 2021 THE LENS WAS EXPLANTED. ADDITIONAL SYMPTOMS REPORTED: TRANSIENT LOSS OF BCVA, PUPIL BLOCK, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND UNKNOWN: "POSSIBLE FAILURE OF PI'S." H6-CLINICAL CODE 4581 [PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL]. HEALTH IMPACT CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION-UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493426&pc=MTA | 2023826-2022-01674 | 2022-04-23 | Injury | STAAR SURGICAL COMPANY | 2022-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -12.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WERE PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, SAME MODEL AND POWER LENS. IT WAS REPORTER THAT THE EXCHANGE RESOLVED THE PROBLEM. HOWEVER, IT WAS ALSO REPORTED THAT THE VAULT IS STILL HIGH. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET, COMFORTABLE, STABLE, HAPPY PATIENT." IN THE REPORTERS OPINION THE DEVICE AND PATIENT-RELATED FACTOR WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493429&pc=MTA | 2023826-2022-01673 | 2022-04-23 | Injury | STAAR SURGICAL COMPANY | 2022-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WERE PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, SAME MODEL AND POWER LENS. IT WAS REPORTER THAT THE EXCHANGE RESOLVED THE PROBLEM. HOWEVER, IT WAS ALSO REPORTED THAT THE VAULT IS STILL HIGH. FOR STATUS OF THE EYE THE REPORTER STATED " QUIET, COMFORTABLE, GOOD VISION, HAPPY." IN THE REPORTERS OPINION THE DEVICE AND PATIENT-RELATED FACTOR IS THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WAS PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND PROBLEM RESOLVED. STATUS OF THE EYE WAS REPORTED AS "QUIET, COMFORTABLE, GOOD VISION, HAPPY". HOWEVER IT WAS ALSO REPORTED THE VAULT IS STILL HIGH. IN THE REPORTER'S OPINION THE DEICE AND PATIENT RELATED FACTOR IS THE CAUSE OF THIS EVENT." H6- HEALTH EFFECT: IMPACT CODE- 4627 SHOULD BE CORRECTED TO 4629. D6B- (B)(6) 2022 CORRECTED TO (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM MICL13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, ANGLE CLOSURE, ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION. ADDITIONAL MEDICAL INTERVENTION (AQUEOUS SUPPRESSANTS AND DIOMOX) WAS PRESCRIBED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND PROBLEM RESOLVED. STATUS OF THE EYE WAS REPORTED AS "QUIET, COMFORTABLE, GOOD VISION, HAPPY". HOWEVER IT WAS ALSO REPORTED THE VAULT IS STIL HIGH. IN THE REPORTER'S OPINION THE DEICE AND PATIENT RELATED FACTOR IS THE CAUSE OF THIS EVENT." H6- HEALTH EFFECT: IMPACT CODE- 4627. D6B- (B)(6) 2022. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509197&pc=MTA | 2023826-2024-00227 | 2023-10-16 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK,ANGLE CLOSURE, UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY WAS REPORTED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. MILD CORTICAL LENS OPACITY WAS OBSERVED ON (B)(6) 2023. IT WAS REPORTED THAT A MONOFOCAL LENS WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED HOWEVER, IT WAS ALSO REPORTED THAT IT DID NOT RESOLVE. ADDITIONAL INFORMATION AND CLARIFICATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581:SIGNIFICANT REDUCTION OF ICA, ANGLE CLOSURE,UNREACTIVE/FIXED PUPIL, IRIS ATROPHY. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727938&pc=MTA | 2023826-2015-00471 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE - 50MM HG (PAPILLARY BLOCK), NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT'S VISION WAS BLURRY. ANOTHER PERIPHERAL IRIDECTOMY WAS PERFORMED BUT WITH NO IMPROVEMENT. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2015. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15, THE PATIENT IS STABLE AND DOING WELL. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - ANTERIOR CHAMBER ANGLE CLOSURE, PUPILLARY BLOCKAGE AND INCREASED IOP ARE CONSEQUENCES OF EXCESSIVE VAULTING. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY THE EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE. HOWEVER, NO DEFINITE ROOT CAUSE DETERMINED. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305169&pc=MTA | 2023826-2015-01689 | 2015-11-01 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS PROBABLY DUE TO SIZING ISSUE. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE REPORTER STATED, THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. SINCE THE EXCHANGE WITH THE SMALLER SIZE LENS RESOLVED THE ISSUE, ALTHOUGH NOT CONFIRMED, IT IS LIKELY THE CAUSE OF THE EVENT IS DUE TO LENS SIZING. THE PRODUCT WAS NOT RETURNED TO STAAR FOR AN EVALUATION. MEDICAL REVIEW - REPORTEDLY, MICL( 13.2MM) WAS EXPLANTED/EXCHANGED APPROXIMATELY ONE AND A HALF MONTHS POSTOPERATIVELY FOR A SHORTER SIZE MICL (12.6 MM) TO ADDRESS HIGH VAULT AND SUBSEQUENT COMPLICATIONS (PUPILLARY BLOCK, ELEVATED IOP, BLURRY VISION). ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2015, THE CAUSE OF THE EVENT WAS UNKNOWN BUT WAS PROBABLY DUE TO SIZING ISSUE. PATIENT WAS UNDER MONITOR AND THE MOST CURRENT UCVA WAS 20/25. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7187682&pc=MTA | 2023826-2018-00043 | 2017-12-12 | Injury | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO SPIKED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: PATIENT'S NAME IS CORRECTED TO (B)(6). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED IOP (PUPILLARY BLOCK), NARROWING OF ANGLES, SHALLOWING OF ACD AND ANGLE CLOSURE. THERE WAS NO PATIENT INJURY. THE PATIENT CURRENTLY HAS BLURRED VISION, AN ENLARGED/ UNRESPONSIVE PUPIL (URRETS-ZAVALIA SYNDROME) AND PHOTOPHOBIA. THE REPORTER STATED THAT THE DOCTOR MIGHT SUGGEST SOFT CONTACTS TO BE USED BY PATIENT IN ORDER TO IMPROVE THE PATIENT CONDITION. (B)(4). CONCLUSION CODE: AS PER THE DFU, THIS LENS IS CONTRAINDICATED FOR IMPLANTATION IN PATIENTS ABOVE THE AGE OF 45 YEARS AND IN THOSE WITH ACD (ANTERIOR CHAMBER DEPTH) LESS THAN 3.0MM CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT, IN A SMALL POUCH TAPED TO A PAPER WITH FACILITY INFORMATION. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES AND PRESENCE OF RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A SMALL POUCH TAPED TO PAPER WITH FACILITY INFORMATION. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES BEING UNABLE TO EVALUATE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710842&pc=MTA | 2023826-2016-00807 | NA | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Prolapse; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK), NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND IRIS PROLAPSE. THE PATIENT EXPERIENCED BLURRED VISION. A SECONDARY SURGERY WAS PERFORMED TO LASER THE PERIPHERAL IRIDOTOMIES AND THE LENS WAS REPOSITIONED BUT THE PATIENT STILL HAD PRESSURE ISSUES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S PUPIL HAS REMAINED LARGE WITH SOME SHALLOWING OF THE ANTERIOR CHAMBER BUT THE ISSUE HAS IMPROVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). RESULTS CODE(S): (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18377615&pc=MTA | 2023826-2023-05714 | 2023-10-11 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -7.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, BLURRED VISOIN AND PROLONGED POSTOPERATIVE ANTERIOR UVEITIS. IT WAS REPORTED THAT THE IOP WAS CONTROLLED WITH MEDICATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10786734&pc=MTA | 2023826-2020-02499 | 2020-03-11 | Injury | STAAR SURGICAL COMPANY | 2020-11-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON 11/MAR/2020. ELEVATED IOP (58.0 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED WITH A STRONGER POWER, SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, SURGEON "INADVERTENTLY IMPLANTED LENS PREPARED FOR OD." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE LENS WAS REMOVED AND EXCHANGED" SHOULD BE CORRECTED TO "THE LENS WAS EXPLANTED AND IN A SEPARATE SURGERY WAS REPLACED". D9: DATE ENTERED IN SUPPLEMENTAL #1 SHOULD BE REMOVED AND DATE ENTERED IN D9 IN INITIAL MEDWATCH REPORT SHOULD BE KEPT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED GLARE/HALOES AND AC IRRIGATION. THIS SHOULD HAVE BEEN INCLUDED. H6 - 4625 - AC IRRIGATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308925&pc=MTA | 2023826-2021-00320 | 2021-01-05 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP (28MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2020), PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, GLARE/HALOES, IRIS ATROPHY, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "EXCESSIVE VAULT CAUSING PIGMENTARY DISPERSION AND INCREASED EYE PRESSURE." REPORTEDLY, "VAULT IS NOW BETTER, ALONG WITH IOP. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- REPLACEMENT LENS (B)(6) DID NOT RESOLVE THE PROBLEM AND WAS EXPLANTED ON (B)(6)2022. PER EMAIL ON (B)(6) 2022, THE "REASON FOR EXPLANT- PATIENT COMPLAINED OF DISCOMFORT". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# : (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688983&pc=MTA | 2023826-2016-00253 | 2015-12-15 | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.5 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED IOP, REDUCTION OF ICA, UNREACTIVE PUPIL, SHALLOW CHAMBER, AND GLARES/HALOES WERE REPORTED. A LENS EXCHANGE IS PLANNED BUT HAS NOT YET BEEN PERFORMED. PATIENT IS CURRENTLY ON ANTIGLAUCOMA AND ANTI-INFLAMMATORY EYE DROPS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 2015/12/15. ELEVATED INTRAOCULAR PRESSURE (IOP), REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE PUPIL, SHALLOW CHAMBER, AND GLARES/HALOES WERE REPORTED. THE SURGEON PERFORMED ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS AND INTRAOCULAR INJECTION (MEDICATION). PATIENT WAS ON ANTIGLAUCOMA AND ANTI-INFLAMMATORY EYE DROPS. THE LENS WAS EXPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7480272&pc=MTA | 2023826-2018-00657 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2, -12.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (DATE ASSESSED: (B)(6) 2018), UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND BLURRED VISION. ADDITIONAL SURGICAL INTERVENTION ( ENLARGEMENT OF PI; YAG) AND MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WERE PERFORMED/PRESCRIBED. THE REPORTED STATED THAT THE CAUSE OF THIS EVENT WAS UNKNOWN. "FIXED PUPIL" WAS REPORTED FOR STATUS OF THE EYE AT PATIENT'S LAST FOLLOW UP. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: AS OF (B)(6) 2018 BEST CORRECTED ACUITY 20/30, IOP IS ELEVATED TO 27UM. PATIENT IS BEING MONITORED FOR OCULAR HYPERTENSION. CORRECTIONS: PATIENT DID NOT HAVE INTRAOCULAR INJECTIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) . WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7480527&pc=MTA | 2023826-2018-00662 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM MICL13.2, -9.5, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, GLARE/HALOES, AND BLURRED VISION. ADDITIONAL SURGICAL INTERVENTION (YAG) AND MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WERE PERFORMED/PRESCRIBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE STATUS OF THE EYE WAS REPORTED TO BE WITHIN NORMAL LIMITS. PATIENT IS BEING MONITORED FOR OCULAR HYPERTENSION. CORRECTIONS: PATIENT DID NOT HAVE INTRAOCULAR INJECTIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. CORRECTED DATA: CORRECT DATE IS (B)(6) 2018, (B)(6) 2018 IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546111&pc=MTA | 2023826-2017-00770 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -8 DIOPTER LENS. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307839&pc=MTA | 2023826-2019-00222 | 2019-01-10 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP AND GLARE/HALOS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED FOR "ADVERSE EVENT" TO "PRODUCT PROBLEM" FOR INITIAL MDR. TYPE OF REPORTABLE EVENT: CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL PATIENT CODE ADDED, 3191 - NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES) PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9745394&pc=MTA | 2023826-2020-00378 | 2019-11-25 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, ELEVATED IOP AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 3001: GLARE/HALOES. METHOD CODE 4110: LENS WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105253&pc=MTA | 2023826-2020-01077 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, PUPIL BLOCK WITH ELEVATED IOP, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835113&pc=MTA | 2023826-2018-01317 | 2018-02-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 , -125 DIOPTER WAS IMPLANTED AND REMOVED ON (B)(6) 2018 DUE TO HIGH IOP AND UNHAPPY VISION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: TWO SIMILAR COMPLAINT TYPE EVENT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP (DATE ASSESSED: (B)(6) 2018 AND (B)(6) 2018), AND UNREACTIVE PUPIL. ENLARGEMENT OF PI (YAG AND SURGICAL) WAS REPORTED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND SHORTER LENGTH LENS. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION REPORTED: PATIENT HAD ICL OD WITHOUT COMPLICATIONS (B)(6) 2018. PATIENT HAD PRK OU WITHOUT COMPLICATION (B)(6) 2018. REPORTER STATED THAT POSSIBLE CAUSE OF EVENT WAS UGH SYNDROME DESPITE PATENT PI. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: PATIENT CODE: SECONDARY SURGICAL INTERVENTION,EXPLANT. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631451&pc=MTA | 2023826-2018-00930 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Consequences Or Impact To Patient; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, AND GLARE/HALOES. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT AND ADDITION OF NEW PI) WERE PERFORMED. ADDITIONAL INFORMATION REPORTED; DECREASED CELL AND FLARE + REMOVAL. PATIENT REPORTS DECREASE IN GLARE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A THERMO SCIENTIFIC CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN. CORRECTED DATA: PRE-EXISTING CONDITION: AXIAL LENGTH GREATER THAN 29. CONDITION SHOULD HAVE BEEN MENTIONED IN INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4) SECONDARY SURGICAL PROCEDURE, ENLARGEMENT OF PI, ADDITION OF NEW PI, EXPLANT. PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842749&pc=MTA | 2023826-2020-00509 | 2020-01-01 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Iritis; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL LENS DUE TO PIGMENT DISPERSION, BLURRED VISION, AND IRITIS. REPORTEDLY THE INFLAMMATION WAS TREATED WITH MEDICATION. THE SURGEON STATES IT LOOKED LIKE "AN INFLAMMATORY CELL WAS ATTACHED TO THE ICL." THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760983&pc=MTA | 2023826-2015-00536 | 2014-02-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Loss of Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED PATIENT HAD A MICL 13.6 MM IMPLANTABLE COLLAMER LENS IMPLANTED IN THE PATIENT'S RIGHT EYE ON (B)(6) 2012. AS PART OF THE POST MARKET STUDY, THE PATIENT REPORTED EXPERIENCING HAVING LOST ACUITY IN HIS NEAR VISION AFTER THE SURGERY AND HAVE DIFFICULTY FOCUSING WITHOUT THE USE OF CONTACTS OR GLASSES TO BALANCE HIS EYESIGHT. ALSO BEING TOLD BY HIS OPTOMETRIST THAT SHE CAN SEE THE BEGINNINGS OF CATARACTS. THE PHYSICIAN'S OFFICE WAS CONTACTED FOR FURTHER INFORMATION. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2014. THE PATIENT IS NOW HYPEROPIC. STATED ADVERSE EVENT IS RELATED TO THE DEVICE. PROGNOSIS IS GOOD. VISUAL ACUITY POST-TREATMENT BCDVA 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). ICL REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328084&pc=MTA | 2023826-2015-01749 | 2014-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED THE VISION IN HER RIGHT EYE HAS GOTTEN WORSE. THE FACILITY REPORTED THE PATIENT HAD VIRAL CONJUNCTIVITIS IN THE LEFT EYE, WHICH WAS RESOLVED AND WAS NOT RELATED TO THE DEVICE. THE PATIENT'S VA POST-OP WAS 20/30 AND THE PROGNOSIS IS GOOD. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: PER MEDICAL REVIEW: ON (B)(6) 2015 THE PATIENT COMPLETED 48 MONTH PATIENT FOLLOW-UP FORM AND REPORTED THAT "VISION HAS GOTTEN WORSE" FOLLOWING ICL IMPLANTATION. NO REPORTED SECONDARY SURGICALLY INTERVENTION. ACCORDING TO DCR, DATED ON (B)(6) 2015, THE CAUSE OF THE REPORTED EVENT WAS VIRAL CONJUNCTIVITIS NOT RELATED TO THE DEVICE. THE PATIENT WAS MEDICALLY TREATED (TOBRADEX) AND THE ISSUE RESOLVED WITH GOOD PROGNOSIS. VA WAS 20/30 (UNKNOWN IF THAT WAS CORRECTED OR UNCORRECTED). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: CONCLUSION CODE(S): (DEVICE FAILURE RELATED TO PATIENT CONDITION): VIRAL CONJUNCTIVITIS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151923&pc=MTA | 2023826-2015-01257 | 2015-09-15 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. POST-OP. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS. SEE MFR #2023826-2015-01256 FOR RIGHT EYE. Manufacturer Narrative: INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(4) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(4) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. CLAIM # 708141 Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -05.00. EVALUATION: VISION, LOSS OF, (NEW INCISION), (SECONDARY SURGERY). (LENS EXPLANTED). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4) | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327972&pc=MTA | 2023826-2015-01747 | NA | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED SHE HAD LOST VISION DURING HER PREGNANCY AND THE VISION HAS NOT RETURNED. THE PATIENT IS NOW WEARING CONTACTS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PER MEDICAL REVIEW: PATIENT COMPLETED 36 MONTH FOLLOW-UP FORM FOLLOWING ICL IMPLANTATION STATING THAT SHE HAD EXPERIENCED DECREASE IN VISUAL ACUITY DURING PREGNANCY SO SHE HAD TO WEAR CONTACT LENS AGAIN. NO REPORT OF ANY SECONDARY SURGICALLY OR MEDICALLY INTERVENTION PERFORMED AND NO REPORT OF LOSS OF BCVA. ACCORDING TO DATA CLARIFICATION REQUEST, RECEIVED ON (B)(6) 2015, ICL WAS IMPLANTED BILATERALLY IN 2012 AND THE PATIENT WAS LAST SEEN ON (B)(6) 2013 WITH NO NOTED ISSUES (VA OU: 20/20). NO ADDITIONAL INFORMATION WAS RECEIVED TO DATE FROM ANY OTHER DOCTOR. IT SHOULD BE NOTED THAT HORMONES PRODUCED DURING PREGNANCY AND/OR BREASTFEEDING CAN CHANGE EYESIGHT SO PER DFU: "THE VISIAN ICL IS CONTRAINDICATED IN PATIENTS WHO ARE PREGNANT OR NURSING". A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). LENS IMPLANTED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327977&pc=MTA | 2023826-2015-01748 | NA | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DURING HER PREGNANCY AND THE VISION HAS NOT RETURNED. THE PATIENT IS NOW WEARING CONTACTS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: PATIENT COMPLETED 36 MONTH FOLLOW-UP FORM FOLLOWING ICL IMPLANTATION STATING THAT SHE HAD EXPERIENCED DECREASE IN VISUAL ACUITY DURING PREGNANCY SO SHE HAD TO WEAR CONTACT LENS AGAIN. NO REPORT OF ANY SECONDARY SURGICALLY OR MEDICALLY INTERVENTION PERFORMED AND NO REPORT OF LOSS OF BCVA. ACCORDING TO DATA CLARIFICATION REQUEST, RECEIVED ON 11/19/2015, ICL WAS IMPLANTED BILATERALLY IN 2012 AND THE PATIENT WAS LAST SEEN ON (B)(6) 2013 WITH NO NOTED ISSUES (VA OU:20/20). NO ADDITIONAL INFORMATION WAS RECEIVED TO DATE FROM ANY OTHER DOCTOR. IT SHOULD BE NOTED THAT HORMONES PRODUCED DURING PREGNANCY AND/OR BREASTFEEDING CAN CHANGE EYESIGHT SO PER DFU: "THE VISIAN ICL IS CONTRAINDICATED IN PATIENTS 3. WHO ARE PREGNANT OR NURSING". A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). LENS IMPLANTED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578311&pc=MTA | 2023826-2015-00256 | 2014-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOSS OF BCVA (PREOP 20/25, POST-OP 20/40). THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. PATIENTS LAST BCVA ON (B)(6) 2014 WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218719&pc=MTA | 2023826-2015-01482 | 2015-08-11 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). THERE WAS A LOSS OF BEST-CORRECTED VISUAL ACUITY WHICH WAS RESOLVED WITH AN EXCHANGED LENS MODEL. ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS WAS DEFORMED AND RETURNED IN PIECES - IT WAS UNABLE TO BE EVALUATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5189312&pc=MTA | 2023826-2015-01448 | 2015-09-21 | Injury | STAAR SURGICAL COMPANY | 2015-10-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY AND THERE WERE POST-OPERATIVE DEPOSITS ON THE CRYSTALLINE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS -(EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664097&pc=MTA | 2023826-2016-00660 | 2016-04-07 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED AND IT IS PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA) WAS NOT REPORTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4749511&pc=MTA | 2023826-2015-00501 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY | 2015-05-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED, AS WELL AS LENS ROTATION AND LOSS OF VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOW BCVA WAS 20/28 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: (OTHER) A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918795&pc=MTA | 2023826-2019-01491 | 2019-03-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | No Consequences Or Impact To Patient; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY, FOR ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: HALOES WERE ALSO REPORTED. UDI#: N/A. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11547918&pc=MTA | 2023826-2021-00739 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENT LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE IRIS GOT EXTRUDED OUT THE WOUND DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING 2ND SURGERY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS ALSO INDICATED THE PATIENT EXPERIENCED PAIN; NON CONFORMITY THAT WASN'T SPECIFIED WHAT TYPE; LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED, LENS WAS EXPLANTED ON 03-MARCH-2021 AND THE DAY AFTER WAS RE-IMPLANTED. H6: 4581- LOSS OF BCVA, IRIS PROLAPSE. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11557439&pc=MTA | 2023826-2021-00740 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Defective Device; Unintended Movement | Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT WAS REPORTED WITH HYPEROPIA AND ASTIGMATISM, WHICH CAUSED BAD AND BLURRED VISION. ALSO REPORTED PERMANENET LOSS OF BCVA. THE LENS WAS ALSO REPORTED AS HAVING FLIPPED. THE PATIENT HAD A SMALL EYE, THERE WAS PUPILAR CONSTRICTION AND THE PUPIL BEGAN TO CLOSE DURING THE SURGERY. THE LENS WAS EXPLANTED AND RE-IMPLANTED ON (B)(6) 2021. PERIBULBAR BLOCK WAS ADMINISTERED DURING THE 1ST SURGERY (INITIAL) AND 2ND SURGERY (RE-IMPLANT). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. TWO HOURS AFTER THE RE-IMPLANT SURGERY, ANTERIOR CHAMBER A LITTLE BIT NARROW AND PUPIL DILATED. THE PATIENT IS HAPPY WITH THE RESULTS OF THE SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS ALSO INDICATED THAT THE PATIENT EXPERIENCED PAIN. REPORTER INDICATED THERE WAS PRODUCT NONCONFORMITY BUT THIS WAS NOT SPECIFIED ON. LENS WAS IMPLANTED UPSIDE DOWN/FLIPPED. IT WAS LATER CLARIFIED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021, AND THE NEXT DAY RE-IMPLANTED. H6 - 4581 - LOSS OF BCVA. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5/H10- IT WAS LATER CLARIFIED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND PROPERLY REIMPLANTED, INTRAOPERATIVELY AND NOT IN A SEPARATE SURGERY. THIS HAD RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917237&pc=MTA | 2023826-2015-00799 | 2015-06-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO NO LENS VAULT. THERE WAS NO LENS OPACITY. THE PATIENT EXPERIENCED BLURRY VISION AND EYE PAIN. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY INDICATED THE DATE OF INITIAL SURGERY WAS (B)(6) 2015. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. USE OF THE WRONG LENS WAS MOST LIKELY THE CAUSE OF LACK OF VAULT. (B)(4). Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION - IMPLANTED DATE - (B)(6) 2015. METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THE CONSEQUENCE OF INADEQUATE VAULTING COULD BE CATARACT THAT WARRANT LENS EXCHANGE TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074717&pc=MTA | 2023826-2016-01575 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS REPOSITIONED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THAT LENS ALSO IS VAULTED. THE PATIENT'S POST-OP BSCVA WAS 20/30 +1, WITH BLURRY VISION AND PAIN. Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6775540&pc=MTA | 2023826-2017-01298 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPILLARY BLOCK (ELEVATED IOP), BLURRED VISION, AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A CASE/VIAL. THE LENS VIAL WAS RETURNED TAPED TO LAF CARD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146624&pc=MTA | 2023826-2017-01945 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Pain; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -7 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE, GLARES, HALOES, INDUCED ASTIGMATISM, AND CILIARY GROOVE PAIN. THE LENS WAS EXCHANGED FOR A DIFFERENT LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285658&pc=MTA | 2023826-2020-01486 | 2020-06-22 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 LENS, -17.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LENS ROTATION, LENS DISLOCATION/SUBLUXATION, CILIARY PAIN, LENS DECENTRATION, TYNDALL 2+, AND DOUBLE IMAGE. TOPICAL STERIODS WERE ALSO ADMINISTERED. REPOSITIONING DID NOT RESOLVE THE PROBLEM. REPORTEDLY, ON (B)(6) 2020 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER 1 BROKEN HAPTIC WAS DISCOVERED UPON EXPLANT WHICH COULD BE A CONTRIBUTING CAUSE. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE. PAIN PERSISTS. REFERENCE FILE# (B)(4) FOR REPLACEMENT LENS INFORMATION. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385347&pc=MTA | 2023826-2019-02271 | 2019-10-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE: 3191 - NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN MANUFACTURER NARRATIVE, BUT CODE SHOULD HAVE ALSO BEEN SELECTED IN H6. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385361&pc=MTA | 2023826-2019-02272 | 2019-10-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Pain; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -08-50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PATIENT REPORTED, "ISSUES WITH BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILIARY BODY." LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (CONTRACTION OF THE CILIARY BODY) WAS REPORTED IN THE MANUFACTURER NARRATIVE, BUT SHOULD HAVE ALSO BEEN SELECTED IN H6 SECTION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487527&pc=MTA | 2023826-2017-00618 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pupillary Block; Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, AND THE LENS WAS NOTED AS HAVING AN EXCESSIVE VAULT. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2017 IN THE PATIENT'S RIGHT EYE (OD), -11.0 DIOPTER. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HYPEROPIC SURPRISE, EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND ANGLE CLOSURE. A YAG PERIPHERAL IRIDOTOMY WAS PERFORMED TO RELIEVE PUPIL BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS POST-OP BSCVA WAS 20/25. (B)(4). CORRECTED DATA: THE IMPLANTED DATE IS (B)(6) 2017 NOT (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL, IN A BIOHAZARD BAG. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: CONCLUSION CODE: (VISIAN® IMPLANTABLE COLLAMER® LENSES ARE INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. THIS PATIENT WAS (B)(6) YEARS OF AGE). (B)(4) Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816245&pc=MTA | 2023826-2016-00976 | 2016-02-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Vertigo; Blurred Vision; Loss of Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOSS OF VISION AND VERTIGO. THE LENS WAS EXCHANGED WITH ANOTHER SAME MODEL LENS WITH A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. Manufacturer Narrative: (B)(4). REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, ACD BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS DRIED, DISCOLORED MATERIAL. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO UNEXPECTED REFRACTIVE OUTCOME AND VERTIGO. THE LENS WAS EXCHANGED WITH ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/30. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073304&pc=MTA | 2023826-2020-03183 | 2000-08-21 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -17.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO " PATIENTS VISION WAS YELLOW". IT WAS REPORTED THAT THE PROBLEM RESOLVED, PATIENTS VISION RETURNED TO NORMAL AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12644941&pc=MTA | 2023826-2021-04020 | 2021-07-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6_13.2,-6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT HAS PROBLEMS WITH NEAR VISION, LENS REMAINS IMPLANTED. THE REPORTER STATED "THE PATIENT DOESN'T WANT ANY FURTHER TREATMENT FOR THE TIME BEING." FOR PATIENTS CURRENT CONDITION " PROBLEMS WITH READING" WAS REPORTED. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182716&pc=MTA | 2023826-2022-00027 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS THE LENS IS STILL NOT CENTERED AND THE PATIENT IS UNCOMFORTABLE WITH THE GLARE/HALOS. SEE MFR REP# 2023826-2021-04895 FOR INITIAL LENS. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923026&pc=MTA | 2023826-2015-00829 | 2015-06-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WILL BE EXCHANGED FOR ANOTHER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. METHOD - WORK ORDER SEARCH RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE MANUFACTURE DATE: 03/19/2015. (B)(4). Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VILE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. CONCLUSION: PATIENT IS UNDER THE AGE OF 21 AND HAS AN ACD OF 2.9. AS PER DFU FOR THIS MODEL FOR THE PATIENTS UNDER THE AGE OF 21 YEARS AND IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158574&pc=MTA | 2023826-2016-01737 | 2016-11-11 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. THE LENS REMAINS IMPLANTED AT THE TIME OF THIS REPORT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7329352&pc=MTA | 2023826-2018-00356 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.0 DIOPTER ON (B)(6) 2017 IN THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTED THAT "THE PATIENT IS SUFFERING FROM SEVERE GLARE AND HE BELIEVES THAT IT IS FROM THE CENTRAL PORTS OF THE LENSES IN HER EYES". THE REPORTER STATED THAT THE PATIENT STILL HAS "GLARES/ DECENTRATION". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE REPORTER BUT NONE HAVE BEEN FORTHCOMING. A SUPPLEMENTAL MDR WILL BE SUBMITTED IF NEW INFORMATION IS RECEIVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7304830&pc=MTA | 2023826-2018-00300 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE PATIENT COMPLAINED OF GLARE AND FOUND IT HARD TO DRIVE. THE SURGEON PLANS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023228&pc=MTA | 2023826-2021-04895 | 2021-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, LENS ROTATION, AND GLARE/HALOS. SEE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "LENGTH OF THE ICL." Manufacturer Narrative: B5- CORRECTED TO : "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULTS, LENS ROTATION, AND GLARE/HALOS. SEE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "LENGTH OF THE ICL." CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16044661&pc=MTA | 2023826-2022-04707 | 2022-10-21 | Injury | STAAR SURGICAL COMPANY | 2022-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -6.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. REPORTEDLY, "LENS TOO BIG SO HIGH VAULT BUT ALSO GLARE AT NIGHT SO NOT GOING TO EXCHANGE. PATIENT HAD DRY EYE BUT THIS HAS CLEARED SO HAVE DECIDED TO LASER INSTEAD." STATUS OF THE EYE IS SAID, "LENS HAS BEEN REMOVED AND WILL LASER INSTEAD." IT WAS ALSO STATED, "PATIENT HAD DRY EYE WHEN FIRST SEEN SO DECIDED UPON ICL, HOWEVER THE PATIENT HAS A HIGH VAULT AND THE LENSES NEED REMOVING. SINCE THE PATIENT NO LONGER HAS DRY EYE AND IS GETTING GLARE AT NIGHT WHERE SHE CANNOT GO OUT AFTER 4PM, HAVE DECIDED TO LASER RATHER THAN EXCHANGE." PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC BROKEN. RECEIVED TWO LENSES IN ONE BAG AND ONE VIAL,DUE TO SIGNIFICANT LENS DAMAGE, UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16044665&pc=MTA | 2023826-2022-04705 | 2022-10-21 | Injury | STAAR SURGICAL COMPANY | 2022-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -7.0 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS THE DEVICE. REPORTEDLY, "LENS TOO BIG SO HIGH VAULT BUT ALSO GLARE AT NIGHT SO NOT GOING TO EXCHANGE. PATIENT HAD DRY EYE BUT THIS HAS CLEARED SO HAVE DECIDED TO LASER INSTEAD." STATUS OF THE EYE IS SAID, "LENS HAS BEEN REMOVED AND WILL LASER INSTEAD." IT WAS ALSO STATED, "PATIENT HAD DRY EYE WHEN FIRST SEEN SO DECIDED UPON ICL, HOWEVER THE PATIENT HAS A HIGH VAULT AND THE LENSES NEED REMOVING. SINCE THE PATIENT NO LONGER HAS DRY EYE AND IS GETTING GLARE AT NIGHT WHERE SHE CANNOT GO OUT AFTER 4PM, HAVE DECIDED TO LASER RATHER THAN EXCHANGE." PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. RECEIVED TWO LENSES IN ONE BAG AND ONE VIAL, THUS UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM. CLAIM# (B)(4). Manufacturer Narrative: PT INFO:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742913&pc=MTA | 2023826-2023-01257 | 2023-03-03 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) -2023 DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED DOUBLE VISION AND GLARE AT NIGHT. THE PATIENT WAS PRESCRIBED MEDICATION BUT THE SYMPTOMS DID NOT IMPROVE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC SPLIT. THE HAPTIC WAS BROKEN AND BENT. UNABLE TO EVALUATE THE LENS DUE TO LENS WAS RETURNED IN PIECES. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595988&pc=MTA | 2023826-2017-00841 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER VISION HAD WORSENED. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE DATE OF THE LAST VISIT WAS (B)(6) 2014. THE REPORTER STATED THE CAUSE OF THE EVENT WAS RESIDUAL ASTIGMATISM, THE PRE-OP ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH THE ICL. THE VA AT PATIENT'S 6 MONTH VISIT WAS 20/20 AND AT 12 MONTH VISIT IT WAS 20/25. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL EVALUATION OF COMPLAINT EVENTS FOUND NO SERIOUS INJURY TO PATIENT REQUIRING SURGICAL OR MEDICAL INTERVENTION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE DATE OF THE LAST VISIT WAS (B)(6) 2014. THE REPORTER STATED THE CAUSE OF THE EVENT WAS RESIDUAL ASTIGMATISM, THE PRE-OP ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH THE ICL. THE VA AT PATIENT'S 6 MONTH VISIT WAS 20/20 AND AT 12 MONTH VISIT IT WAS 20/25. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429111&pc=MTA | 2023826-2020-01845 | 2020-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-19 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) OF 2015. THE SURGEON REPORTS DOUBLE VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: DATE IS APPROXIMATE. PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11070813&pc=MTA | 2023826-2020-03179 | 2020-08-15 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1,-7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXPLANTED, THE REPORTER STATED "THE PATIENT WAS DISSATISFIED WITH THE QUALITY OF VISION AND REQUESTED TO BE REMOVED THE LENS." IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15411920&pc=MTA | 2023826-2022-03050 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "MACULAR HOLE 3 DAYS AFTER ICL." Manufacturer Narrative: B5 - VITRECTOMY WAS PERFORMED ON (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND REMOVED ON (B)(6) 2022, DUE TO A MACULAR HOLE 3 DAYS AFTER ICL. CAUSE OF THE EVENT IS UNKNOWN. PROBLEM WAS NOT RESOLVED. REPORTEDLY, "PATIENT TRANSFERRED TO GENERAL HOSPITAL, RECENT DETAILS ARE UNAVAILABLE." CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11073301&pc=MTA | 2023826-2020-03182 | 2020-11-14 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511816&pc=MTA | 2023826-2020-01994 | 2020-07-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1M VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT REQUESTED TO CHANGE THE LENS TO A LOWER POWER BECAUSE HE FELT THE VISUAL ACUITY WAS TOO GOOD. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS (-06.00) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL,DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524373&pc=MTA | 2023826-2021-03721 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT DIOPTER LENS DUE TO LOW VAULT AND DISCOMFORT WITH VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT) ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911484&pc=MTA | 2023826-2022-00962 | 2020-01-09 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Distortion | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS DIOPTER -17.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE SAME DIOPTER LENS DUE TO POSTOPERATIVE VISUAL DISTORTION AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. AFTER EXCHANGE IT IS REPORTED THAT, "SELF-COMPLAINT VISUAL DISTORTION IMPROVED SLIGHTLY." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). . CLAIM# (B)(4). Manufacturer Narrative: "THE REPORTER MARKED THE CAUSE OF THE EVENT AS PATIENT RELATED FACTOR AND STATED 'THE PATIENT HAS HIGH MYOPIA, POSTERIOR SCLERAL STPHYLOMA' AS THEIR CAUSE ANALYSIS." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082456&pc=MTA | 2023826-2021-04965 | 2021-05-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558556&pc=MTA | 2023826-2021-03809 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS GLARE/HALOS AND PILOCARPINE ADMINISTRATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856855&pc=MTA | 2023826-2023-01586 | 2023-03-06 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022 . THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED GLARE AND THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT ACCEPT THE GLARE AND REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18297786&pc=MTA | 2023826-2023-05508 | 2023-08-24 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18297821&pc=MTA | 2023826-2023-05509 | 2023-08-24 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916944&pc=MTA | 2023826-2023-04390 | 2023-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT COMPLAINED OF GLARE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK . D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. AT ONE MONTH POST-OP THE PATIENT COMPLAINED OF PERSISTENT GLARE AND THE LENS REMAINED IMPLANTED. POST-OP, THE PROBLEM IS GETTTING BETTER. THE CAUSE OF THE EVENT IS PATIENT-RELATED FACTOR, THE PATEINT IS A CAREFUL OBSERVER. CLAIM # 433379. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498032&pc=MTA | 2023826-2019-02401 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498117&pc=MTA | 2023826-2019-02400 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634639&pc=MTA | 2023826-2020-00204 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT THE PATIENT HAD SOME ISSUES OF GLARE AT NIGHT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. REFERENCE MFR# 2023826-2020-00202 FOR INITIAL LENS. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634784&pc=MTA | 2023826-2020-00205 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-01-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE LENS RESOLVED THE PROBLEM BUT THE PATIENT HAD SOME ISSUES OF GLARE AT NIGHT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. REFERENCE MFR# 2023826-2020-00203 FOR INITIAL LENS. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599727&pc=MTA | 2023826-2021-03909 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -6.5 INTO THE PATIENT'S EYE ON (B)(6) 2021. THE SURGEON REPORTS HALOS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT - UNK. INITIAL REPORTER -ESTABLISHMENT NAME/ADDRESS: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492552&pc=MTA | 2023826-2022-01382 | NA | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY GLARES/HALOES OBSERVED. TATTOO INK WAS INJECTED INTO CORNEAL POCKETS AND ANTERIOR CHAMBER, STAINING THE LENS. THE LENS WAS REMOVED ON (B)(6) 2022 AND REPLACED IN A SEPARATE SURGERY WITH AN IDENTICAL LENS. CAUSE IS DUE TO THE DEVICE - ICL WAS PERFECT. PATIENT HAD GLARE FROM PI, TATTOO WAS INJECTED INTO CORNEAL POCKETS. Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636051&pc=MTA | 2023826-2023-01065 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED ON (B)(6), 2023 DUE TO THE PATIENT COMPLAINED OF GLARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK. IMPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16638118&pc=MTA | 2023826-2023-01066 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311165&pc=MTA | 2023826-2023-05091 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF A -15.0 DIOPTER INTO THE PATIENT'S EYE. PATIENT SUBJECTIVE DISTURBANCE OF "GLARE AND HALOS" WAS REPORTED. THERE IS NO VISUAL ACUITY LOSS OR LOSS OF BCVA REPORTED AND THE LENS REMAINS IMPLANTED. THE ICL DIRECITONS FOR USE IDENTIFIES GLARE AND HALO VISUAL DISTURBANCE AS A KNOWN POTENTIAL ADVERSE EVENT. PATIENT RISK FACTORS INCLUDE PRE-EXISTING/CURRENT OCULAR STATUS AND MESOPIC PUPIL SIZE THAT IS GREATER THAN THE ICL OPTIC DIAMETER. BASED ON THE INFORMATION RECEIVED, THE ROOT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323156&pc=MTA | 2023826-2015-01725 | NA | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WILL BE EXPLANTED DUE TO THE LENS HAS BEEN SHIFTING AND THE PATIENT HAS BEEN EXPERIENCING VISUAL DISTURBANCE. THE SURGEON PLANS TO EXCHANGE THE LENS FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE SHIFTING OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. PER MEDICAL REVIEW: REPORTEDLY, MICL (13.2 MM) CHANGED THE POSITION ("SHIFTED") FOLLOWING IMPLANTATION IN (B)(6) 2015 CAUSING VISUAL DISTURBANCES TO PATIENT. THE SURGEON WAS PLANNING TO EXCHANGE THE LENS, IN (B)(6) 2015, FOR A LONGER LENGTH ICL BUT HE PERFORMED ONLY MANUAL ROTATION INSTEAD. ACCORDING TO EMAIL REPORT DATED ON (B)(6) 2015, THE MICL SHIFTED AGAIN SO THE SURGEON WAS PLANNING LENS EXCHANGE FOR MICL (13.7 MM) TO PREVENT FUTURE COMPLICATIONS. DESPITE MULTIPLE ATTEMPTS NO CONFIRMATION ABOUT LENS EXCHANGE WAS RECEIVED TO DATE. THE SERIOUSNESS IS BASED ON INFORMATION ABOUT SECONDARY SURGICALLY INTERVENTION PERFORMED (ROTATION).DEVICE CAUSALITY IS BASED ON INFORMATION THAT AT THE TIME OF THE SURGERY PATIENT WAS (B)(6) AND PER DFU :"INDICATIONS: THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187453&pc=MTA | 2023826-2023-05231 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. REFRACTION IS PLANO -0.25 BUT ISSUES WITH NIGHT VISION ARE REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16177977&pc=MTA | 2023826-2023-00153 | 2022-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED NIGHT TIME GLARE CAUSED BY LIGHTS. THE VA HAS BEEN GOOD AND PATIENT IS DOING WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT ALSO EXPERIENCED KERATOCONUS. H6: HEALTH EFFECT- CLINICAL CODE: 4581-KERATOCONUS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178186&pc=MTA | 2023826-2023-00154 | 2022-10-12 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED NIGHT TIME GLARE CAUSED BY LIGHTS. THE VA HAS BEEN GOOD AND PATIENT IS DOING WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- "THE PATIENT ALSO EXPERIENCED KERATOCONUS." H6- HEALTH EFFECT- CLINICAL CODE: 4581- KERATOCONUS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313440&pc=MTA | 2023826-2023-02767 | 2023-06-23 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS) ON(B)(6) 2020. IRIS TRANSILLUMINATION DEFECTS ARE REPORTED. CAUSE OF THE EVENT WAS THE DEVICE. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "ADVERSE EVENT STARTED 23JUNE23, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- IRIS TRANSILLUMINATION DEFECTS" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED IRIS TRANSILLUMINATION DEFECTS. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "ADVERSE EVENT STARTED (B)(6) 23, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE". ADDITIONAL INFORMATION PROVIDED: 31AUG23, THE TID SEEN AT THE 36-MONTH VISIT WAS VERY MILD. IT WAS DIFFICULT TO EVEN PHOTOGRAPH. CLINICIAN DID NOT NOTICE IT AT THE 24-MONTH VISIT OR OTHER PREVIOUS VISITS. NO OTHER CHANGES WERE NOTED IN SUBJECT'S HEALTH, MEDS, VISION AND REFRACTION WERE STABLE." THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313441&pc=MTA | 2023826-2023-02766 | 2023-06-23 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IRIS TRANSILLUMINATION DEFECTS ARE REPORTED. CAUSE OF THE EVENT WAS THE DEVICE. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "ADVERSE EVENT STARTED (B)(6) 2023, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- IRIS TRANSILLUMINATION DEFECTS" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED IRIS TRANSILLUMINATION DEFECTS. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "ADVERSE EVENT STARTED (B)(6) 2023, NOTED AS MILD AND RELATED TO DEVICE. (B)(6) 2023 WAS THE LAST PROTOCOL REQUIRED STUDY VISIT, SUBJECT EXITED SAME DAY AND WILL BE FOLLOWED PER STANDARD OF CARE". ADDITIONAL INFORMATION PROVIDED: "(B)(6) 2023 THE TID SEEN AT THE 36-MONTH VISIT WAS VERY MILD. IT WAS DIFFICULT TO EVEN PHOTOGRAPH. CLINICIAN DID NOT NOTICE IT AT THE 24-MONTH VISIT OR OTHER PREVIOUS VISITS. NO OTHER CHANGES WERE NOTED IN SUBJECT'S HEALTH, MEDS, VISION AND REFRACTION WERE STABLE." THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772254&pc=MTA | 2023826-2019-01177 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Discomfort | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 REPORTER INDICATED THAT LOW VAULT WITH TEMPORAL/INFERIOR DECENTRATION AND WITH "DOUBLING IMAGE AND DISCOMFORT WAS REPORTED. DOCTOR AND PATIENT DO NOT WANT MORE INTERVENTIONS AT THIS TIME." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT : LENS ROTATION (30 DEGREES) "AGAIN", LENS LOCATED OUTSIDE THE PUPIL, OFF-CENTER AND PATIENT "FEELS IMAGE UNFOLDING AGAIN." FOUND NO ZONULAR DEFECT AND "SURGEON WANTS TO EXPLANT FOR FEAR OF MAJOR PROBLEMS." CLAIM# (B)(4). Manufacturer Narrative: LENS ROTATION WAS ALSO REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852586&pc=MTA | 2023826-2020-02687 | 2020-09-12 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Discomfort | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT WAS UNCOMFORTABLE AND EXPERIENCED VISION FATIGUE. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279016&pc=MTA | 2023826-2023-02617 | 2022-08-06 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -01.25 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO PATIENT DISCOMFORT AND GLARE ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1IMPLANTABLE COLLAMER LENS OF DIOPTER -1.25 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED DISCOMFORT/ GLARE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "NO. THERE IS NO PLAN TO IMPLANT ANOTHER LENS." CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521049&pc=MTA | 2023826-2017-00696 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AND HALOS, WORSE IN AM AND AT NIGHT. THE PATIENTS PUPILS ARE 6.5MM IN DIM ILLUMINATION, ARE ESSENTIALLY EQUAL IN SIZE AND THERE IS NO EVIDENCE OF TRAUMATIC MYDRIASIS. THE PATIENT IS BEING TREATED WITH DROPS AND THE LENS REMAINS IMPLANTED. THE PATIENTS VISION IS 20/20 UNCORRECTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6521056&pc=MTA | 2023826-2017-00695 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT WAS EXPERIENCING GLARE AND HALOS, WORSE IN AM AND AT NIGHT. THE PATIENTS PUPILS ARE 6.5MM IN DIM ILLUMINATION, ARE ESSENTIALLY EQUAL IN SIZE AND THERE IS NO EVIDENCE OF TRAUMATIC MYDRIASIS. THE PATIENT IS BEING TREATED WITH DROPS AND THE LENS REMAINS IMPLANTED. THE PATIENTS VISION IS 20/20 UNCORRECTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020633&pc=MTA | 2023826-2017-01701 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAS BEEN BOTHERED BY PERSISTENT GLARE AND HALOS TWO MONTHS AFTER THE SURGERY. THE PATIENT HAS TRIED GLASSES AND MIOTICS BUT WITH NO IMPROVEMENT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 6 MONTHS. Manufacturer Narrative: CORRECTED DATA: THE EVENT IS CORRECTED TO "PRODUCT PROBLEM" FROM "ADVERSE EVENT". TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO "MALFUNCTION" FROM "SERIOUS INJURY". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020681&pc=MTA | 2023826-2017-01702 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAS BEEN BOTHERED BY PERSISTENT GLARE AND HALOS TWO MONTHS AFTER THE SURGERY. THE PATIENT HAS TRIED GLASSES AND MIOTICS BUT WITH NO IMPROVEMENT. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 6 MONTHS. Manufacturer Narrative: CORRECTED DATA: THE EVENT IS CORRECTED TO "PRODUCT PROBLEM" FROM "ADVERSE EVENT". TYPE OF REPORTABLE EVENT: IT IS CORRECTED TO "MALFUNCTION" FROM "SERIOUS INJURY". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930935&pc=MTA | 2023826-2020-02813 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930952&pc=MTA | 2023826-2020-02812 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738945&pc=MTA | 2023826-2021-01199 | 2020-09-29 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON 29/SEP/2020. REPORTEDLY, LENS TOO LONG, SURGEON PREFERENCE TO CHANGE. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH, HIGHER POWER LENS. PER PATIENT'S (B)(6) 2021 VISIT, GLARE AND HALOS ARE BEING EXPERIENCED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7945358&pc=MTA | 2023826-2018-01501 | 2018-08-14 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS), THE HAPTICS TORE DURING INJECTION. THE SURGEON STATES THE LENS IS "TILTED AND DISLOCATED" AND "THE LENS WENT OUT WITH A BROKEN HAPTIC." THE LENS WAS IMPLANTED ON (B)(6) 2018 AND REPORTEDLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS LISTED AS "LENS TORN." ADDITIONALLY, THE SURGEON REPORTS LENS DISLOCATION/SUBLUXATION AND GLARE/HALOS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: SERIOUS INJURY CORRECTED TO MALFUNCTION IN SUPPLEMENTAL #1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7578068&pc=MTA | 2023826-2018-00864 | NA | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6, -10.5 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HALOS AND GLARE. IT WAS REPORTED THAT THE HALOS AND GLARES WERE NOTICED ON THE SAME DAY AS THE IMPLANT SURGERY. THERE IS NO PLANS TO IMPLANT ANOTHER LENS. PATIENT IS STABLE BUT STILL HAS GLARE FROM PI. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717807&pc=MTA | 2023826-2020-00333 | 2020-01-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SEVERE GLARE AND HALOES. THE REPORTER STATED THAT THE SEVERE HALOES MAY BE DUE TO VAULT PUSHING PUPIL WIDER. THE LENS REMAINS IMPLANTED. THE DOCTOR PLANS ON EXCHANGING THE LENS FOR A SHORTER LENGTH LENS BUT THE PATIENT IS CURRENTLY PREGNANT, SO HE WILL HAVE TO WAIT UNTIL SHE DELIVERS LATER THIS YEAR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717808&pc=MTA | 2023826-2020-00332 | 2020-01-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND SEVERE GLARE AND HALOES. THE REPORTER STATED THAT THE SEVERE HALOES MAY BE DUE TO VAULT PUSHING PUPIL WIDER. THE LENS REMAINS IMPLANTED. THE DOCTOR PLANS ON EXCHANGING THE LENS FOR A SHORTER LENGTH LENS BUT THE PATIENT IS CURRENTLY PREGNANT, SO HE WILL HAVE TO WAIT UNTIL SHE DELIVERS LATER THIS YEAR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403835&pc=MTA | 2023826-2019-00347 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED A 12.6MM, MICL12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS EYE. THE SURGEON REPORTS EXCESSIVE VAULT, GLARE, AND HALOS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091302&pc=MTA | 2023826-2020-03208 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO THE LENS HAD A HIGHER VAULT AND WAS CAUSING BAD HALO AND GLARE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON 30-OCT-2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CORRECTED DATA: B3: DATE OF EVENT - (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636085&pc=MTA | 2023826-2018-00952 | 2018-04-10 | Malfunction | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE PATIENT COMPLAINED OF GLARE AND HALOS AND COULD SEE THE BORDER OF THE LENS. THE LENS REMAINS IMPLANTED. SEE MFR REPORT # 2023826-2018-00951 FOR INITIAL LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223261&pc=MTA | 2023826-2023-05166 | 2023-09-28 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223321&pc=MTA | 2023826-2023-05167 | 2023-09-28 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195876&pc=MTA | 2023826-2022-02607 | 2021-10-01 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE PATIENT DID NOT WANT THE LENS REPLACED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195903&pc=MTA | 2023826-2022-02608 | 2021-10-01 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE PATIENT DID NOT WANT THE LENS REPLACED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921079&pc=MTA | 2023826-2022-04454 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921081&pc=MTA | 2023826-2022-04455 | 2022-09-06 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295992&pc=MTA | 2023826-2023-02835 | 2023-03-21 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295993&pc=MTA | 2023826-2023-02836 | 2023-03-21 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722680&pc=MTA | 2023826-2023-03959 | 2023-03-24 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715576&pc=MTA | 2023826-2023-03874 | 2023-06-24 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715689&pc=MTA | 2023826-2023-03875 | 2023-06-24 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18115810&pc=MTA | 2023826-2023-04768 | 2023-09-26 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167101&pc=MTA | 2023826-2021-02514 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MANUFACTURER: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167100&pc=MTA | 2023826-2021-02513 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: G3: CORRECTED TO 21-JUN-2021 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16579641&pc=MTA | 2023826-2023-00988 | 2023-01-09 | Injury | STAAR SURGICAL COMPANY | 2023-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS GLARE AND HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS NOTED "AFTER IMPLANTATION THE PATIENT SHOWED SYMPTOMS OF GLARE; AND THE CLINICIAN TRIED TO COMMUNICATE WITH HIM FOR A MONTH; BUT STILL COULD NOT ADAPT; THE PATIENT STRONGLY REQUESTED REMOVAL". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16580026&pc=MTA | 2023826-2023-00989 | 2023-02-09 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS GLARE AND HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS NOTED "AFTER IMPLANTATION THE PATIENT SHOWED SYMPTOMS OF GLARE; AND THE CLINICIAN TRIED TO COMMUNICATE WITH HIM FOR A MONTH; BUT STILL COULD NOT ADAPT; THE PATIENT STRONGLY REQUESTED REMOVAL". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18508980&pc=MTA | 2023826-2024-00216 | 2023-11-30 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Distortion; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 30-NOV-2023 THE LENS WAS EXCHANGED FOR A LONGER LENS OF THE SAME MODEL AND POWER DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750667&pc=MTA | 2023826-2021-01280 | 2018-09-07 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 14 AUG 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751109&pc=MTA | 2023826-2021-01279 | 2018-09-07 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307835&pc=MTA | 2023826-2019-00226 | 2019-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS GLARE/HALOS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273785&pc=MTA | 2023826-2021-00245 | 2019-09-11 | Injury | STAAR SURGICAL COMPANY | 2021-02-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO _13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. PER THE REPORTER ON (B)(6) 2021, ANOTHER LENS WILL NOT BE IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385986&pc=MTA | 2023826-2022-00343 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13386000&pc=MTA | 2023826-2022-00342 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10819682&pc=MTA | 2023826-2020-02645 | 2020-07-19 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE SURGEON PLANS ON IMPLANTING A REPLACEMENT LENS BUT THE DATE IS NOT FIXED YET. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10819835&pc=MTA | 2023826-2020-02644 | 2020-07-19 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE SURGEON PLANS ON IMPLANTING A REPLACEMENT LENS BUT THE DATE IS NOT FIXED YET. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112700&pc=MTA | 2023826-2021-02210 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF EVENT. IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. THE REPORTER STATED "AFTER THE OPERATION, THE PATIENTS VISION IS NOT ADAPTED AND CANNOT BE RELIEVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507468&pc=MTA | 2023826-2021-00669 | 2020-08-31 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY IMMEDIATELY AFTER THE ICL SURGERY THE PATIENT HAD GOOD VISION. HOWEVER, HE COMPLAINED OF POSTOPERATIVE HALOS ESPECIALLY IN THE DARK. ONE MONTH AFTER SURGERY THE SITUATION WAS WORSE. THE PATIENT HAD DIFFICULTY GOING OUTSIDE AND DRIVING AT NIGHT. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507480&pc=MTA | 2023826-2021-00668 | 2020-08-31 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY IMMEDIATELY AFTER THE ICL SURGERY THE PATIENT HAD GOOD VISION. HOWEVER, HE COMPLAINED OF POSTOPERATIVE HALOS ESPECIALLY IN THE DARK. ONE MONTH AFTER SURGERY THE SITUATION WAS WORSE. THE PATIENT HAD DIFFICULTY GOING OUTSIDE AND DRIVING AT NIGHT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507478&pc=MTA | 2023826-2021-00673 | 2020-09-19 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN WITH RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507482&pc=MTA | 2023826-2021-00674 | 2020-09-19 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE, RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820901&pc=MTA | 2023826-2021-01399 | 2020-09-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6). 2020. GLARES AND HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. PATIENT UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. GLARES AND HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. PER EMAIL FROM (B)(6) 2021 IT WAS REPORTED TO "PLEASE ACCEPT THE FORM AS 'RESOLVED', PATIENT WILL KEEP ICL FOR NOW." Manufacturer Narrative: B3 - CORRECTION: UNK. B5 - DATE WHEN THE LENS WAS IMPLANTED CORRECTED TO SAY (B)(6) 2021. B5 - ADDITIONAL INFORMATION RECEIVED: CLAIM CAN BE ACCEPTED AS "RESOLVED" PER EMAIL FROM (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: DATE RECD BY MFR - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11758251&pc=MTA | 2023826-2021-01305 | 2021-02-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOS. THE PATIENT WANTED THE LENS TO BE REMOVED AND ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11758266&pc=MTA | 2023826-2021-01304 | 2021-02-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF GLARE/HALOS. THE PATIENT WANTED THE LENS TO BE REMOVED AND ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT : UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11853624&pc=MTA | 2023826-2021-01564 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VICM5_12.6, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. FOR ADDITIONAL INFORMATION "PATIENT WEARS CONTACTLENS AND GLASSES" WAS REPORTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233836&pc=MTA | 2023826-2021-02868 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 17/DEC/2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. REPORTEDLY, THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233910&pc=MTA | 2023826-2021-02869 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. REPORTEDLY, THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477684&pc=MTA | 2023826-2022-03177 | 2022-07-25 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT COMPLAINED OF GLARE/HALOES. THE PROBLEM WAS RESOLVED AND THE PATIENT WAS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477694&pc=MTA | 2023826-2022-03176 | 2022-07-25 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT COMPLAINED OF GLARE/HALOES. THE PROBLEM WAS RESOLVED AND THE PATIENT WAS HAPPY. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562170&pc=MTA | 2023826-2022-03519 | 2022-08-01 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY AND WISHES NO FURTHER IMPLANTATION. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379628&pc=MTA | 2023826-2023-03112 | 2023-04-24 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075195&pc=MTA | 2023826-2023-04727 | 2023-07-15 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT REQUESTED LENS TO BE REMOVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17978639&pc=MTA | 2023826-2023-04558 | 2023-08-26 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852200&pc=MTA | 2023826-2019-01330 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED GLARE AND HALOS AND THE PLAN IS TO EXPLANT THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965149&pc=MTA | 2023826-2020-00905 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT COMPLAINED OF GLARE AND HALOS. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM IN INITIAL MDR. B2 - N/A IN INITIAL MDR. H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13656840&pc=MTA | 2023826-2022-00649 | NA | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT COMPLAINED OF GLARE/HALOES AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6)2022 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, IMPLANT DATE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664267&pc=MTA | 2023826-2022-00650 | NA | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT COMPLAINED OF GLARE/HALOES AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17620296&pc=MTA | 2023826-2023-03602 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2022. GLARE/HALOES WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H4 - DEVICE MANUFACTURE DATE CORRECTED TO 18-JAN-2022. CLAIM #(B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17620390&pc=MTA | 2023826-2023-03604 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2022. GLARE/HALOES WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 : TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18079189&pc=MTA | 2023826-2023-04799 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. GHOSTING IMAGES ARE REPORTED AND PATIENT IS BEING MONITORED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18079674&pc=MTA | 2023826-2023-04800 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. GHOSTING IMAGES ARE REPORTED AND PATIENT IS BEING MONITORED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16091922&pc=MTA | 2023826-2023-00045 | 2022-12-03 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY, "MULTIPLE VISUAL DISTURBANCES OVER THE WEEKEND." Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "SEVERE NON-TOLERABLE GLARE AND HALOS. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS STATED AS, "GLARE AND HALOS RESOLVED FOLLOWING EXPLANTATION." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. CLAIM#(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16028488&pc=MTA | 2023826-2022-04635 | 2022-10-18 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARE/HALOES AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. HOWEVER, THE REPORTER STATED THAT THE LENS FAILED TO PERFORM AS INTENDED. PER UPDATED INFORMATION THE REPORTER STATES"THE PATIENT HAD CALIPER W-W 12.1 MM OU, QUITE DEEP AC-S (3.4-3.5MM), AND WAS SUGGESTED BY OCOS, AND I WAS EXPECTING IT TOO, THAT THE SIZE 13.2MM ICL SHOULD FIT THE SULCUS PERFECTLY. BUT, OBVIOUSLY DUE TO SOME POSTERIOR SULCUS ANATOMY VARIATIONS, IT HAD EXCESSIVE VAULT AND DID NOT WORK AS EXPECTED." Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16029118&pc=MTA | 2023826-2022-04636 | 2022-10-18 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -2.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARE/HALOES AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. HOWEVER, THE REPORTER STATED THAT THE LENS FAILED TO PERFORM AS INTENDED. PER UPDATED INFORMATION THE REPORTER STATES"THE PATIENT HAD CALIPER W-W 12.1 MM OU, QUITE DEEP AC-S (3.4-3.5MM), AND WAS SUGGESTED BY OCOS, AND I WAS EXPECTING IT TOO, THAT THE SIZE 13.2MM ICL SHOULD FIT THE SULCUS PERFECTLY. BUT, OBVIOUSLY DUE TO SOME POSTERIOR SULCUS ANATOMY VARIATIONS, IT HAD EXCESSIVE VAULT AND DID NOT WORK AS EXPECTED." Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO THE PATIENT'S LEFT EYE (OS). CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924575&pc=MTA | 2023826-2023-01803 | 2023-03-20 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES AND MILD DILATED PUPIL. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581:SIGNIFICANT REDUCTION OF ICA, MILD DILATED PUPIL. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14655726&pc=MTA | 2023826-2022-01785 | NA | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT,SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARE/HALOES. THE EXCHANGE RESOLVED THE PROBLEM. INCISION ENLARGEMENT WAS PERFORMED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK . CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787767&pc=MTA | 2023826-2023-01374 | 2023-02-20 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. AT 3 WEEKS POST-OP, THE PATIENT COMPLAINED OF DYSPHOTOPSIA AND GLARE/HALOES. THIS WAS VERY DISTRESSING TO THE PATIENT BOTH DAY AND NIGHT. THE DOCTOR FELT THIS WAS RELATED TO THE CENTRAFLOW HOLE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. POST-OP, THE PATIENT IS BACK TO HIS PRE-OP STATUS. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN GAUZE, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTICS TORN. THERE WAS DEBRIS AND RESIDUE THROUGHOUT THE LENS. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: A4: UNKNOWN. A5: UNKNOWN. A6: UNKNOWN. D6B: UNKNOWN. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DYSPHOTOPSIA. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15910698&pc=MTA | 2023826-2022-04371 | 2022-07-27 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, DIZZINESS AND EYE DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15716593&pc=MTA | 2023826-2022-03885 | 2022-02-01 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER FROM THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO PATIENT WAS DISSATISFIED WITH VISION DUE TO GLARE/HALOES AND THE PROBLEM WAS RESOLVED. POST-OP, THE STATUS OF THE EYE WAS NORMAL. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATIION: B5- " ON (B)(6)2022 THE LENS WAS EXPLANTED." CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15717531&pc=MTA | 2023826-2022-03886 | 2022-02-01 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, FROM THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO PATIENT WAS DISSATISFIED WITH VISION DUE TO GLARE/HALOES AND THE PROBLEM WAS RESOLVED. POST-OP, THE STATUS OF THE EYE WAS NORMAL. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATIION: B5- " ON (B)(6) 2022 THE LENS WAS EXPLANTED." CLAIM# (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11384756&pc=MTA | 2023826-2021-00535 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, MICL12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS LENS WAS GOING TO BE EXPLANTED DUE EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: IT WAS REPORTED THAT A 12.6MM, MICL12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO BORDERLINE EXCESSIVE VAULT. PATIENT-RELATED FACTOR WAS REPORTED FOR CAUSE OF EVENT, REPORTER STATED " PATIENT UNHAPPY WITH VISION AFTER SURGERY (HALO/GLARE) AT NIGHT TIME. DISSATISFIED WITH INTRAOP DISCOMFORT." IT WAS REPORTED THAT THE PATIENT DECIDED ON ICL REMOVAL, WANTED TO WEAR GLASSES AGAIN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11384762&pc=MTA | 2023826-2021-00534 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, MICL12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS LENS WAS GOING TO BE EXPLANTED DUE EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: IT WAS REPORTED THAT A 12.6MM, MICL12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO BORDERLINE EXCESSIVE VAULT. PATIENT-RELATED FACTOR WAS REPORTED FOR CAUE OF EVENT, REPORTER STATED " PATIENT UNHAPPY WITH VISION AFTER SURGERY (HALO/GLARE) AT NIGHT TIME. DISSATISFIED WITH INTRAOP DISCOMFORT." IT WAS REPORTED THAT THE PATIENT DECIDED ON ICL REMOVAL, WANTED TO WEAR GLASSES AGAIN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091314&pc=MTA | 2023826-2020-03209 | 2020-12-14 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS HAD A HIGHER VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR (HIGH VAULT CAUSING BAD HALO/GLARE). CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7896040&pc=MTA | 2023826-2018-01415 | 2018-05-22 | Injury | STAAR SURGICAL COMPANY | 2018-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -7 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, EXPLANT. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE AND A TORN/ BENT HAPTIC. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824059&pc=MTA | 2023826-2018-01286 | 2018-07-30 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B) (6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6507577&pc=MTA | 2023826-2017-00678 | 2016-03-19 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARE/ HALOES, AND UNDERWENT ADDITIONAL REFRACTIVE TREATMENT. ON (B)(6) 2017 THE LENS WAS EXPLANTED. ATTEMPTS TO CONTACT REPORTER FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ON (B)(6) 2017 THE V4C LENS WAS EXPLANTED AND EXCHANGED FOR THE V4 LENS. THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: LENS WAS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6624647&pc=MTA | 2023826-2017-00912 | 2016-12-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT REPORTED EXPERIENCING GLARES AND HALOS. THE LENS CURRENTLY REMAINS IMPLANTED BUT THE DOCTOR IS PLANNING TO EXPLANT THE LENS. THE PATIENT'S NEXT VISIT IS SCHEDULED ON (B)(6) 2017 AND THE DOCTOR WILL SCHEDULE THE SURGERY DATE TO EXPLANT THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE INITIAL LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2017. THE PATIENT'S POST-OP INDICATES THAT THE PATIENT HAS INTRACTABLE PERIPHERAL GLARE IN SPITE OF EXCELLENT ICL PLACEMENT AND IS UNABLE TO DRIVE AT NIGHT. H3-DEVICE EVALUATION: THE LENS SPECIMEN WAS RETURNED IN A SPECIMEN CUP AND THERE WAS SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: LENS REMAINS IMPLANTED. PLANNED EXPLANT DATE IS UNDETERMINED. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6625059&pc=MTA | 2023826-2017-00913 | 2016-12-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A MICL12.6, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT REPORTED EXPERIENCING GLARES AND HALOS. THE LENS CURRENTLY REMAINS IMPLANTED BUT THE DOCTOR IS PLANNING TO EXPLANT THE LENS. THE PATIENT'S NEXT VISIT IS SCHEDULED ON (B)(6) 2017 AND THE DOCTOR WILL SCHEDULE THE SURGERY DATE TO EXPLANT THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE INITIAL WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2017. THE PATIENT'S POST-OP INDICATES THAT THE PATIENT HAS INTRACTABLE PERIPHERAL GLARE IN SPITE OF EXCELLENT ICL PLACEMENT AND IS UNABLE TO DRIVE AT NIGHT. THE LENS WAS RETURNED IN A SPECIMEN CUP AND IT HAD SOME MOISTURE ON IT. VISUAL INSPECTION FOUND A PIECE OF THE LENS HAPTIC MISSING. (B)(4). Manufacturer Narrative: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300927&pc=MTA | 2023826-2018-00290 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED DUE TO GLARE/HALOES. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955569&pc=MTA | 2023826-2017-01584 | 2017-03-29 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT COMPLAINING OF EXTREME GLARE. NO ADDITIONAL DETAILS ARE KNOWN AT THIS TIME. Manufacturer Narrative: ADDITIONAL SYMPTOM OF HALOS REPORTED. REPORTEDLY, SYMPTOMS BEGAN ON (B)(6) 2017 AND BECAME WORSE AT NIGHT. THE SYMPTOMS WERE RESOLVED WITH THE LENS EXPLANT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND PIECES OF THE HAPTIC MISSING. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955572&pc=MTA | 2023826-2017-01583 | 2017-03-29 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT COMPLAINING OF EXTREME GLARE. NO ADDITIONAL DETAILS ARE KNOWN AT THIS TIME. Manufacturer Narrative: ADDITIONAL SYMPTOM OF HALOS REPORTED. REPORTEDLY, SYMPTOMS BEGAN ON (B)(6) 2017. SYMPTOMS WERE WORSE AT NIGHT. THE PROBLEM WAS RESOLVED WITH EXPLANT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281442&pc=MTA | 2023826-2018-00256 | 2017-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO GLARE/HALOES. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND DEBRIS ON LENS. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456990&pc=MTA | 2023826-2018-00596 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO GLARES AND HALOES. IT WAS REPORTED THAT THE PATIENT HAS RECOVERED AND THERE IS NO MORE PROBLEMS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7580116&pc=MTA | 2023826-2018-00857 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890304&pc=MTA | 2023826-2017-01472 | NA | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT NOT BEING HAPPY WITH THE HALOES AND GLARES. THERE WAS NO PATIENT INJURY AND THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. NO OTHER INFORMATION WAS PROVIDED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: SECONDARY SURGERY, LENS EXPLANTED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7578161&pc=MTA | 2023826-2018-00865 | NA | Injury | STAAR SURGICAL COMPANY | 2018-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -12.5 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HALOS AND GLARES. IT WAS REPORTED THAT THE HALOS AND GLARES WERE NOTICED ON THE SAME DAY AS THE IMPLANT SURGERY. THERE IS NO PLANS TO IMPLANT ANOTHER LENS. PATIENT IS STABLE BUT STILL HAS GLARE FROM PI. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041209&pc=MTA | 2023826-2016-01089 | 2014-10-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARES/HALOES AND BLURRED VISION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE DEVICE WAS WITHIN SPECIFICATIONS. (B)(4).- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6233468&pc=MTA | 2023826-2017-00036 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. THE LENS WAS EXCHANGED FOR ANOTHER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DEVICE CODE: (B)(4) (LENS EXCHANGED FOR ANOTHER LENS) TO (B)(4) - PREVIOUS CODE IS NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE LISTED UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253843&pc=MTA | 2023826-2017-00079 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS RETURNED DRY AND BENT IN THE LENS CONTAINER. VISUAL INSPECTION FOUND DISCOLORED FIBERS ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701668&pc=MTA | 2023826-2017-01153 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE/HALOES, SLOW PUPILLARY REACTION, OVERCORRECTION, AND SEES "CENTERFLOW" REFLECTIVES EVEN ON DAYTIME. THE CUSTOMER STATED THAT THE SURGEON PLANNED TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6004348&pc=MTA | 2023826-2016-01389 | 2016-08-18 | Injury | STAAR SURGICAL COMPANY | 2016-10-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP ((B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/12. DEVICE EVALUATION - THE LENS WAS RETURNED BENT AND DRY AND BENT IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE CORRECT REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172601&pc=MTA | 2023826-2016-01749 | 2016-10-29 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253621&pc=MTA | 2023826-2017-00076 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IRIS, ENLARGED PUPIL, AND GLARE/HALOES. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS. PROBLEM IS RESOLVED. AS OF THE DATE OF THE SUPPLEMENTAL SUBMISSION, THE LENS HAS NOT BEEN RETURNED TO THE MFG FOR EVALUATION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2017, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN/BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6253661&pc=MTA | 2023826-2017-00077 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IRIS, ENLARGED PUPIL, AND GLARE/HALOES. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF THIS SUPPLEMENTAL, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT ALSO EXPERIENCED DIFFICULTIES WITH NEAR VISION. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6629954&pc=MTA | 2023826-2017-00928 | 2017-04-08 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), GALRE/HALOES, AND IRIS ATROPHY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. ON (B)(6) 2017, THE PATIENT LAST VISIT DATE, THE SURGEON COMMENTED THAT THE PROBLEM WAS RESOLVED, BUT THE PATIENT STILL HAS MYDRIASIS INTERMEDIATE, IRIS ATROPHY, AND MILD PHOTOPHOBIA. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. ATTEMPTS TO CONTACT TO THE CUSTOMER HAVE BEEN ON-GOING. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636064&pc=MTA | 2023826-2018-00951 | 2018-03-14 | Injury | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT STILL COMPLAINED OF GLARE AND HALOS AND COULD SEE THE BORDER OF THE LENS. SEE MFR REPORT # 2023826-2018-00952 FOR EXCHANGED LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL TESTING WAS PERFORMED ON THE LENS AND THE LENS PASSED SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7440966&pc=MTA | 2023826-2018-00569 | 2018-03-23 | Injury | STAAR SURGICAL COMPANY | 2018-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PIGMENT DISPERSION, AND GLARE/HALOS. REPORTEDLY, THE PATIENT SEES HALOS AND GLARE AT DUSK AND CAN'T DRIVE. THE SURGEON IS CONTINUING TO MONITOR THE PATIENT; THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULT SIGNIFICANT REDUCTION OF ICA PIGMENT DISPERSION AND GLARES/ HALOS. REPORTEDLY THE PATIENT SEES GLARE AND HALOS AT DUSK AND CAN'T DRIVE. THE SURGEON IS CONTINUING TO MONITOR THE PATIENT. THE REPORTER ALSO STATED THE IRIS IS REACTIVE AND THAT IN THE PICTURE OF THE NON-DILATED PUPIL; THE LENS LOOKS TILTED ON ONE SIDE. THE LENS REMAINS IMPLANTED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495924&pc=MTA | 2023826-2018-00714 | 2018-03-23 | Injury | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PIGMENT DISPERSION, AND GLARE/HALOS. REPORTEDLY, THE PATIENT SEES HALOS AND GLARE AT DUSK AND CAN'T DRIVE. THE SURGEON IS CONTINUING TO MONITOR THE PATIENT; THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740006&pc=MTA | 2023826-2018-01143 | 2018-06-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT AND LENS SURFACE. THERE WAS THICK RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6535419&pc=MTA | 2023826-2017-00724 | 2017-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, GLARE/HALOES, AND DECENTERED LENS. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.0MM). (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229463&pc=MTA | 2023826-2019-00027 | 2018-11-24 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT COMPLAINED OF GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546104&pc=MTA | 2023826-2017-00769 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -6.5 DIOPTER. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE PATIENT DID NOT EXPERIENCE A VISIBLE DISTORTION TO THE IRIS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912815&pc=MTA | 2023826-2019-01510 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. REPORTEDLY, THE PATIENT EXPERIENCED "EXTREMELY BAD" GLARES AND HALOS THAT MADE HER NOT ABLE TO DRIVE AT NIGHT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: "03-JUL-2019" SHOULD BE CORRECTED TO "07-AUG-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912812&pc=MTA | 2023826-2019-01513 | 2019-07-09 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, AND SIGNIFICANT REDUCTION OF ICA. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DATE OF THIS REPORT: 26-JUL-2019 SHOULD BE CORRECTED TO 29-JUL-2019. DATE RECEIVED BY MANUFACTURER: 26-JUL-2019 SHOULD BE CORRECTED TO 29-JUL-2019 IN INITIAL MEDWATCH REPORT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) (CORRECTED FROM CLAIM# PREVIOUSLY REPORTED). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962458&pc=MTA | 2023826-2019-01585 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, UNREATIVE (FIXED) PUPIL AND GLARE/HALOS. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. AT ONE DAY POST-OP, IOP WAS 16 AND THERE WAS SOME EDEMA. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THE LENS WAS RETURNED IN A VIAL, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203031&pc=MTA | 2023826-2019-01952 | 2019-09-02 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH LENS ROTATION AND GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY SURGEON NOTED "HALOS/GLARE PRESENT BOTH WITH SMALL AND DILATED PUPIL. RELATED IN MY OPINION TO CENTRAL HOLE." CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE(B)(6): SMALL DILATED PUPIL SHOULD HAVE BEEN INCLUDED IN INITIAL MDR ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LONGER LENGTH LENS ON 14-JAN-2020. THE PROBLEM WAS RESOLVED. REPORTEDLY, 'GLARE AND HALOS ARE GONE.' H6 - PATIENT CODE (B)(6): SECONDARY SURGICAL INTERVENTION - LENS EXCHANGE CLAIM# 714773. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233832&pc=MTA | 2023826-2019-00073 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 AND WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, LENS DISLOCATION/SUBLUXATION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. AFTER ICL EXCHANGE THE INTRAOCULAR PRESSURE REMAINED ELEVATED. A SECONDARY YAG WAS PERFORMED FOR ENLARGEMENT OF THE PERIPHERAL IRIDECTOMIES. THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS, CAUSING ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). SEE MFR REPORT # 2023826-2019-00071 FOR EXCHANGED LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842714&pc=MTA | 2023826-2020-00517 | 2020-02-18 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOS. THE PROBLEM WAS RESOLVED HOWEVER, PATIENT'S CONDITION AT FINAL VISIT IS MARKED AS "MODERATE CORNEAL EDEMA." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR AND "ANATOMICAL." REFERENCE MFR FLE# (B)(4) FOR 2ND IMPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518762&pc=MTA | 2023826-2018-00735 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC MISSING. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7518794&pc=MTA | 2023826-2018-00734 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON 06/21/2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9247857&pc=MTA | 2023826-2019-02032 | NA | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: PER PATIENT EMAIL HAD AN OD IMPLANT DATE OF "ON OR AROUND (B)(6) 2017". PATIENT STATES "SYMPTOMS CURRENTLY ON RIGHT EYE - FOGGINESS, RIPPLES OF WATER EFFECT, GHOSTING, HALOS, LOSS OF NIGHT VISION. SEEKING HELP TO REFUND REMOVAL OF RIGHT ICL LENS." INFORMATION RECEIVED FORM THE IMPLANTING SURGEON INDICATED THAT A 13.2MM, MICL13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. FOR CAUSE OF THIS EVENT THE SURGEON REPORTED "NO ADVERSE EVENT SURGICALLY". PRK TREATMENT WAS PERFORMED ON (B)(6) 2018. THE LENS REMAINS IMPLANTED. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS UCVA WAS 20/20-3. CLEAR CORNEA, SUPERIOR PI, CLEAR ICL WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). ADDITONAL INFORMATION STATES, "PATIENT C/O GLARE/HALOS SINCE ORIGINAL ICL IMPLANTED > AT NIGHT". Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: 2993 SHOULD BE CORRECTED TO CODE 1401 AND 3001. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552450&pc=MTA | 2023826-2019-00654 | 2019-02-18 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427043&pc=MTA | 2023826-2019-02297 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO GLARE/HALOES. PER REPORTER, "PATIENT EXPERIENCED CONCENTRIC CIRCLES OF LIGHT THRU LENS HOLD - INTOLERABLE." PATIENT WILL NOT BE PROCEEDING WITH ANOTHER LENS." Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT) WAS REPORTED IN INITIAL MEDWATCH REPORT MANUFACTURER NARRATIVE BUT WAS NOT SELECTED IN H6. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138814&pc=MTA | 2023826-2020-01191 | 2020-05-22 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED AFTER IMPLANTABLE COLLAMER LENS SURGERY. STARBURST AND FLOATERS ARE BEING OBSERVED. REPORTEDLY, HIGH REFRACTIVE POWER (AROUND -16 D) AND PATIENT NOTED STARBURST, FLOATERS AND HALOS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VIICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE REPORTER STATED THAT THE PATIENT "HAD HIGH REFRACTIVE POWER AROUND -16 DIOPTER, AND THE SURGEON TOLD HIM THAT FLOATERS ARE NORMAL. BUT HE WAS WORRIED ABOUT STARBURST AS BECAUSE OF IT HE COULD NOT DRIVE IT SEEMS." THE CLINIC SUGGESTED TO USE PILO (MEDICATION) FOR A MONTH (DURING NIGHT TIME) AND HE SAID FOR 3 DAYS HE WAS FINE AND STARTED HAVING THE SAME PROBLEM BEYOND THAT. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H6 - HEALTH EFFECT IMPACT CODE 4644 - PILO MEDS. H6 - MEDICAL DEVICE PROBLEM CODE 1494- OFF-LABEL (AGE< 21 YEARS). H6 - MEDICAL DEVICE PROBLEM CODE 3189 IS NOT APPLICABLE IN INITIAL MDR. H6 - INVESTIGATION FINDINGS CODE 213 IS CORRECTED TO 114 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609206&pc=MTA | 2023826-2019-00840 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND "ACHE." REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE LENS HAPTIC IS BENT. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609210&pc=MTA | 2023826-2019-00839 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND "ACHE." REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7943031&pc=MTA | 2023826-2018-01475 | 2018-09-03 | Injury | STAAR SURGICAL COMPANY | 2018-10-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. PATIENT COMPLAINED OF CIRCLES OVER THE LIGHT. ON (B)(6) 20218 THE LENS WAS EXPLANTED. POST OP 1 WEEK PATIENT'S EYE IS QUITE. NO COMPLAINTS OF HALOS POST OPERATIVE. Manufacturer Narrative: CORRECTED DATA: DEVICE MANUFACTURE DATE CORRECTED TO: "11-JAN-2018". (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: PT INFO: - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255304&pc=MTA | 2023826-2019-00106 | 2018-09-14 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS. THE PATIENT WAS SEEING THE EDGE OF THE LENS DUE TO PUPIL SIZE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/SAME LENGTH AND DIOPTER LENS, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895490&pc=MTA | 2023826-2019-01452 | 2019-03-02 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM WAS RESOLVED, PATIENT CONDITION WAS STABLE. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593229&pc=MTA | 2023826-2019-00751 | 2019-03-26 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593245&pc=MTA | 2023826-2019-00753 | 2019-04-02 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND TEARS ON THE LENS HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8754986&pc=MTA | 2023826-2019-01137 | 2019-05-16 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO GLARE/HALOES. THE LENS WAS LATER REPLACED AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9458068&pc=MTA | 2023826-2019-02387 | 2019-05-16 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS SATISFIED AND HAPPY, AND HAS NO PROBLEM WITH GLARES ANYMORE." Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191 SHOULD HAVE BEEN SELECTED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9458110&pc=MTA | 2023826-2019-02388 | 2019-05-16 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS SATISFIED AND HAPPY, AND HAS NO PROBLEM WITH GLARES ANYMORE." Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191 SHOULD HAVE BEEN SELECTED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188673&pc=MTA | 2023826-2019-01948 | 2019-07-06 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND POWER, DIFFERENT MODEL LENS DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE; THE REPORTER STATES, "SEVERE HALOS DUE TO AQUAPORT..." Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912810&pc=MTA | 2023826-2019-01500 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON EXPLANTED THE LENS DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED, IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WERE TORN. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912824&pc=MTA | 2023826-2019-01501 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON EXPLANTED THE LENS DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279653&pc=MTA | 2023826-2019-02062 | 2019-09-09 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND POWER, DIFFERENT MODEL LENS DUE TO GLARE/HALOS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254208&pc=MTA | 2023826-2020-01452 | 2020-02-24 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CLAIM #(B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033109&pc=MTA | 2023826-2020-01030 | 2020-04-24 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO GLARE/ HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS WEARING CONTACT LENSES AGAIN." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033110&pc=MTA | 2023826-2020-01029 | 2020-04-24 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO GLARE/ HALOES. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS WEARING CONTACT LENSES AGAIN." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965138&pc=MTA | 2023826-2020-00904 | NA | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. TWO MONTHS POST-OP, THE PATIENT COMPLAINED OF GLARE AND HALOS. THE SURGEON PERFORMED A SECOND SURGERY AND PLACED THE LENS VERTICALLY, HOPING THE HALOS WOULD DISAPPEAR, BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282312&pc=MTA | 2023826-2019-02077 | 2019-10-07 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED GLARE/HALOS. THE PATIENT ALSO HAD AN UNREACTIVE IRIS AND DILATED (FIXED) PUPIL. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10162874&pc=MTA | 2023826-2020-01295 | 2019-10-21 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED, REPORTER STATED " AFTER ICL IMPLANTATION SURGERY THE PATIENT'S VISION AND EYE CONDITIONS WERE VERY GOOD. HOWEVER, THE PATIENT COMPLAINED OF SEVERE GLARE AND HALOES AND COULD NOT ACCEPT IT. THE PATIENT WAS UNWILLING TO ACCEPT OTHER TREATMENT METHODS". IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER LENS EXPLANT. PATIENT RELATED FACTOR WAS REPORTED AS CAUSE OF THIS EVENT. Manufacturer Narrative: ADDITONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGICAL INTERVENTION, EXPLANT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10162981&pc=MTA | 2023826-2020-01296 | 2019-10-21 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED, REPORTER STATED " AFTER ICL IMPLANTATION SURGERY THE PATIENT'S VISION AND EYE CONDITIONS WERE VERY GOOD. HOWEVER, THE PATIENT COMPLAINED OF SEVERE GLARE AND HALOES AND COULD NOT ACCEPT IT. THE PATIENT WAS UNWILLING TO ACCEPT OTHER TREATMENT METHODS". IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER LENS EXPLANT. PATIENT RELATED FACTOR WAS REPORTED AS CAUSE OF THIS EVENT. Manufacturer Narrative: CORRECTED DATA : H1- TYPE OF REPORTABLE EVENT IN INITIAL MDR AND SUPPLEMENTAL MDR: MALFUNCTION IS CORRECTED TO SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149320&pc=MTA | 2023826-2018-01799 | 2018-10-25 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE LENS "IS TOO BIG" AND IS PLANNING TO EXCHANGE IT FOR A SHORTER LENS. HOWEVER, UNABLE TO OBTAIN ADDITIONAL INFORMATION DESPITE REPEATED REQUESTS.Event Description: THE REPORTER INDICATED THAT ON (B)(6)2018 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. ON (B)(6)2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR, STATING "THE IMPLANT PLACED WAS TOO BIG." Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED IN THE LEFT (OS) EYE. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE LENS WAS RETURNED IN LIQUID AND TAPED TO AN EMAIL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS FOUND. (B)(4). Manufacturer Narrative: EVENT-UPDATED EVENT STATEMENT. PLEASE DISREGARD ALL PREVIOUS EVENT STATEMENTS. PATIENT CODES ADDED. DEVICE CODE ADDED. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241101&pc=MTA | 2023826-2019-00099 | 2018-11-15 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, MICL12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED, SURGEON PLANNED TO EXCHANGED LENS ON (B)(6) 2019 BUT THE PATIENT RESCHEDULED SURGERY. ON (B)(6) 2019 IS THE EXPECTED DATE OF SURGERY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, MICL12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018.ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS DOING GREAT. SECONDARY SURGICAL INTERVENTION, EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SMALL PLASTIC CONTAINER. VISUAL INSPECTION FOUND A TORN HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9226984&pc=MTA | 2023826-2019-02020 | 2019-05-14 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Known Impact Or Consequence To Patient; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 AND LOW VAULT WAS OBSERVED. PATIENTS CURRENT UCVA WAS REPORTED TO BE 20/20, PATIENT IS DOING GOOD. LENS REMAINS IMPLANTED. LENS EXCHANGE IS PLANNED FOR JANUARY 2020. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 2227: HALOS, PATIENT CODE 2140: GLARE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ADVERSE EVENT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: THE REPORTER INDICATED THAT THE SURGEON. IMPLANTED A 12.6MM VICM5 12.6 OF -10.00 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A DIFFERENT MODEL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTEVENTION; LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 OF -10.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 16-01-2019. LOW VAULT AND GLARE/HALOES WERE REPORTED. THE LENS WAS EXCHANGED ON 12-12-2019 FOR A DIFFERENT MODEL; LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 OF -10.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT AND GLARE/HALOES WERE REPORTED. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A DIFFERENT MODEL; LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. POST-OP INFORMATION PROVIDED."PATIENT STILL HAS GLARE." REFERENCE MFR# 2023826-2020-00696 FOR REPLACEMENT LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227163&pc=MTA | 2023826-2019-02021 | 2019-05-14 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Known Impact Or Consequence To Patient; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 AND LOW VAULT WAS OBSERVED. PATIENTS CURRENT UCVA WAS REPORTED TO BE 20/20, PATIENT IS DOING GOOD. LENS REMAINS IMPLANTED. LENS EXCHANGE IS PLANNED FOR JANUARY 2020. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 OF -11.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED WITH GLARE/HALOES. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A DIFFERENT MODEL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. EXPLANT DATE: (B)(6) 2019 PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. PATIENT CODE 2227: HALOS, PATIENT CODE 2140: GLARE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE:3001 GLARE/HALOES CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 OF -11.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON 16-JAN-2019. LOW VAULT WAS OBSERVED WITH GLARE/HALOES. THE LENS WAS EXCHANGED ON 12-DEC-2019 FOR A DIFFERENT MODEL AND LONGER LENGTH LENS. THE EXCHANGE DID NOT RESOLVED THE PROBLEM. POST-OP INFORMATION PROVIDED."PATIENT STILL HAS HALO." REFERENCE MFR# 2023826-2020-00697 FOR REPLACEMENT LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15466007&pc=MTA | 2023826-2022-03249 | 2022-08-02 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -07.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT IS EXPERIENCING LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -07.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING; GLARE; LENS DISLOCATION/SUBLUXATION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THIS SAME DAY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "LOW VAULT - COULD HAVE BEEN MIS-MEASURED OR ABNORMAL RATIO OF DIAMETER TO WHITE TO WHITE. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15466080&pc=MTA | 2023826-2022-03250 | 2022-08-02 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -07.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -07.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING; GLARE; LENS DISLOCATION/SUBLUXATION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THIS SAME DAY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "LOW VAULT - COULD HAVE BEEN MIS-MEASURED OR ABNORMAL RATIO OF DIAMETER TO WHITE TO WHITE. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16486390&pc=MTA | 2023826-2023-00730 | 2022-08-26 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION, GLARE/HALOES, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE REPORTED AS OTHER. REPORTEDLY, "THE ICL WAS TOO SMALL". Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2022. ON (B)(4) 2023, THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION, GLARE/HALOES, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE REPORTED AS OTHER. REPORTEDLY, "THE ICL WAS TOO SMALL."" IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7207524&pc=MTA | 2023826-2018-00108 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. WITHIN THE FIRST WEEKS FOLLOWING SURGERY AS THE EYE HEALED AND THE VISION NOT IMPROVING ACCORDINGLY, IT WAS DETERMINED THAT THE ICL HAD BEEN INSERTED UPSIDE DOWN. THE PATIENT UNDERWENT SUCCESSFUL REMOVAL AND IMPLANTATION OF A REPLACEMENT LENS ON (B)(6) 2018. IT IS REPORTED THAT SURGEON ERROR WAS THE CAUSE OF THE EVENT. Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172483&pc=MTA | 2023826-2016-01759 | NA | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED DECREASE OF DISTANCE VISION AND PRK SURGERY WAS PERFORMED. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY CONFIRMED THE PATIENT'S REPORT AND INDICATED THE EVENT WAS NOT RELATED TO THE DEVICE. NO MEDICATIONS OR TREATMENTS WERE REQUIRED. AT THE LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S VISUAL ACUITY WAS 20/20. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802406&pc=MTA | 2023826-2017-01351 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE GLARE. THE LENS WAS REPLACED WITH THE SAME MODEL NUMBER AND DIOPTER. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850671&pc=MTA | 2023826-2017-01402 | 2017-08-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -14.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING GLARE/HALOS. REPORTEDLY, THE PATIENT WAS UNHAPPY BECAUSE OF HALOS AND SO THE SURGEON DECIDED TO EXPLANT THE LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (FELLOW EYE (B)(4)). (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850674&pc=MTA | 2023826-2017-01403 | 2017-08-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -14.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING GLARE/HALOS. REPORTEDLY, THE PATIENT WAS UNHAPPY BECAUSE OF HALOS AND SO THE SURGEON DECIDED TO EXPLANT THE LENS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND THE LENS RETURNED DRY IN A LENS CASE/ VIAL AND CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION IN SUPPLEMENTAL MDR #1: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (FELLOW EYE #(B)(4)). (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773012&pc=MTA | 2023826-2018-01189 | 2018-04-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO MODERATELY HIGH VAULTING AND NIGHT GLARE. THE LENS WAS REPLACED WITH ANOTHER ICL LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773025&pc=MTA | 2023826-2018-01188 | 2018-04-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO MODERATELY HIGH VAULTING AND NIGHT GLARE. THE LENS WAS REPLACED WITH ANOTHER ICL LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244072&pc=MTA | 2023826-2017-00053 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED DISTORTED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMEASUREMENT OF SULCUS. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND A SCRATCH AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244098&pc=MTA | 2023826-2017-00052 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED DISTORTED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMEASUREMENT OF SULCUS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429837&pc=MTA | 2023826-2017-00488 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13MM ICM130V4 IMPLANTABLE COLLAMER LENS, -20 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. THE PATIENT IS EXPERIENCING EXCESSIVE VAULT AND IRIS ATROPHY. THE PATIENT ALSO EXPERIENCES MID-DILATED AND FIXED PUPIL, GLARE AND HALOES. IN 2006 PATIENT UNDERWENT PUPILLOPLASTY. IN 2006 AND 2008 PATIENT UNDERWENT AN IRIDECTOMY. PATIENT COMPLAINS OF GLARE AND HALOS "AS THE HAPTIC IS EXPOSED." AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LAST FOLLOW UP EXAM WAS ON (B)(6) 2016. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. CONCLUSION CODE: IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION, PATIENT HAD CORNEAL GRAFT. THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) -NO CODE AVAILABLE (SECONDARY SURGERY-PUPILLOPLASTY AND IRIDECTOMY (X2), IRIS ATROPHY, FIXED PUPIL) (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7304920&pc=MTA | 2023826-2018-00301 | NA | Injury | STAAR SURGICAL COMPANY | 2018-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE PATIENT COMPLAINED OF GLARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA WORK ORDER SEARCH: CORRECTED TO, "ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT." (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT THE PROBLEM WAS NOT PRODUCT RELATED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887948&pc=MTA | 2023826-2019-01416 | 2019-07-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 20/JUN/2019. PIGMENT DISPERSION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. PER REPORTER, ON PATIENTS (B)(6) 2019 VISIT, "BCVA WAS 20/20, VAULT 867, IOP 17 MMHG, AND STATUS OF THE EYE PIGMENT DISPERSION. I WOULD LIKE TO EXCHANGE THE ICL." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE 114 NOT APPLICABLE IN INITIAL MDR. RESULT CODE 213 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE PATIENT IS 'ASYMPTOMATIC'. THE BCVA WAS RECORDED AS 20/20 WITH A VAULT OF 801 UM. THE SURGEON HAS DECIDED NOT TO REPLACE THE LENS. THE LENS REMAINS IMPLANTED. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194491&pc=MTA | 2023826-2018-01904 | 2018-12-10 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: REPORTEDLY, THE PATIENT IS SUFFERING FROM DIPLOPIA IN THE RIGHT (OD) EYE AND LENS DISLOCATION. THE LENS REMAINS IMPLANTED. THIS IS A SPHERICAL 12.6MM ICL, -12.0 DIOPTER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 WAS IMPLANTED INTO THE PATIENT'S LEFT (OD) EYE. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON NOTING PIGMENT DISPERSION AND LENS DISLOCATION WHICH CAUSED DIPLOPIA. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MFR DATE: UNK. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM-PLEASE DISREGARD PREVIOUS DESCRIBE EVENT OR PROBLEM STATEMENT. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (PIGMENT DISPERSION, LENS EXPLANT, SECONDARY SURGERY). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140763&pc=MTA | 2023826-2020-01231 | 2019-11-24 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED, REPORTER STATED " THE PATIENT FELT THAT THE GLARE WAS OBVIOUS AFTER SURGERY, AND IT WAS DIFFICULT TO ADAPT, SO HE WAS REQUIRED TO REMOVE IT". EXPLANTING THE LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7329320&pc=MTA | 2023826-2018-00357 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-03-09 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER ON (B)(6) 2017 IN THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTED THAT "THE PATIENT IS SUFFERING FROM SEVERE GLARE AND HE BELIEVES THAT IT IS FROM THE CENTRAL PORTS OF THE LENSES IN HER EYES". THE REPORTER STATED THAT THE PATIENT STILL HAS "GLARES/ DECENTRATION". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE REPORTER BUT NONE HAVE BEEN FORTHCOMING. A SUPPLEMENTAL MDR WILL BE SUBMITTED IF NEW INFORMATION IS RECEIVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8447848&pc=MTA | 2023826-2019-00480 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem; No Apparent Adverse Event | Visual Disturbances; No Known Impact Or Consequence To Patient | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY THE SURGEON "WILL DETERMINE IF THE LENS NEEDS TO BE EXCHANGED DUE TO VAULT". LENS REMAINS IMPLANTED. "THERE ARE NO COMPLICATIONS WITH PATIENT AS OF NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMAITON; SERIAL# UPDATED FROM SERIAL NUMBER (B)(4) TO SERIAL NUMBER (B)(4). EXPIRATION DATE UPDATED FROM 31-AUG-2020 TO 31-JUL-2019. UNIQUE IDENTIFIER (UDI)# UPDATED FROM (B)(4) TO (B)(4). DEVICE MANUFACTURE DATE UPDATED FROM 14-SEP-2018 TO 11-AUG-2017. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THERE IS CONCERN FOR ASSESSMENT OF VAULTING AND THE SURGEON WILL MONITOR TO DETERMINE IF AN EXCHANGE IS NECESSARY. THE PATIENT IS DOING WELL WITH MILD COMPLAINTS OF GLARE. THE LENS REMAINS IMPLANTED. MICL 13.7 SHOULD BE ADDED TO PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9886889&pc=MTA | 2023826-2020-00696 | 2020-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICM_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS. IT IS REPORTED THAT " THE PATIENT STILL HAS GLARE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. REFERENCE MFR# 2023826-2019-02020 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4960431&pc=MTA | 2023826-2015-00853 | 2014-12-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Contamination with Chemical or Other Material | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND ON (B)(6) 2014 THE SURGEON OBSERVED A WHITE MATERIAL ADHERED ON THE LENS. A DECREASE IN UVCA WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11058380&pc=MTA | 2023826-2020-03119 | 2020-06-22 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Impairment | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT COMPLAINED OF UNSATISFACTORY VISUAL ACUITY AND ASKED TO BE REPLACED WITH A TICL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. THE LENS WAS RETURNED SIGNIFICANTLY DAMAGED WHICH CAN CAUSE DIFFICULTY WHEN CENTERING THE LENS DURING FUNCTIONAL INSPECTION. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH AND LENGTH WAS IN SPECIFICATION BUT REFRACTION (FUNCTIONAL) VERIFICATION FAILED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5040390&pc=MTA | 2023826-2015-01107 | 2015-07-11 | Injury | STAAR SURGICAL COMPANY | 2015-08-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Unintended Movement | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS TORE WHILE BEING INSERTED AND THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE. THE LENS HAS SINCE DECENTERED AND THE EDGE OF THE LENS IS IN THE PATIENT'S VISION. THE SURGEON'S PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4966256&pc=MTA | 2023826-2015-00943 | NA | Injury | STAAR SURGICAL COMPANY | 2015-08-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED A LOSS OF UCVA AND THE LENS WAS EXCHANGED FOR TICM. THIS RESOLVED THE PROBLEM. PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS IN SPECIFICATION. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206084&pc=MTA | 2023826-2015-01470 | NA | Injury | STAAR SURGICAL COMPANY | 2015-11-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT WAS UNDER-CORRECTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PATIENT IS NOW 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - REPORTEDLY, ICL (-5.5 D;12.6 MM) WAS EXPLANTED/EXCHANGED, THREE WEEKS POSTOPERATIVELY, FOR A DIFFERENT POWER ICL OF TO ADDRESS PATIENT RESIDUAL REFRACTIVE ERROR ("UNDERCORRECTION"). NO REPORTED PROBLEM WITH THE LENS POSITION PRIOR TO THE LENS EXCHANGE. ACCORDING TO THE REPORT, BY A DOCTOR, DATED ON (B)(6) 2015, THE REPLACEMENT LENS RESOLVED THE ISSUE AND UVA WAS 20/20. GIVEN THE INFORMATION THAT A DIFFERENT POWER LENS WAS IMPLANTED AS A REPLACEMENT AND THE REFRACTIVE ERROR WAS CORRECTED IT IS VERY LIKELY THAT USER ERROR (MISCALCULATION) WAS THE CAUSE OF THE EVENT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, THE MOST LIKELY ROOT CAUSE OF THE EVENT WAS USER ERROR (MISCALCULATION). (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION, WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE SURGERY FACILITY REPORTED THAT THERE WAS NO PROBLEM WITH THE LENS. AS REPORTED ON THE MEDICAL REVIEW, SINCE A DIFFERENT POWER LENS WAS IMPLANTED AS A REPLACEMENT AND THE REFRACTIVE ERROR WAS CORRECTED, IT IS VERY LIKELY THE USER ERROR (MISCALCULATION) WAS THE CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5131893&pc=MTA | 2023826-2015-01231 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-10-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE PATIENT HAD A HYPEROPIC REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP UCVA WAS 20/25 ON (B)(6) 2015. SEE MFR # 2023826-2015-01311 FOR LEFT EYE. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS CODE(S): (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S) BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT HAD A HYPEROPIC REFRACTIVE SURPRISE. TWO MONTHS POSTOPERATIVELY, THE SURGEON EXCHANGED THE LENS FOR ANOTHER OF THE SAME MODEL BUT DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. AFTER THE EXCHANGE, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4509422&pc=MTA | 2023826-2015-00115 | NA | Injury | STAAR SURGICAL COMPANY | 2015-02-11 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT'S VISION WAS HYPEROPIC. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH NO COMPLICATIONS.Event Description: ADDITIONAL INFORMATION RECEIVED - THERE WERE NO RELATED COMPLICATIONS TO THE EVENT AND THE ICL WAS EXPLANTED WITH NO REPORTED INJURY. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. AS REPORTED ON THE MEDICAL REVIEW, THE MEDICAL DEVICE WAS UNLIKELY THE CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF LENS HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5395179&pc=MTA | 2023826-2016-00072 | NA | Injury | STAAR SURGICAL COMPANY | 2016-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE PATIENT COULD NOT SEE WELL WITH THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE ICL LENS WAS EXPLANTED DUE TO TOO SMALL/FLAT SO IT WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2015. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT DID WELL AFTER THE EXCHANGE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW: REPORTEDLY, MICL (-15D; 12.1 MM) WAS EXPLANTED POSTOPERATIVELY TO ADDRESS PATIENT RESIDUAL REFRACTIVE ERROR ("COULD NOT SEE WELL WITH THE LENS"). NO REPORTED POSTOPERATIVE SEQUELAE. DESPITE MULTIPLE ATTEMPTS NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540585&pc=MTA | 2023826-2017-00741 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Vitreous Floaters; Headache; Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED IRIS DISTORTION, DISCOMFORT, HEADACHES, GLARE/HALOES, BLURRED VISION, DOUBLE VISION OR "GHOSTING", "VITREOUS FLOATERS", PUPIL REMAINS DILATED AND NONRESPONSIVE, AND OTHER MISCELLANEOUS VISUAL DISTURBANCES. THE PATIENT IS NOT HAPPY WITH THE RESULT OF THE IMPLANT. ON (B)(6) 2017 THE PATIENT BEGAN PHYLOCARPINE DROPS TREATMENT AND THE DR. IS MONITORING THE OCCURENCE. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, AND IRIS ATROPHY. (B)(6). (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. "PHYLOCARPINE" PILOCARPINE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6341778&pc=MTA | 2023826-2017-00244 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Vitreous Floaters; Visual Disturbances | P030016 | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THEIR RIGHT EYE (OD). THE PATIENT REPORTED HAD FLOATERS AND WAS BLURRY SOMETIME. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THEY CONFIRMED THE PATIENT REPORT AND THE DATE OF LAST VISIT WAS (B)(6) 2014. IT WAS UNKNOWN IF THE EVENT WAS CAUSED BY THE DEVICE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP PROGNOSIS WAS GOOD AND VISUAL ACUITY WAS 20/30-2. Manufacturer Narrative: CONCLUSION: IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU "INDICATIONS THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE". THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483346&pc=QCB | 2023826-2020-01947 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENTS EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6). EXPIRATION DATE: 30-APR-2023 D6A- UNK. D6B-UNK. H1- SERIOUS INJURY. H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00 DIOPTER INTO THE PATIENT RIGHT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A: ON (B)(6) 2020, D6B: ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC TORN WITH THREADS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483351&pc=QCB | 2023826-2020-01948 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: REPORTER INDICATED SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B1- ADVERSE EVENT B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON AN UNKNOWN DATE FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON. D4- MODEL#: VICMO 13.2. SERIAL #: (B)(6),. EXPIRATION DATE: 30-APR-2023 H1- SERIOUS INJURY H6-IMPACT CODE 4629: LENS WAS EXCHANGED. H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATE THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50 DIOPTER INTO THE PATIENT LEFT EYE ON (B)(6) 2020. REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT IS REPORTED AS WRONG CALCULATION IN OCOS BY SURGEON D6A:10-AUG-2020. D6B: 31-AUG-2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL WIDTH VERIFICATION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. H6 -TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778982&pc=QCB | 2023826-2021-01323 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.19D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #719586. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779013&pc=QCB | 2023826-2021-01322 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6 -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "PRESCRIBED DISTANCE GLASSES." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM6 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PATIENT WAS PRESCRIBED DISTANCE GLASSES. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION ON THE OPTOCRAFT FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VAULES MEASURED AT THE TIME OF MANUFACTURING; THE SPHERE WAS OUT OF TOLERANCE BY 0.08D. H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H3 - RESULT CODE: 114 IN SUPPLEMENTAL MDR #3 CLAIM # (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140273&pc=QCB | 2023826-2021-02413 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 TH LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM,VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " ON THE FIRST DAY AFTER OPERATION, THE VISUAL ACUITY WAS 0.6 AND THE CORNEAL EDEMA WAS MILD." NO SPECIAL TREATMENT WAS NEEDED FOR THE EDEMA. THE EDEMA DISAPPEARED THE NEXT DAY, AND THERE WAS NO SIGNIFICANT CHANGE BETWEEN CORNEAL ENDOTHELIUM AND PREOPERATIVE EXAMINATION. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13160463&pc=QCB | 2023826-2022-00077 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -14.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME WERE ASSESSED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17542184&pc=QCB | 2023826-2023-03515 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -2.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION (PT. RIGHT EYE FIELD IS BLURRED) AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45 S OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882069&pc=QCB | 2023826-2023-04269 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6-CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882376&pc=QCB | 2023826-2023-04270 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF-4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON 07-JUL-2023. BLURRED VISION; SEVERE DISTENDING SENSATION;PAIN; DIZZINESS WAS REPORTED AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS REPOSITIONED THEN ON SAME DAY REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLINICAL CODE 4581- SEVERE DISTENDING SENSATION; DIZZINESS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=MTA | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16981458&pc=QCB | 2023826-2023-01954 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2;-5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THE LENS HAS SOME DIFFUSE WHITISH PLAQUES/FILM ON THE LENS, WHICH LOOKS LIKE IT'S ON THE POSTERIOR SURFACE OF THE EVO LENS AND NOT THE LENS OR CORNEA. THE SURGEON ALSO INDICATED THERE IS A BIT MORE INFLAMMATION THEN TYPICAL, THE WHITISH FILM/PLAQUES APPEAR STABLE, NEITHER INCREASING OR DECREASING AND THERE SEEMS TO BE NO CONTINUED INFLAMMATION. THE SURGEON ALSO REPORTS GLARE/HALOES, BLURRED VISION, AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY, A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT MODEL/POWER WAS IMPLANTED. IT WAS NOTED "VISION IMPROVED SIGNIFICANTLY WITH EXCHANGE". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "SUSPECT THIS WAS INFLAMMATORY DEBRIS/FIBRING BUILDING UP ON THE LENS". Manufacturer Narrative: ADDITIONAL INFORMATION: INITIAL MDR IS NOT APPLICABLE AS THERE WAS NO DEVICE PROBLEM, REPORTER LATER CLARIFIED THE ISSUE WAS DIRECTLY PATIENT RELATED. CLAIM#: (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE: 4581- WHITISH PLAQUE/FILM. H6- TYPE OF INVESTIGATION: 4110- LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved