## IOP Issue-1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code...37 | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8970000&pc=QCB | 2023826-2019-01541 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (26 MMHG) AND LOW VAULT WERE ASSESSED ON (B)(6) 2019. EYE STATUS ON (B)(6) 2019, "AC QUIET WITH MINIMAL CORNEAL EDEMA. NEAR VA J8." LENS REMAINS IMPLANTED. Manufacturer Narrative: "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SAME LENGTH, SAME MODEL LENS. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED IN THE INITIAL MDR. PATIENT CODE 3191: SECONDARY SURGERY- EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON A FOLLOW UP VISIT ON (B)(6) 2019, IT WAS REPORTED THAT THE PATIENT HAS "SLIGHT ANTERIOR CAPSULE OPACITY, ANTERIOR CHAMBER QUIET AND OPEN ANGLES". SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. EXPLANTED DATE: (B)(6) 2019. PATIENT CODE: 1791, PATIENT CODE: 1766, PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE), DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), DEVICE CODE 2017: IMPROPER OR INCORRECT PROCEDURE OR METHOD. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874226&pc=QCB | 2023826-2020-00634 | 2020-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN-HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUES EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK; ELEVATED IOP AND INFLAMMATION WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE ON (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN- HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUED EVALUATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- INCLUDED "ELEVATED IOP AND INFLAMMATION" TO STATEMENT. H6-PATIENT CODE 1937- ELEVATED IOP. 1932- INFLAMMATION. H6-DEVICE CODE: 1401-BLURRED VISION. H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107370&pc=QCB | 2023826-2019-01751 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Inflammation; Intraocular Pressure Increased; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -18.00/+1.0/177 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED. THE LENS WILL BE EXPLANTED BUT TO DATE REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1937: INFLAMMATION. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: LENS IMPLANTED UPSIDE DOWN. PATIENT CODE 1937 IN PREVIOUS MDR CORRECTED TO 1932. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16261376&pc=QCB | 2023826-2023-00326 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER (PATIENT CURRENT OBSERVING DOCTOR) STATES THAT A SURGEON AT A DIFFERENT CLINIC IMPLANTED A 13.7MM VTICMO 3.7 IMPLANTABLE COLLAMER LENS OF -7.00/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE CURRENT OBSERVING DOCTOR SAW PATIENT FOR THE FIRST TIME ON (B)(6) 2022, WHERE HE UNFORTUNATELY BECAME AWARE OF SEVERAL MEDICAL ERRORS. "TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT." COMPLAINT QUESTIONNAIRE PROVIDED REPORTS CATARACT (ASC); ANGLE CLOSURE WITH ELEVATED IOP; LENS DISLOCATION/ SUBLUXATION; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; BLURRED VISION; AND IRIS ATROPHY WAS REPORTED. CURRENT STATUS OF THE EYE: "EYE ALREADY CALM, CORNEA CLEAR, ANTERIOR CHAMBER SHALLOW, PUPIL WIDE DILATED, ICL IN SITUATION ROTATED COUNTER CLOCKWISE TO INTENDED POSITION, ANTERIOR SUBCAPS CAT. CATARACT FORMATION." CAUSE REPORTED AS USER ERROE; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE IMPLANTING SURGEON HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538715&pc=QCB | 2023826-2020-02049 | 2020-07-08 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Cataract; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; Vomiting; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/1.50/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE SURGEON ASSESSED AN ELEVATED IOP OF 22MMHG. THIS WAS CONTROLLED WITHIN THE FIRST 24 HOURS WITH MEDICATION. ALSO REPORTED IS PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, DILATED PUPIL, GLARE/HALOES, IRIS ATROPHY, PUPIL BLOCK, AND ANTERIOR SUBCAPSULAR OPACITIES. THE PATIENT ALSO EXPERIENCED NASEAU AND EXCESSIVE MOVEMENT DUE TO VOMITING CAUSED BY THE ANESTHESIA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE PUPILLARY DILATION WAS "SLOWLY CORRECTED WITHIN THE FIRST MONTH, LEAVING AN ANISOCORIA OF APPROX. 2.5MM AND PRESENTING SECTORAL ATROPHY OF THE IRIS." THE PATIENT ALSO REPORTS "DECREASE IN THE INTENSITY OF THE COLORS." THE PATIENT STILL EXPERIENCES GLAUKOMFLECKEN, IRIS PIGMENT DISPERSION, AND GLARES AT NIGHT, BUT IS OVERALL "REALLY SATISFIED." THE CAUSE OF THE EVENT IS REPORTED AS "REMNANT OF OVD BEHIND THE ICL." Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8807833&pc=QCB | 2023826-2019-01275 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -11.00/1.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE DAY AFTER OPERATION THE EYE HAD UNDERGONE MYDRIASIS, WHICH WAS IMPOSSIBLE TO NARROW. ON (B)(6) 2019 THE LENS WAS EXPLANTED. THE SURGEON STATED " THE EVENT WAS DUE TO AN UNKNOWN PROBLEM. I CAN ONLY SUPPOSE, THAT DIAMETER ICL WAS BIGGER. WE STILL DONT KNOW WHAT HAPPENED. CATARACT (BEHIND CAPSULA ANTERIOR) WAS OBSERVED AFTER IMPLANTATION. SHE STILL HAS PARALYTIC MYDRIASIS (5-6 MM).'' PATIENT CONDITION IS STILL THE SAME, ON (B)(6) 2019 THE DIAMETER OF THE PUPIL WAS 5-6 MM AND IT DIDN'T RESPOND TO THE LIGHT. PATIENT USED PILLOCARPIN 1% THE LAST 2 MONTHS. Manufacturer Narrative: DATE OF THIS REPORT: 20-JUN-2019. CONCOMITANT MEDICAL PRODUCTS: CARTRIDGE MODEL-SFC-45; LOT#-UNK SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT'S DATE OF BIRTH: (B)(6) 1996. PATIENT CODE 3191: NO CODE AVAILABLE (UNREACTIVE FIXED PUPIL) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 26 AUG 2019. OVD USED: EYEVISC PFS 2 ML (HYDROXYPROPYL METHYLCELLULOSE 2% W/V) OPHTALMIC SOLUTION USP FROM BIOTECH IOP WAS 50.0MM THE DAY AFTER SURGERY, WHEN MYDRIASIS WAS NOTICED. THE ANGLES WERE OPEN. MEDICATIONS USED TO TREAT PATIENT: DEXAMETASON-INIECTION 3 DAYS ONE TIME PER DAY, BETOPTIC AND AZOPT IN DROPS FOR 3 DAYS AND THEN THE IOP WAS OK -14 MM. PATIENT CODE 3191: MEDICAL INTERVENTION. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620830&pc=QCB | 2023826-2022-03493 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/084, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR (ASC) CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. POST-OP, THE PUPIL IN MYDRIASIS NON REACTIVE AND SLIGHT EPITHELIAL HAZE. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EPITHELIAL HAZE, MYDRIATIC PUPIL SHOULD HAVE BEEN INCLUDED. B5: POSTOP COMMENTS ADDITIONAL NOTED PUPIL IN MYDRISASI NON REACTIVE; SLIGHT EPITHELIAL HAZE. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13698235&pc=QCB | 2023826-2022-00761 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Conjunctivitis; Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.00/+2.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT INTRAOCULAR PRESSURE INCREASED AND COMPLAINED OF BLURRY VISION AND PAIN ON THE DAY OF SURGERY; MEDICATION WAS PRESCRIBED AND THE SURGEON PERFORMED IRRIGATION BEFORE SENDING HER HOME. ONE DAY POST OP THE PATIENT HAD HIGH IOP AND CORNEAL EDEMA CENTRALLY WITH CONJUNCTIVAL REDNESS. PUPILLARY BLOCK 3+ CELLS/FLARES NOTED. ON THIS DAY THE SURGEON REPORTED NOTICING A THIN MEMBRANE OVER THE CENTRAL HOLE OF THE ICL; WHICH THEY PERFORMED AN ADDITIONAL IRRIGATION WITH ADDITIONAL MEDICATION PRESCRIBED. IT WAS INDICATED AFTER ANOTHER POSTOP VISIT SURGEON PERFORMED AN ADDITIONAL IRRIGATION AND THE MEMBRANE WASHED OUT DURING THIS AND THE PATIENT'S IOP BECAME STABLE. IT WAS FURTHER INDICATED 1 WEEK POSTOP THE IOP IS STABLE; NO CORNEAL EDEMA BUT PATIENT HAS PUPIL MID-DILATED DUE TO THE PUPILLARY BLOCK AND 1.2 CT VAULTING. THE PATIENT WILL CONTINUE BEING MONITORED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840739&pc=QCB | 2023826-2021-04397 | 2021-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Blurred Vision; Discomfort; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO CORNEAL DECOMPENSATION. REPORTEDLY, "HOURS AFTER UNEVENTFUL SURGERY, PAIN, DECREASED VISION; ON EXAMINATION: CORNEAL DECOMPENSATION, NO DETAILS VISIBLE." THE PROBLEM IS MARKED AS RESOLVED AFTER EXPLANT HOWEVER THERE WAS STILL SOME CORNEAL EDEMA AT 7H. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE PATIENT STARTED HAVING DISCOMFORT A FEW HOURS AFTER SURGERY AND THE LENS HAD TO BE REMOVED TWO DAYS LATER. SYMPTOMS ABATED AFTER EXPLANT. THE SURGEON STATES THE ICL LOOKED FINE WITHOUT VISIBLE DEFECTS AND HE REMAINS UNCERTAIN ABOUT THE EXACT CAUSE OF THE EVENT. CLAIM# (B)(4). Manufacturer Narrative: D4: CORRECTED MODEL#, SERIAL#, AND EXPIRATION DATE. D6A: CORRECTED TO ON (B)(6) 2021. D6B: CORRECTED TO ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634551&pc=QCB | 2023826-2020-02295 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+5.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION (CORNEAL EDEMA) AND UNREACTIVE (FIXED) PUPIL. THE IOP WAS CONTROLLED WITH AGGRESSIVE TREATMENT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO AN OVERSIZED LENS. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON 12-OCT-2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS ASYMPTOMATIC, THE ANTERIOR SEGMENT WAS OK, THERE WAS SOME PIGMENT DISPERSION AND THE PUPIL WAS MYDRIATIC. THE PATIENT WAS HAPPY WITH THE RESULTS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558781&pc=QCB | 2023826-2021-00704 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200 AND PUPIL IS DILATED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SIGNIFICANT REDUCTION OF ICA, CORNEAL DECOMPENSATION, CORNEAL EDEMA ALSO REPORTED. THE CAUSE IS REPORTED AS RETAINED VISCOELASTIC BEHIND THE ICL AND "ICL STILL A BIT OVERSIZED." REPORTEDLY THE PATIENT IS RECOVERED, VISION IS GOOD. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558806&pc=QCB | 2023826-2021-00699 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS -15.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTEDLY EXPLANTED THE LENS AND REPLACED IT WITH A SHORTER LENS BECAUSE THE LENS WAS "TOO BIG." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: EXCESSIVE VAULT-TOUCHING THE CORNEA, ELEVATED IOP, 4+ CORNEAL EDEMA, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION REPORTED. CAUSE IS REPORTED AS DEVICE: ICL TOO LARGER PER CALCULATOR. PROBLEM MARKED AS RESOLVED. HOWEVER, SEE MFR REP# 2023826-2021-00704 FOR REPLACEMENT LENS. ORIGINAL LENS IMPLANTED, REMOVED, AND REPLACED WITHIN THE SAME SURGERY. CLAIM #(B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUIED IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741585&pc=QCB | 2023826-2021-04219 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -16.0/2.5/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED AND AT A LATER DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699716&pc=QCB | 2023826-2019-00980 | NA | Injury | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/5.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK WITH ELEVATED IOP (50MMHG), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, AND IRIS ATROPHY WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INTRAOCULAR INJECTION AND THE LENS WAS REMOVED ON (B)(6) 2019. PER THE REPORTER, ON (B)(6) 2019 PATIENT'S EYE STATUS, "IRIS ATROPHYIC AND PUPIL DILATED. REPORTER NOTED CORNEA BECAME CLEAR AFTER 1 MONTH. EYE IS CALM. CAUSE OF THE EVENT IS REPORTED AS "URETS-ZAVALIA SYNDROME DUE TO OCCLUSION OF THE AQUAPORT FIRST WITH VISCOELASTIC THEN WITH IRIS PIGMENTS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 1937: ELEVATED IOP. DEVICE CODE: 1401: MISFOCUSING; PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1401 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "DEVICE CODE 3191: PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984533&pc=QCB | 2023826-2019-01677 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.5/+3.0/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, CORNEAL DECOMPENSATION AND UNREACTIVE (FIXED) PUPIL. THERE WAS AN ADDITION OF NEW PERIPHERAL IRIDECTOMY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PATIENTS POST-OP CONDITION WAS NORMAL CORNEAL THICKNESS, NO EDEMA, IOP 10MMHG, CELL COUNT (B)(6). Manufacturer Narrative: METHOD CODE: 4111 SHOULD HAVE BEEN INCLUDED IN INITAL MEDWATCH REPORT. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12242256&pc=QCB | 2023826-2021-02943 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Glaucoma; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.5/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK; (IN (B)(6)) HYPHEMA ON THE ANTERIOR AND POSTERIOR SURFACE OF THE ICL AND WAS PRESCRIBED MEDICATION WHICH THE EYE RECOVERED TO 1.5 CORRECTED VISUAL ACUITY AND 1.2 NAKED EYE VISION. AT A FOLLOW UP DATE HYPHEMA AND MILD CORNEAL EDEMA WAS NOTED WITH ADDITIONAL MEDICATION; AS WELL AS CONDITION OF SECONDARY GLAUCOMA ASSOCIATED WITH HYPHEMA CONTINUED. THE PATIENT IS BEING OBSERVED. THE PROBLEM WAS REPORTED AS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17935878&pc=QCB | 2023826-2023-04537 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Corneal Edema; Headache; Blurred Vision; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -12.5/+1.00/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION, THE SURGEON ALSO NOTED HEADACHES, CORNEAL EDEMA, BLURRY VISION AND GLARES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR - PREOP AXIS WAS NOT ACCURATE BASED ON THE TOPOGRAPHY AND AUTO-REFRACT AXIS OF ABOUT 49 DEGREES AND IT WAS A PRE-OP MEASUREMENT ERROR NOT AN ISSUE WITH THE LENS ITSELF. Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THAT THE PATIENT HAD EYE PAIN AND MEDICATION WAS USED, TIMOLOL AND DIAMOX. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390681&pc=QCB | 2023826-2023-00595 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Corneal Edema; Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. CAUSE WAS "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION." STATUS OF THE EYE IS PSEUDOPHAKIA. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- ANGLE CLOSURE; PIGMENT DISPERSION; IRIS CHAFING. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/21 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED INFLAMMATION; IRIS CHAFING; PIGMENT DISPERSION; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACOEMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PSEUDOPHAKIA". CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: H6-HEALTH EFFECT- CLINICAL CODE: "1791-MILD CORNEAL EDEMA" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/21 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 INFLAMMATION, IRIS CHAFING, PIGMENT DISPERSION, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED. REPORTEDLY "INFLAMMATORY REACTION DUE TO IRIS CHAFING AND PIGMENT DISPERSION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACOEMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "MILD CORNEAL EDEMA". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334443&pc=QCB | 2023826-2021-00358 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP MILD CORNEAL EDEMA. THE LENS WAS DISCARDED. Manufacturer Narrative: B5 - PATIENT IS USING ARTIFICIAL TEARS EVERY ONCE IN A WHILE AS NEEDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005314&pc=QCB | 2023826-2021-01926 | 2021-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). ANOTHER LENS WAS NOT IMPLANTED. THE IOP WAS FIRST LOWERED WITH PARACENTESIS AND TAMPONADE ICL/REFILL A/C-FLUID WITH IOP INCREASING AFTER 1 HOUR. POST-OP CORNEAL EDEMA, DILATED PUPIL, OPTIC NERVE NOT PALE AT THIS TIME BUT CONCERNED AT LONG TERN CORNEAL AND OPTIC NERVE HEALTH. PATIENT WAS TAKING KRATOM-CHOLINERGENIC BLOCKER ALPHA AGONIST; POSSIBLE OPIOD EFFECTS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND WITH RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14645301&pc=QCB | 2023826-2022-01780 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO SIZING ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. FOR CAUSE DETAILS THE REPORTER STATED "AT THE DAY 1 POSTOPERATIVE VISIT, THE UNCORRECTED VISUAL ACUITY MEASURED 20/320 OS. INTRAOCULAR PRESSURE WAS 45 MMHG OS. EXAMINATION DEMONSTRATED SEVERE CORNEAL EDEMA OS WITH NARROW ANGLES OU." A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - PER EMAIL ON (B)(6) 2023, PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE, EXPLANT DATE:UNK TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473796&pc=QCB | 2023826-2024-00131 | 2023-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129041&pc=QCB | 2023826-2018-01773 | 2018-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.5/+2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY, THE PATIENT PRESENTED WITH BLURRY VISION, CORNEAL EDEMA, HIGH IOP AND EXCESSIVE VAULT. THE SURGEON PERFORMED A YAG PI AND PATIENT WAS PLACED ON IOP MEDICATIONS. CURRENTLY, THE PATIENT'S SYMPTOMS ARE RESOLVING AND VISION HAS IMPROVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9709011&pc=QCB | 2023826-2020-00272 | 2019-12-17 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. POST-OP, THERE WAS CORNEAL EDEMA AND THE INTRAOCULAR PRESSURE (IOP) INCREASED, WHICH WAS MANAGED WITH WOUND BURPING, PO DIAMOX, PREDNEFRINE FORTE AND TOPICAL G COMBIGAN. THE CORNEAL EDEMA AND THE IOP HAD SETTLED DOWN IN LESS THAN 2 WEEKS. POST-OP, THE PUPIL WAS REPORTED AS OVOID (OVAL SHAPED). THE LENS REMAINS IMPLANTED AND THE SURGEON AND PATIENT ARE WAITING FOR THE CORNEA TO SETTLE DOWN FOR A COUPLE OF MONTHS BEFORE ANY DECISIONS ARE MADE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930510&pc=QCB | 2023826-2020-02837 | 2020-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Corneal Edema; Intraocular Pressure Increased; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THAT IMMEDIATELY POSTOP THERE WAS ELEVATED IOP, PAIN, CORNEAL EDEMA, AND A VAULT OF 1200. DIAMOX WAS ADMINISTERED. DAY 21 POSTOP THE VAULT IS 1000, THE PUPIL DOESN'T TIGHTEN WELL, NO PAIN OR EDEMA, IOP 18. Manufacturer Narrative: B5-UPDATED INFORMATION: LENS WAS EXPLANTED. CLAIM# (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: ON 18-NOV-2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. VAULTING WAS RESOLVED WITH EXCHANGE, HOWEVER SEE MFR CLAIM#: (B)(4) FOR RESIDUAL SYMPTOMS AFTER EXCHANGE. CAUSE OF THE EVENT IS REPORTED AS A PATIENT RELATED FACTOR: "THE MEASUREMENT OF THE WTW WAS CORRECT AND ICL CALCULATION WAS COHERENT WITH IT. BUT APPARENTLY, THE ICL OF 13.2MM WAS TOO BIG AND DIDN'T FOUND THE GREAT POSITION OF IMPLANTATION IN THE SULCUS. WE HAD TO REMOVE IT FOR A SMALLER ONE." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120422&pc=QCB | 2023826-2019-01744 | 2012-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -2.5/+0.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, AND CORNEAL EDEMA. AN ENLARGEMENT OF THE PI-YAG WAS PERFORMED. THE PROBLEM IS RESOLVED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER INFORMATION RECEIVED ALSO STATES THAT THE ACD WAS NOT MEASURED PER DFU INSTRUCTIONS (ACD ASSUMED TO BE 3.2MM). INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120423&pc=QCB | 2023826-2019-01743 | 2013-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -1.0/+1.0/142 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, AND CORNEAL EDEMA. AN ENLARGEMENT OF THE PI-YAG WAS PERFORMED. THE PROBLEM IS RESOLVED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER INFORMATION RECEIVED ALSO STATES THAT THE ACD WAS NOT MEASURED PER DFU INSTRUCTIONS (ACD ASSUMED TO BE 3.2MM). INFORMATION INITIALLY RECEIVED VIA AN ARTICLE PUBLISHED IN CLINICAL OPHTHALMOLOGY ENTITLED 'A PROSPECTIVE PILOT STUDY USING A LOW POWER PIGGY-BACK TORIC IMPLANTABLE COLLAMER LENS TO CORRECT RESIDUAL REFRACTIVE ERROR AFTER MULTIFOCAL IOL IMPLANTATION.' Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780421&pc=QCB | 2023826-2019-01173 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/4.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA (PRE-OP BCVA = 20/80 POST-OP BCVA = 20/500) AND ELEVATED IOP (40 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, IOP 38 MMHG, CORNEAL EDEMA, MYDRIASIS, PHOTOPHOBIA, BLURRED VISION, PAIN. REPORTER INDICATED ON (B)(6) 2019, THAT PATIENT RECEIVED TREATMENT FOR HIGH IOP, DILATING DROPS WITH STEROIDS AND NON-STEROIDS. "NO MORE CORNEAL EDEMA AFTER" LENS REMOVAL. AFTER LENS REMOVAL, SURGEON CHOSE CLE (CLEAR LENS EXTRACTION) TO RESOLVE THE REFRACTIVE ISSUE. THE PATIENT IS FINE. THE REPORTER INDICATED THAT CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. CORRECTED TO NO, DEVICE EVALUATION ANTICIPATED BUT NOT YET BEGUN, IN INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546404&pc=QCB | 2023826-2022-00529 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -8.00/+1.00/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THERE WAS REPORT OF PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE; ON (B)(6) 2022 IT WAS ALSO NOTED PATIENT EXPERIENCED MID DILATED PUPIL; UNCORRECTED VISUAL ACUITY; CORNEAL EDEMA AND A YAG PI SURGERY WAS PERFORMED WHICH RELIEVED THE PRESSURE AND WHEN THE PATIENT RETURNED (B)(6) 2022 THEY HAD IMPROVED VISUAL ACUITY/REDUCED INTRAOCULAR PRESSURE AND MID-DILATED PUPIL. THE PROBLEM WAS MARKED AS RESOLVED BUT NOTED PATIENT IS SCHEDULED TO COME IN FOR A 1 MONTH FOLLOW UP. Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT "EXAMINATION ON (B)(6) 2022- DEMONSTRATED REDUCED UNCORRECTED VISUAL ACUITY OF 20/32; MILD CORNEAL EDEMA; A MID-DILATED PUPIL AND NON-PATENT PERIPHERAL IRIDOTOMY OD. REPEAT LASER IRIDOTOMY OD IMMEDIATELY RELIEVED THE PRESSURE; AND THE SUBJECT RETURNED THE FOLLOWING DAY (B)(6) 2022, WITH IMPROVED VISUAL ACUITY; REDUCED INTRAOCULAR PRESSURE AND PERSISTENT MID-DILATED PUPIL." THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION :LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334355&pc=QCB | 2023826-2021-00359 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP MILD CORNEAL EDEMA. THE LENS WAS DISCARDED. Manufacturer Narrative: B5 - THE PATIENT STATED SEEING A GLARE THE LEFT EYE (OS) WHILE NIGHT DRIVING AND USES ARTIFICIAL TEARS OCCASIONALLY AS NEEDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498160&pc=QCB | 2023826-2019-02408 | 2019-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.50/+3.5/065 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND THE PATIENT FELT PAIN. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND DEVICE. Manufacturer Narrative: POSTOPERATIVE REPORT STATED PATIENT'S EYE PRESSURE RETURNED TO NORMAL AND IS HEALING WITH CORNEAL EDEMA. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860009&pc=QCB | 2023826-2019-01291 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Pain; Red Eye(s); Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/3.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION, "PAINFUL, AND RED EYE" WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS "THE SIZES OF THE LENS PROBABLY TOO BIG." PER PATIENTS LAST VISIT ON (B)(6) 2019, "REDUCED VISION LEFT EYE STILL RIGID BIG PUPIL, IRIS DEFECT AS 13.00, CORNEAL EDEMA 201090410, RIGID PUPIL HAS NOT RETURNED TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191 - IRIS DEFECT AND FIXED PUPIL. DEVICE CODES 3001 AND 1583 THAT WERE PREVIOUSLY REPORTED ARE NOT APPLICABLE. COMMON DEVICE NAME/ PRODUCT CODE CORRECTED. MTA PRODUCT CODE IN INITIAL MDR NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350551&pc=QCB | 2023826-2019-02189 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Ulcer; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/081 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE WAS INITIATED AND A THERAPEUTIC CONTACT LENS INSERTED. PAIN WAS REPORTED AND CORNEAL ULCER. THE LENS REMAINS IMPLANTED. PATIENT CODE: 3191 MEDICATION TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE PRESCRIBED, THERAPEUTIC CONTACT LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983283&pc=QCB | 2023826-2022-01033 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.5/+2.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6: CLINICAL CODE 4581 (ROUND PUPIL). CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983288&pc=QCB | 2023826-2022-01032 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.5/+1.0/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6-CLINICAL CODE 4581 (ROUND PUPIL). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136571&pc=QCB | 2023826-2023-04906 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT WAS FOUND TO HAVE ENDOPHTHALMITIS AND PERMANENT LOSS OF BCVA. INTRAOCULAR PRESSURE (IOP) CHECK AND ANTERIOR SEGMENT IMAGING WERE PERFORMED ON (B)(6) 2023. PATIENT WAS PRESCRIBED MEDICATIONS FOR INCREASED IOP. PATIENT HAD HYPOPYON, FIBRIM NET, CELLS IN AC, WITH REDNESS, SEVERE LIGHT SENSITIVITY AND PAIN, INITIALLY DIAGNOSED AS TASS DUE TO FAST COURSE OF SIGNS AND CLEAR B-SCAN BY DOCTOR AND RETINAL SPECIALIST. PATIENT HAD VITREOUS BIOPSY WITH INTRAVITREAL INJECTION OF VANCOMYCIN AND CEFTAZIDIME TO RULE OUT AND TREAT ENDOPHTHALMITIS AS WELL AS SUBTENON KENALOG DUE TO INITIAL FAST COURSE OF SIGNS OF TASS. THE NEXT DAY, DUE TO CELLS ON VITREOUS B-SCAN, THE PATIENT HAD BILATERAL PPV (PARS PLANA VITRECTOMY) WITH INTRAVITREAL INJECTION OF CEFTAZIDIME, VANCOMYCIN AND VITREOUS CULTURE. VITREOUS CULTURE GREW PSEUDOMONA AERUGINOSA AND CURRENTLY THE PATIENT IS HOSPITALIZED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT IS RECEIVING IV MEROPENEM, TOPICAL ANTIBIOTICS AND GLAUCOMA MEDS. IN THE LAST EXAMINATION BY RETINAL SPECIALIST, SHE WAS STILL HAVING PAIN IN BOTH EYES AND PERIORBITAL INFLAMMATION THAT HAD BEEN IMPROVED SINCE INITIAL HOSPITALIZATION EVALUATED BY CT AND MRI OF THE BRAIN AND ORBITS. CURRENT UCVA IS NLP OU (B)(6) 2023). H6: HEALTH EFFECT IMPACT CODE: 4644 - MEROPENEM. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, D6B: UNK, H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PERMANENT LOSS OF BCVA, REDNESS, LIGHT SENSITIVITY. H6: HEALTH EFFECT IMPACT CODE: 4644 - CEFTAZIDIME, VANCOMYCIN. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PARS PLANA VITRECTOMY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136578&pc=QCB | 2023826-2023-04905 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT WAS FOUND TO HAVE ENDOPHTHALMITIS AND PERMANENT LOSS OF BCVA. INTRAOCULAR PRESSURE (IOP) CHECK AND ANTERIOR SEGMENT IMAGING WERE PERFORMED ON (B)(6) 2023. PATIENT WAS PRESCRIBED MEDICATIONS FOR INCREASED IOP. PATIENT HAD HYPOPYON, FIBRIM NET, CELLS IN AC, WITH REDNESS, SEVERE LIGHT SENSITIVITY AND PAIN, INITIALLY DIAGNOSED AS TASS DUE TO FAST COURSE OF SIGNS AND CLEAR B-SCAN BY DOCTOR AND RETINAL SPECIALIST. PATIENT HAD VITREOUS BIOPSY WITH INTRAVITREAL INJECTION OF VANCOMYCIN AND CEFTAZIDIME TO RULE OUT AND TREAT ENDOPHTHALMITIS AS WELL AS SUBTENON KENALOG DUE TO INITIAL FAST COURSE OF SIGNS OF TASS. THE NEXT DAY, DUE TO CELLS ON VITREOUS B-SCAN, THE PATIENT HAD BILATERAL PPV (PARS PLANA VITRECTOMY) WITH INTRAVITREAL INJECTION OF CEFTAZIDIME, VANCOMYCIN AND VITREOUS CULTURE. VITREOUS CULTURE GREW PSEUDOMONA AERUGINOSA AND CURRENTLY THE PATIENT IS HOSPITALIZED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT IS RECEIVING IV MEROPENEM, TOPICAL ANTIBIOTICS AND GLAUCOMA MEDS. IN THE LAST EXAMINATION BY RETINAL SPECIALIST, SHE WAS STILL HAVING PAIN IN BOTH EYES AND PERIORBITAL INFLAMMATION THAT HAD BEEN IMPROVED SINCE INITIAL HOSPITALIZATION EVALUATED BY CT AND MRI OF THE BRAIN AND ORBITS. CURRENT UCVA IS NLP OU ((B)(6) 2023). H6: HEALTH EFFECT IMPACT CODE: 4644 - MEROPENEM. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, D6B: UNK, H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PERMANENT LOSS OF BCVA, REDNESS, LIGHT SENSITIVITY. H6: HEALTH EFFECT IMPACT CODE: 4644 - CEFTAZIDIME, VANCOMYCIN. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PARS PLANA VITRECTOMY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4) | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980501&pc=QCB | 2023826-2020-02926 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PAIN, PHOTOPHOBIA, AND FIXED PUPIL. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED/ADDITIONAL INFORMATION: AN ADDITIONAL SUTURE WAS ALSO PERFORMED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE VAULT WAS RESOLVED. REFERENCE MFR FILE# (B)(4) FOR ADDITIONAL SYMPTOMS WITH THE REPLACEMENT LENS. H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136765&pc=QCB | 2023826-2022-02524 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. INTERMITTENT EYE PAIN IS REPORTED. THE LENS REMAINS IMPLANTED. THE REPORTER STATES CAUSE WAS DUE TO THE DEVICE. CAUSE DETAILS PROVIDED BY THE REPORTER DESCRIBE MULTIPLE PATIENT RELATED FACTORS THAT POTENTIALLY CONTRIBUTE TO THE INTERMITTENT EYE PAIN. THE REPORTER STATES, "THIS SUBJECT'S SYMPTOMS ARE COMPATIBLE WITH INTERMITTENT SUBACUTE ANGLE CLOSURE GLAUCOMA PRECIPITATED BY ACCOMMODATIVE MIOSIS DURING COMPUTER WORK. HOWEVER, THERE ARE FEATURES WHICH ARGUE AGAINST THIS DIAGNOSIS, INCLUDING THE ABSENCE OF ANY SIGNS OR SYMPTOMS FOLLOWING PHARMACOLOGIC DILATION, PATENT PIS AND NORMAL ICL VAULT. THE HISTORY OF MIGRAINE AND THE CONSTELLATION OF SYMPTOMS IN THIS SUBJECT SUGGESTS THAT A DIAGNOSIS OF BENIGN EPISODIC UNILATERAL MYDRIASIS MAY ALSO BE ENTERTAINED. NEVERTHELESS, THIS DIAGNOSIS MUST ONLY BE MADE IF INTERMITTENT ANGLE CLOSURE IS FIRST RULED OUT." Manufacturer Narrative: (B)(4). Manufacturer Narrative: B1: CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. B5: "REPORTEDLY, AFTER ORIGINAL TMICL SURGERY SUBJECT BEGAN EXPERIENCING EYE PAIN, FEELING OF PRESSURE, NAUSEA, AND PULSATING VISION WITH HEAR BEAT, INTERMITTENTLY OVER THE PAST SEVERAL MONTHS. EPISODES WOULD HAPPEN WHEN READING AND WHEN IN DIM LIGHTING. SUBJECT HAD OS TMICL REPLACED ON (B)(6) 2023 WHICH DECREASED THE VAULT OF THE ICL OVER THE CRYSTALLINE LENS. SUBJECT WILL BE FOLLOWED UNTIL HER EYE PAIN HAS BEEN SHOWN TO BE RESOLVED. " SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO: SERIOUS INJURY IN THE INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 2199 SHOULD BE REMOVED FROM THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5- IT WAS ALSO REPORTED, "SEVERITY: GRADE 3/SEVERE, OUTCOME: UNKNOWN AND RELATIONSHIP TO DEVICE: UNKNOWN". H6-HEALTH EFFECT- IMPACT CODE: 4629 SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- HEALTH EFFECT- CLINICAL CODE: 4581- INTERMITTENT MYDRIASIS. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED INTERMITTENT EYE PAIN AND INTERMITTENT MYDRIASIS. REPORTEDLY, "AFTER ORIGINAL TMICL SURGERY THE SUBJECT BEGAN EXPERIENCING EYE PAIN, FEELING OF PRESSURE, NAUSEA, AND PULSATING VISION WITH HEARTBEAT, INTERMITTENTLY OVER THE PAST SEVERAL MONTHS. EPISODES WOULD HAPPEN WHEN READING AND WHEN IN DIM LIGHTING". IT WAS ALSO REPORTED, "SEVERITY: GRADE 3/SEVERE, OUTCOME: UNKNOWN AND RELATIONSHIP TO DEVICE: UNKNOWN". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: " "(B)(6) 2022 SUBJECT COMPLAINED OF INTERMITTENT EPISODES OVER THE PAST SEVERAL WEEKS OF "FEELING EYE PRESSURE" OS WITH ACCOMPANYING OCULAR PAIN, HAZY VISION, OCCASIONAL NAUSEA AND A LARGER PUPIL THAN FELLOW EYE. UNCORRECTED VISUAL ACUITY MEASURED 20/20 AND INTRAOCULAR PRESSURE WAS 14 MMHG, AND THE REMAINDER OF THE EXAMINATION WAS UNREMARKABLE. THIS ADVERSE EVENT WAS NOTED AS RESOLVED 06APR2022. ON (B)(6) 2022 THE SUBJECT PRESENTED FOR AN UNSCHEDULED VISIT REPORTING EYE PAIN WHICH STARTED (B)(6) 2022. VISUAL ACUITY MEASURED 20/16, INTRAOCULAR PRESSURE WAS 16 AND LENS VAULT MEASURED 616 [?]M OS. THE TMICL LENS WAS EXCHANGED ON (B)(6) 2023 FOR VTICM5 12.1, -12.5/1.0/093 EVO+ ICL LENS. THE POSTOPERATIVE COURSE WAS UNEVENTFUL. ON 03MAR2023, THE SUBJECT REPORTED THERE HAD BEEN NO FURTHER EPISODES OF PAIN." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6: HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942788&pc=QCB | 2023826-2023-01890 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.5/011 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT COMPLAINT OF PAIN AROUND THE EYE AND THE SURGEON NOTES THE PATIENT HAD MUSCAE VOLITANTES (EYE FLOATERS). THE LENS WAS EXPLANTED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - MUSCAE VOLITANTES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387537&pc=QCB | 2023826-2023-05730 | 2023-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Eye Pain; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT WAS EXPERIENCING ORBITAL PAIN DUE TO PSYCHOLOGICAL EFFECTS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18389058&pc=QCB | 2023826-2023-05729 | 2023-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Eye Pain; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT WAS EXPERIENCING ORBITAL PAIN DUE TO PSYCHOLOGICAL EFFECTS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778992&pc=QCB | 2023826-2021-01327 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Swelling/ Edema | N/A | NA | 1 | Event Description: HE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE SUSPENSORY LIGAMENT OF THE EYE LENS THAT WAS FOUND TO BE RELAXED, AND PATIENT REPORTED SEVERE EYE PAIN, SWELLING AND DISTENSION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE INTRAOCULAR LENS WAS IMPLANTED INTO THE EYE TO SINK. AFTER THE OPERATION, THE PATIENT HAD SEVERE EYE PAIN, SWELLING AND DISTENSION, WHICH AFFECTED HIS LIFE. THEREFORE, THE LENS WAS STRONGLY REQUIRED TO BE REMOVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779023&pc=QCB | 2023826-2021-01328 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Eye Pain; Swelling/ Edema | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. SEVERE EYE PAIN, SWELLING AND DISTENSION WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. REPORTEDLY, "DURING THE OPERATION, THE SUSPENSORY LIGAMENT OF THE LENS WAS FOUND TO BE RELAXED, AND THE LENS WAS IMPLANTED INTO THE EYE TO SINK. AFTER THE OPERATION, THE PATIENT HAD SEVERE EYE PAIN, SWELLING AND DISTENSION, WHICH AFFECTED HIS LIFE. THEREFORE, THE LENS WAS STRONGLY REQUIRED TO BE REMOVED." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16622823&pc=QCB | 2023826-2023-01088 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO 09-FEB-2023. H6: 4629 SHOULD BE CORRECTED TO 4627. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS.IT IS UNCLEAR WHETHER LENS REMOVAL AND REPLACEMENT OCCURRED INTRAOPERATIVELY. RESIDUAL CYLINDER WAS PRESENT AFTER THE EXCHANGE. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 HYPEROPIC/SHOWING CYL. TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715562&pc=QCB | 2023826-2023-03867 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715631&pc=QCB | 2023826-2023-03868 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868572&pc=QCB | 2023826-2020-02721 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/+2.50/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND GLAUCOMA. IT WAS REPORTED THAT PATIENT'S CURRENT CONDITION IS GOOD AND GLAUCOMA WAS RESOLVED THROUGH MEDICATION. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10868645&pc=QCB | 2023826-2020-02720 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/+2.50/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND GLAUCOMA. IT WAS REPORTED THAT PATIENT'S CURRENT CONDITION IS GOOD AND GLAUCOMA WAS RESOLVED THROUGH MEDICATION. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: UNK. UNK. UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453104&pc=QCB | 2023826-2021-00590 | 2021-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -4.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENT'S EYE DUE TO CLOSED ANGLE GLAUCOMA BEING OBSERVED. NO PATIENT INJURY. ADDITIONAL SUTURE TO CLOSE THE WOUND. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN CONTACT LENS CASE, DRY, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND STRETCHED OUT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142126&pc=QCB | 2023826-2023-00099 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. POST-OP, THERE WAS REMAINING PUPIL DISSIPATION AFTER AN ACUTE GLAUCOMA ATTACK. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00098. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766002&pc=QCB | 2023826-2023-01455 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.00+1.50/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION; GLAUCOMA AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED THE PATIENT IS ON ANTI-GLAUCOMA DROPS FOR NOW. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION . TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773437&pc=QCB | 2023826-2023-01461 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -12.00/+1.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION; GLAUCOMA AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED THE PATIENT IS ON ANTI-GLAUCOMA DROPS FOR NOW. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766109&pc=QCB | 2023826-2023-01337 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Glaucoma; Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)-2023, YAG WAS PERFORMED DUE TO TRANSIENT LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY, "PATIENT HAD INTERMITTENT TRANSIENT VISION LOSS RESOLVED WITH LASER PI. HER LATEST IOP IS 12MMHG. POSSIBLE INTERMITTIENT RESULTS OF PUPILARY BLOCK PRIOR TO LASER PI IN THIS CASE. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(6). Manufacturer Narrative: CORRECTED DATA: B5- THE PATIENT ALSO EXPERIENCED GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. H6- HEALTH EFFECT- CLINICAL CODE: 4581- GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, YAG WAS PERFORMED, ELEVATED IOP: 24 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, PERIORBITAL EYE PAIN, MEDICATION WAS PRESCRIBED, GLAUCOMA, INTERMITTENT PUPILLARY BLOCK, AND HEADACHES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE PATIENT DEVELOPED OS SECONDARY GLAUCOMA POST ICL, PATIENT DECIDED TO REMOVE BOTH EYE ICL". CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153715&pc=QCB | 2023826-2022-02531 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Glaucoma; Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS GRADUAL LOSS OF VISION AND DIAGNOSIS OF BLINDNESS ((B)(6) 2022). BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION (ANTI GLAUCOMA AND ANTI INFLAMMATION) . THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- " DATE OF (B)(6) 2014 SHOULD BE CORRECTED TO (B)(6) 2014." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM 12.5 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LOSS OF VISION; BLINDNESS DUE TO GLAUCOMA (DIAGNOSED ON (B)(6) 2022; ELEVATED IOP; SECONDARY YAG PERFORMED; LENS REPOSITIONING, LENS ROTATION; INFLAMMATION; PIGMENT; AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE LENS REMAINS IMPLANTED. THE IMPLANTED LENSES ARE NOT LENSES SOLD BY STAAR JAPAN. THE LENSES WERE PURCHASED BY PHYSICIANS THROUGH PRIVATE IMPORTATION." THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE WAS REPORTED AS "BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION." THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN." H6- HEALTH IMPACT-CLINICAL CODE: 4581- PIGMENT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434301&pc=QCB | 2023826-2022-03078 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND BLURRED VISION. "PAIN IN RIGHT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "OCCASIONAL CLOUDINESS RIGHT EYE ;MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: CLINICAL CODE 4581: PIGMENT DISPERSION. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434626&pc=QCB | 2023826-2022-03079 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND GLARE/HALOES. "PAIN IN LEFT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE; PHOTOSENSITIVITY. PIGMENTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6-CLINICAL CODE 4581- PIGMENT DISPERSION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9672545&pc=QCB | 2023826-2020-00246 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Pain; Vomiting; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.0/+1.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP. THE REPORTER STATED THAT THE PATIENT DEVELOPED DENSE AC FIBRINOUS INFLAMMATORY REACTION, SECONDARY GLAUCOMA AND RIGHT FROM IMPLANTATION COMPLAINED OF SEVERE PAIN, WATERING AND VOMITING. PATIENT WAS UNDER STEROIDS. IT WAS REPORTED THAT HE PROBLEM HAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC, RESIDUE AND DEBRIS ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE, STEROIDS). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9672785&pc=QCB | 2023826-2020-00247 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Pain; Vomiting; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.0/+1.5/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP. THE REPORTER STATED THAT THE PATIENT DEVELOPED DENSE AC FIBRINOUS INFLAMMATORY REACTION, SECONDARY GLAUCOMA AND RIGHT FROM IMPLANTATION COMPLAINED OF SEVERE PAIN, WATERING AND VOMITING. PATIENT WAS UNDER STEROIDS. IT WAS REPORTED THAT HE PROBLEM HAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE, STEROIDS). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586880&pc=QCB | 2023826-2020-00065 | 2019-12-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/3.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE, EXCESSIVE VAULT, AND HEADACHE WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HEADACHE RELIEVED BY WEARING GLASSES PRESCRIBED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586884&pc=QCB | 2023826-2020-00064 | 2019-12-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.00/3.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE, EXCESSIVE VAULT WAS OBSERVED, WITH HEADACHE AND DIFFICULTY DRIVING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "FIRST CASE TO DEVELOP HYPEROPIA POST-OP WHEN SUPPOSED TO BE, THE PATIENT SHOULD BE SLIGHTLY MYOPIC POST-OP. HEADACHE WITHOUT GLASSES. PATIENT GIVEN GLASSES TO EASE HEADACHE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039558&pc=QCB | 2023826-2023-01879 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Headache | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCING SEVERE AND PERSISTENT HEADACHE AND MIGRAINE. PROBLEM RESOLVED. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. PATIENT USED HER NORMAL PAINKILLERS USED FOR MIGRAINE(NAPROXEN, SUMITRAPTAN). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039559&pc=QCB | 2023826-2023-01880 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Headache | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. LENS WAS EXPLANTED ON (B)(6)2023 DUE TO PATIENT EXPERIENCING SEVERE AND PERSISTENT HEADACHE AND MIGRAINE. PROBLEM RESOLVED. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. PATIENT USED HER NORMAL PAINKILLERS USED FOR MIGRAINE(NAPROXEN, SUMITRAPTAN). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980949&pc=QCB | 2023826-2021-04799 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -1.00/0.50/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND HEADACHES IN THE NIGHT ON THE PATIENT'S LEFT SIDE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641115&pc=QCB | 2023826-2022-00645 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -01.50/+1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PATIENT HAD A HEADACHE AT NIGHT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14294131&pc=QCB | 2023826-2022-01424 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -17.50/3.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. IT REPORTED THE CAUSE OF THE EVENT WAS "OTHER" AND NOTED "ICL TOO LARGE - PATIENT WITH HIGH HTO NON-MEASURABLE) AT D+1 POST OP VISIT WITH VOMITING; SEVERE HEADACHE". THE LENS WAS EXPLANTED ON (B)(6) 2022. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022, A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316632&pc=QCB | 2023826-2022-02946 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/1.50/064 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCPLANTED DUE TO ANISEIKONIA AND HEADACHE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). EXPIRATION DATE: 04/30/2024. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749935&pc=QCB | 2023826-2020-02521 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749941&pc=QCB | 2023826-2020-02520 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701923&pc=QCB | 2023826-2023-03896 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701924&pc=QCB | 2023826-2023-03895 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916646&pc=QCB | 2023826-2023-01725 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, INFLAMMATION, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), THE PATIENT IMPLANTED THE LENS IN JORDAN BY (B)(6) ON (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR (B)(6) INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO JORDAN FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP.". Manufacturer Narrative: CORRECTED DATA: TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/95 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(6). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916648&pc=QCB | 2023826-2023-01726 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND INFLAMATION WITH ELEVATED IOP WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), TE PATIENT IMPLANTED THE LENS IN (B)(6) BY (B)(6) IN (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR SAMER MULKI INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO (B)(6) FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP." Manufacturer Narrative: A4-A6:UNK H6: PATIENT RELATED FACTOR. CLAIM#(B)(4). Manufacturer Narrative: H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6- HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/78 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173083&pc=QCB | 2023826-2022-00057 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/172, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON 2021 DUE TO EXCESSIVE VAULTING, SHALLOW CHAMBER, SIGNIFICANT REDUCTION OF ANGLE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF A HEADACHE, WHICH WAS RESOLVED WITH MEDICATION (DIAMOX AND TOPICAL DRUGS). IOP REMAINS STABLE UNDER TOPICAL BETA-BLOCKER. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH ANOTHER ICL LENS WITHIN THE SAME SURGERY ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: (B)(4). HEALTH EFFECT IMPACT CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274802&pc=QCB | 2023826-2022-01399 | 2022-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.0/+1.0/14 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE FIXED PUPIL. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS ADDITIONALLY REPORTED THAT THERE IS NO PLANS TO REPLACE THE LENS AND THE PATIENT EXPERIENCED "GLAUCOMA SYMPTHOMATIC LIKE A HEADACHE AND VOMITING; HE DID NOT CONTACT THE CLINIC". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) ANGLE CLOSURE, FIXED PUPIL, VOMITING. TYPE OF INVESTIGATION: (B)(4) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433012&pc=QCB | 2023826-2023-00707 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -16.5/1.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; HEADACHES; SHALLOW ANTERIOR CHAMBER. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR AND NOTED "PATIENT HAS EXTREME FLOPPY IRIS WHICH IS CAUSING EXTREME BOWING OF IRIS TISSUE FORWARDS LEADING TO REDUCTION OF IRIDOCORNEAL ANGLE; LEFT EYE WORSE THAN RIGHT. IOP SPIKES PRESENT CAUSING HEADACHE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE, SHALLOW ANTERIOR CHAMBER. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433013&pc=QCB | 2023826-2023-00708 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -17.0/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; HEADACHES; SHALLOW ANTERIOR CHAMBER. THE PROBLEM IS NOT RESOLVED; AS THE SURGEON NOTED THE PATIENT STILL HAS HIGH VAULT WITH EXPLANT OR YAG PI IN CONSIDERATION. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR AND NOTED "PATIENT HAS EXTREME FLOPPY IRIS WHICH IS CAUSING EXTREME BOWING OF IRIS TISSUE FORWARDS LEADING TO REDUCTION OF IRIDOCORNEAL ANGLE; LEFT EYE WORSE THAN RIGHT. IOP SPIKES PRESENT CAUSING HEADACHE. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE, SHALLOW ANTERIOR CHAMBER. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16947904&pc=QCB | 2023826-2023-01916 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.50/6.0/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPOSITIONED/ROTATED, PER REPORTER THE PATIENT PRESENTED WITH ELEVATED INTRAOCULAR PRESSURE, HEADACHE AND NAUSEA FOLLOWING REPOSITIONING. FIBRIN PLUG OVER CENTRAFLOW HOLE WAS OBSERVED, ONCE RELEASED, INTRAOCULAR PRESSURE RETURNED TO NORMAL AND PROBLEM RESOLVED. Manufacturer Narrative: H6: HEALTH EFFECT- CLINICAL CODE 4581: NAUSEA, FIBRIN. H6: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781744&pc=QCB | 2023826-2020-00390 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.0/+2.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO MORE HEADACHES. COMPLIANT WITH DROPS." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200458&pc=QCB | 2023826-2021-00183 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/1.5/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED PAIN, HEADACHE, LIGHT SENSITIVITY, AND BLURRY VISION. PUPIL BLOCK WITH ELEVATED IOP WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (YAG PI) AND MEDICAL INTERVENTION (TROPICAMIDE, PHENYLEPHRINE, COSOPT PF, DIAMOX, BROMINIDINE) WERE PERFORMED/PRESCRIBED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, STATUS OF EYE "NORMAL". IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: B3 - (B)(6) 2020. B5 - EVALUATION ON 19-DEC-2020: EXAMINATION DEMONSTRATED CORNEAL EDEMA, A MID-DILATED PUPIL, ELEVATED IOP, ELEVATED ICL VAULT AND IOP. YAG PERFORMED. AQUEOUS TAP DONE (B)(6) 2020. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145766&pc=QCB | 2023826-2022-01205 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -11.5/+1.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 A PI-YAG WAS PERFORMED, THE SURGEON REPORTING TRANSIENT LOSS OF BCVA, PERIORBITAL EYE PAIN, HEADACHE, ELEVATED IOP, AND POSSIBLE INTERMITTENT BOUTS OF PUPILLARY BLOCK PRIOR TO LASER PI. THE LENS REMAINS IMPLANTED. REPORTEDLY THE TRANSIENT VISION LOSS WAS RESOLVED WITH THE LASER PI; LATENT IOP IS 12MMHG. Manufacturer Narrative: B5 LENS WAS EXPLANTED ON (B)(6) 2023. STATUS OF EYE IS REPORTED AS "PATIENT DEVELOPED LEFT EYE SECONDARY GLAUCOMA POST ICL IMPLANTATION ASSOCIATED WITH HEADACHE. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. "PATIENT COMPLAINTED HEADACHE ASSOCIATED WITH EYE PAIN." CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, THIS CLAIM IS FOR THE RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9677189&pc=QCB | 2023826-2020-00248 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -7.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE REPORTED. THE REPORTER ALSO STATED THAT "IRIS HAS BEEN DAMAGED AT THE LIMBAL REGION, WHERE YOU CAN SEE THAT THE PIGMENT LAYER HAS BEEN DISSOLVED." SECONDARY SURGICAL INTERVENTION (YAG) WAS PERFORMED ON POST-OP DAY 1 "BECAUSE THE CONDITION OF THE EYE AND PATIENT'S RAPPORT (HEADACHE AND VOMITING DEBUTING CA 8 H AFTER SURGERY) SUGGESTED THAT IOP HAD BEEN HIGH." IN THE SURGEONS OPINION THE CAUSE OF THE EVENT WAS DUE TO ICL TOO LARGE/ ACTUAL SULCUS DIAMETER SMALLER THAN EXPECTED FROM PREOP MEASUREMENTS. FOR PATIENTS' CURRENT CONDITION THE SURGEON REPORTED " PATIENT IS VERY PLEASED WITH THE OUTCOME, WEARS NO LENSES OR GLASSES AND WOULD LIKE TO KEEP EVERYTHING AS IT CURRENTLY IS; SO CLOSE TO EMMETROPIA OA. HOWEVER, THERE IS RATHER LARGE AMOUNTS OF PIGMENTATION ON ICL, IRIS ATROPHY AND THE VAULT IS STILL CA 4 OA". THE SURGEON PLANS TO REMOVE THE ICL WITH NO ICL IMPLANTATION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT CLINICAL CODE 1937- INTRAOCULAR PRESSURE INCREASE SHOULD BE IN THE INITAL MDR. H6- INVESTIGATION CONCLUCIONS CODE 4310 SHOULD BE IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9677374&pc=QCB | 2023826-2020-00249 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Vomiting; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_13.7, -7.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE REPORTED. SECONDARY SURGICAL INTERVENTION (YAG) WAS PERFORMED ON POST-OP DAY 1 "BECAUSE THE CONDITION OF THE EYE AND PATIENT'S RAPPORT (HEADACHE AND VOMITING DEBUTING CA 8 H AFTER SURGERY) SUGGESTED THAT IOP HAD BEEN HIGH." IN THE SURGEONS OPINION THE CAUSE OF THE EVENT WAS DUE TO ICL TOO LARGE/ACTUAL SULCUS DIAMETER SMALLER THAN EXPECTED FROM PREOP MEASUREMENTS. FOR PATIENTS' CURRENT CONDITION THE SURGEON REPORTED " PATIENT IS VERY PLEASED WITH THE OUTCOME, WEARS NO LENSES OR GLASSES AND WOULD LIKE TO KEEP EVERYTHING AS IT CURRENTLY IS; SO CLOSE TO EMMETROPIA OA. HOWEVER, THERE IS RATHER LARGE AMOUNTS OF PIGMENTATION ON ICL, IRIS ATROPHY AND THE VAULT IS STILL CA 4 OA". THE SURGEON STATED "I PLAN TO EXCHANGE THE ICL FOR ONE TWO SIZES DOWN." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT CLINICAL CODE 1937- INTRAOCULAR PRESSURE INCREASED SHOULD BE IN THE PREVIOUS MDR. H6- HEALTH EFFECT CLINICAL CODE 4851- IRIS ATROPHY SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616940&pc=QCB | 2023826-2021-00926 | 2021-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Headache; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A LENS WITH THE WRONG AXIS WAS IMPLANTED. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A1: PATIENT IDENTIFIER - (B)(6). B3: DATE OF EVENT - (B)(6) 2021. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND A HEADACHE. THERE WAS NO IOP ELEVATION AND NORMAL VAULT. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS STABLE AND VISION IS GOOD. THE REPORTER INDICATED THE MANIFEST REFRACTION AXIS WAS TYPED WRONG ON THE ORIGINAL ICL FORM (15 INSTEAD OF 115). D6A: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941304&pc=QCB | 2023826-2019-01572 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "HEADACHE." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE OPTICE AND HAPTIC WERE TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941310&pc=QCB | 2023826-2019-01573 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "HEADACHE." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275294&pc=QCB | 2023826-2023-00263 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+3.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO COMPLAINED OF PAIN AND PRESSURE SENSATION IN THE EYE, SO THE LENS WAS REMOVED. POST-OP, THE PATIENT HAS BEEN SEEING THE SURGEON EVERY 2 WEEKS, SOMETIMES COMPLAINING OF PAIN IN THE EYE, SOMETIMES OF PRESSURE AND HEADACHES. SHE HAS ALREADY VISITED THE EMERGENCY ROOM SEVERAL TIMES (WITHOUT RESULTS). H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PRESSURE SENSATION. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. PROCODE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275298&pc=QCB | 2023826-2023-00262 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/087(SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT IDENTIFIER: UNK. PROCODE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO COMPLAINED OF PAIN AND PRESSURE SENSATION IN THE EYE, SO THE LENS WAS REMOVED. POST-OP, THE PATIENT HAS BEEN SEEING THE SURGEON EVERY 2 WEEKS, SOMETIMES COMPLAINING OF PAIN IN THE EYE, SOMETIMES OF PRESSURE AND HEADACHES. SHE HAS ALREADY VISITED THE EMERGENCY ROOM SEVERAL TIMES (WITHOUT RESULTS). H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PRESSURE SENSATION. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182322&pc=QCB | 2023826-2018-01885 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.00/+4.0/033 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED SEVERE HEADACHE. THE REPORTER INDICATED THE RELATED SYMPTOM MAY BE DUE TO SHALLOW ANTERIOR CHAMBER (MID PERIPHERAL IRIDOCORNEAL TOUCH IN TURN CAUSING HEADACHE AND EYE PAIN). THE LENS REMAINS IMPLANTED. THE REPORTER STATED THE CAUSE OF THE EVENT IS THE DEVICE. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE EXCHANGE LENS DID NOT RESOLVE THE ISSUE. REFERENCE MFR REPORT # 2023826-2019-00435 FOR FURTHER DETAILS ON THE EXCHANGE LENS. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND BLUE RESIDUE ON LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070995&pc=QCB | 2023826-2022-02369 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070996&pc=QCB | 2023826-2022-02370 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986875&pc=QCB | 2023826-2023-04601 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSTATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548122&pc=QCB | 2023826-2020-02091 | 2020-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Headache; Vomiting | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PRESSURE 16MMHG, OPEN ANGLE, WERE OBSERVED. HEADACHE AND VOMIT WERE OCCSSIONALLY FELT BY PATIENT. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "SERGEON IS GOING TO REVIEW PATIENT IN 1 MONTH. NO SPECIFICT ACTION, ONLY OBSERVATION FOR NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - "REPORTEDLY, THE SURGEON THINKS THE CAUSE OF THE HEADACHE AND VOMITING IS DUE TO THE PATIENT'S ADAPTATION OF THE VISION." SHOULD BE ADDED TO THE INITIAL MDR. CORRECTED DATA: B5 - "...IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020." CORRECTED TO "...IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020." IN THE INITIAL MDR. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10548144&pc=QCB | 2023826-2020-02092 | 2020-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Headache; Vomiting | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.00/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. PRESSURE 16MMHG, OPEN ANGLE, WERE OBSERVED. HEADACHE AND VOMIT WERE OCCASIONALLY FELT BY PATIENT. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "SURGEON IS GOING TO REVIEW PATIENT IN 1 MONTH. NO SPECIFIC ACTION, ONLY OBSERVATION FOR NOW." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5 - "REPORTEDLY, THE SURGEON THINKS THE CAUSE OF THE HEADACHE AND VOMITING IS DUE TO THE PATIENT'S ADAPTATION OF THE VISION." SHOULD BE ADDED TO THE INITIAL MDR. B5 - "IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2020." CORRECTED TO "IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2020." IN THE INITIAL MDR. CLAIM#: 717258 Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147542&pc=QCB | 2023826-2021-02274 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Inflammation; Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 13.2MM TMICL13.2, -11.50/3.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE ON (B)(6) 2021. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ELEVATED IOP, INFLAMMATION AND PIGMENT DEBRIS OBSERVED. THE LENS WAS REMOVED AND CATARACT SURGERY WAS PERFORMED AND A NEW LENS AS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE ICL WAS IMPLANTED UPSIDE DOWN. THE PATIENT ALSO DISPLAYED A TEAR IN THE ANTERIOR CAPSULE AND CATARACT FORMATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14058808&pc=QCB | 2023826-2022-01151 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2022 SECONDARY YAG WAS PERFORMED. PER UPDATED INFORMATION AS OF (B)(6) 2022 INTRAOCULAR PRESSURE WAS 42MMHG, PATIENT IS BEING MONITORED WITH MEDICATION. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2022. THE REPORTER STATED, "PERSISTENT INFLAMMATION, HIGH IOP, VAULTING ½ CORNEA THICKNESS. EXPLANTED TO CONTROL INFLAMMATION. AND TO BETTER MANAGE IOP". CLAIM (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14062269&pc=QCB | 2023826-2022-01150 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.1, -8.00/+1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, INFLAMMATION AND PIGMENT DISPERSION. MEDICAL INTERVENTION WAS PRESCRIBED, IT WAS REPORTED THAT THE CONDITION OF THE PATIENT IS IMPROVING. PATIENT IS BEING MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741439&pc=QCB | 2023826-2023-03979 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND PIGMENT DISPERSION. PATIENT DEVELOPED CHRONIC INFLAMMATION WITH HIGH IOP. MEDICATION WAS ADMINISTERED - ANTIGLAUCOMA; COMBIGAN; AZOPT; XALATAN. POST-OP, THE ANTERIOR CHAMBER HAD MILD INFLAMMATION WITH PIGMENT CELL. THE PATIENT STILL CONTINUES WITH 2 TYPES OF ANTIGLAUCOMA - TIMOLOL BD AND AZOPT BD. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6: HEALTH EFFECT IMPACT CODE: 4644 - ANTIGLAUCOMA MEDS - COMBIGAN, AZOPT, XALATAN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549694&pc=QCB | 2023826-2023-03497 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A,12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -18.0/1.00/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "UP TO DATE THE EYE CONTINUES WITH THE INFLAMMATORY PROCESS." CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER STATED, "DUE TO A PATIENT INFLAMMATORY RESPONSE." Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984492&pc=QCB | 2023826-2019-01675 | 2018-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Iritis; Red Eye(s); No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED 2 WEEKS AFTER THE INITIAL SURGERY DUE TO POSSIBLE INFLAMMATION. ELEVATED INTRAOCULAR PRESSURE WAS ALSO REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS HAD LOW VAULTING AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE PATIENT'S LEFT EYE (OS) WAS RED AND HE CAME TO THE HOSPITAL FOR DIAGNOSIS ON (B)(6) 2018. THE DIAGNOSIS WAS IRITIS. THE PATIENT WAS TREATED WITH ANTI-INFLAMMATION AND INTRAOCULAR PRESSURE LOWERING MEDS. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE PATIENT'S CONDITION IS STABLE. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15336964&pc=QCB | 2023826-2022-02893 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | Inflammation; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. NARROW ANGLES AND IRIS ELEVATED REPORTED DUE TO LENS IMPLANTED UPSIDE DOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2022 THE LENS WAS REMOVED AND THE SAME LENS WAS RE-IMPLANTED IN THE CORRECT ORIENTATION. AFTER THE LENS WAS RE-IMPLANTED;. SUBSEQUENTLY, IOP INCREASED, INFLAMMATION-LIKE REACTION, MYDRIASIS. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. PUPIL BLOCK WITH ELEVATED IOP NOT MEASURABLE. ANGLE CLOSURE WITH ELEVATED IOP NOT MEASURABLE. ANTERIOR CHAMBER CONTENTS PRESENT, 8/6 IOP DIFFICULT TO MEASURE. THE STATUS OF THE PATIENT'S EYE WAS REPORTED AS " ON AUGUST 6, IOP ROSE (UNMEASURABLE) AND ANTERIOR CHAMBER CONTENTS APPEARED, SO THE ANTERIOR CHAMBER WAS RINSED. WITH EYE DROPS AND ORAL MEDICATION, THE IOP DECREASED TO 10 ON (B)(6), BUT THERE WAS RESIDUAL CONTENT, POOR PUPIL SHAPE, AND CORNEAL EDEMA. STEROID EYE DROPS WERE CONTINUED AND THE PATIENT'S SYMPTOMS IMPROVED. CURRENTLY, SHE CONTINUES TO USE STEROIDS AND PILOCARPINE IN THE LEFT EYE. PUPIL SHAPE HAS IMPROVED, BUT THE LEFT EYE TENDS TO CONTRACT, AND THERE IS CURRENTLY A DIFFERENCE IN PUPIL DIAMETER BETWEEN THE LEFT AND RIGHT EYES. THE PROBLEM WAS RESOLVED. CAUSE IS USER ERROR; AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D6B- (B)(6) 2022. H6- CLINICAL CODE: 4581: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ANGLE CLOSURE. ANTERIOR CHAMBER CONTENTS PRESENT. POOR PUPIL SHAPE. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022." D6A - CORRECTED TO: IMPLANT DATE (B)(6) 2022 . CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE: 30-APR-2025, CLINICAL SIGNS CODE (B)(4) - NARROW ANGLES; IRIS ELEVATED. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136620&pc=QCB | 2023826-2023-05112 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS NOTED THAT "OD PAIN EYELID EDEMA - INFLAMMATION DESPITE ANTI-INFLAMMATORY EYE DROPS FOR 3 DAYS - TOBRADEX = DEXAMETHASONE AND TOBRAMYCINE. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EYELID EDEMA. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10556172&pc=QCB | 2023826-2020-02123 | NA | Injury | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Decreased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE SURGEON REPORTS A HIGH VAULT OF 1006UM AND AN IOP<5MMHG. HOWEVER, BECAUSE OF HYPOTONY AND IRIS INCISION AND AFTER CORNEAL SUTURE, THE VAULT IS 720UM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD- BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORD(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." SHOULD HAVE BEEN IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/080(SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT, IRIS PROLAPSE AND ADDITION OF SUTURE, GLARE AND HALOES WERE OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2020. A LENS OF A SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WERE "NO COMPLAINTS" FOR THE STATUS OF THE EYE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619462&pc=QCB | 2023826-2020-02273 | 2020-09-01 | Malfunction | STAAR SUGICAL COMPANY | 2020-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.0/73 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. IT WAS REPORTER THAT THE PROBLEM RESOLVED, REPORTER STATED " IOP WENT DOWN. LENS REMAINS IN THE EYE. PATIENT IS HAPPY." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619517&pc=QCB | 2023826-2020-02272 | 2020-09-01 | Malfunction | STAAR SUGICAL COMPANY | 2020-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/2.0/102 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. IT WAS REPORTER THAT THE PROBLEM RESOLVED, REPORTER STATED " IOP WENT DOWN. LENS REMAINS IN THE EYE. PATIENT IS HAPPY." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138575&pc=QCB | 2023826-2020-01183 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -3.00/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (40 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR. PER THE REPORTER ON (B)(6) 2020, "PREDISPOSITION TO SUFFER FROM OPEN ANGLE GLAUCOMA. THE ANGLE IS CONFIRMED OPEN BY OCT OF THE ANTERIOR POLE. CURRENTLY INTRAOCULAR PRESSURE IS CONTROLLED BY AZARGA AND MONOPROST." Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - VTICH12.6 DIOPTER CORRECTED TO "+3.00/2.5/082" IN INITIAL MDR. H6 - DEVICE CODE: 1494, OFF-LABEL USE, ACD < 3.0MM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515444&pc=QCB | 2023826-2020-02020 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP (40MMHG) WITHOUR PUPIL BLOCK AND ANGLE CLOSURE OBSERVED ON (B)(6) 2020, AS WELL AS EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537754&pc=QCB | 2023826-2020-01978 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS LOST AND WILL NOT BE RETURNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537776&pc=QCB | 2023826-2020-01977 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND WITH RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538983&pc=QCB | 2023826-2020-02161 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10539564&pc=QCB | 2023826-2020-02162 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10601013&pc=QCB | 2023826-2020-02236 | 2020-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (18.5 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE BEING OBSERVED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "BIGGER SIZE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635478&pc=QCB | 2023826-2020-02299 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.00/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. RESIDUAL REFTRACTIVE ERROR WAS OBSERVED AND ROTATION. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "IT DOES NOT SEEM LIKE THE LENS ROTATED FAR FROM PLANNED ORIENTATION, AND FURTHER ROTATION IS NOT ADVISED." PATIENT WILL BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635480&pc=QCB | 2023826-2020-02298 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.00/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. RESIDUAL REFTRACTIVE ERROR WAS OBSERVED AND ROTATION. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "IT DOES NOT SEEM LIKE THE LENS ROTATED FAR FROM PLANNED ORIENTATION, AND FURTHER ROTATION IS NOT ADVISED." PATIENT WILL BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11060109&pc=QCB | 2023826-2020-03151 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT CYLINDER AND AXIS LENS DUE TO ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013551&pc=QCB | 2023826-2020-02984 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. A WEEK AFTER LENS WAS IMPLANTED, THERE WAS AN INTRAOCULAR PRESSURE SPIKE AND PATIENT WAS PRESCRIBED ALPHAGAN DROPS. TO DATE THE MEDICATION HAS BEEN TAPERED OFF AND LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11013716&pc=QCB | 2023826-2020-02983 | 2020-09-28 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. A WEEK AFTER LENS WAS IMPLANTED, THERE WAS AN INTRAOCULAR PRESSURE SPIKE AND PATIENT WAS PRESCRIBED ALPHAGAN DROPS. TO DATE THE MEDICATION HAS BEEN TAPERED OFF AND LENS REMAINED IMPLANTED. THE PATIENT IS HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851811&pc=QCB | 2023826-2020-02665 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/4.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2020, THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER ON (B)(6) 2020, SURGEON SAID PATIENT MIGHT HAVE HAD SYNECHIAE RETRO-IRIS, THEREFORE LENS DID NOT SIT PROPERLY. "SITUATION CALMED DOWN IMMEDIATELY AFTER THE EXPLANT, IOP BECAME NORMAL AND EXCEPT FEW DAYS AFTER EXPLANT SURGERY, CORNEA BECAME ALSO PERFECTLY TRANSPARENT. NO SIGNIFICANT EC LOS, NO SECONDARY CATARACT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON PRODUCT LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN THE PREVIOUS MDR. "LENS IS CURVED." SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135907&pc=QCB | 2023826-2021-00040 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT AND ELEVATED IOP, ENHANCED PIS AND ADMINISTRATION OF IOP DROPS AND DIAMOX. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN A PLASTIC BAG WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.5/+4.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 2020 (B)(6). THE SURGEON REPORTS ELEVATED IOP, ENHANCED PIS AND ADMINISTRATION OF IOP DROPS AND DIAMOX. ON 2021 (B)(6) THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11169226&pc=QCB | 2023826-2021-00069 | 2020-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, TICML12.1, 14.00/3.0/086, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD). ACCOMMODATIVE SPASMS, REFRACTIVE CHANGE OVERTIME, AND FOGGY VISION WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "AT (B)(6) 2020 VISIT, PATIENT VOICED COMPLAINT OF BLURRY VISION ONLY AFTER WORK WHICH WOULD PERSIST FOR MINUTES TO HOURS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TICML12.1 IMPLANTABLE COLLAMER OF DIOPTER -14.00/3.0/086 (SPHERE/CYLINDER/AXIS) LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ACCOMMODATIVE SPASMS, REFRACTIVE CHANGE OVERTIME, AND FOGGY VISION WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY ON (B)(6) 2020, "THE PATIENT VOICED A COMPLAINT OF BLURRY VISION ONLY AFTER WORK WHICH WOULD PERSIST FOR MINUTES TO HOURS." THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11175631&pc=QCB | 2023826-2021-00141 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE " LENS TOO LARGE". Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206965&pc=QCB | 2023826-2021-00109 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." MANUFACTURER NARRATIVE: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206967&pc=QCB | 2023826-2021-00110 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS." MANUFACTURER NARRATIVE: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405310&pc=QCB | 2023826-2021-00547 | NA | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A HIGH PRESSURE SPIKE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405316&pc=QCB | 2023826-2021-00548 | NA | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A HIGH PRESSURE SPIKE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467925&pc=QCB | 2023826-2021-00630 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -15.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTED AS CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467943&pc=QCB | 2023826-2021-00631 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -14.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTED AS CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -14.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. "LENGTH OF IMPLANT" WAS REPORTER AS CAUSE OF EVENT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11548140&pc=QCB | 2023826-2021-00736 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S EYE IS STABLE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE - THE LENS WAS TOO LARGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. EXPIRATION DATE: UNKNOWN. DEVICE MANUFACTURE DATE : UNKNOWN. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654112&pc=QCB | 2023826-2021-00967 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVER SIZE LENS." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721500&pc=QCB | 2023826-2021-01160 | 2020-10-13 | Injury | STAAR SURGICAL COMPANY | 2021-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # ((B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721925&pc=QCB | 2023826-2021-01038 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. A NEW PI AND YAG WAS PERFORMED ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP(34MMHG ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, "PATIENTS SULCUS IS AN ANATOMICAL VARIANET FROM THE NORMAL WHICH RESULTED IN HIGHER VAULT THAN EXPECTED WITH THE FDA NOMOGRAM." Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -12/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. A NEW PI AND YAG WAS PERFORMED ON (B)(6) 2021.THE LENS WAS EXPLANTED ON (B)(6) DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP(34MMHG ASSESSED ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED (B)(6) 2021, IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, "PATIENTS SULCUS IS AN ANATOMICAL VARIANET FROM THE NORMAL WHICH RESULTED IN HIGHER VAULT THAN EXPECTED WITH THE FDA NOMOGRAM." CORRECT INFORMATION: H6 - HEALTH EFFECT-IMPACT CODE - REPORTED AS 4629 - CORRECT TO 4627. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE PUPIL ISSUE WAS A RESULT OF THE INITIAL SURGERY NOT THE REPLACEMENT. THE PATIENT IS DOING WELL WITH ALL OTHER SIDE EFFECTS BESIDES THE PUPIL RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822933&pc=QCB | 2023826-2021-01437 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -14.50/0.50/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. YES BECAUSE THE INITAL LENS WAS TOO LONG AND THE VAULT TOO HIGH". Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12018124&pc=QCB | 2023826-2021-01921 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.5/2/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP. ADDITIONAL MEDICAL INTERVENTION (BROMINIDINE AND TIMOLOL) WAS NECESSARY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS CAUSED BY BOTH PROCEDURE AND DEVICE. Manufacturer Narrative: B5- PER UPDATED INFORMATION "CAUSED BY BOTH PROCEDURE AND DEVICE. INCREASED IOP SAME DAY POSTOPERATIVE VISIT DUE TO RETAINED VISCOELASTIC." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12058246&pc=QCB | 2023826-2021-02040 | NA | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ELEVATED IOP. REPORTER STATED " SHE WAS SEEN IN THE ER MULTIPLE TIMES AND WAS PLACED ON GLAUCOMA DROPS. IT WAS REPORTED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND A SHORTER LENGTH LENS WAS LATER IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12059229&pc=QCB | 2023826-2021-02039 | NA | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT EXPERIENCED ELEVATED IOP. REPORTER STATED " SHE WAS SEEN IN THE ER MULTIPLE TIMES AND WAS PLACED ON GLAUCOMA DROPS. IT WAS REPORTED THAT THE LENS WAS REMOVED ON (B)(6) 2021 AND A SHORTER LENGTH LENS WAS LATER IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072258&pc=QCB | 2023826-2021-02069 | 2021-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, IT WAS REMOVED DUE TO ELEVATED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: A SHORTER LENS WAS IMPLANTED INTO THE PATIENT'S EYE AFTER THE INITIAL LENS WAS REMOVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097194&pc=QCB | 2023826-2021-02208 | 2021-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -6.5/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND THE PROBLEM RESOLVED. USER ERROR WAS THE CAUSE OF THIS EVENT, IT WAS REPORTED THAT PERIPHERAL IRRIDECTOMIES ARE MISSING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102434&pc=QCB | 2023826-2021-02115 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/6.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. TRANSIENT LOSS OF BCVA AND ELEVATED IOP (30MMHG) WITHOUT PUPIL BLOCK,ANGLE CLOSURE, AND PIGMENT DISPERSION, WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PRESENCE OF INFLAMMATION COMPOSED OF CELLS AND IRIS PIGMENTS IN THE ANTERIOR CHAMBER. OCT WILL BE DONE TO SEE IF THERE IS ADHERENCE OF LENS TO IRIS ANTI INFLAMMATORY DROPS GIVEN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12150435&pc=QCB | 2023826-2021-02532 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/118 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID SOME ADDITIONAL LASER SURGERY TO THE IRIDOTOMY (ENLARGED) AND THAT RESOLVED THE ISSUE. THE PATIENT IS FINE AND THERE WAS NO NEED TO EXPLANT THE LENS, THE LENS REMAINED IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12203583&pc=QCB | 2023826-2021-02557 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -12.5/1.0/131 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE AND NOTED "EXCESSIVE VAULTING CAUSING ANGLES TO NARROW". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE PATIENT WAS SEEN FOR AN EVALUATION, PATIENT HAD HIGH VAULTING BUT NORMAL IOP ON HYPOTENSIVE MEDICATION AND WAS INSTRUCTED TO DISCONTINUE MEDICATION. THEN ON (B)(6) 2021 ANOTHER EVALUATION WAS PERFORMED AND THE PATIENT HAD INCREASED IOP OFF MEDICATION AND A DECISION WAS MADE TO REPEAT/ENLARGE LPIS TO ASSESS IF ANY IMPROVEMENT IN IOP. IT WAS ALSO INDICATED IF NO IMPROVEMENT WITH IOP OCCURS A LENS EXCHANGE IS PLANNED. H6 - HEALTH EFFECT IMPACT CODE: 4625 ADDITIONAL SURGERY - LPI REPEAT/ENLARGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT'S DIURNAL IOP MEASUREMENT AND IOP IS WITHIN NORMAL LIMITS/STABLE AFTER REPEAT LPI AND OFF HYPOTENSIVE MEDICATIONS, AND THEY WOULD NOT BE PROCEEDING WITH A LENS EXCHANGE. PATIENT WILL CONTINUED BEING MONITORED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204759&pc=QCB | 2023826-2021-02558 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -17.0/2.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE AND NOTED "EXCESSIVE VAULTING CAUSING ANGLES TO NARROW". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE PATIENT WAS SEEN FOR AN EVALUATION, PATIENT HAD HIGH VAULTING BUT NORMAL IOP ON HYPOTENSIVE MEDICATION AND WAS INSTRUCTED TO DISCONTINUE MEDICATION. THEN ON (B)(6) 2021, ANOTHER EVALUATION WAS PERFORMED AND THE PATIENT HAD INCREASED IOP OFF MEDICATION AND A DECISION WAS MADE TO REPEAT/ENLARGE LPIS TO ASSESS IF ANY IMPROVEMENT IN IOP. IT WAS ALSO INDICATED IF NO IMPROVEMENT WITH IOP OCCURS A LENS EXCHANGE IS PLANNED. H6 - HEALTH EFFECT IMPACT CODE: 4625 ADDITIONAL SURGERY - LPI REPEAT/ENLARGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT'S DIURNAL IOP MEASUREMENT AND IOP IS WITHIN NORMAL LIMITS/STABLE AFTER REPEAT LPI AND OFF HYPOTENSIVE MEDICATIONS, AND THEY WOULD NOT BE PROCEEDING WITH A LENS EXCHANGE. PATIENT WILL CONTINUED BEING MONITORED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214302&pc=QCB | 2023826-2021-02675 | 2021-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP (45 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2021. CONCOMITANT MEDICATIONS WERE PERSCRIBED. LENS REMAINS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250513&pc=QCB | 2023826-2021-02972 | 2021-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE PATIENTS INTRAOCULAR PRESSURE WAS ELEVATED DURING THE SURGERY. THE LENS WAS REPLACED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308398&pc=QCB | 2023826-2021-03083 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.00/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (54MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308417&pc=QCB | 2023826-2021-03084 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (60MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12407186&pc=QCB | 2023826-2021-03328 | 2021-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(32MMHG) (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12510969&pc=QCB | 2023826-2021-03705 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/1.00/22 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IT WAS REPORTED THAT A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS OTHER AND IT WAS REPORTED THAT THE DEVICE FAIL TO PERFORM AS INTENDED. CAUSE DETAILS " HIGH WTW MEASUREMENT " WAS REPORTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587047&pc=QCB | 2023826-2021-03864 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. IOP WAS STILL HIGH ON THE 1 AND 3 DAY FOLLOW-UP POST LENS REMOVAL. PATIENT IS RESPONDING TO ANTI-GLAUCOMA MEDICATIONS AND THE IOP IS CONTROLLED. THE PATIENT WILL BE CLOSELY FOLLOWED UP. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587050&pc=QCB | 2023826-2021-03863 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+4.0/109 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND LARGE VAULT. IOP WAS STILL HIGH ON THE 1 AND 3 DAY FOLLOW-UP POST LENS REMOVAL. PATIENT IS RESPONDING TO ANTI-GLAUCOMA MEDICATIONS AND THE IOP IS CONTROLLED. THE PATIENT WILL BE CLOSELY FOLLOWED UP. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5- AS PER THE COMPLAINT QUESTIONNAIRE, THE REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN, HOWEVER SUBSEQUENT COMMUNICATION VIA EMAIL ON (B)(6) 2021 STATED THAT "THE DOCTOR TOLD ME TODAY THAT IN HIS OPINION THIS IS STEROID INDUCED IOP". CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS ARE REQUIRED BUT CANNOT BE PERFORMED DUE TO THE LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685440&pc=QCB | 2023826-2021-04063 | 2021-09-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED AFTER EXPLANT, NO FURTHER SURGICAL INTERVENTION IS PLANNED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "PREOPERATIVELY IOP WAS 23MMHG, BCVA 20/32; AFTER IMPLANTATION IOP WAS 60MMHG (AT 5PM), PRESSURE VIA PARACENTESIS REDUCED IOP TO 20MMHG (18 MINUTES AFTER 5PM); AT 6PM IOP INCREASED TO 50MMHG AGAIN SO EXPLANTATION WAS PERFORMED IMMEDIATELY." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685432&pc=QCB | 2023826-2021-04062 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+0.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THERE WAS NO REDUCTION OF IOP UNDER ORAL MEDICATION BY GLAUPAX. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685468&pc=QCB | 2023826-2021-04061 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/034 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THERE WAS NO REDUCTION OF IOP UNDER ORAL MEDICATION BY GLAUPAX. THE PROBLEM WAS RESOLVED AFTER EXPLANT. SEE MFR REP# 2023826-2021-04063 FOR REPLACEMENT LENS. Manufacturer Narrative: H3: DEVICE EVALUATON - LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014953&pc=QCB | 2023826-2021-04810 | 2018-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+5.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT -RELATED FACTOR.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+5.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT -RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054368&pc=QCB | 2023826-2021-04738 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065424&pc=QCB | 2023826-2021-04743 | 2021-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: D6: (B)(6) 2024. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115357&pc=QCB | 2023826-2021-04768 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/088 (SPHERE/CYLINDER/AXIS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256067&pc=QCB | 2023826-2022-00117 | 2021-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.5/3.0/091 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256668&pc=QCB | 2023826-2022-00119 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297269&pc=QCB | 2023826-2022-00186 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 24MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00185. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 24MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00185. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM # (B)(6). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM # (B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297379&pc=QCB | 2023826-2022-00185 | 2021-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00186 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297459&pc=QCB | 2023826-2022-00184 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED ON (B)(6) 2021, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 29MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00183. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL SURGERY WAS PERFORMED ON (B)(6) 2021, WITH THE ADDITION OF NEW PIS BY YAG BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. AT LAST POST-OP VISIT THE IOP WAS 29MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00183. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297565&pc=QCB | 2023826-2022-00183 | 2021-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00184 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE IOP INCREASED CONTINUOUSLY AFTER OPERATION. SEE MFR. REPORT # 2023826-2022-00184 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13361414&pc=QCB | 2023826-2022-00300 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/6.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/6.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423349&pc=QCB | 2023826-2022-00253 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -11.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (28MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED 03/JAN/2022) EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. REPORTEDLY EYE QUIET AND PATIENT HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13481675&pc=QCB | 2023826-2022-00320 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509882&pc=QCB | 2023826-2022-00333 | 2021-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.50/3.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577875&pc=QCB | 2023826-2022-00499 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -8.00/2.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -8.00/2.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577879&pc=QCB | 2023826-2022-00500 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -6.00/1.50/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (47MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726672&pc=QCB | 2023826-2022-00758 | 2021-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE IS VERY GOOD. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13905945&pc=QCB | 2023826-2022-00907 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.6; -6.00/+2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXPLANTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(4) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14116520&pc=QCB | 2023826-2022-01194 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.00/+1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE AND PATIENT-RELATED FACTOR. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14282382&pc=QCB | 2023826-2022-01213 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. "PATIENT WITH ABNORMAL CILIARY BODY MORPHOLOGY AND NARROWING OF THE ATRIAL ANGLE ARE NOT SUITABLE FOR SURGERY." Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14282387&pc=QCB | 2023826-2022-01212 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS EXPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. "PATIENT WITH ABNORMAL CILIARY BODY MORPHOLOGY AND NARROWING OF THE ATRIAL ANGLE ARE NOT SUITABLE FOR SURGERY." Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14403295&pc=QCB | 2023826-2022-01524 | NA | Injury | STAAR SURGICAL COMPANY | 2022-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A REPLACEMENT LENS OF SPHERICAL AND SHORTER LENGTH WAS IMPLANTED. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. SEE MFR # 2023826-2022-01552 FOR REPLACEMENT LENS. Manufacturer Narrative: CORRECTED DATA: H6-TYPE OF INVESTIGATION CODE : 10 IN PREVIOUS MDR SHOULD BE CORRECTED TO 4110. ADDITIONAL INFORMATION: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14403301&pc=QCB | 2023826-2022-01525 | NA | Injury | STAAR SURGICAL COMPANY | 2022-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/111(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED. A REPLACEMENT LENS WAS LATER IMPLANTED WITH A SPHERICAL MODEL AND SHORTER LENGTH. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. SEE MFR # 2023826-2022-01553 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-TYPE OF INVESTIGATION CODE : 10 IN PREVIOUS MDR SHOULD BE CORRECTED TO 4110. ADDITIONAL INFORMATION: H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM #: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632891&pc=QCB | 2023826-2022-01661 | 2022-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2, -9.0/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, LENS WAS NOT PROPERLY SIZED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14687899&pc=QCB | 2023826-2022-01855 | 2022-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. LATER ON THE SAME DATE IN A SECOND SURGERY THE LENS WAS EXPLANTED AND NOTED "THE PATIENT WOULDN'T IMPLANTED WITH THE LENS". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829910&pc=QCB | 2023826-2022-01877 | 2022-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -5.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829913&pc=QCB | 2023826-2022-01878 | 2022-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -5.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983249&pc=QCB | 2023826-2022-02206 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED. IT WAS NOTED "THE PATIENT IS UNDER CONSIDERATION; AND AT THE TIME OF THE SECOND OPERATION IS UNCERTAIN". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4): CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042197&pc=QCB | 2023826-2022-02165 | 2022-02-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -10.5/3.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117753&pc=QCB | 2023826-2022-02455 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/129 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP) (STEROID RESPONDER). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE PATIENT IS DOING WELL AND FEELS GOOD. THE IOP IS GOOD. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136362&pc=QCB | 2023826-2022-02502 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.0/3.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS HIGH IOP. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567405&pc=QCB | 2023826-2022-03565 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT . THE LENS WAS LATER REPLACED WITH A SPHERICAL LENS OF THE SAME LENGTH DIFFERENT POWER AND THE PROBLEM RESOLVED. FOR ADDITIONAL INFORMATION THE REPORTER STATED " BECAUSE OF THE HIGH INTRAOCULAR PRESSURE AT THE INITIAL STAGE AFTER SURGERY, THE DOCTOR ORIGINALLY DECIDED TO CUSTOMIZE THE ASTIGMATIC LENS WITH SPECIAL AXIAL POSITION FOR REPLACEMENT AFTER FACE-TO-FACE CONSULTATION. AFTER REPEATED EXAMINATIONS, THE PATIENT AGAIN DECIDED TO CUSTOMIZE THE NON-ASTIGMATIC LENS FOR REPLACEMENT, AND THE ASTIGMATISM DEGREE WAS CORRECTED BY LASER SURGERY." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6: IMPACT CODE: 4628 ADDED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801805&pc=QCB | 2023826-2022-04085 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE INTRAOCULAR PRESSURE (IOP) WAS REPORTED AS BEING ELEVATED AND COMBIGAN BID AND PRED FORTE BID WERE INITIATED. THE ELEVATED IOP WAS RESOLVED AND WAS MEASURED WITHIN NORMAL RANGE ON (B)(6) 2022. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811021&pc=QCB | 2023826-2022-03961 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/057(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. CAUSE OF THE EVENT IS THE DEVICE. THE REPORTER STATED THAT THE LENS WAS, "OVERSIZED." THE PROBLEM IS NOT RESOLVED. REPORTEDLY, "EXPLANTED ICL, IOP NOT CONTROLLED AFTGER ICL REMOVAL. 30MMHG 2 WEEKS POST-OP." Manufacturer Narrative: D4: CORRECTED TO UNIQUE IDENTIFIER (UDI) # 00841542131032 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15903005&pc=QCB | 2023826-2022-04235 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "2 MONTHS AFTER THE OPERATION, THE IOP HAS BEEN BETWEEN 26-29 MMHG, THE PATIENT WAS ANXIOUS AND ASKED TO REMOVE THE LENS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE: SHOULD BE CORRECTED TO 31-MAY-2025. H4-DEVICE MANUFACTURE DATE:SHOULD BE CORRECTED TO 13-JUN-2022. CLAIM#(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15903054&pc=QCB | 2023826-2022-04234 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "2 MONTHS AFTER THE OPERATION, THE IOP HAS BEEN BETWEEN 26-29 MMHG, THE PATIENT WAS ANXIOUS AND ASKED TO REMOVE THE LENS." Manufacturer Narrative: OFF LABEL AGE<21, PRE-EXISTING CONDITION: PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000610&pc=QCB | 2023826-2022-04626 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000620&pc=QCB | 2023826-2022-04625 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+2.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269394&pc=QCB | 2023826-2023-00310 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269398&pc=QCB | 2023826-2023-00311 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. REPORTEDLY, "REMOVE AND NO LONGER IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." IN THE INITAL MDR. CLAIM #(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275247&pc=QCB | 2023826-2023-00258 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275257&pc=QCB | 2023826-2023-00257 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500236&pc=QCB | 2023826-2023-00819 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500270&pc=QCB | 2023826-2023-00818 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00/+0.5/100 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500271&pc=QCB | 2023826-2023-00817 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.00/+0.5/045 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500966&pc=QCB | 2023826-2023-00820 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587673&pc=QCB | 2023826-2023-00952 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/109 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). AN ADDITION OF NEW PI WAS PERFORMED BY YAG ON (B)(6) 2022 BUT THE PRESSURE DID NOT GO DOWN. THE PATIENT ENDED UP HAVING A CATARACT LENS IMPLANTED. Manufacturer Narrative: PATIENT INFO: UNK. HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION OF NEW PI. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16600700&pc=QCB | 2023826-2023-01008 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.5/003 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 DUE TO REFRACTIVE SUPRISE. REPORTEDLY, "POSTOPERATIVE DAY 1 IOP AND I/A IS PERFORMED. CURRENTLY, DOCTOR IS TAKING A WAIT-AND-SEE APPROACH." THE CAUSE OF THE EVENT WAS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16721439&pc=QCB | 2023826-2023-01332 | 2022-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF --14.0/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPANTED AND REMOVED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: UNK. ACD <3.0. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792688&pc=QCB | 2023826-2023-01434 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WERE OBSERVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. REPORTEDLY, "HIGH IOP WITH MEDICATIONS. PLAN TO EXCHANGE TO A SMALLER SIZE (12.6). Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON 21 APR 2023 WITH A SHORTER LENGTH LENS BUT PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. LENS RETURNED IS NOT THE SIZE STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792797&pc=QCB | 2023826-2023-01433 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND ANGLE CLOSURE WITH ELEVATED IOP WERE OBSERVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. REPORTEDLY, "HIGH IOP WITH MEDICATIONS. PLAN TO EXCHANGE TO A SMALLER SIZE (12.6)." Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS BUT PROBLEM WAS NOT RESOLVED. CLAIM# (B)(6). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16917117&pc=QCB | 2023826-2023-01760 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -13.00/+1.0/39 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216365&pc=QCB | 2023826-2023-02549 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). ELEVATED IOP AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT TO BE OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216368&pc=QCB | 2023826-2023-02550 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). ELEVATED IOP AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT TO BE OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253137&pc=QCB | 2023826-2023-02657 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.00/1.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. DEVICE WAS MENTIONED AS THE CAUSE OF THE EVENT. REPORTEDLY, "ANTERIOR CATARACT. FIXED BLOWN PUPIL WITH SYNECHAIE." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322467&pc=QCB | 2023826-2023-02901 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/129 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT DECIDED TO REMOVE IT AND IS AFRAID TO IMPLANT IT AGAIN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES-UNABLE TO EVALUATE. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322480&pc=QCB | 2023826-2023-02848 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/126 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT DECIDED TO REMOVE IT AND IS AFRAID TO IMPLANT IT AGAIN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES-UNABLE TO EVALUATE. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17465253&pc=QCB | 2023826-2023-03310 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LENS ROTATION AND IOP SPIKES. IT WAS INDICATED THERE WILL BE NOT REPLACEMENT LENS. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17465254&pc=QCB | 2023826-2023-03311 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.0/002 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION AND IOP SPIKES. IT WAS INDICATED THERE WILL BE NOT REPLACEMENT LENS. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17914527&pc=QCB | 2023826-2023-04273 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136876&pc=QCB | 2023826-2023-05060 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/0.5/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136885&pc=QCB | 2023826-2023-05061 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223147&pc=QCB | 2023826-2023-05376 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.00/3.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258193&pc=QCB | 2023826-2023-05318 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -7.5/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ELEVATED IOP 34MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED. OPHTHALMIC DROPS PRESCRIBED(BROMFENAC SODIUM HYDRATE AND LATANOPROST). NO OTHER SYSTEMS WERE OBSERVED. PROBLEM RESOLVED. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK (B)(4) | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18287965&pc=QCB | 2023826-2023-05442 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC EVO ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS EYE. AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. H6: HEALTH EFFECT IMPACT CODE: 4644 - IOP LOWERING MEDICATIONS. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. THE LENS REMAINED IMPLANTED. AT THE MOST RECENT POST-OP, PRESSURE HAS STABILIZED ON MEDICATIONS AND THE PATIENT IS BEING MONITORED CLOSELY. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298068&pc=QCB | 2023826-2023-05435 | 2023-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346344&pc=QCB | 2023826-2023-05691 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.00/2.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER AND INDICATED THAT THE DEVICE PERFORMED AS INTENDED BUT THE LENS WAS TOO LARGE. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238020&pc=QCB | 2023826-2021-00168 | 2020-12-18 | Malfunction | STAAR SURGICAL COMPANY AG | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/0.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2013. REFRACTIVE CHANGE OVERTIME AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE - (B)(6) 2013. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537728&pc=QCB | 2023826-2020-02067 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON ENLARGED/ADDITION OF PERIPHERAL IRIDOTOMY, BY YAG. A SUTURE WAS REQUIRED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS VISION IS GOOD BUT HAS MILD GLARE/HALO AT NIGHT. Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566551&pc=QCB | 2023826-2020-02177 | 2020-07-18 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HUGE VAULT AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON PERFORMED A COLOBOMA ON (B)(6) 2020, WHICH CORRECTED THE SITUATION A LITTLE BIT, BUT ANGLES WERE STILL CLOSED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10566577&pc=QCB | 2023826-2020-02176 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HUGE VAULT AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON PERFORMED A COLOBOMA ON (B)(6) 2020, WHICH CORRECTED THE SITUATION A LITTLE BIT, BUT ANGLES WERE STILL CLOSED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE LENS HAS BEEN EXPLANTED ON (B)(6) 2020. A NEW LENS WAS LATER IMPLANTED. THE PATIENT WAS SATISFIED WITH OPERATION; AND THE DOCTORS WAS SATISFIED WITH RESULTS. D6B. (B)(6) 2020 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE LENS HAS BEEN EXPLANTED ON AN UNKNOWN DATE. FURTHER INFORMATION HAS BEEN REQUESTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D6B. UNK H3- DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT. HAPTIC TORN/BROKEN. VISUAL OBSERVATION CONCLUDED NO DIMENSIONAL INSPECTION CAN BE PERFORMED DUE LENS RETURNED IN PIECES. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +6.00/+2.0/132 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS HAS BEEN EXPLANTED ON (B)(6) 2020 AND ON THE SAME DAY A NEW LENS OF SAME MODEL BUT LONGER LENGTH WAS IMPLANTED. THIS RESOLVED THE PROBLEM. REPORTER STATES CAUSE AS "THE DOCTOR SUPPOSE THAT THE SIZE WAS SMALLER THAN THEY NEED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634500&pc=QCB | 2023826-2020-02281 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND INTERMITTENT ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2020. ADDITIONAL SURGERY WAS PERFORMED, ENLARGEMENT OF PI BY YAG PROCEDURE AND THERE WAS AN ADDITIONAL SUTURE. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; ENLARGEMENT OF PI BY YAG, SUTURE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634519&pc=QCB | 2023826-2020-02294 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+5.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND UNREACTIVE (FIXED) PUPIL. THE IOP WAS CONTROLLED WITH AGGRESSIVE TREATMENT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO AN OVERSIZED LENS. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ANTERIOR SEGMENT WAS OK AND THERE WAS A LITTLE BIT OF PIGMENT DISPERSION BUT STILL OK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE CORRECTED TO 4627. B5 - IT WAS INDICATED THAT THE LENS WAS EXPLANTED ON (B)(6) 2020, AND LATER ON THE SAME DAY IN A SEPERATE SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. POSTOPERATIVE REPORT INDICATED THAT THE ANTERIOR SEGMENT WAS OK AND THERE WAS A LITTLE BIT OF PIGMENT DISPERSION BUT STILL OK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635452&pc=QCB | 2023826-2020-02242 | 2020-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/2.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT AND ELEVATED IOP (22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE THAT WAS ASSESSED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "OVER SIZE LENS." REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED SHALLOW CHAMBER. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW CHAMBER. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALAUTION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11099005&pc=QCB | 2023826-2020-03218 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND SHALLOW CHAMBER. THERE WAS AQUEOUS MISDIRECTION, WHICH WAS RELIEVED WITH VITREOUS TAPS. THE LENS WAS EXCHANGED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD AN EXCESSIVE VAULT. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. SEE MFR. REPORT # 2023826-2020-03220 FOR REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - HEALTH EFFECT CLINICAL CODE : 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5 - "THE CORNEA WAS CLEAR, CENTERED ICL WITH NOTE OF IRIS PIGMENTS ON LENS" SHOULD BE REMOVED AS IT IS FOR THE REPLACEMENT LENS. H6 - HEALTH IMPACT CLINICAL CODE: 4581 IN H10 DESCRIPTION - SHOULD BE CORRECTED TO "HEALTH EFFECT CLINICAL CODE: 4581", AND IRIS PIGMENT SHOULD BE REMOVED AS IT IS FOR REPLACEMENT LENS. H6 - HEALTH EFFECT CLINICAL CODE : 4581 - MEDICATION, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091484&pc=QCB | 2023826-2020-03221 | 2020-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00/+3.0/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091943&pc=QCB | 2023826-2020-03217 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND SHALLOW CHAMBER. THERE WAS AQUEOUS MISDIRECTION, WHICH WAS RELIEVED WITH VITREOUS TAPS. THE LENS WAS EXCHANGED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS STILL HAD AN EXCESSIVE VAULT. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. SEE MFR. REPORT # 2023826-2020-03219 FOR REPLACEMENT LENS. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH IMPACT - CLINICAL CODE: 4581 - IRIS PIGMENTS SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED, THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICIATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHALLOW CHAMBER, IRIS PIGMENTS HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY - AQUEOUS MISDIRECTION/VITREOUS TAPS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091947&pc=QCB | 2023826-2020-03220 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING HIGH VAULTING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, THE CORNEA WAS CLEAR, ICL WAS CENTERED, WITH IRIS PIGMENTS ON LENS. THIS WAS THE REPLACEMENT LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2020-03218. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT PATIENT ALSO EXPERIENCED ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM RESOLVED AND NOTED "PATIENT STILL ON AZARGA AND LUMIGAN". H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, MEDICATION. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11092035&pc=QCB | 2023826-2020-03219 | 2020-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING A HIGH VAULT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. MEDICATION WAS PRESCRIBED, CLEAR CORNEA, CENTERED ICL, WITH IRIS PIGMENTS ON LENS. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-03217. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. IT WAS NOTED THE PATIENT WAS STILL ON AZARGA AND LUMIGAN. HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BENT, WITH DEBRIS ON THE LENS. CORRECTED DATA: G6 - DISREGARD PREVIOUS SUPPLEMENTAL #2. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135881&pc=QCB | 2023826-2021-00039 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.0/+4.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT AND ELEVATED IOP, WHICH RESPONDED TO ENHANCED PIS. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: UNREACTIVE (FIXED) PUPIL AND ADDITION OF NEW PI-YAG ALSO REPORTED. ON (B)(6) 2020, THE LENS WAS EXPLANTED. AFTER EXPLANT THE EYE STATUS IS LISTED AS DEEPENED AC, GOOD IOP, PUPIL STILL SLUGGISH AT 5.5MM. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: SULCUS DIAMETER IN VIVO APPEARS MUCH SMALLER THAN THE NOMOGRAM-RECOMMENDED 13.2MM TICL. H6-CLINICAL CODE 4581: [UNREACTIVE (FIXED) PUPIL, DEEPENED AC]. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343216&pc=QCB | 2023826-2021-00407 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IIDO-CORNEAL ANGLES AND AC SHALLOW IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS SHALLOW AC. PER THE REPORTER ON (B)(6) 2021, THERE WILL NOT BE ANOTHER LENS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343218&pc=QCB | 2023826-2021-00406 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2, -7.50/2.0/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IIDO-CORNEAL ANGLES AND AC SHALLOW IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS SHALLOW AC. PER THE REPORTER ON (B)(6) 2021, THERE WILL NOT BE ANOTHER LENS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366548&pc=QCB | 2023826-2021-00478 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "TOO BIG LENGTH OF ICL 13.7 FOR THIS EYE. DISSOCIATION BETWEEN THE REAL SIZE OF POSTERIOR CHAMBER AND PREDICTED BY CALCULATOR." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11389451&pc=QCB | 2023826-2021-00545 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF ICA. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478839&pc=QCB | 2023826-2021-00639 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM WAS RESOLVED. THE SURGEON IMPLANTED THE EXCHANGED ICL LENS VERTICALLY TO REDUCE VAULT AND AN INCISION WAS USED TO REDUCE ASTIGMATISM. THE PATIENT WAS SATISFIED WITH THE RESULT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616998&pc=QCB | 2023826-2021-00922 | 2021-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TOPICAL ANTIHYPERTENSIVE TREATMENT (KRYTANTEK) WAS ADMINISTERED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021. POST-OP VISUAL ACUITY WAS 20/30 UVA; 20/25 CVA. PIO: 18MMHG WITH NO ANTI GLAUCOMA TREATMENT. SLIT LAMP - NO INFLAMMATORY FINDINGS. VAULT - 386UM, ANGLE:20 DEGREES. THE PATIENT AND DOCTOR ARE HAPPY WITH THE RESULTS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11624856&pc=QCB | 2023826-2021-00934 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM012.6, -17.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PIGMENT DISPERSION WAS OBSERVED. ON PATIENTS LAST VISIT DATE (B)(6) 2021 FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) " VISION IS EXCELLENT, IOP 12MM/HG WITH MEDS BUT THERE IS A IRIS DISPERSION FOR 3 MONTHS POST-OP ASSOCIATED WITH TRANSPARENT IRIS AT THE SIDE OF LENS HAPTICS" WAS REPORTED. FOR ADDITIONAL INFORMATION "HOWEVER THE LENS VAULT IS VERY LOW, THERE IS A LIGHT TRANSPARANCY AT THE SIDE OF THE HAPTICS" WAS REPORTED. ON (B)(6) 2021 THE REPORTER STATED " I LIKE TO ACKNOWLEDGE THAT PATIENT IS GETTING BETTER WEEK AFTER WEEK. DISPERSION HAS ALREADY STOPPED BUT THE ILLUMINATION IS EXISTING". LENS REMAINS IMPLANTED IT WAS REPORTED THAT NO INTERVENTION HAS BEEN PERFORMED FOR THE EYE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER STATED THAT "2 WEEKS AFTER SX HE COME WITH HIGH IOP OF 37. AND ANT.CHAMBER REACTION OF 2+. GIVEN TOPICAL ANTI GLAUCOMA AND REDUCED STEROID". THIS SHOULD HAVE BEEN ADDED IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11648783&pc=QCB | 2023826-2021-00923 | 2021-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TOPICAL ANTIHYPERTENSIVE TREATMENT (KRYTANTEK) WAS ADMINISTERED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021. POST-OP VISUAL ACUITY WAS 20/30 UVA; 20/25 CVA. PIO: 18MMHG WITH NO ANTI GLAUCOMA TREATMENT. SLIT LAMP: NO INFLAMMATORY FINDINGS. VAULT 439UM; ANGLE 21 DEGREES. THE PATIENT AND THE DOCTOR ARE HAPPY WITH THE RESULTS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK EXPLANT DATE: UNK. REPORT SOURCE: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11722521&pc=QCB | 2023826-2021-01209 | 2021-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/+0.5/164 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, -11.50 DIOPTER, SPHERICAL LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND UNREACTIVE PUPIL (PUPIL IS REACTIVE TO LIGHT BUT VERY DILATED). FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "AWAITING FOLLOW UP" WAS REPORTED." DEVICE" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " SURGEON INTERPRETATION IS THAT THE ANGLE IS OPEN BUT THAT THE CONFIGURATION AND VAULT OF ICL IS SUCH THAT THE PUPIL CANNOT CONSTICT OVER THE LIP OF THE OPTIC RESULTING IN THE PUPIL BEING MECHANICALLY PREVENTED FROM RETURNING TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT IT WAS AN IRIS CONFIGURATION ISSUE AND NOT AN ISSUE WITH VAULT OR THE ICL. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11780620&pc=QCB | 2023826-2021-01371 | 2021-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+6.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (THE IRIS WAS TOUCHING THE CORNEA). ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS CALCULATION ERROR IN SUGGESTED LENS LENGTH. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11780627&pc=QCB | 2023826-2021-01370 | 2021-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00/+6.0/116 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (THE IRIS WAS TOUCHING THE CORNEA). ANOTHER ICL LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS CALCULATION ERROR IN SUGGESTED LENS LENGTH. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11796745&pc=QCB | 2023826-2021-01390 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-05-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -16/2/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND ANGLE CLOSURE. ADDITIONAL SURGICAL INTERVENTION (AQUEOUS TAP AND ENLARGEMENT OF PI YAG LASER) AND MEDICAL INTERVENTION (IOP LOWERING MEDICATION) WERE PERFORMED/PRESCRIBED PRIOR TO ICL EXPLANTATION BUT DIDN`T RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANT " IOP AND ANGLE GRADE RETURNED TO NORMAL RANGE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: PATIENT INFORMATION :UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11796748&pc=QCB | 2023826-2021-01391 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-05-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -17.5/3.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND ANGLE CLOSURE. ADDITIONAL SURGICAL INTERVENTION (AQUEOUS TAP AND ENLARGEMENT OF PI YAG LASER) AND MEDICAL INTERVENTION (IOP LOWERING MEDICATION) WERE PERFORMED/PRESCRIBED PRIOR TO ICL EXPLANTATION BUT DIDN`T RESOLVE THE ISSUE. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXPLANT " IOP AND ANGLE GRADE RETURNED TO NORMAL RANGE." IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: PATIENT INFORMATION :UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855173&pc=QCB | 2023826-2021-01511 | 2021-04-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP(24MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "LENS TOO BIG FOR EYE." REPORTEDLY, SURGEON TO REPLACE WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855239&pc=QCB | 2023826-2021-01522 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/1.0/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP(57MMHG), ANGLE CLOSURE WITH ELEVATED IOP(57MMHG), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, TRANSIENT LOSS OF BCVA, EXCESSIVE VAULT, IRIS ATROPHY, BLURRED VISION, AND UNREACTIVE(FIXED) PUPIL BEING OBSERVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED STUCK IN A GAUZE, WITH MOISTURE, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, WITH RESIDUE ON SURFACE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894540&pc=QCB | 2023826-2021-01710 | NA | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY, INCREASED LASER ENERGY WAS APPLIED AND A 2ND MID PERIPHERAL PI WAS PERFORMED. PROBLEM RESOLVED. THE ELEVATED IOP WAS RELATED TO A NON PATENT SUPER PERIPHERAL PI. Manufacturer Narrative: B5: UPDATED INFO: REPORTEDLY THE PRESSURE IS UP AND THE ANGLE IS SHALLOW. THE DOCTOR BELIEVES THAT THE LENS MAY BE UPSIDE DOWN. HIGH VAULT IS REPORTED. THE DOCTOR WOULD LIKE TO EXCHANGE THE LENS. H6-HEALTH CLINICAL CODE 4581 (SHALLOW ANGLE). CLAIM# (B)(4). Manufacturer Narrative: LENS WAS EXPLANTED AND EXCHANGED AND TO THE REPORTER'S BEST KNOWLEDGE, PATIENT HAS DONE QUITE WELL. #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938946&pc=QCB | 2023826-2021-01644 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP(75MMHG,ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938953&pc=QCB | 2023826-2021-01645 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP (75MMHG,ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966819&pc=QCB | 2023826-2021-01819 | 2021-05-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH3.2, -10.00/3.0/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ANGLE CLOSURE WITH ELEVATED IOP(64MMHG ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT WERE OBSERVED. ADDITION OF NEW PI, AND YAG PERFORMED ON (B)(6) 2021. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR, LENS SLIGHTLY TOO LARGE FOR PATIENT'S EYE. PATIENT IS HIGHLY HYPEROPIC HENCE SMALLER EYE. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966836&pc=QCB | 2023826-2021-01818 | 2021-05-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH3.2, -10.00/2.0/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ANGLE CLOSURE WITH ELEVATED IOP(76MMHG ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT WERE OBSERVED. ENLARGEMENT OF PI, ADDITION OF NEW PI, AND YAG PERFORMED ON (B)(6) 2021. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR, LENS SLIGHTLY TOO LARGE FOR PATIENT'S EYE. PATIENT IS HIGHLY HYPEROPIC HENCE SMALLER EYE. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170834&pc=QCB | 2023826-2021-02603 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.5/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170835&pc=QCB | 2023826-2021-02602 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178448&pc=QCB | 2023826-2021-02701 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -16.0/+1.0/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "DEVICE BEING TOO LARGE CAUSED ANTERIOR VAULT, RESULTING IN NARROWING ANTERIOR CHAMBER AND INCREASED IOP. THE PROBLEM IS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: PER EMAIL ON (B)(6) 2021, THE REPORTER STATED THAT "REPLACEMENT LENS WERE IMPLANTED ON (B)(6) 2021. RIGHT EYE: (B)(6) & LEFT EYE: (B)(6). ALL PROBLEMS HAVE BEEN RESOLVED AT THIS TIME." CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214185&pc=QCB | 2023826-2021-02636 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(26.7MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT CLINICAL CODE (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214348&pc=QCB | 2023826-2021-02679 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/3.5/044 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. PUPIL BLOCK WITH ELEVATED IOP (65MMHG) WAS ASSESSED ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. MEDICATION PRESCRIBED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT RELATED FACTOR. REPORTEDLY, NON-REACTIVE PUPIL IS RESPONDING TO MYOSIS WITH THE ADMINISTRATION OF EYE DROPS (PILOCARPINE). SURGEON IS CONSIDERING DISCONTINUING DROPS AND ASSESS IF THE PUPIL REACTS TO LIGHT WITHOUT MEDICATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -2.50/3.5/044 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK; ANGLE CLOSURE AND ELEVATED IOP REPORTED. A WEEK FOLLOWING TILTED LENS; HIGH VAULT; AND FIBRIN IN AC, AND REFRACTIVE SURPRISE WAS OBSERVED. MEDICATION WAS ADMINISTERED. IRIS ATROPHY REPORTED FOLLOWINGLY. LENS REMAINS IMPLANTED. UPDATE ABOUT PATIENT: RESIDUAL ASTIGMATISM 0.75 D; NON-REAGENT PUPIL; PUPIL DIAMETER 4,5/5MM; EYE IS QUIET; AND IOP AS 24MMHG BUT FOLLOWINGLY WAS 34 MMHG. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12228684&pc=QCB | 2023826-2021-02916 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/+3.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND IT WAS NOTED VISCOELASTIC WAS DRAINED IN A SEPARATE SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND INDICATED "SINCE LENS WAS TORIC AND WTW WAS IN THE LIMIT FOR BOTH SIZE; WE CHOSE THE BIGGER SIZE". Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE VAULT STILL BEING HIGH. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12264140&pc=QCB | 2023826-2021-03041 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13/+1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)-2021. THE SURGEON REPORTS EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND TOPICAL TREATMENT (PROSTAGLANDIN DROPS) WERE USED. THE SURGEON NOTED A REPLACEMENT LENS WAS ORDERED AS A SOLUTION. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406884&pc=QCB | 2023826-2021-03327 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/1.0/095 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478337&pc=QCB | 2023826-2021-03586 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.3MM, TMICL13.2, -6.5/2.5/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTOT THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED. IT WAS REPPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS REQUIRED. AC WASHOUT, VRT TAP AND SUTURE WOUND WAS REPORTED TO BE PERFORMED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND IT WAS REPORTED THAT IT IS UNSURE YET IF THE PROBLEM RESOLVED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): THE REPORTER STATED " 2+ CELL, FORMED AC, LENS EDGE VISIBLE SUP AND INF. ANGLE NARROW AT NASAL & TEMP. OPEN BUT SHALOW, DILATED IRIS." IT WAS ALSO REPORTED THAT LENSES MAY NEED TO BE REPLACED WITH SMALLER SIZE. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502723&pc=QCB | 2023826-2021-03578 | NA | Injury | STAAR SURGICAL COMPANY | 2021-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/088, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. POST-OP, AFTER LENS WAS EXPLANTED, PATIENT HAD FIXED DILATED PUPIL WITH IRIS ATROPHY (URRETS ZAVALIA). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558563&pc=QCB | 2023826-2021-03805 | NA | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS INTERMITTENT ANGLE CLOSURE WITH INCREASED IOP, PLATEAU IRIS, AND LOW-DOSE PILO ADMINISTRATION. REPORTEDLY THE VISION IS GOOD AND PATIENT DOES NOT WANT TO REMOVE THE LENS; LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12910019&pc=QCB | 2023826-2021-04438 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -12.50/1.50/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS DEVICE RELATED TO HIGH VAULT WITH 13.2 LENS SIZE. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12993610&pc=QCB | 2023826-2021-04882 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385990&pc=QCB | 2023826-2022-00337 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.0/+1.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES. THE PROBLEM IS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.0/+1.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES. THE PROBLEM IS RESOLVED. Manufacturer Narrative: REPORTEDLY THE LENS WAS TOO LONG, CAUSING A SHALLOW ANGLE AND IOP ISSUES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385991&pc=QCB | 2023826-2022-00339 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. REPORTEDLY, "THE PATIENT HAS EYE PRESSURE IN SPITE OF ALL OCULAR TREATMENT." AFTER REMOVAL, THE EYE PRESSURE IS STABLE BUT THE PUPIL IS DILATED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. PRODUCT EVALUATION PERFORMED EVEN THOUGH INSPECTION WAS NOT REQUIRED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423368&pc=QCB | 2023826-2022-00265 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2021 DUE TO ELEVATED IOP (21.3MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13518185&pc=QCB | 2023826-2022-00479 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.00/+1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5- PER ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2022: THE PROBLEM WAS REPORTED AS RESOLVED. "PATIENT DOESN'T WISH FURTHER SURGERY PATIENT WEAR GLASSES/CONTACTLENSES". CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13538841&pc=QCB | 2023826-2022-00512 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.0/057 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE; ELEVATED IOP; AND ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.0/057 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE; ELEVATED IOP; AND ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H3 DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND WITH FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552706&pc=QCB | 2023826-2022-00437 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13581966&pc=QCB | 2023826-2022-00593 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; ANGLE CLOSURE AND INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588934&pc=QCB | 2023826-2022-00598 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995866&pc=QCB | 2023826-2022-00933 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (23.3 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14058325&pc=QCB | 2023826-2022-01149 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/+3.5/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. IT WAS REPORTER THAT INTRAOCULAR PRESSURE WAS TAPER DOWN TO 12MMHG. PER UPDATED INFORMATION INTRAOCULAR PRESSURE IS CONTROLLING AND IMPROVING AND DOCTOR CONTINUES TO MONITOR PATIENT. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "IT WAS ALSO REPORTED HTA THE PATIENT HAD A GRAND MAL SEIZURE IN THE IMMEDIATE POT OPERATIVE PERIOD ON A BACKGROUND OF A HISTORY OF SEIZURE FIVE YEARS PREVIOUSLY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161436&pc=QCB | 2023826-2022-01083 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(331MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2022), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439471&pc=QCB | 2023826-2022-01368 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/5.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT; ELEVATED IOP AND SIGNIFICANT REDUCTION OF ICA WERE REPORTED. ANTI GLAUCOMA MEDICATIONS ADMINISTERED TO MAINTAIN NORMAL IOP. ENLARGED PUPIL REACTIVE TO LIGHT; IRIS BOWING; PIGMENTS IN ANTERIOR CHAMBER. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697153&pc=QCB | 2023826-2022-01766 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.00/3.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (45MMHG ASSESSED ON (B)(6) 2022). ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS VERY HAPPY WITH THE NEW VISION. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15168658&pc=QCB | 2023826-2022-02589 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.5/+2.5/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE/FIXED PUPIL AND IRIS ATROPHY. IT WAS REPORTED THAT THE EXPLANT DID NOT RESOLVED THE ISSUE. THE REPORTED STATED " AFTER REMOVAL OF THE LENS, THE PUPIL DID NOT CONSTRICT. NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE: UNDER 21 YRS OF AGE AT TIME OF INITIAL IMPLANTATION. | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178277&pc=QCB | 2023826-2022-02601 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+2.0/039 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. AT 3 WEEKS POST-OP IOP WAS ELEVATED TO MID 20'S AND PATIENT WAS STARTED ON TRAVATAN EYEDROP. AT 5 WEEKS POST-OP PRESSURE NORMALISED BUT PATIENT HAD OCULAR SURFACE SIDE EFFECTS FROM EYEDROP. PATIENT STARTED ON TREATMENT FOR OCULAR SURFACE DISEASE. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT WAS DIAGNOSED WITH PERIPHERAL ANTERIOR SYNECHIAE. PHACO WAS PERFORMED AFTER REMOVAL OF THE ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179016&pc=QCB | 2023826-2022-02575 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -16.0/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282214&pc=QCB | 2023826-2022-02651 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP(51 MMHG) ASSESSED ON (B)(6) 2022, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477821&pc=QCB | 2023826-2022-03263 | 2022-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.0/0.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INRAOPERATIVELY DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506720&pc=QCB | 2023826-2022-03359 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15630190&pc=QCB | 2023826-2022-03698 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.00/4.50/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE, THE REPORTER STATED " THE TICL WAS LARGE IN SIZE". FOR ADDITIONAL INFORMATION " EYE IS MONITORED FOR ELEVATED IOP BUT ANGLE IS STILL CLOSED BY THE IRIS AND THE TICL SIZE IS LARGE." WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659600&pc=QCB | 2023826-2022-03713 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS OF THE SAME LENGTH DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ELEVATED IOP WITHOUT PUPIL BLOCK/ANGLE CLOSURE WAS ALSO REPORTED. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN,BROKEN AND BENT TO THE LENS. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764262&pc=QCB | 2023826-2022-04078 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; AND ADDITIONALLY NOTED "QUIET EYE; UNREACTIVE FIXED DILATED PUPIL/BORDERLINE IOP; SLIGHTLY DISCORIC PUPIL/PATIENT COMPLAINS ALSO FROM THE PUPIL SIZE; THE EXCESSIVE VAULTING VAULT CAUSED IRIS CHAFFING WITH SEVERE PIGMENT DISPERSION AND PIGMENT DEPOSITS ON THE ICL'S OPTIC. ADDITIONALLY IT CAUSED DISCORIC PUPIL AND THE PUPIL BECAME MID-DILATED AND UNRESPONSIVE; ALSO IRIS ATROPHY CAUSE BY IRIS CHAFING". ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE PROBLEM IS REPORTED AS NOTE RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED EXCESSIVE VAULTING USING THE MANUFACTURED SUGGESTED ICL SIZE; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15910184&pc=QCB | 2023826-2022-04341 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+5.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE EYE WAS BURNING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO ADVERSE EVENT: B2: SHOULD BE CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. D4: UNIQUE IDENTIFIER NUMBER: (B)(4). H1: SHOULD BE CORRECTED TO SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+5.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE EYE WAS BURNING. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993011&pc=QCB | 2023826-2022-04420 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -4.00/2.5/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. "PATIENT DID NOT WANT TO UNDERGO ANOTHER LENS IMPLANTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993013&pc=QCB | 2023826-2022-04419 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.00/1.5/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. "PATIENT DID NOT WANT TO UNDERGO ANOTHER LENS IMPLANTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000862&pc=QCB | 2023826-2022-04423 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.5/1.0/129 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000874&pc=QCB | 2023826-2022-04424 | NA | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT DOES NOT WANT THE LENS TO BE REMOVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.0/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. REPORTEDLY THE LENS HAS BEEN EXPLANTED AT AN UNKNOWN DATE. H6:4627 ADDED AND 2199 OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT 13.2MM VTICM5 13.2 IMPLANTABLE COLLAR LENS OF -13.0/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ELEVATED IOP AND NARROW ANGLES WERE REPORTED. REPORTEDLY THE LENS HAS BEEN EXPLANTED ON (B)(6)2023. THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. PROBLEM RESOLVED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309076&pc=QCB | 2023826-2023-00446 | 2023-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.5/2.0/050 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS PIGMENT DISPERSION AND ELEVATED INTRAOCULAR PRESSURE AND MEDICATION IS BEING USED TO LOWER IOP. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND NOTED "ICL-IRIS CHAFE". Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) - PIGMENT DISPERSION. CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE LENS WAS EXPLANTED AND REPLACED WITH A VTICM5 12.6 WITH A DIOPTER OF -7.5/2.0/070. REPORTEDLY, " I WAS WITH DR. SKINNER FOR HIS EXPLANT RELATED TO THIS COMPLAINT. HE IMPLANTED A SMALLER SIZE AND IS WATCHING THE PATIENT."" SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343041&pc=QCB | 2023826-2023-00480 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+1.5/27 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND LENS REPOSITIONING WAS PERFORMED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16439865&pc=QCB | 2023826-2023-00603 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.50/005 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULT; ELEVATED IOP AND UNREACTIVE FIXED PUPIL. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.50/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULT; ELEVATED IOP AND UNREACTIVE FIXED PUPIL. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLINICAL CODE 4581: UNREACTIVE (FIXED) PUPIL. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773438&pc=QCB | 2023826-2023-01462 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.5/1.0/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED "REMOVED THE LENS AS THE IOP WENT UP WITH LOT OF PIGMENT DISPERSION". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773439&pc=QCB | 2023826-2023-01463 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.0/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ADDITIONALLY NOTED "REMOVED THE LENS AS THE IOP WENT UP WITH LOT OF PIGMENT DISPERSION". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791590&pc=QCB | 2023826-2023-01471 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP; 30MMHG. ON (B)(6) 2023 A REPLACEMENT LENS TWO SIZES SHORTER IN LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN AND THE REPORTER STATED IT WAS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE 4581- ANGLE CLOSURE (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16873134&pc=QCB | 2023826-2023-01681 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/2.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THERE IS ICL DEPOSITS ON THE LENS, BUT NO REFRACTIVE OUTCOMES OR LOSS OF VISION ACUITY OR VISION QUALITY. BUT LENS TRANSPARENCY SEEMS TO BE AFFECTED; THERE IS A RISE IN IOP AND MEDICATION HAS USED AS TREATMENT. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION. B5 - IT WAS LATER INDICATED THAT THE SURGEON EXPLANTED/REPLACED THE LENS AND THE PATIENT'S POSTOP IS INDICATED AS EVERYTHING IS FINE BUT A 1 MONTH POSTOP IS SCHEDULED SOON. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DEPOSITS ON LENS, LENS TRANSPARENCY. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16962964&pc=QCB | 2023826-2023-01910 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/003 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/003 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. H6 - HEALTH EFFECT CLINICAL CODE: 1937 (INTRAOCULAR PRESSURE RISE). CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976342&pc=QCB | 2023826-2023-01871 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF DOUBLE VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 14-JUN-2023. G3: 14/AUG/2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DOUBLE VISION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16988828&pc=QCB | 2023826-2023-01872 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2022. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT COMPLAINED OF DOUBLE VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED AND THE OPTIC TORN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DOUBLE VISION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17014968&pc=QCB | 2023826-2023-02065 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.0/+3.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. CLAIM#:(B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON 06-JUN-2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17014969&pc=QCB | 2023826-2023-02066 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.5/+2.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091758&pc=QCB | 2023826-2023-02289 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+0.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ANGLE CLOSURE WITH ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE ALSO NOTED. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT CLINICAL CODE 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342540&pc=QCB | 2023826-2023-02889 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) . THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE MYDRIASIS CONTINUED AFTER GLAUCOMA ATTACK. SEE MFR. REPORT # 2023826-2023-02890 FOR REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON 27-APR-2023. THE LENS WAS REMOVED ON 01-JUN-2023 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE MYDRIASIS CONTINUED AFTER GLAUCOMA ATTACK. SEE MFR. REPORT # 2023826-2023-02890 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D6B: 01-JUN-2023. CLAIM # (B)(4). Manufacturer Narrative: A3: UNK. A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17492163&pc=QCB | 2023826-2023-03432 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.50/2.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. OCULAR HYPER TENSION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.5/2.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. A SHORTER LENGTH DIFFERENT POWERED ( SPHERE/CYLINDER/AXIS) LENS WAS EXCHANGED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP 21MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2023) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. H4 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: OCULAR HYPER TENSION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION PROVIDED WITHIN REPORT. CORRECTION: H4 REPORTED AS BLANK - CORRECTED TO 16/JUN/2022. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17502541&pc=QCB | 2023826-2023-03460 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/+1.0/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IDIRO-CORNEAL ANGLES, ADDITIONALLY MEDICAL INTERVENTION OF COMBIGAN AND COSOPT 3 TIMES A DAY. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE SURGEON NOTED THERE IS A PLANNED LENS EXCHANGE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS EXPLANTED AND REPLACED IN A DIFFERENT SURGERY ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLEAR CORNEA, FORMED AC, ICL IN PLACE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17502604&pc=QCB | 2023826-2023-03461 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.00/+0.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IDIRO-CORNEAL ANGLES, ADDITIONALLY MEDICAL INTERVENTION OF COMBIGAN AND COSOPT 3 TIMES A DAY. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE SURGEON NOTED THERE IS A PLANNED LENS EXCHANGE. Manufacturer Narrative: B5 - LENS EXPLANTED AND REPLACED IN A DIFFERENT SURGERY ON (B)(6)2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLEAR CORNEA, FORMED AC, ICL IN PLACE. (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669068&pc=QCB | 2023826-2023-03734 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -8.0/3.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP 21MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2023). WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. H6 - WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17913515&pc=QCB | 2023826-2023-04272 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/1.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL DUE TO ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916405&pc=QCB | 2023826-2023-04520 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2023 AN ADDITION OF NEW PI (YAG) WAS PERFORMED. LENS REMAINS IMPLANTED. A FOLLOW UP WILL BE PERFORMED WAS NOTED BY THE SURGEON. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2023. THIS RESOVED THE PROBLEM. . H6 - CODE 4581 - ANGLE CLOUSRE, SHOULD BE REMOVED. H6- CODE 4625 - YAG, SHOULD BE REMOVED. CLAIM #: (B)(4). Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT THE PROBLEM WAS NOT RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916406&pc=QCB | 2023826-2023-04521 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/+2.5/121 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND ON (B)(6) 2023 AN ADDITION OF NEW PI (YAG) WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948893&pc=QCB | 2023826-2023-04442 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, -3.0/3.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(30MMHG)(ASSESED ON (B)(6) 2023), AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, NARROW ANGEL IN MYOPIC PATIENT. EPISODES OF HIGH PREESURE WHEN DILATED PUPIL. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- "-3.0/3.0/101" SHOULD BE CORRECTED TO "+3.0/3.0/101". H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970807&pc=QCB | 2023826-2023-04632 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE AND PIGMENT DISPERSION. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) IT WAS REPORTED THAT THE IOP (INTRAOCULAR PRESSURE) IS NORMAL WITH SINGLE DRUG THERAPY. THE REPORTER ALSO STATED " WAITING FOR IOP TO LOWER FURTHER. THE ANGLES ARE OPEN NOW PIGMENT DISPERSION STILL PRESENT". Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) . | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17998693&pc=QCB | 2023826-2023-04775 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.50/+2.5/129 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO A PRIOR ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED. IT IS IMPLIED THE PREVIOUS PROBLEMS STILL PERSIST. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999332&pc=QCB | 2023826-2023-04699 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THERE WAS A SMALL TEAR IN THE ICL DURING LOADING. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE. H6 - WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048081&pc=QCB | 2023826-2023-04728 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -7.0/1.0/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED IOP(69MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: LENS TOO BIG. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 (IRIS COMPROMISE). H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097675&pc=QCB | 2023826-2023-05045 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.00/+1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, ON (B)(6) 2023 AN ADDITIONAL PI WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED THE PATIENT HAD HIGH IOP AND WAS ON GLAUCOMA DRUGS. PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4) | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337478&pc=QCB | 2023826-2023-05678 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, 10.00/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). AT ONE MONTH POST-OP, THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP), WHICH WAS BEING TREATED WITH IOP LOWERING MEDICATIONS. THE PATIENT HAD NO SYMPTOMS OF ANGLE CLOSURE AND NO SIGNS OF INFLAMMATION. CURRENT POST-OP FINDS THE PRESSURE HAS STABILIZED ON MEDICATION. THE PATIENT IS BEING WATCHED CLOSELY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - IOP LOWERING MEDICATIONS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537395&pc=QCB | 2023826-2020-02066 | 2020-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS BEING TOO LARGE AND WILL BE REMOVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: ADDITIONAL SURGERY WAS PERFORMED - ENLARGEMENT OF PI BY YAG. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP PATIENT EXPERIENCED HALO, GLARE AND BLURRED VISION. SEE MFR#: 2023826-2020-02282. THE CAUSE WAS DUE TO USER ERROR. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; YAG - ENLARGEMENT OF PI. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE BLA NUMBER: NA. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135926&pc=QCB | 2023826-2021-00050 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H10 - CODE 4581: ANGLE CLOSURE, SHOULD HAVE BEEN SPECIFIED IN PREVIOUS SUPPLEMENTAL. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED WITH RESIDUE ON THE LENS. CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616797&pc=QCB | 2023826-2021-00840 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WILL BE EXPLANTED AND EXCHANGED DUE TO OVER VAULT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -06.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PIS WERE ENLARGED, WITH AN ADDITION OF PI BY YAG ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGEMENT/ADDITION OF PIS. CORRECTED DATA: B3: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060666&pc=QCB | 2023826-2021-01987 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, EYE IS QUIET. AC WAS FINE AFTER LENS REMOVAL. PATIENT NOT WILLING TO EXCHANGE AT PRESENT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CLAIM# (B)(4) Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITH ANGLE CLOSURE. CLAIM# (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT CLINAL CODE 1937 - ELEVATED IOP. H6: HEALTH EFFECT CLINAL CODE 4581 - ANGLE CLOSURE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161445&pc=QCB | 2023826-2022-01085 | 2022-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/2.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14647638&pc=QCB | 2023826-2022-01781 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.50/3.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO SIZING ISSUES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.5/+3.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. FOR CAUSE DETAILS THE REPORTER STATED "AT THE DAY 1 POSTOPERATIVE VISIT, THE UNCORRECTED VISUAL ACUITY MEASURED 20/32 OD. INTRAOCULAR PRESSURE WAS 18 MMHG OD. NARROW ANGLES OU." A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, INCREASED IOP >= 30 MMHG OR > 10 MMHG ABOVE PREOP IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER EMAIL ON 21 JUN 2023, PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REPORTEDLY, INCREASED IOP >= 30 MMHG OR > 10 MMHG ABOVE PREOP IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY AS REMOVAL. PROBLEM RESOLVED." SHOULD BE OMITTED FOR CORRECTION. H6: 1937 SHOULD BE OMITTED FOR CORRECTION. CLAIM (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308771&pc=QCB | 2023826-2021-00318 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (32MMHG) OBSERVED ON (B)(6) 2020, REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -15.00/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION; REFRACTIVE SURPRISE; PUPIL BLOCK; ELEVATED IOP; AND ANGLE CLOSURE WERE REPORTED. THE LENS WAS EXPLANTED AND LATER REPLACED FOR A LONGER LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. THE REPORTER ALSO NOTES "STABLE ICL VAULT, IOP, VISION, TEMPORAL SUTURE IN SITU." CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11787674&pc=QCB | 2023826-2021-01388 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ELEVATED IOP. LENS REMAINS IMPLANTED. FOR STATUS OF THE EYE " PATIENT CANNOT READ ANYMORE " WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER REPORTED " HIGH IOP OS: PROBLEM RESOLVED WITH DIAMOX AND PRESSURE WAS AFTER USING DIAMOX 20 MM HG. PROBABLY SOME VISCO REMAINED IN THE EYE. BUT THE PUPIL IN THIS EYE IS STILL ASYMMETRIC? THEY DONT KNOW WHAT IS THE CAUSE OF THIS REFRACTIVE SURPRISE. DOCTOR WANTS TO DO A LASER TO SOLVE THE REFRACTIVE SURPRISE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14507171&pc=QCB | 2023826-2022-01692 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ONE HAPTIC WAS NOTED TO BE TORN AFTER IMPLANTATION. PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, SUBLUXATION AND LESS VISION. MANAGEMENT WAS GIVEN FOR ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. FOR ADDITIONAL INFORMATION THE REPORTER STATED, "CORNEA CLEAR, VAULT- TEMPORARILY SUBLAXATED DUE TO LESS SUPPORT (HAPTIC TEAR)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED. EVERYTHING WENT WELL, PATIENT AND THE SURGEON ARE HAPPY WITH RESULTS. VA 6/9, VAULT OK, IOP 13MMHG. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962122&pc=QCB | 2023826-2023-04570 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/6.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING KERATOCONUS ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE IS "STABLE." CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE LENS NOT PERFORMING AS INTENDED. REPORTEDLY, "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/6.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND ELEVATED IOP. ON (B)(6) 2023, THE LENS WAS EXPLANTED BUT THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "STABLE". REPORTEDLY "THE ICL WAS REMOVED BECAUSE IT WAS TO BIG FOR THE EYE OF THE PATIENT SO THE EYE PRESSURE INCREASED". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE". H6: HEALTH EFFECT- CLINICAL CODE: "1937 SHOULD BE ADDED. CLAIM#: (B)(6). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL/PATIENT RELATED FACTOR: PRE-EXISTING KERATOCONUS. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875424&pc=QCB | 2023826-2020-02715 | 2020-07-18 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A12.6MM, VTICMO12.6, -16.0/+1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO "TILTED ICL", ELEVATED IOP, PUPIL BLOCK, LENS DISLOCATION/SUBLUXATION AND BLURRED VISION. ADDITION OF NEW PI WAS PERFORMED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2020. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11070811&pc=QCB | 2023826-2020-03177 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND IOP (INTRAOCULAR PRESSURE), UNRESPONSIVE TO MAXIMUM MEDICAL THERAPY. DOCTOR EXPLANTED THE LENS AS HE COULD NOT REDUCE THE IOP ADEQUATELY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -15.50/1,0/118, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA AND ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H3- DEVICE EVALUATION : LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN A SPECIMEN CUP. VISULA INSPECTION FOUND ONE OF THE HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308736&pc=QCB | 2023826-2021-00317 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BENT AND DEFORMED WITH RESIDUE AND DEBRIS ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG), LENS ROTATION AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT HAS "STABLE VAULT, STABLE IOP, CLEAR CORNEA, AND TEMPORAL SUTURE IN SITU." PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# : (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366547&pc=QCB | 2023826-2021-00485 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.5/+1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PERIPHERAL ANTERIOR SYNECHIA (PAS), AND A "LITTLE" MYDRIASIS. AN ADDITION OF PI-YAG WAS PERFORMED, PILOCARPINE EYE DROPS ADMINISTERED, IRRIGATION AND ASPIRATION (IA) PERFORMED. THE PAS WAS RELEASED BY PEELING OFF THE ADHESIONS OF THE ANGLES BY A ROUND NEEDLE. THE PROBLEM IS NOTED AS RESOLVED THOUGH THE PUPIL IS REPORTED AS SLIGHTLY DILATED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939022&pc=QCB | 2023826-2021-01669 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.00/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(32MMHG, ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: ELEVATED IOP AND BLURRINESS WAS NOTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939023&pc=QCB | 2023826-2021-01670 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.50/2.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP (25MMHG, ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: BLURRY VISION AND ANGLE CLOSURE OBSERVED. H10: CODE 4581 - ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222320&pc=QCB | 2023826-2021-02802 | 2021-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SECTORAL IRIS ATROPY. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THE PROBLEM IS NOT RESOLVED THOUGH THE LENS HAS BEEN EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS . CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/3.50/95 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH LENS. "THE SPECIALIST REPORTS THAT THE AGREEMENT IS FINE, WITH A LITTLE BLURRED VISION IN THE IMMEDIATE POST-OPERATIVE CONTROL, IT HAS OPEN ANGLES OF 2 PER GONIOSCOPY, VAULT OF 700 MICRONS, STABLE IOP OF 10/10 MMHG. THE PATIENT FINDS HIMSELF BETTER WITH THE EXCHANGE, SURGERY WITHOUT COMPLICATIONS." H6- HEALTH EFFCET CLINICAL CODE; 4581-OPEN ANGLES OF 2 PER GONIOSCOPY. CLAIM# (B)(4) Manufacturer Narrative: WEIGHT, ETHNICITY, RACE -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874494&pc=QCB | 2023826-2021-04515 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.5/2.5/62, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR STATUS OF THE EYE "BLURRY VISION" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B3: 05-OCT-2021. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061327&pc=QCB | 2023826-2022-01130 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -13.0/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO BLURRED VISION, IRIS SPHINCTER LOSS AT 5 O'CLOCK, HEAVILY PIGMENTS TM, KIV EXPLANT, AND IRREGULAR PUPIL. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- ELEVATED IOP SHOULD HAVE BEEN REPORTED AS AN ADVERSE EVENT IN INITIAL MDR. H6- INTRAOCULAR PRESSURE INCREASED . CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15221446&pc=QCB | 2023826-2022-02760 | 2020-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15556104&pc=QCB | 2023826-2022-03511 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.00/6.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HIGH INTRAOCULAR PRESSURE WITH ANGLE CLOSURE AND CHANGE IN REFRACTION. CLOUDY CORNEA WAS ALSO REPORTED, HOWEVER THE REPORTER STATED, " REGARDING THE SYMPTOMS RELATED TO THE EXPLANTED LENS, EVERYTHING RESOLVED." IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600761&pc=QCB | 2023826-2023-03684 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -3.00/+3.00/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17608279&pc=QCB | 2023826-2023-03689 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.00/+5.50/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701985&pc=QCB | 2023826-2023-03775 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.00/0.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED WITH BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. PATIENT WAS LOST TO FOLLOW-UP. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17705120&pc=QCB | 2023826-2023-03951 | 2021-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE - 4581 - PUPIL ANOMALY. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885010&pc=QCB | 2023826-2023-04315 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS HAD A LOW VAULT. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION AND BLURRED VISION. OCT PERFORMED ON (B)(6) 2023 REVEALED RNFL DAMAGE AND TEMPORAL SCOTOMA. MEDICATION WAS PRESCRIBED. THE LENS REMAINED IMPLANTED. POST-OP, GOOD IOP DURING 6 HOURS AFTER DROPS. PROSTAGLANDIN INHIBITOR WAS ADDED AT NIGHT. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, RETINAL NERVE DAMAGE, TEMPORAL SCOTOMA H6: HEALTH EFFECT IMPACT CODE: 4644 - HYPOTENSIVE DROPS - S. COSOPT (DORZOLAMIDE/TIMOLOL) X2, S. LUXFEN (BIMATOPROST) X2 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+0.5/69 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS, FOGGY VISION, BLURRED VISION, LASER YAG PIS PERFORMED AND MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6-HEALTH EFFECT- CLINICAL CODE: 4581- "FOGGY VISION" AND "OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS" SHOULD BE ADDED. H6-HEALTH IMPACT CODE: 4625- "'LASER YAG PI' WAS PERFORMED" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525781&pc=QCB | 2023826-2024-00170 | 2023-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6, -11.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ANTERIOR CHAMBER IRRIGATION/VACUATION OF VISCO/FLUIDS AND ANTERIOR CHAMBER WASH ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED IOP(40MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(DATE ASSESSED (B)(6) 2023 & (B)(6) 2023), PIGMENT DISPERSION, AND BLURRED VISION. THIS DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618161&pc=QCB | 2023826-2019-00798 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 15.5/+2.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ICL WAS TOO BIG. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282596&pc=QCB | 2023826-2019-02091 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.5/1.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND DEBRIS ON THE LENS BUT NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558794&pc=QCB | 2023826-2021-00700 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200. PUPIL IS DILATED AND AR IS +2.75. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, HYPEROPIC SURPRISE +2.00, GLARE/HALOS, BLURRED VISION. REPORTEDLY VISION IS NOW "OK" AT +1.00 HYPEROPIC. PILOCARPINE USED TO CONSTRICT PUPIL AND SET THE ICL IN THE SULCUS. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE/FIXED PUPIL). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664235&pc=QCB | 2023826-2022-00658 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. ADDITIONAL SYMPTOMS REPORTED: ELEVATED IOP, GLARE/HALOS, PIGMENT DISPERSION, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. TRACE PIGMENT IS PRESENT ON THE ENDOTHELIUM AT 6 O'CLOCK AFTER EXCHANGE. HOWEVER, THE EVENT IS UNLIKELY TO BE CAUSED BY THE DEVICE AND NO INTERVENTION IS PLANNED. AN ENLARGEMENT OF PI-YAG WAS PERFORMED AFTER LENS EXPLANT AND PRIOR TO REPLACEMENT LENS IMPLANT. H6-CLINICAL CODE 4581 (PIGMENT DISPERSION). CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. IMPLANT DATE, EXPLANT DATE: UNK. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=QCB | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4) | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112743&pc=QCB | 2023826-2019-01626 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION, GLARE/HALOES, AND BLURRED VISION. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES. ADDITIONAL INFORMATION: PRE-EXISTING CORNEAL OR REFRACTIVE SURGERY. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140318&pc=QCB | 2023826-2021-02397 | 2021-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. GOOD POSITION AND VAULT, SYMPTOMS (GLARE) RESOLVED AND IOP NORMAL. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416136&pc=QCB | 2023826-2021-00506 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.50/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP(44MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, GLARE/HALOES, BLURRED VISION, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938959&pc=QCB | 2023825-2021-01652 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.50/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP(50MMHG), ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, IRIS ATOROPHY, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, "UNFORTUNATELY DURING THE ICL EXCHANGE BECAUSE OF THE LIMITED DISTANCE BETWEEN THE ICL AND THE ENDOTHELIUM SHE SUSTAINED A 1MM TEAR IN DESCEMET'S MEMBRANE. AS OF NOW WE ARE WAITING TO SEE IF HER CORNEA WILL CONTINUE TO CLEAR. WHEN THE CORNEA CLEARS ENOUGH WE INTEND TO ASSESS THE STATUS OF THE ENDOTHELIUM TO SEE HOW MUCH DAMAGE SHE HAS SUSTAINED BY THE THREE PROCEDURES ON THIS EYE. CURRENTLY HER ANGLES ARE ADEQUATE IN BOTH EYES AND THE IOP IS ACCEPTABLE." Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US, (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14809062&pc=QCB | 2023826-2022-01953 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/150/66 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PATIENT HAS EXPERIENCED EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOES AND BLURRED VISION. FOR PATIENTS CURRENT CONDITION THE REPORTER STATED " IRIS PERIPHERAL ANTERIOR SYNACHIAE WHICH REQUIRES REMOVAL OF CURRENT LENS SIZING OF 13.2MM TO 12.6MM SIZING. THE INITIAL LENS IS STILL IN PATIENT'S EYES." FOR ADDITIONAL INFORMATION THE REPORTER STATED " PATIENT HAS IRIS ATROPHY DUE TO HIGH LENS VAULTING BUT DR HAS BEEN MONITORING THIS CASE FOR 2 YEARS. NEEDS A SECONDARY INTERVENTION AS WELL DUE TO IRIS DISTORTION AND CAUSING PATIENT TO HAVE VISUAL DISTURBANCE ESPECIALLY AT NIGHTS." IN THE REPORTS OPINION THE CAUSE OF THIS EVENT IS USER ERROR. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON 28-JUN-2022 AND REPLACED WITH A SHORTER LENGTH LENS ON THE SAME DATE BUT DIFFERENT SURGERY AND THE PROBLEM RESOLVED. THE REPORTER STATED "ANTERIOR SYNECHIAE AND IRIS ATROPHY STILL PERSIST BUT VAULTING REDUCED AND EYE LOOKED BETTER, LESS GLARE." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#:(B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17060362&pc=QCB | 2023826-2023-02031 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO TRANSIENT LOSS OF BCVA (ASSESSED ON (B)(6)2023, ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2023). PIGMENT DISPERSION. GLARE, HALOES, AND BLURRED VISION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PARACENTESIS DONE AND DIAMOX GIVEN TO REDUCE IOP. PATIENT DOES NOT WANT NEW LENS IMPLANTED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6: HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6:WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT - IMPACT CODE: "4625- PARACENTESIS DONE." B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-13.5/+1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, DILATED PUPIL, MEDICATION WAS PRESCRIBED. REPORTEDLY "PARACENTESIS DONE AND GAVE DIAMOX TO REDUCE IOP." ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ALSO REPORTED, "IOP REDUCED AND PUPIL STILL DILATE AFTER REMOVAL. PATIENT NOT KEEN TO IMPLANT ANOTHER ICL ANYMORE." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393935&pc=QCB | 2023826-2015-00168 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/+3.5/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS HAS AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF THE IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOS, BLURRED VISION AND IRIS ATROPHY. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. EXPLANTED DATE UNKNOWN. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED AND OBSERVED NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). RESUBMISSION OF INITIAL MDR REQUESTED BY FDA. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528011&pc=QCB | 2023826-2019-02495 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 6.50/+2.0/068 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH? NEEDS SMALLER LENS?) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CONTACT CASE, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS H6 - PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECTED TO (B)(6) 2019 IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, THE "PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT YET FULLY RESOLVED." THE PATIENT STATES "VISION HAS IMPROVED SINCE PROCEDURE OVERALL." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528024&pc=QCB | 2023826-2019-02494 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 7.00/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH NEEDS SMALLER LENS) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT CASE, WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS. PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: IT WAS LATER INDICATED BY THE REPORTER THAT AS OF 25/FEB/2020 "THE PATIENT'S PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT FULLY HEALED. THE PATIENT'S VISION HAS IMPROVED SINCE THE PROCEDURE OVERALL THOUGH" CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140634&pc=QCB | 2023826-2020-01200 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, CONTINUOUS GRAVITATIONAL PAIN-FEELING SOMETHING PUSHING ON THE EYE, HALOS, CIRCLES, AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14957809&pc=QCB | 2023826-2022-02108 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/+3.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14958057&pc=QCB | 2023826-2022-02109 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.0/+4.0/65 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908414&pc=QCB | 2023826-2019-01486 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/106 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT REPORTED DISCOMFORT. THE LENS WILL BE EXPLANTED BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913315&pc=QCB | 2023826-2019-01487 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT REPORTED DISCOMFORT. THE LENS WILL BE EXPLANTED BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454729&pc=QCB | 2023826-2023-03241 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO HIGH VAULT CAUSING PAIN AND HIGH PRESSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED ON THE SAME DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271533&pc=QCB | 2023826-2023-05502 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE) ANS PAIN. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278375&pc=QCB | 2023826-2021-03031 | NA | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS OTHER - ICL LENS LENGTH CALCULATION ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS EXCHANGED ON (B)(6) 2021 WITH A DIFFERENT MODEL, SIMILAR POWER, SHORTER LENGTH LENS WHICH RESOLVED THE ISSUE. REPORTEDLY, DOCTOR AND PATIENT HAPPY WITH THE FINAL RESULT. THE REPORTER MARKED THE CAUSE OF THE EVENT AS PATIENT-RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 721078. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15469467&pc=QCB | 2023826-2022-03140 | 2022-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Iritis | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF -16.50/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO IRITIS/IOP ELEVATION. LENS WAS NOT REPLACED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15474202&pc=QCB | 2023826-2022-03244 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Iritis | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF --17.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO IRITIS/HIGH IOP. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SEPIMENT CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16479533&pc=QCB | 2023826-2023-00722 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Intraocular Pressure Increased; Iritis | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION AND IRITIS WERE OBSERVED. CAUSE OF THE EVENT IS BOTH A UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITITATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS.". Manufacturer Narrative: B5 - "REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT.'" SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2023 SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION, PRESSURE SPIKES, AND IRITIS ARE OBSERVED. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT PROBLEM: PATIENT RELATED FACTOR. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16585378&pc=QCB | 2023826-2023-00994 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Intraocular Pressure Increased; Iritis | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION AND IRITIS WERE OBSERVED. CAUSE OF THE EVENT IS BOTH A UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY, THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. ADVERSE EVENT PROBLEM: PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B: EXPLANTATION DATE: 28-MAR-2023 ADDED FOR CORRECTION. H1: SERIOUS INJURY SHOULD BE ADDED AND MALFUNCTION OMITTED FOR CORRECTION. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE OF A PATIENT WITH PRE-EXISTING GLAUCOMA ON (B)(6) 2022. PIGMENT DISPERSION, PRESSURE SPIKES, AND IRITIS ARE OBSERVED. REPORTEDLY THE STATUS OF THE EYE IS, "MOSTLY QUIET" AND "TAPERING PREDNISOLONE INITIATED LATANOPROST IN MY OPINION BEST IRRESPECTIVE OF PIGMENT DISPERSION, IRITIS." REPORTEDLY, 'BOTH ICLS WERE REMOVED TUESDAY (B)(6) 2023 FOR PERSISTENT IOP ISSUES FROM PIGMENT DISPERSION. MAXIMAL MEDICAL THERAPY HAD BEEN TRIED, AS WELL AS SLT LASER FOR THE LEFT. THE ONLY THING WE DID AFTER SURGERY WAS PERFORM A GONIOTOMY ON THE LEFT EYE AT THE TIME OF THE REMOVAL.' CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969886&pc=QCB | 2023826-2020-02884 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Iritis; Pain; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/018 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT DEVELOPED AQUEOUS MISDIRECTION AND EXPERIENCED PAIN. THE SURGEON ENLARGED THE PI BY SURGICAL PROCEDURE. ANOTHER ICL LENS WAS NOT IMPLANTED. THE PATIENT IS IMPROVING, WITH SOME RESIDUAL IRITIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405559&pc=QCB | 2023826-2021-00562 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT AFTER LENS REPOSITIONING THE PATIENT'S IOP INCREASED, ALTHOUGH THE PROBLEM WAS STILL RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241308&pc=QCB | 2023826-2021-02895 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK, SEX: UNK, WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK IMPLANT DATE: UNK, EXPLANT DATE: UNK, DEVICE MFR DATE: UNK. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241986&pc=QCB | 2023826-2021-02896 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK, DEVICE MFR DATE: UNK. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161442&pc=QCB | 2023826-2022-01084 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/2.0/046 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15227548&pc=QCB | 2023826-2022-02765 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO A PATIENT'S EYE IN (B)(6) 2022. THE SURGEON NOW REPORTS A PLANNED LENS EXCHANGE FOR A SHORTER LENGTH LENS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -3.50/+1.00/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15674217&pc=QCB | 2023826-2022-03632 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS OF -18/3.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY. CAUSE OF THE EVENT IS THE DEVICE. THE REPORTER STATED THAT THE LENS WAS, "OVERSIZED." THE PROBLEM IS NOT RESOLVED. REPORTEDLY, "EXPLANTED ICL, AWAITING PUPIL DILATION TO RESOLVE BEFORE RE-IMPLANTING." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "VTICM 12.6" SHOULD BE CORRECTED TO "VTICMO 12.6" CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4582- ANGLE CLOSURE; IRIS ATROPHY; UNREACTIVE (FIXED) PUPIL. ADDITIONAL DATA: B5: IT WAS LATER REPORTED "THE LENS WAS NOT REPLACED." CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15763937&pc=QCB | 2023826-2022-04065 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409310&pc=QCB | 2023826-2019-00378 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -16.0/+3.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND ANGLE CLOSURE. THE STATUS OF THE EYE IS "GOOD" AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(6). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475395&pc=QCB | 2023826-2019-00520 | 2019-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, PUPIL BLOCK, AND ANGLE CLOSURE. IN ADDITION, THE SURGEON PERFORMED ENLARGEMENT OF THE PI. NO ADDITIONAL INFORMATION CURRENTLY AVAILABLE REGARDING PATIENT'S CONDITION AFTER LENS EXCHANGE. Manufacturer Narrative: PREVIOUSLY STATED A LENS SPHERE/CYLINDER/AXIS VALUE OF -8.5/+2.5/103. ACTUAL VALUE IS -8.5/+2.5/108. THE LENS EXCHANGE RESOLVED THE PROBLEM. "EXCELLENT OUTCOME." CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PUPIL BLOCK, ANGLE CLOSURE). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489348&pc=QCB | 2023826-2019-00555 | 2018-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/1.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501145&pc=QCB | 2023826-2019-00505 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/2.0/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (35MMHG) WITHOUT PUPIL BLOCK AND ANGEL CLOSURE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8501252&pc=QCB | 2023826-2019-00506 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/1.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (26MMHG) WITHOUT PUPIL BLOCK AND ANGEL CLOSURE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511516&pc=QCB | 2023826-2019-00597 | 2019-03-14 | Injury | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, - 6.0/+4.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), ANGLE CLOSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. PRIOR TO EXPLANT, GAVE THE PATIENT DIAMOX, TOPICAL OCULAR ANTIHYPERTENSIVE MEDICATIONS, THEN IV MANNITOL ATTEMPTING TO LOWER IOP, THEN ATTEMPTED TAPPED ANTERIOR CHAMBER AT SLIT LAMP ATTEMPTING TO LOWER IOP. THESE MEASURES WERE UNSUCCESSFUL AND IOP CONTINUED TO RISE. GONIOSCOPY DEMONSTRATED ANGLE CLOSURE, AND THE DECISION WAS MADE TO EXPLANT. THE SURGEON PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED LENS. THE PATIENT CONDITION AT ONE DAY POST-OP, MILD CORNEAL EDEMA AND ANGLE OPEN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519014&pc=QCB | 2023826-2019-00610 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. IN ADDITION, THE SURGEON PERFORMED ENLARGEMENT OF THE PI. THE PROBLEM WAS RESOLVED. THE SURGEON NOTES "EXCHANGE WITHOUT COMPLICATION" AND "EXCELLENT 20/20 OD." Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND RESIDUE ON THE LENS SURFACE AND THE OPTIC WAS SCRATCHED. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8620926&pc=QCB | 2023826-2019-00871 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -7.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. THE SURGEON AND PATIENT ARE CONSIDERING IMPLANTING A NEW LENS AFTER THREE OR FOUR MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621026&pc=QCB | 2023826-2019-00872 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -7.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 27 (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. THE SURGEON AND PATIENT ARE CONSIDERING IMPLANTING A NEW LENS AFTER THREE OR FOUR MONTHS. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 3189 SHOULD BE CORRECTED TO 2993 IN INITIAL MEDWATCH REPORT. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED DRY WITH SURGICAL RESIDUE IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633170&pc=QCB | 2023826-2019-00886 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.0/+1.0/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION, ELEVATED IOP, AND LENS DISLOCATION/SUBLUXATION INFERIORLY. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. THE REPORTER STATES, "LARGER SULCUS HORIZONTAL THAN ESTIMATED BY WTW CAUSING ICL TO DROP INFERIORLY (HOLE AT INFERIOR PUPIL EDGE)." REFERENCE CLAIM# (B)(4) FOR REPLACEMENT LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAGGIE, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8634346&pc=QCB | 2023826-2019-00804 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/177 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS RIGHT EYE (OD) DURING THE SAME SURGERY. THE LENS WAS REMOVED DUE TO EXCESSIVE VAULTING, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACITVE (FIXED) PUPIL AND IRIS ATROPHY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO LENS SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652099&pc=QCB | 2023826-2019-00933 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL DRY WITH SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731290&pc=QCB | 2023826-2019-01084 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -9.50/4.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULT THE NEXT DAY FOLLOWING SURGERY, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ELEVATED IOP (31 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). DATE OF EVENT IS (B)(6) 2019 CONCOMITANT MEDICAL PRODUCT: INJECTOR LOT# 1431485. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.5/4.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULT WITH ROTATION THE NEXT DAY FOLLOWING SURGERY; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 AND REPLACED FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. DATE RECEIVED BY MANUFACTURER: DATE RECEIVED IN ORIGINAL MDR IS NOT APPLICABLE; ADDITIONAL INFORMATION WAS RECEIVED ON 12-JULY-2019. PATIENT CODE:3191 SECONDARY SURGERY (EXPLANT); SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.5/4.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON OBSERVED EXCESSIVE VAULTING THE NEXT AT FOLLOWING SURGERY; LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8731314&pc=QCB | 2023826-2019-01101 | NA | Injury | STAAR SURGICAL COMPANY | 2019-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/4.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULT THE NEXT DAY FOLLOWING SURGERY, LOSS OF BCVA AND ELEVATED IOP (31 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -9.0/4.5/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULTING THE NEXT AT FOLLOWING SURGERY; LOSS OF BCVA; AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -9.0/4.5/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS). THE SURGEON OBSERVED EXCESSIVE VAULT WITH ROTATION THE NEXT DAY FOLLOWING SURGERY; AND ELEVATED IOP (31MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND LATER REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. DEVICE CODE 1401 IN ORIGINAL MDR IS NOT APPLICABLE. DEVICE CODE 2923 ADDED FOR ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736583&pc=QCB | 2023826-2019-01126 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 WITHIN THE SAME SURGERY AS THE REMOVAL OF THE ORIGINAL LENS. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K #: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809460&pc=QCB | 2023826-2019-01259 | 2019-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS CONDITION AT LAST VISIT WAS THE PUPIL WAS BIGGER THAN NORMAL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889701&pc=QCB | 2023826-2019-01418 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.00/4.0/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (40 MMHG), AND EXCESSIVE VAULT WERE OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SMALLER LENGTH LENS. REPORTER NOTED "VAULT IS STILL GENEROUS BUT SAFE." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS RECOMMENDED LENS TOO LARGE. PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP, STILL GENEROUS VAULT BUT SAFE EXAM & IOP, ANGLES OPEN." Manufacturer Narrative: SEX: FEMALE SHOULD BE CORRECTED TO MALE. RETURNED TO MANUFACTURER: 08-AUG-2019 SHOULD BE CORRECTED TO 30-JUL-2019. (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889704&pc=QCB | 2023826-2019-01419 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (50 MMHG), AND EXCESSIVE VAULT WERE OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SMALLER LENGTH LENS. REPORTER NOTED "VAULT IS STILL GENEROUS BUT SAFE." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "RECOMMENDED LENS TOO LARGE." PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP, STILL GENEROUS VAULT BUT SAFE EXAM & IOP, ANGLES OPEN." Manufacturer Narrative: EVENT DESCRIPTION: "PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, BCVA 20/50, "EYE IS SAFE, UNINFLAMED, GOOD IOP" SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. INFORMATION IS NOT APPLICABLE. ON (B)(6) 2019 THE LENS WAS REMOVED SHOULD BE CORRECTED TO (B)(6) 2019 THE LENS WAS REMOVED, IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894304&pc=QCB | 2023826-2019-01441 | 2019-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918840&pc=QCB | 2023826-2019-01497 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF HAPTIC TORN OFF. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8998364&pc=QCB | 2023826-2019-01665 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT LENS A 12.6MM, VTICM5_12.6, -13.00/+2.00/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE). IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, PATIENT CONDITION IS WELL. NO PLAN TO IMPLANT OTHER LENS. Manufacturer Narrative: REFERENCE MFR# 2023826-2019-01615 FOR INITIAL LENS. DEVICE EVALUATION: DIMENSIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. INVESTIGATION FOUND THAT IT IS LIKELY THAT THE CUSTOMERS SWITCHED THE INITIAL AND EXCHANGE LENSES WHEN PACKAGING THEM FOR RETURN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. DEVICE CODE 1494: OFF-LABEL USE- AGE < 21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: H6 - METHOD CODE 3331: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9070787&pc=QCB | 2023826-2019-01707 | 2019-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/1.0/091 SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: EVENT DESCRIPTION: INITIAL MDR SHOULD HAVE STATED: ELEVATED IOP WAS ALSO REPORTED. IT WAS REPORTED THAT AT NINE DAYS POST-OP, 'HIS IOP AFTER MEDICATIONS IS 27MMHG.' (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120432&pc=QCB | 2023826-2019-01735 | 2019-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.5/+2.5/127 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH EXPLANT, NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120434&pc=QCB | 2023826-2019-01736 | 2019-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -13.5/+1.5/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH EXPLANT, NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9194673&pc=QCB | 2023826-2019-01927 | 2019-07-30 | Injury | STAAR SURGICAL COMPANY | 2019-10-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). ANOTHER ICL LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9202984&pc=QCB | 2023826-2019-01911 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/5.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216530&pc=QCB | 2023826-2019-01987 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO 13.2, -8.5/1.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE, AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DESCRIPTION OF EVENT OR PROBLEM- "-8.5/1.5/069" SHOULD BE CORRECTED TO "-8.5/1.5/068" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279655&pc=QCB | 2023826-2019-02069 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292244&pc=QCB | 2023826-2019-02051 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS AND EXCESSIVE VAULT WAS OBSERVED. ELEVATED IOP (38 MMHG) WAS ASSESSED ON (B)(6) 2019. LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDDOTHIAL CELL LOSS AND EXCESSIVE VAULT WAS OBSERVED." CORRECTED TO: "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT WERE OBSERVED." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299931&pc=QCB | 2023826-2019-02078 | 2019-09-07 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/051 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS DISCARDED BY THE HOSPITAL AND WILL NOT BE RETURNED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306729&pc=QCB | 2023826-2019-02050 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -14.0/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306739&pc=QCB | 2023826-2019-02049 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -14.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312711&pc=QCB | 2023826-2019-02079 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312732&pc=QCB | 2023826-2019-02080 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317888&pc=QCB | 2023826-2019-02129 | 2019-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336755&pc=QCB | 2023826-2019-02173 | 2019-10-18 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -10.0/+3.0/175 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IOP VALUE WENT FROM 45MMHG TO 15MMHG AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A CONTACT LENS CONTAINER. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336761&pc=QCB | 2023826-2019-02172 | 2019-10-18 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -10.0/+3.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. IOP VALUE WENT FROM 59MMHG TO 12MMHG AFTER EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A CONTACT LENS CONTAINER. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9335805&pc=QCB | 2023826-2019-02149 | 2019-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2, -11.00/1.0/085 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (23 MMHG) .THIS RESOLVED THE PROBLEM. AN ENLARGEMTNE OF PI AND YAG WAS PERFORME. PER THE REPORTER, CAUSE OF THE EVENT IS REPORTED AS "STEEP VAULT.". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE IN A LENS VIAL INSIDE A ZIPLOC BAG. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338940&pc=QCB | 2023826-2019-02120 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/087 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6)2019. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS IN THE LEFT EYE. THE PATIENT IS DOING WELL. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "A REPLACEMENT LENS OF A SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2020. THE REPLACEMENT LENS RESOLVED THE ISSUE. PATIENT IS NOW FINE." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338965&pc=QCB | 2023826-2019-02119 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/102 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, INCISION ENLARGEMENT. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND NARROW ANGLES. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE VAULT IS VERY GOOD AND THE PATIENT CONDITION IS FINE. CLAIM # (B)(4). Manufacturer Narrative: (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN THE INITIAL MDR. G4-(B)(6) -2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. WEIGHT UNK. AGE: UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498077&pc=QCB | 2023826-2019-02407 | NA | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.5/+6.0/084 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, IRIS ATROPHY, UNREACTIVE (FIXED) PUPIL AND ANGLE CLOSURE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. POST-OP THE PATIENT STILL HAS IRIS ATROPHY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512598&pc=QCB | 2023826-2019-02423 | 2019-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -15.5/+2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. IMMEDIATE POSTOP RESULTS WERE GOOD FOLLOWING THE EXCHANGE. PRIOR TO EXCHANGE THE SURGEON HAD PERFORMED A RELEASE OF FLUID VIA PARACENTESIS AND HAD THE PATIENT ON "MAXIMAL MEDICAL THERAPY." Manufacturer Narrative: CORRECTED DATA: A2: PATIENT'S AGE: ON (B)(6) 1983. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512611&pc=QCB | 2023826-2019-02422 | 2019-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -14.5/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. IMMEDIATE POSTOP RESULTS WERE GOOD FOLLOWING THE EXCHANGE. PRIOR TO EXCHANGE THE SURGEON HAD PERFORMED A RELEASE OF FLUID VIA PARACENTESIS AND HAD THE PATIENT ON "MAXIMAL MEDICAL THERAPY." Manufacturer Narrative: PATIENT'S AGE: (B)(6) 1983. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# 430811 | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516563&pc=QCB | 2023826-2019-02448 | 2019-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. ELEVATED IOP (50MMHG) WAS ASSESSED ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565909&pc=QCB | 2023826-2020-00062 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B3- DATE OF EVENT: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. B5- THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM. D6- IMPLANT DATE: (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/-1.50/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP AND UNREACTIVE (FIXED) PUPIL. THE EXCHANGE RESOLVED THE PROBLEM. H6- PATIENT CODE 3191- UNREACTIVE (FIXED) PUPIL. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639933&pc=QCB | 2023826-2020-00201 | 2019-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.0/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET. THE SURGEON AND PATIENT ARE HAPPY. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED ELEVATED IOP. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9656163&pc=QCB | 2023826-2020-00212 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT HAS NO PAIN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689251&pc=QCB | 2023826-2020-00250 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (20 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER STATES "PATIENT IS HAPPY". PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690171&pc=QCB | 2023826-2020-00278 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). AN ADDITION OF A PI-YAG WAS ALSO PERFORMED. ON THE SAME DATE IN A SEPARATE SURGERY, A SHORTER LENGTH LENS WAS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING CALCULATION ERROR." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED ON (B)(6) 2019 YAG WAS PERFORMED. POSTOP PATIENT RESULTS ARE "SO FAR EVERYTHING SEEMS FINE". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781511&pc=QCB | 2023826-2020-00296 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT OTHER LENS DUE TO CONSTANT ELEVATED IOP. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S CURRENT CONDITION IS GOOD, IOP WAS TREATED WITH MEDICATION AND THE SURGEON KEEP MONITORING THE PATIENT'S IOP. AFTER REMOVING THE LENS, ALL THE PROBLEMS WERE RESOLVED AND THE DOCTOR PRESCRIBED GLASSES." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781512&pc=QCB | 2023826-2020-00297 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.00/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT AND ELEVATED IOP (27 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "THERE ARE NO PLANS TO IMPLANT OTHER LENS DUE TO CONSTANT ELEVATED IOP. PER REPORTER ON (B)(6) 2020, "PATIENT IN GOOD CONDITION, IOP WAS TREATED WITH MEDICATION AND THE SURGEON KEEP MONITORING THE PATIENT'S IOP, AND AFTER REMOVING THE LENS, ALL THE PROBLEMS WERE RESOLVED AND THE DOCTOR DESCRIBED GLASSES." Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842477&pc=QCB | 2023826-2020-00547 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THIS RESOLVED THE PROBLEM, PATIENT CURRENT CONDITION IS GOOD. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842582&pc=QCB | 2023826-2020-00548 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/5.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. LENS EXPLANT RESOLVED THE PROBLEM, PATIENT CURRENT CONDITION IS GOOD. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1937 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842748&pc=QCB | 2023826-2020-00505 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -9.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. THE SURGEON REPORTS "SIZE SEEMS TO BE BIG, WITH PERSISTENT ELEVATED IOP REFRACTIVE TO MEDICAL TREATMENT." THE SURGEON PLANS TO INCREASE THE SIZE OF THE PI AND WILL DETERMINE IF AN EXCHANGE IS NEEDED. IT IS UNKNOWN AT THIS TIME AS TO WHETHER THE PI WAS INCREASED OR THE EXCHANGE OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, YAG- ENLARGEMENT OF PI). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842787&pc=QCB | 2023826-2020-00506 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -9.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. THE SURGEON REPORTS "SIZE SEEMS TO BE BIG, WITH PERSISTENT ELEVATED IOP REFRACTIVE TO MEDICAL TREATMENT." THE SURGEON PLANS TO INCREASE THE SIZE OF THE PI AND WILL DETERMINE IF AN EXCHANGE IS NEEDED. IT IS UNKNOWN AT THIS TIME AS TO WHETHER THE PI WAS INCREASED OR THE EXCHANGE OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, YAG- ENLARGEMENT OF PI). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9891498&pc=QCB | 2023826-2020-00597 | 2020-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (GREATER THAN 50MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (GREATER THAN 50MMHG) WERE ASSESSED ON (B)(6) 2020, AND EXCESSIVE VAULT WAS OBSERVED. ADDITION OF NEW PI AND YAG WERE PERFORMED. THE LENS WAS REMOVED ON (B)(6) 2020. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE (EXCESSIVE VAULT). REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: PATIENT CODE- 3191- SECODARY YAG NEEDS BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913332&pc=QCB | 2023826-2020-00737 | 2020-01-14 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THREE DAYS LATER, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED AGAIN. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON IS NOT PLANNING TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953427&pc=QCB | 2023826-2020-00872 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953577&pc=QCB | 2023826-2020-00874 | 2018-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/3.0/96 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- TYPE OF INVESTIGATION CODE 10 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953691&pc=QCB | 2023826-2020-00875 | 2018-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/3.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(60 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953821&pc=QCB | 2023826-2020-00882 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957585&pc=QCB | 2023826-2020-00847 | 2020-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.00/0.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ANGLE CLOSURE WITH ELEVATED IOP (20.5MMHG) BEING OBSERVED ON (B)(6) 2020 AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY - ALL IS FINE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957586&pc=QCB | 2023826-2020-00846 | 2020-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO ANGLE CLOSURE WITH ELEVATED IOP (28.5MMHG) BEING OBSERVED ON (B)(6) 2020 AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY - ALL IS FINE." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CORRECTED DATA: H6-PATIENT CODE: 3191-ANGLE CLOSURE SHOULD BE INCLUDED IN INITIAL MDR. PATIENT CODE 1937- ELEVATED IOP SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960150&pc=QCB | 2023826-2020-00856 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991666&pc=QCB | 2023826-2020-00950 | 2019-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025642&pc=QCB | 2023826-2020-01002 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/045 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AS WELL AS THE OPTIC AND HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- DEVICE EVALUATION: "WITH MOISTURE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032664&pc=QCB | 2023826-2020-01022 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -3.00/2.5/017 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ANGLE CLOSURE WITH ELEVATED IOP (20 MMHG) ON (B)(6) 2020 AND EXCESSIVE VAULT WERE OBSERVED. ADDITION OF NEW PI AND YAG WAS PERFORMED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - -3.00/2.5/107 SHOULD BE CORRECTED TO +3.00/2.5/107 FOR LENS (SPHERE/CYLINDER/AXIS). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. IT WAS ALSO REPORTED THE PATIENT WAS PRESCRIBED MEDICATION. PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044695&pc=QCB | 2023826-2020-01053 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "..THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT." CORRECTED TO "..THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP OF 40MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044899&pc=QCB | 2023826-2020-01054 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-14.0/1.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE 40 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10044986&pc=QCB | 2023826-2020-01055 | 2020-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.1,-12.0/1.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : B5-THE REPORTER INDICATED THAT A 13.7MM, VTICMO 13.7, -12.0/1.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM# : (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242445&pc=QCB | 2023826-2020-01423 | 2020-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392960&pc=QCB | 2023826-2020-01672 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND DEBRIS ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398745&pc=QCB | 2023826-2020-01707 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, CAUSING ELEVATED INTRAOCULAR PRESSURE (IOP). THE INCISION WAS ENLARGED TO REMOVE THE LENS AND A SUTURE WAS REQUIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. AN ADDITIONAL SURGERY WAS PERFORMED - ENLARGEMENT/ADDITION OF PI BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. H6: HEALTH EFFECT IMPACT CODE: 4625, ADDITIONAL SURGERY - ENLARGEMENT/ADDITION OF PI BY YAG. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416135&pc=QCB | 2023826-2020-01714 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020 AND WAS EXCHANGED FOR A SHORTER LENS. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED AC SHALLOWING. CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZING ISSUE". H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PARTIAL AC SHALLOWING, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10417803&pc=QCB | 2023826-2020-01752 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -4.5/1.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP, IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER LENS EXPLANT. THE REPORTER STATED " THE BASIC INTRAOCULAR PRESSURE OF THE PATIENT WAS HIGH, AND POST-OP WAS HIGH. THE IOP WAS STILL GREATER THAN 26MMHG AFTER MEDICATION. CONSIDERING THE HIGH RISK OF GLAUCOMA, THE LENS WAS REMOVED, AND THEN THE CORNEAL LASER OPERATION WAS PERFORMED. AFTER LASER SURGERY THE PATIENTS UCVA REACHED OU1.2 AND THE IOP IS NORMAL (17MMHG)". Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - HEALTH IMPACT: 4644 - PATIENT WAS PRESCRIBED MEDICATION, SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPLACT: 4625 - LASER REFRACTIVE SURGERY, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440601&pc=QCB | 2023826-2020-01816 | 2020-07-29 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A COUPLE OF WEEKS LATER, THE VAULT HAD COME DOWN, SO THE SURGEON DECIDED TO LEAVE THE LENS IN THE EYE AND MONITOR THE PATIENT. THE PATIENT WAS TAKING DROPS FOR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445835&pc=QCB | 2023826-2020-01815 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/010 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476011&pc=QCB | 2023826-2020-01938 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.00/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476114&pc=QCB | 2023826-2020-01934 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2020-01935 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO EXCESSIVE VAULTING. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. IT IS LIKELY THAT THESE LENSES WERE INADVERTENTLY SWAPPED DURING RETURN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL WIDTH INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. DIMENSIONAL OVERALL LENGTH INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416144&pc=QCB | 2023826-2020-01715 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/064 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WA EXPLANTED ON (B)(6) 2020 AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CORRECTED DATA: B5 - IT WAS ALSO REPORTED THE PATIENT EXPERIENCED PARTIAL AC SHALLOWING. CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZING ISSUE". H6 - HEALTH EFFECT CLINICAL CODE: 4581: PARTIAL AC SHALLOWING SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12957740&pc=QCB | 2023826-2021-04643 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.00/+4.0/117 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/ EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A SPECIMEN CUP WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND STRETCHED OUT WITH RESIDUE THROUGHOUT THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965310&pc=QCB | 2023826-2021-04644 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/065 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/065 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED IOP. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/ FLUIDS AND THE WOUND WAS BURPED, CONFIRMING THERE WAS NO VISCO REMAINING. THE PIS WERE ENLARGED BY YAG. THE PROBLEM WAS RESOLVED AND PATIENT IS NO LONGER IN PAIN. THE IOP HAS DECREASED SIGNIFICANTLY AND ANGLES ARE OPEN. THE CAUSE OF THE EVENT WAS DUE TO THE ICL WAS TOO BIG, CAUSING PERIPHERAL ANGLE CLOSURE. Manufacturer Narrative: H3: DEVICE EVALUATION : LENS WAS RETURNED IN AA SPECIMEN CUP DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182702&pc=QCB | 2023826-2022-00022 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.5/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SHALLOW AC AND ANGLES, PAIN, AND VOMITING. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. HEALTH EFFECT (B)(4). CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182711&pc=QCB | 2023826-2022-00023 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -14.0/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SHALLOW AC AND ANGLES, PAIN, AND VOMITING. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970222&pc=QCB | 2023826-2023-04631 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) -2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE AND PIGMENT DISPERSION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED (VAULT). Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, PIGMENT DISPERSION, PAIN, BLURRED VISION AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "IOP NORMAL WITH SINGLE DRUG GLAUCOMA THERAPY". ADDITIONAL INFORMATION PROVIDED - "WAITING FOR IOP TO LOWER FURTHER. THE ANGLES ARE OPEN NOW. PIGMENT DISPERSION STILL PRESENT". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "RAISED IOP DUE TO ANGLE CLOSER AND HIGH VAULT AND PIGMENT DISPERSION". H6: HEALTH EFFECT- CLINICAL CODE: "2137" AND "1994" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "4644" SHOULD BE ADDED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378095&pc=QCB | 2023826-2023-03101 | NA | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; Pain; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO HIGH VAULTING, CAUSING PAIN AND HIGH PRESSURE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611930&pc=QCB | 2023826-2019-00782 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -10.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SHALLOW ANTERIOR CHAMBER AND LOW VAULT WAS OBSERVED. PATIENT HAD ELEVATED IOP AND PAIN. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS REPORTED AS "CALCULATOR AND REP RECOMMENDED LENGTH OF LENS DUE TO AGE (13.7MM). 13.2MM IMPLANTED AFTER EXPLANT" OF LENS. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THAT A 13.7MM TMICL 13.7 -10.00/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2019. SHALLOW ANTERIOR CHAMBER AND EXCESSIVE VAULT WAS OBSERVED. PATIENT HAD ELEVATED IOP AND PAIN. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9034019&pc=QCB | 2023826-2019-01672 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -12.50/3.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE "LENS WAS REMOVED DUE TO SEVERE PAIN, LENS WAS WRONG LENGTH. SURGEON REMOVED LENS AND DID NOT IMPLANT ANY OTHER LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SUPPLEMENTAL #1 MEDWATCH REPORT IN THE EVENT DESCRIPTION SHOULD BE CORRECTED TO: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2, -12.5/3.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (50 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED LENS. PER PATIENT'S LAST VISIT ON "(B)(6) 2019, BVCA 20/30, IOP 20 MMHG, EYE STATUS IS WHITE AND QUIET. IOP REMAINED ELEVATED 42 MMHG (B)(6) 2019 DESPITE MAX MEDICAL TREATMENT AND AQUEOUS RELEASE." CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 1937. PATIENT CODE 3191: NO CODE AVAILABLE (MEDICAL INTERVENTION). CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019 SHOULD BE REMOVED AND REPLACED WITH UNK AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (52 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED PER REPORTER, PER PATIENT'S LAST VISIT ON (B)(6) 2019, BCVA: 20/40, IOP: 18 MMHG, EYE STATUS WHITE AND QUIET." (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9034043&pc=QCB | 2023826-2019-01671 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE "LENS WAS REMOVED DUE TO SEVERE PAIN, LENS WAS WRONG LENGTH. SURGEON REMOVED LENS AND DID NOT IMPLANT ANY OTHER LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (52 MMHG) WITHOUT PUPIL BLOCK WAS OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS OVERSIZED PER REPORTER, PER PATIENT'S LAST VISIT ON (B)(6) 2019, BCVA: 20/40, IOP: 18 MMHG, EYE STATUS WHITE AND QUIET." CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: 16-SEP-2019 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. PATIENT CODE: 1937. PATIENT CODE: 3191 NO CODE AVAILABLE (MEDICAL INTERVENTION). PATIENT RECEIVED MEDICAL INTERVENTION AND ELEVATED IOP STILL REMAINED. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. .WEIGHT:UNK. ETHNICITY- UNK. RACE: UNK. -DATE IF EVENT:UNK EXPLANTED: UNK. .IMPLANTED: UNK. (B)(4).WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13240884&pc=QCB | 2023826-2022-00156 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL_12.6; -10.50/+1.00/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; ACUTE ANGLE CLOSURE SECONDARY TO PUPILLARY BLOCK FOLLOWING PHARMACOLOGIC MYDRIASIS; PAIN; DECREASED VISUAL ACUITY; NAUSEA; PUPIL DILATION. IT WAS REPORTED THAT "SUBJECT TREATED WITH ORAL ACETAZOLAMIDE AND TOPICAL TIMOLOL; PILOCARPINE; LATANOPROST AND BRINZOLAMIDE OVER THE COURSE OF APPROXIMATELY 2 HOURS. SUBJECT DISCHARGED WITH INSTRUCTIONS TO USE TIMOLOL; BRIMONIDINE; LATANOPROST AND ACETAZOLAMIDE AT HOME AND RETURN FOR FOLLOW-UP EXAM THE NEXT DAY". POST OP NOTES INDICATED THAT "THE SUBJECT RETURNED FEELING MUCH BETTER; WITHOUT PAIN OR HEADACHE. INTRAOCULAR PRESSURE MEASURED 8 MMHG; AND THE CORNEA WAS CLEAR OS. UNCORRECTED VISUAL ACUITY IMPROVED TO 20/16 OS. BRIMONIDINE WAS TAPERED; AND ALL OTHER MEDICATIONS DISCONTINUED". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS NOTED AS THE DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5: "ON (B)(6) 2022 LENS OPACITY (CATARACT) ADC WAS OBSERVED. ANTERIOR SUBCAPSULAR OPACITY NOTED AS GRADE1/MILD. NO ACTION TAKEN, PATIENT REMAINS ON STUDY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200420&pc=QCB | 2023826-2023-02472 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; Pupillary Block; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE - THE LENS HAD NOT CENTRAL HOLE. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS (THE PERFORATING HOLE WAS NOT PRESENT). DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: THE INVESTIGATIVE RESULTS DETERMINED THE MILLING TOOL NOT BEING REPLACED AFTER ITS MAX LIFE WAS REACHED IS MOST LIKELY THE ROOT CAUSE TO WHY THE CENTRAL HOLE WAS NOT PRESENT ON THE LENS. AUGUST 2022, IT HAS BEEN MADE A COMMON PRACTICE ACROSS MANUFACTURING TO CHANGE THE PRIMARY TOOL PRIOR TO MAX LIFE BEING REACHED TO MINIMIZE CHIPPED EDGES/HOLES. THIS PRACTICE WILL ALSO HELP PREVENT FURTHER OCCURRENCES OF A MISSING CENTRAL HOLE. IN ADDITION, RETRAINING FOR SOP & WORK INSTRUCTION ON DETECTION OF MISSING CENTRAL HOLE AND/OR LENS DEFECTS HAVE BEEN PERFORMED. SPECIFIC OPERATORS RESPONSIBLE HAVE BEEN MADE AWARE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. PRODUCT NON-CONFORMITY WAS NOTED AS "THE ICL LENS HAVE NO CENTRAL HOLE".THE PATIENT EXPERIENCED ELEVATED IOP WITH PUPIL BLOCK, PAIN AND DISCOMFORT. REPORTEDLY, "THE PATIENT FELT PAIN AND UNCOMFORTABLE".ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME PARAMETERS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "THE PATIENT HAS JUST COMPLETED THE SURGERY AND IS STILL IN FOLLOW-UP. THE PATIENT'S EYE CONDITION IS GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "THE ICL LENS HAVE NO CENTRAL HOLE". H6: HEALTH EFFECT- CLINICAL CODE: "2330" AND "1994" SHOULD BE ADDED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478860&pc=QCB | 2023826-2021-00635 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS BEING TOO BIG, THE CORNERS OF THE ANTERIOR CHAMBER ARE CLOSED, INTRAOCULAR PRESSURE IS ELEVATED, DESPITE REGULAR USE OF ANTIHYPERTENSIVE DRUGS AND THE PATIENT COMPLAINS OF SQUEEZING PAIN FROM TIME TO TIME. THERE IS ALSO A RETINAL DETACHMENT. THE CAUSE OF THE EVENT IS UNKNOWN. THE PATIENTS VISION IS NORMAL AND THERE ARE NO STRONG COMPLAINTS FROM PATIENT. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - YAG PI, MEDICAL INTERVENTION OUTSIDE STANDARD PROTOCOL. H6 - HEALTH EFFECT CLINICAL CODE: 2047 SHOULD BE REMOVED. B5 - IT WAS LATER CONFIRMED BY THE REPORTER THAT THE PATIENT DID NOT EXPERIENCE RETINAL DETACHMENT. ALSO ADDITIONAL YAG SURGERIES WERE PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11349059&pc=QCB | 2023826-2021-00497 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/1.0/023 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP . THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS "LENGTH ICL" Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942691&pc=QCB | 2023826-2021-01763 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -16.0/+1.0/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFO: "DESPITE ALL CALCULATIONS INDICATING 12.6 OR EVEN 13.2 THE LENS WAS CLEARLY OVERSIZED IN THE EYE AND A 12.1 WAS NEEDED. THERE WAS NOTHING IN THE PREOPERATIVE EVALUATION THAT SHOULD HAVE INDICATED THAT RESULT." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057192&pc=QCB | 2023826-2021-04911 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS SLOWLY RECOVERING AND THE PUPIL IS GETTING BETTER. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, THE PROBLEM IS RESOLVED. THE PATIENT IS HAPPY. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14654881&pc=QCB | 2023826-2022-01797 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO HIGH VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE IOP WAS 12 MMHG, VAULT WAS 750UM AND BCVA 20/25. THE OVERSIZED LENS CAUSED THE PUPIL BLOCK AND HIGH PRESSURE. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1583 SHOULD BE CORRECTED TO 2993. B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON 06-MAY-2022. THE LENS WAS EXPLANTED ON 13-MAY-2022 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) WITH PUPIL BLOCK. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE IOP WAS 12 MMHG, VAULT WAS 750UM AND BCVA 20/25. THE OVERSIZED LENS CAUSED THE PUPIL BLOCK AND HIGH PRESSURE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760468&pc=QCB | 2023826-2022-01928 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.00/+1.00/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.00/+1.00/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE STATUS OF THE EYE IS "ALL FINE AND PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11042998&pc=QCB | 2023826-2020-03085 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-9.5/1,5/109, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LOW VAULT AND PUPIL BLOCK WITH ELEVATED IOP 26MMHG (ASSESSED: (B)(6) 2020). ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: N5- "ANGLE CLOSURE WITH ELEVATED IOP WAS REPORTED."- SHOULD BE IN INITIAL MDR. H6: CLINICAL CODE 4581 - ANGLE CLOSURE- SHOULD BE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515299&pc=QCB | 2023826-2021-00653 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO PUBIL BLOCK WITH ELEVATED IOP(34MMHG), ANGLE CLOSURE WITH ELEVATED IOP((34MMHG) ASSESSED ON (B)(6) 2021), AND EXCESSIVE VAULT, ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515298&pc=QCB | 2023826-2021-00652 | 2021-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021, AND EXCESSIVE VAULT, ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7, -16.00/4.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO PUBIL BLOCK WITH ELEVATED IOP(39MMHG), ANGLE CLOSURE WITH ELEVATED IOP(39MMHG) ASSESSED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12270523&pc=QCB | 2023826-2021-03056 | 2021-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TICM12.6; -14.0/1.0/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE (FIXED) PUPIL; AND IT WAS INDICATED THAT ENLARGEMENT OF PI (YAG) WAS PERFORMED ON (B)(6) 2021. THE PROBLEM IS NOT RESOLVED BUT IT WAS NOTED A PLANNED LENS REPLACEMENT IS TAKING PLACE ON (B)(6) 2021. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12686036&pc=QCB | 2023826-2021-04108 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP. PROBLEM RESOLVED. AT A LATER UNKNOWN DATE "SOME IRIS ATROPHY + 3MM PUPIL" WAS REPORTED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN, " PATIENT ANATOMY/ OVERSIZE MANUFACTURED ICL" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT MODEL AND SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP AND THE PROBLEM RESOLVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS ALSO REPORTED AS ADDITIONAL INTERVENTION/TREATMENT. "SOME IRIS ATROPHY + 3MM PUPIL, DEVELOPED DURING AC WASH" WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER STATED, "PUPIL STILL 3MM BUT BETTER FUNCTION AND REACTS TO LIGHT (SLOWER THAN OD)." CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN, "PATIENT ANATOMY/OVER MANUFACTURED ICL" WAS REPORTED FOR CAUSE DETAILS. H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286583&pc=QCB | 2023826-2022-00174 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -7.00/+3.50/100 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED IOP; MEDICATION WAS PRESCRIBED (DIAMOX); ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT IT WAS LATER REPORTED THAT THE ISSUE WAS NOT RESOLVED AND THE PATIENT HAS PUPIL BLOCK/FIXED PUPIL; PRESSURE SPIKES AND PRK SURGERY WAS PERFORMED DUE TO THE VISION. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER UNEXPECTED SHORT CILIARY SULCUS SPACE". SEE MDR # 2023826-2022-00176 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15936540&pc=QCB | 2023826-2022-04499 | 2022-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -14.0/1.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; FIXED (UNREACTIVE) PUPIL; ENLARGEMENT OF PI - YAG WHICH WAS PERFORMED ON (B)(6) 2022; AND ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142929&pc=QCB | 2023826-2023-00098 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), PUPIL BLOCK, ANGLE CLOSURE, UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY. AN ADDITIONAL PI WAS PERFORMED ON (B)(6) 2022 AND THERE WAS ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-00099 FOR REPLACEMENT LENS. REMAINING PUPIL DISSIPATION AFTER AN ACUTE GLAUCOMA ATTACK. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNFICANT LENS HAPTIC DAMAGE. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14116256&pc=QCB | 2023826-2022-01185 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: VAULT WAS REPORTER IN THE RIGHT EYE (OD) OF A PATIENT WITH AN IMPLANTABLE COLLAMER LENS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/128 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PUPIL BLOCK WITH ELEVATED IOP, PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY, WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS REPORTED AS OTHER, THE REPORTER STATED "FIXED DILATED PUPIL, TWO WEEKS POST ICL. IOP SLIGHTLY ELEVATED, THOUGHT MIGHT BE PUPILLARY BLOCK. PIGMENT DISPERSION NOTED, THOUGHT MIGHT BE DUE TO PRESSURE SPIKES NOT IN OFFICE. NOW THINK MOXI-DEX (5MG/CC) MIGHT BE CASE." ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI YAG, ADDITIONAL SUTURE AND ATTEMPT TO EXCHANGE) WAS PERFORMED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PHAKIC WITHOUT ICL IMPLANTED. IRIS TRAUMA TEMPORALLY, SUTURE IN PLACE, NOW WITH MR OF -12.75 +1.75X119. THREE DIOPTER SHIFT. WAITING TO SETTLE TO SEE IF PROCEED WITH SECONDARY IMPLANT. FOR ADDITIONAL INFORMATION THE REPORTER STATED "FIRST CASE OF IRIS ATROPHY. WAS CONCERNED IT WAS IOP RELATED TO HIGH VAULT. ADDITIONAL PI ENLARGEMENT NO HELP. PROGRESSIVE TID DECIDED TO EXCHANGE ICL. VERY DIFFICULT CASE, WITH INDUCTION LEVEL ANESTHESIA NEEDED, STILL NOT ABLE TO KEEP PATIENT STILL. LUCKY TO GET LENS OUT. IRIS TRAUMA/PROLAPSE. ATTEMPTED TO PLACED, THEN PATIENT BEGAN TO ASPIRATE, FORTUNATE TO BE DONE WITH CASE. CONTEMPLATING SECONDARY ICL UNDER GENERAL ANESTHESIA BUT WANT REFRACTION TO STABILIZE. ATTEMPTED TO IMPLANT TMICL 12.6 -11.0/1.5/109 (ONLY ONE AVAIL URGENTLY) SN: (B)(6)." H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION, UNREACTIVE/FIXED PUPIL, IRIS ATROPHY, TRANSILLUMINATION IRIS DEFECTS. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DILATED PUPIL, IRIS TRAUMA/PROLAPSE. SHOULD HAVE BEEN INCLUDED. (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14505054&pc=QCB | 2023826-2022-01687 | 2021-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -11.0/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; REFRACTIVE SURPRISE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK; ANGLE CLOSURE; ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE (FIXED) PUPIL; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ON (B)(6) 2021 A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION-LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794474&pc=QCB | 2023826-2022-01950 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794499&pc=QCB | 2023826-2022-01951 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE, THE REPORTER STATED "THE DOCTOR THINKS THAT THE PATIENT'S MYOPIA PROBLEM HAS NOT BEEN SOLVE." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440465&pc=QCB | 2023826-2024-00046 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OP DUE TO EXCESSIVE VAULT; REFRACTIVE SURPRISE; PUPIL BLOCK WITH ELEVATED IOP; AND ANGLE CLOSURE WITH ELEVATED IOP. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE. SEE MFR# 2023826-2024-00047 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-CLINICAL CODE 4581: ANGLE CLOSURE WITH ELEVATED IOP. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12662188&pc=QCB | 2023826-2021-04050 | 2021-09-12 | Injury | STAAR SURGICAL COMPANY | 2021-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/074 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND PUPIL BLOCK WITH INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245039&pc=QCB | 2023826-2020-01403 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (34 MMHG) WAS OBSERVED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS (ASSESSED ON (B)(6) 2020), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE (FIXED) PUPIL, LENS OPACITY AND EXCESSIVE VAULT WERE OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INCISION ENLARGEMENT, ADDITIONAL SUTURE AND THE LENS WAS REMOVED ON (B)(6) 2020. REPORTEDLY THERE ARE NO PLANS TO REPLACE THE LENS. PER REPORTER ON (B)(6) 2020, PRESCRIPTION OF IOP LOWERING DROPS PROVIDED AND AS OF (B)(6) 2020 IOP WAS 18MMHG AND PUPILS ARE STILL UNREACTIVE. Manufacturer Narrative: CORRECTED DATA: B5 - 15.50/2.5/100 (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -15.00/3.0/100 (SPHERE/CYLINDER/AXIS). H6 - HEALTH EFFECT - CLINICAL CODE: 4581 (UNREACTIVE FIXED PUPIL) AND (LENS OPACITY) AND (SIGNIFICANT ENDOTHELIAL CELL LOSS) SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPACT - DEVICE CODE: 4625 ADDITIONAL SURGERY - INCISION ENLARGEMENT, ADDITIONAL SUTURE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. SHOULD HAVE BEEN INCLUDED. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN WITH FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245040&pc=QCB | 2023826-2020-01402 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) WAS OBSERVED ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE (FIXED) PUPIL, AND EXCESSIVE VAULT WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INCISION ENLARGEMENT, ADDITIONAL SUTURE AND THE LENS WAS REMOVED ON (B)(6) 2020. REPORTEDLY THERE ARE NO PLANS TO REPLACE THE LENS. PER REPORTER ON (B)(6) 2020, PRESCRIPTION OF IOP LOWERING DROPS PROVIDED AND AS OF (B)(6) 2020 IOP WAS 19MMHG AND PUPILS ARE STILL UNREACTIVE. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT - CLINICAL CODE: 4581 (UNREACTIVE FIXED PUPIL). SHOULD HAVE BEEN INCLUDED. H6 - HEALTH IMPACT - DEVICE CODE: 4625 ADDITIONAL SURGERY - INCISION ENLARGEMENT, ADDITIONAL SUTURE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. SHOULD HAVE BEEN INCLUDED. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN WITH FIBER ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279904&pc=QCB | 2023826-2020-01454 | 2020-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED THE SAME DAY ON (B)(6) 2020 DUE TO EXCESSIVE VAUTLING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON HAD ENLARGED THE PI BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (HIGH VAULT). H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A PLASTIC BAG, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6: HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6: HEALTH EFFECT IMPACT CODE: 4625 - SECONDARY SURGICAL INTERVENTION; PI, YAG. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990641&pc=QCB | 2023826-2022-04529 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS. THE SURGEON REPORTED LATE POSTOPERATIVE PUPILLARY BLOCK WITH FIBRIN OCCLUDING THE PUPIL. THE SURGEON PERFORMED A PI, WHICH RESOLVED THE HIGH IOP. CULTURE OF THE OCCLUSION WAS NEGATIVE AND THE PATIENT DID WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 -PARS PLANE VITRECTOMY. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTED LATE POSTOPERATIVE PUPILLARY BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE (FIXED) PUPIL AND FIBRIN OCCLUDING THE PUPIL. THE SURGEON ENLARGED THE PI, WHICH RESOLVED THE HIGH IOP. THE PATIENT HAD A RETINAL DETACHMENT MACULA ON (B)(6) 2022, WHICH REQUIRED A PARS PLANE VITRECTOMY SURGERY. CULTURE OF THE OCCLUSION WAS NEGATIVE AND THE PATIENT DID WELL. THE LENS REMAINED IMPLANTED. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM#: (B)(4). Manufacturer Narrative: AGE AT TIME OF EVENT, DATE OF BIRTH :UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL: EXPIRATION DATE UNK. IF IMPLANTED, GIVE DATE: UNK. IF EXPLANTED, GIVE DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10956516&pc=QCB | 2023826-2020-02851 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON (B)(6) 2020 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.0/2.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON 06 NOV 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON 19 NOV 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). HOWEVER, SIGNS/SYMPTOMS WERE NOT BEYOND THAT OF TYPICAL SURGERY. IT WAS REPORTED THAT IRRITATION AND BLURRY VISION RESOLVED. PATIENT IS NOW 20/20. H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17432754&pc=QCB | 2023826-2023-03243 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Red Eye(s); Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.50/+2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARES, REDNESS, MEDICATION USED AS WELL AS GONIOSCOPY PERFORMED TO OPEN ANGLES. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9039904&pc=QCB | 2023826-2019-01703 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Retinal Detachment; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICMO, -18.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION, ELEVATED IOP, REFRACTIVE SURPRISE, AND RETINA DETACHMENT WAS OBSERVED. SURGICAL INTERVENTION (REPOSITIONING, RETINAL REPAIR AND GAS INFUSION) AND MEDICAL INTERVENTION (ANTIBIOTICS, STEROIDS, GLAUCOMA MEDICATION) WERE PERFORMED/ PRESCRIBED. LENS CURRENT STATUS WAS REPORTED TO BE IN POSITION BUT VAULT WAS APPROXIMATELY 100 MICRONS. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED VIA EMAIL. THE LENS WAS REPOSITIONED A SECOND TIME ON (B)(6) 2019. "PATIENT HAS GAS LEAKAGE IN THE EYE. SO AGAIN RETINA SURGEON HAS GAS FILLED UP AND THEY WILL WAIT FOR FOUR WEEK FOR VISUAL RECOVERY UPDATE." CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE ON (B)(6) 2019. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271368&pc=QCB | 2023826-2021-03013 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; UGH (Uveitis Glaucoma Hyphema) Syndrome; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO UGH SYNDROME CAUSED IRIS ATROPHY AND CONTINUOUS PIGMENT IN THE ANTERIOR CHAMBER. REMOVAL WENT WELL WITHOUT COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND UNREACTIVE (FIXED) PUPIL/IRIS. THE PIS WERE ENLARGED BY YAG ON 26-FEB-2021. THE CAUSE OF THE EVENT WAS THE DEVICE. RECENT POST-OP VISIT FOUND EYE STILL WITH PIGMENT IN AC AND FIXED IRIS. H3: DEVICE EVALUATION NOT REQUIRED. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL/IRIS H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGE PIS BY YAG CORRECTED DATA: B3: 08-FEB-2021. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE PATIENT HAD A BCVA OF 20/20 AFTER EXPLANT. PER EMAIL RECEIVED ON (B)(6) 2021, " A REPLACEMENT LENS HAS NOT BEEN IMPLANTED YET, AS WE ARE STILL WAITING FOR PIGMENT/IRITIS TO RESOLVE." IN ADDITION THE REPORTER ALSO STATED THAT HIGH ICL VAULT CAUSED UGH SYNDROME. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378904&pc=QCB | 2023826-2023-05715 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.50/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, BLURRED VISOIN AND PROLONGED POSTOPERATIVE ANTERIOR UVEITIS. IT WAS REPORTED THAT THE IOP WAS CONTROLLED WITH MEDICATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112820&pc=QCB | 2023826-2019-01627 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION AND GLARE/HALOES. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS" INJECTOR MODEL-MSI-TF; LOT#-UNK SHOULD HAVE BEEN INCLUDED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980474&pc=QCB | 2023826-2020-02945 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -12.5/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS LEFT (OS) EYE. THE PATIENT REPORTS HALOS, DOUBLE VISION, LIGHT RINGS, POSTOP PAIN, PRESSURE, AND VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. POSTOP PAIN AND PRESSURE RELIEVED BY "PUSHING ONTO THE CUT TO FORCE FLUID OUT." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: 6HRS POST-OP AC FLUID DRAINED VIA EXISTING PARACENTESIS. DIAMOX ADMINISTERED. GHOSTING OF VISION IN LOW LIGHT. ELEVATED IOP DUE TO RETAINED OCCUCOAT. POST-OP HALO/GLARE IS EXPECTED DUE TO SCOTOPIC PUPIL SIZE AND RELATIVELY HIGHT MYOPIC PRESCRIPTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980487&pc=QCB | 2023826-2020-02946 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -13.0/+2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS RIGHT (OD) EYE. THE PATIENT REPORTS HALOS, LIGHT RINGS, AND POSTOP VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. Manufacturer Narrative: B5-ADDITIONAL INFO: ELEVATED IOP 6HRS POST-OP DUE TO RETAINED OCCUCOAT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: CLINICAL CODE: 4581, FOR "GHOSTING VISION". 2227, HALO IS NOT APPLICABLE. DEVICE PROBLEM CODE: 3001, GLARE/HALO IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671674&pc=QCB | 2023826-2019-01020 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: REFERENCE MFR REPORT# 2023826-2019-00886 FOR INITIAL LENS IMPLANT. THE SURGEON EXCHANGED THE INITIAL IMPLANT WITH A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/013 (SPHERE/CYLINDER/AXIS). THE SURGEON REPORTS LENS ROTATION, HALOS, STARBURST, "COMA DOUBLE IMAGE" AND PATIENT IS ON "SYMBIZA+MONOPROST". Manufacturer Narrative: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT DUE TO MEDICATION INTERVENTION. SURGEON REPORTED PER LAST VISIT DATE ON (B)(6) 2019, PATIENT STILL IS EXPERIENCING MILD HALOS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: MALFUNCTION SHOULD BE CORRECTED TO SERIOUS INJURY IN INITIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 1937 SHOULD BE CORRECTED TO PATIENT CODE: 1937, THE PATIENT'S IOP IS MAINTAINED ON TWO IOP LOWERING MEDICATIONS: SIMBRINZA (BRINZOLAMIDE/BRIMONIDINE TARTRATE OPHTHALMIC SUSPENSION) AND MONOPROST. THE ICL REMAINS IMPLANTED AS OF THIS TIME. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140593&pc=QCB | 2023826-2020-01201 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, HALOS AND CIRCLES. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6541708&pc=QCB | 2023826-2016-01557 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.0/97 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: PMA/510(K): NA, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6549621&pc=QCB | 2023826-2016-01558 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-05 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.0/109 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT HAS EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: NA- THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552490&pc=QCB | 2023826-2019-00648 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -15.5/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE VA OF THE EYE AFTER EXPLANT IS REPORTED AS "FOGGY, SMEARED." Manufacturer Narrative: PATIENT CODE 2140: VISUAL DISTURBANCE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030329&pc=QCB | 2023826-2020-01045 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE." SHOULD BE REMOVED FROM SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030336&pc=QCB | 2023826-2020-01044 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -15.00/5.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR." SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 2682 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - INVESTIGATION CODE 4117 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334451&pc=QCB | 2023826-2021-00361 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Iritis; UGH (Uveitis Glaucoma Hyphema) Syndrome | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION AND PATIENT EXPERIENCED IRITIS. THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE PATIENT HAD MILD IRITIS/POSSIBLE UGH (UVEITIS-GLAUCOMA-HYPHEMA), THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE EYE IS GOOD AND IOP NORMAL, ICL STABLE, WITH NO RECURRENCE OF IRITIS. ELECTED TO EXCHANGE ICL DESPITE ADEQUATE VAULT DUE TO ICL ROTATION (60 DEGREES) AND VARIABLE IOP AND POST-OP CELL/FLARE. INCREASING SIZE APPEARS TO HAVE SOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073178&pc=QCB | 2023826-2021-02032 | NA | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. " UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073179&pc=QCB | 2023826-2021-02031 | NA | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS EXPERIENCING PAIN OVER THE INNER EYE FOR NEARLY 10 MONTHS. "UNEVENTFUL SURGERY, GOOD VISION, GOOD POSITIONING AND GOOD VAULTING, CLEAR CORNEA, CLEAN LENS, NO PIGMENT DISPERSION" WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14620821&pc=QCB | 2023826-2022-01794 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639450&pc=QCB | 2023826-2022-01795 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/161 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT WAS COMPLAINING OF PAIN, POOR ABILITY TO READ AND LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022 AND PROBLEM RESOLVED. REPORTEDLY, "ICL HAS BEEN EXPLANTED, PATIENT GOES BACK TO USING REFRACTIVE APPLIANCES. CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704775&pc=QCB | 2023826-2023-03953 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/+2.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON EXPLANTED THE LENS ON (B)(6) 2023, DUE TO THE PATIENT HAVING DISSATISFACTION AND FOREIGN BODY SENSATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PATIENT DISSATISACTION, FOREIGN SENSATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PATIENT ALSO HAD PAIN. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652119&pc=QCB | 2023826-2019-00915 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, BLURRED VISION, PAIN, AND "MINOR TEMPORALLY ICL SHIFT." CORNEA IS CLEAR, NO INJUNCTIVAL INJECTION, IOP AT 14MMHG AT LAST VISIT. PATIENT CLAIMS NO HISTORY OF EYE INJURY. Manufacturer Narrative: PRODUCT PROBLEM: SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE MARKED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. DATE RECEIVED BY MANUFACTURER: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. POST-OP PATIENT UCVA WAS RECORDED AS 20/25 AND THE VAULT WAS 458UM. REPOSITIONING WAS CONDUCTED ON A UNSPECIFIED DATE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454797&pc=QCB | 2023826-2020-01864 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (37MMHG) OBSERVED ON (B)(6) 2020. ADDITIONAL TREATMENT OF ENLARGEMENT OF PI AND A YAG WAS NECESSARY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS NON PATENT PI. REPORTEDLY, PATIENT CONDITION "IMPROVED - NO COMPLAINTS." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED BLURRY VISION WITH EXTREME PAIN. H6 - DEVICE CODE: 2993 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525151&pc=QCB | 2023826-2019-00599 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019 AND PROBLEM WAS RESOLVED. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED WITH ICL EXCHANGE. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. THE PATIENT HAD LASIK OU ON (B)(6) 2019 TO CORRECT RESIDUAL ASTIGMATISM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525163&pc=QCB | 2023826-2019-00598 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED ON 25-FEB-2019 AND EXPLANTED ON 27-MAR-2019. THE PATIENT EXPERIENCED PAIN AND THE PAIN WAS RESOLVED AFTER ICL EXCHANGE.. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT HAD LASIK OU ON 14-MAY-2019 TO CORRECT RESIDUAL ASTIGMATISM. THE CAUSE OF THE EVENT WAS DUE TO SIZING NOMOGRAM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135938&pc=QCB | 2023826-2020-01166 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Pain; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -9.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) IN 2019. THE REPORTER HAS INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS DUE TO THE LENS HAS VAULTED FORWARD AND POSSIBLY RUBBING AGAINST THE IRIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5:THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION AND INTERMITTENT PAIN IN THE DARK. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - PATIENT CODE: 3191 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PIGMENT DISPERSION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478865&pc=QCB | 2023826-2021-00636 | 2021-02-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE PUPIL WAS UNREACTIVE (FIXED) (BLOCKED PUPIL). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE PATIENT IS NOW 8/10 IN THE LEFT EYE AND THE PUPIL IS ALMOST PERFECT BUT STILL NOT 100%. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311393&pc=QCB | 2023826-2022-00176 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Pupillary Block; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739723&pc=QCB | 2023826-2023-01370 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Red Eye(s); Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/1.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. IT WAS INDICATED THAT EVERYONE WAS NORMAL FOR THE 1ST FOUR DAYS AFTER SURGERY BUT ON THE 5TH DAY THE EYE WAS DILATED; WAS PRESCRIBED MEDICATION (PILOCARPINE); BUT PIGMENT DISPERSION STARTED AND RED EYE AND SERVER EYE PAIN. A DIFFERENT SURGEON HAD REPOSITIONED THE LENS BUT THIS DID NOT RESOLVE THE PROBLEM. A 3RD DOCTOR HAD PRESCRIBED ALPHAGAN P. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - DILATED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608173&pc=MTA | 2023826-2017-00900 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/043 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE, AND IRIS ATROPHY. ON (B)(6) 2017 THE PATIENT ALSO EXPERIENCED LENS OPACITY ASCO. THE SURGEON PLANS TO EXCHANGE TO A SHORTER LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. (B)(4). Manufacturer Narrative: (B)(6). THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). H6_WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6220084&pc=MTA | 2023826-2017-00009 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Glaucoma; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DESCRIPTION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/13 AS OF 04/21/2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926670&pc=MTA | 2023826-2016-00374 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED AT FOLLOW-UP VISITS, THE LENS HAPTIC WAS FOLDED AND THE PUPIL WAS OVALISED. THE LENS WAS TILTED ON ONE SIDE AND WAS RE-POSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP #4 FOR MDR # 2023826-2017-00374 WAS SUBMITTED IN ERROR. THE REPORTED INFORMATION WAS MEANT TO BE FOLLOW-UP #4 FOR MDR # 2023826-2016-00374. AS A RESULT, ONLY THE INITIAL MDR # 2023826-2017-00374 PERTAINS TO CLAIM# 710106. FOLLOW-UP #4 FOR MDR # 2023826-2016-00374 WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791893&pc=MTA | 2023826-2018-01244 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/+3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HIGH IOP, MID DILATED PUPIL, CORNEAL EDEMA, CLOSE ANGEL GLAUCOMA AND ANGLE BLOCK. THE SURGEON PERFORMED A PARS PLANA VITRECTOMY, PHACO-EMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION AND IRIS SUTURE. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP VISIT VISUAL ACUITY WAS 20/40 WITHOUT CORRECTION, IOP WAS NORMAL WITH A REGULAR PUPIL. Manufacturer Narrative: CORRECTED DATA: ADDED 'OTHER SERIOUS' TO INITIAL REPORT. REPORTER ALSO INDICATES CORNEAL DECOMPENSATION. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6229803&pc=MTA | 2023826-2017-00020 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/0.96 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND ELEVATED IOP. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND CORNEAL EDEMA. ON (B)(6) 2016 THE LENS WAS EXPLANTED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATION PENDING. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109978&pc=MTA | 2023826-2017-01877 | 2017-10-30 | Injury | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.0/+3.5/171 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, AND CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: REPORTER STATED PATIENT IS OKAY. THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.92 MM)]. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6760836&pc=MTA | 2023826-2017-01145 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Corneal Edema; Edema; Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -4.5/1.5/86 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE NEXT DAY, SIGNS OF TOXIC ANTERIOR SEGMENT SYNDROME, INCLUDE TEARING, PAIN, BLURRY VISION, EYELID EDEMA, HYPEREMIA, CORNEAL EDEMA, FIXED PUPIL AND FIBRIN CLOT IN THE ANTERIOR CHAMBER WAS NOTICED. THERAPY WAS STARTED WITH ORAL PREDNISOLONE (60 MG/DAY) AND ACETAZOLAMIDE 250 MM (1+1+1), TROPICAL PREDNISOLONE 1 DROP PER HOUR, TRAPICAMIDE, OFLOXACINE, AND BROMOFENAC. THE FOLLOWING DAY, THE PATIENT EXPERIENCE LESS PAIN, INTENSE HYPEREMIA, LESS EYELID EDEMA BUT A VERY CLOUDY WHITISH CORNEA. CORNEAL PERIPHERY WAS PARTIALLY PARED AND A BLOOD CLOT WAS SEEN IN THE INFERIOR A.C. AN ANTERIOR CHAMBER WASHING WAS PERFORMED AS WELL AS THE TEMPORAL INCISION WITCH BEING SUTURED WITH 2 10/00 NYLON STICHES AND INTRAVITREOUS INJECTION OF VANCOMICINE 1MG/0.1 ML AND CEFTAZIDIME 1.0 MG/0.1 ML. ACCORDING TO THE SURGEON, THE CORNEA IS CLEARING UP AND WOULD NEED TO MAKE ANOTHER SURGERY TO REMOVE THE MEMBRANE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531883&pc=MTA | 2023826-2017-00727 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Corneal Edema; Inflammation; Intraocular Pressure Increased; Pupillary Block; Uveitis; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED GLARE/HALOS, BLURRED VISION AND INFLAMMATION. ALSO, MILD UVEITIS, PIGMENT DISPERSION AND HAD DEVELOPED URETTS ZAVALIA SYNDROME (PERSISTENT PUPIL DILATION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP SHOWED CORNEAL EDEMA, SHALLOW ANTERIOR CHAMBER AND NORMAL INTRAOCULAR PRESSURE. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594294&pc=MTA | 2023826-2018-00898 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/156 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "SHALLOW AC," DECREASE IN UA, AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER "TOW TIME AC WASH" AND "AC REFORMATION." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN UPDATED INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION: LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.76MM)]. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850665&pc=MTA | 2023826-2017-01401 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/X092 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6)2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANGLE CLOSURE WITH ELEVATED IOP. THE SURGEON REPORTED ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. SURGEON ALSO REPORTS CORNEAL EDEMA/MIDRIASIS. THE SURGEON STATES THAT THE PATIENT STATUS IS OKAY. Manufacturer Narrative: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594288&pc=MTA | 2023826-2018-00897 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/014 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THREE DAYS POSTOP IT IS NOTED THAT THE SITUATION HAS IMPROVED SLIGHTLY, "TOW TIME AC WASH AND AC REFORMATION WITH PI (SURGICAL) AND AIR BUBBLE PLACEMENT." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A "STERIPACK". SPELLING OF "STERIPACK" WAS INCORRECT PRIOR. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER STATED "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION - LENS WAS RETURNED DRY IN A STERIPECLE WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594291&pc=MTA | 2023826-2018-00896 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC," AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER TWO AC WASHES AND SURGICAL PI WITH "AIR BUBBLE IN AC." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC"] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.93MM)] (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221040&pc=MTA | 2023826-2015-01452 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2015-11-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01451 FOR RIGHT EYE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/0.5/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE PUPIL, ANGLE CLOSURE, AND POST-OP PAIN CAUSED BY ELEVATED IOP AND CORNEAL EDEMA. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419929&pc=MTA | 2023826-2018-00537 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (STEROID TREATMENT). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7420014&pc=MTA | 2023826-2018-00535 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+2.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419890&pc=MTA | 2023826-2018-00531 | 2018-03-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; Visual Impairment; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/+2.5/166 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION, UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. A YAG PI WAS PERFORMED ON (B)(6) 2018. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC USED IN SURGERY, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS EXPLANT, THE PATIENT WAS ADVISED TO TAKE PRED/GATI 3 TIMES DAILY. AFTER DISCONTINUING, THE PATIENT DEVELOPED AN INCREASE IN IOP AND INFLAMMATION IN THE ANTERIOR CHAMBER. CURRENTLY, THE PATIENT IS BEING TREATED WITH PREDNISOLONE AND COMBIGAN, AZOPT, AND TRAVATAN Z FOR PRESSURE SPIKES. IT IS BELIEVED THAT THE SPIKES ARE DUE TO "STEROID USE VS INFLAMMATORY CAUSES." REPORTEDLY, NO INFECTION HAS OCCURRED WITHIN THE EYE. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT REPORTS "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA, AND LOSS OF VISION. ALSO STATED IS THAT THE IRIS IS PERMANENTLY FIXED AND DILATED. GLAUCOMA SURGERY PERFORMED. G3-REPORT SOURCE ADDED. H6-ADDITIONAL 3191: (BLOWN IRIS, FIXED IRIS, GLAUCOMA SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, LENS EXPLANT, FIXED PUPIL) OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT (46YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6249746&pc=MTA | 2023826-2017-00051 | 2016-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/098 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), AND DEVELOPMENT OF GLAUCOMA. AS OF THE DATE OF THIS SUBMISSION, THE LENS HAS BEEN RETURNED, BUT THE EVALUATION RESULTS ARE NOT YET AVAILABLE. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS PERFORMED. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419948&pc=MTA | 2023826-2018-00533 | 2018-02-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/+1.0/007 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS PUPIL BLOCK, ANGLE CLOSURE, ELEVATED IOP, AND PIGMENT DISPERSION. THE SURGEON STATED HE THOUGHT THE HIGH IOP WAS DUE TO THE VISCOELASTIC. REPORTEDLY, HE ASSUMES THE AQUA PORT WAS BLOCKED BY THE VISCOELASTIC. A YAG WAS PERFORMED AND THEN IOP RETURNED TO "NORMAL." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PATIENT REPORTS THE LENS WAS EXPLANTED. ALSO REPORTED: "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA AND VISION LOSS, GLAUCOMA SURGERY PERFORMED, IRIS PERMANENTLY FIXED AND DILATED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, ANGLE CLOSURE, YAG) CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.67MM)] (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294520&pc=MTA | 2023826-2018-00267 | 2018-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: MALIGNANT GLACOMA WITH EXPLANT WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5_13.2, -13.5/+1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED 5 HOURS LATER DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL PROCEDURES (ENLARGEMENT OF PI OR ADDITION OF NEW PI AND YAG) WERE REPORTED TO BE PERFORMED. ENTERED STAAR (B)(4)'S ADDRESS DEVICE EVALUATION: LENS WAS RETURNED IN A CONTACT CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600863&pc=MTA | 2023826-2015-00311 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, HEADACHE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199928&pc=MTA | 2023826-2015-01357 | 2015-09-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED INTRAOCULAR PRESSURE (UP TO 25MMHG) STARTED BY THE 4TH DAY. THE PATIENT WAS GIVEN PRESSURE DROPS, THE PRESSURE WENT DOWN TO 22 BUT THEN IT WENT UP TO BETWEEN 32MMHG AND 50MMHG. THE PATIENT'S LAST UCVA WAS 20/20 AND IS SATISFIED WITH THE ICL BUT WAS COMPLAINING OF HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. RECENT IMAGES SHOWS GOOD VAULT AND NORMAL IOP BUT THE PATIENT IS STILL ON MEDICATION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS OBSERVED AND ELEVATED INTRAOCULAR PRESSURE (IOP) STARTED BY THE 4TH DAY. THE SURGEON ESTABLISHED MEDICAL MANAGEMENT AND IOP WAS CONTROLLED TO NORMAL LEVELS WITH ONE GLAUCOMA MEDICATION. PATIENT ALSO REPORTED HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. ON LAST VISIT, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, VAULT WAS NORMAL, AND IOP WAS NORMAL. THE LENS REMAINS IMPLANTED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT IS MARKETED IN THE U.S. DIOPTER: -17.0/+2.0/X093. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444673&pc=MTA | 2023826-2017-00527 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+6.0/085 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED (B)(6) 2016 AND THE SURGEON NOTICES A GRADUAL INCREASE IN IOP. THE PROBLEM IS NOT RESOLVED AND A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112582&pc=MTA | 2023826-2015-01188 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.30/+1.5/075 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT SUFFERED FROM HEADACHE. NO OTHER LENS WAS IMPLANTED. SEE MFR #2023826-2015-01187 FOR RIGHT EYE. Manufacturer Narrative: BASED ON THE DHR REVIEW, THE MANUFACTURING AND PACKAGING PROCESSES WAS PERFORMED ACCORDING TO PROCEDURES AND NO FINDINGS OR DEVIATIONS WERE REPORTED. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED, HOWEVER, THE COMPLAINT IS NOT DEVICE RELATED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THAT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND TWO OF THE FOOTPLATES WERE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.30/+1.5/X075. (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146314&pc=MTA | 2023826-2015-01251 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/X087 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT EXPERIENCING DISCOMFORT, A HEADACHE. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR A FOLLOW-UP. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -17.5/+2.0/X087. LENS NOT RETURNED. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151588&pc=MTA | 2023826-2015-01250 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+2.50/X061 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED DISCOMFORT (HEADACHE) AND THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR FOLLOW UP VISIT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.50/+2.50/X061. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587675&pc=MTA | 2023826-2017-00873 | 2016-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.0/+0.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, IRIS ATROPHY, AND HEADACHE. THE SURGEON COMMENTED THE VAULT ONE MONTH AFTER THE SURGERY IS 600 MICRONS, AND 3 MONTHS AFTER IS 902 MICRONS. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA, IRIS ATROPHY). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298844&pc=MTA | 2023826-2019-00191 | 2018-12-17 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENT COMPLAINED OF A HEADACHE ON LEFT SIDE AFTER SURGERY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. ANOTHER LENS WAS NOT IMPLANTED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION-THE LENS WAS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694745&pc=MTA | 2023826-2017-01119 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN MICRO CENTRIFUGE VIAL WITH LITTLE MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694750&pc=MTA | 2023826-2017-01120 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WITH A CURVED SHAPE. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422330&pc=MTA | 2023826-2017-00444 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDE: ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. ON (B)(6) 2017, THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR THICK AND REDDISH BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDED ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. AFTER LENS EXPLANT, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100 AND THE PATIENT EXPERIENCED A LITTLE INFLAMMATION OF THE ANTERIOR SEGMENT. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. CONCLUSION: OFF-LABEL USE-KERATOCONUS EYE. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721249&pc=MTA | 2023826-2018-01084 | 2018-06-13 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP, ANGLE CLOSURE, INFLAMMATION, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE PATIENT IS "DOING WELL." AZETAZOLAMIDE 250MG GIVEN TO MANAGE SYMPTOMS. THE SURGEON LISTS THE CAUSE OF THE EVENT AS DEVICE. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DATE RECEIVED BY MFR: -IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107278&pc=MTA | 2023826-2018-01723 | 2018-10-31 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination /Decontamination Problem; Unintended Movement | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/114 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UVEITIS, SLIGHTLY DISTORTED PUPIL, CELL DEPOSITS ON THE ICL, AND INFLAMMATION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. PER THE SURGEON, THE LENS "HAS BEEN SLIGHTLY TILTED FROM THE BEGINNING (PATHOLOGY IN THE SULCUS?)." CELL DEPOSITS DISAPPEAR WHEN GIVEN PRED FORTE; REAPPEAR WHEN THE "STERIOD" IS "DISCONSTINUED". REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: EVENT-ADDITIONAL DATA RECEIVED: "THE EXCHANGE RESOLVED THE INFLAMMATION, THE EYE IS CALM. VA SC 0.8, CC 1.0 VAULT 350 MICRON. ICL IN VERTICAL POSITION (AWAY FROM THE NASAL PROBLEM AREA). PUPIL SLIGHTLY NASALLY DISTORTED. THE MD ATTRIBUTES THE INFLAMMATION CLEARLY TO THE OVER VAULTED ICL. EVERYONE IS HAPPY". IMPLANT DATE: (B)(6)2018 DEVICE EVALUATED BY MFR:DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT-CORRECTED IMPLANT DATE INTO THE RIGHT EYE IS: "(B)(6) 2015" IN THE INITIAL MDR. RELEVANT TESTS/LABORATORY DATA-CORRECTED TO: "NA" IN THE SUPPLEMENTAL MDR#1 CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355865&pc=MTA | 2023826-2016-00029 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -18.00/+1.50/X084 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00030 FOR LEFT EYE. Manufacturer Narrative: CONCLUSION: UNABLE TO CONFIRM COMPLAINT - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+1.50/X084. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066047&pc=MTA | 2023826-2015-01174 | NA | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/-1.50/089 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01175 FOR LEFT EYE. Manufacturer Narrative: THERE WAS NO EXCESSIVE VAULTING. THE INTRAOCULAR PRESSURE RISE WAS UNDER CONTROL AND THE PROBLEM WAS RESOLVED. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.50/+1.50/089. (B)(4). METHOD:EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740585&pc=MTA | 2023826-2015-00486 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAME LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/25 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. REDUCTION OF IRIDO-CORNEAL ANGLES. EXPLANTED. VAULTING. DEVICE EVALUATED BY THE MANUFACTURER? NO: LENS WAS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900441&pc=MTA | 2023826-2017-01484 | 2017-08-06 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (CONTROLLED BY MEDICATION). THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SMALLER SIZE LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. EXPLANTED DATE IS CORRECTED TO "N/A" FROM "UNK" IN THE INITIAL; AS THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824489&pc=MTA | 2023826-2018-01252 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355968&pc=MTA | 2023826-2016-00030 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/+3.00/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00029 FOR RIGHT EYE. Manufacturer Narrative: CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+3.00/X083. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066072&pc=MTA | 2023826-2015-01175 | NA | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/2.00/103 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01174 FOR RIGHT EYE. Manufacturer Narrative: SURGEON INDICATED THE IOP WAS UNDER CONTROL. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. LENS IMPLANTED ON (B)(6) 2015. (B)(6). (B)(4). THE PATIENT WAS UNDER THE AGE OF 21 YEARS OLD, AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.00/2.00/103. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185751&pc=MTA | 2023826-2015-01447 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/X067 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (26MMHG). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393718&pc=MTA | 2023826-2023-05723 | 2023-12-01 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.5/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ANGLE CLOSURE WITH ELEVATED IOP 60MMHG(ASSESSED ON (B)(6) 2023). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525773&pc=MTA | 2023826-2024-00171 | 2023-12-01 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.0/043 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ELEVATED IOP(32MMHG) WITH OUR PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON 3RD WEEK POST-OP), PIGMENT DISPERSION WERE OBSERVED. LENS REMAINS IMPLANTED. PATIENT GIVEN STEROIDS AND IOP MEDS. THIS HAS NOT RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16127557&pc=MTA | 2023826-2023-00115 | 2022-12-13 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. B5: "THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "INITIALLY SOME CORNEAL ODEMA AND ENDOTHELIAL CELL LOSS DUE TO TRICKY EXCHANGE (CONSTRICTION OF PUPIL DURING PROCEDURE) WILL COMPLETE ANOTHER COMPLAINT FORM BUT VISION EXCELLENT." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR." CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16128162&pc=MTA | 2023826-2023-00116 | 2022-12-13 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "SITUATION RESOLVED AND PATIENT HAS EXCELLENT VISION." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795682&pc=MTA | 2023826-2017-01320 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/159 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT IOP IS 20.6MMHG, AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN, AND DIAMOX. MULTIPLE ATTEMPS TO GET MORE INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT'S IOP IS 20.6MM HG AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465307&pc=MTA | 2023826-2018-00607 | 2018-02-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+3.0/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2018. TO DATE, THE PATIENT HAS NOT DECIDED WHETHER TO REPLACE WITH ANOTHER LENS OR NOT. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, THE PATIENT DEVELOPED RAISED IOP AFTER TICL IMPLANT, WHICH WAS "REFRACTORY TO TOPICAL ANTI-GLAUCOMA MEDICATIONS." THE TICL WAS EXPLANTED AND A TRABECULECTOMY PERFORMED. THE CORNEA IS CURRENTLY CLEAR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876795&pc=MTA | 2023826-2017-01460 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/142 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876796&pc=MTA | 2023826-2017-01459 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641431&pc=MTA | 2023826-2017-00976 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4725289&pc=MTA | 2023826-2015-00428 | NA | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE LENS ROTATED (NOT ASSOCIATED TO A LOW VAULT). THE INCISION WAS ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED AT THE FIVE MONTH VISIT, THE SURGEON CONFIRMED THE OFF-AXIS ALIGNMENT AND DECIDED TO REPOSITION THE LENS. THE PATIENT EXPERIENCED BLURRED VISION. AFTER THIS INTERVENTION, THE PATIENT EXPERIENCED AN IOP RISE, WHICH WAS SOLVED BY ANTERIOR CHAMBER ASPIRATION AND IRRIGATION. FOR THIS SECONDARY INTERVENTION A NEW INCISION WAS CREATED, LEADING TO INCREASED CORNEAL ASTIGMATISM. THE LENS REMAINS IMPLANTED AND IS STABLE AFTER REPOSITIONING. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS IMPLANTED METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS UNDER-CORRECTING THE CYLINDER AND SURGICAL INDUCED ASTIGMATISM SECONDARY TO A NEW INCISION. THE LENS WAS OFF-AXIS AT ONE DAY POSTOPERATIVELY THEREFORE THIS COULD BE DUE TO SURGICAL MISALIGNMENT OR EARLY MOVEMENTS OF THE LENS. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589772&pc=MTA | 2023826-2017-00878 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS. (B)(6). (B)(4). Manufacturer Narrative: WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6412058&pc=MTA | 2023826-2017-00259 | 2017-01-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON PARTIALLY INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/102 DIOPTER, BUT DUE TO INTRAOCULAR POSITIVE PRESSURE, THE SURGEON COULD NOT IMPLANT THE LENS COMPLETELY. THE REPORTER INDICATED LOOSE ZONULES WERE SUSPECTED AND THE ANTERIOR CHAMBER COLLAPSED. THE LENS WAS REMOVED WITHOUT ANY PATIENT INJURY OR COMPLICATIONS AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN THREE PIECES. THERE WAS RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8077834&pc=MTA | 2023826-2018-01630 | 2018-08-16 | Injury | STAAR SURGICAL COMPANY | 2018-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.0/X084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (13MMHG), WHICH CAUSED THE IRIS TO PROLAPSE (POSTERIOR SYNECHIAE) THROUGH THE WOUND. POST-OP, THE PATIENT STILL HAS ELEVATED IOP AND POSTERIOR SYNECHIAE. IT IS UNKNOWN IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THE SURGEON FELT THE MEASUREMENTS MAY HAVE BEEN INCORRECT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976037&pc=MTA | 2023826-2018-01513 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 28 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976051&pc=MTA | 2023826-2018-01512 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/125 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 38 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215221&pc=MTA | 2023826-2019-00021 | 2018-11-21 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 15.50/+3.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED AFTER 1 MONTH POST-OP, THE SURGEON NOTICED RAISED INTRAOCULAR PRESSURE (IOP), WITH OPEN ANGLE BUT SHALLOW CHAMBER. THERE WAS SUSPECTED STEROID RESPONSE, AND THE SURGEON STARTED ANTIGLAUCOMA. AFTER 2 WEEKS, THERE WAS NO IMPROVEMENT TO IOP, SO THE SURGEON PERFORMED LASER PI. THE IOP SHOT UP AFTER 1 WEEK TO 32MMHG AFTER LASER PI. STARTED ANTIGLAUCOMA, BOTH TOPICAL AND SYSTEMIC. AT PRESENT ONLY TOPICAL ANTIGLAUCOMA AND IOP IS 22MMHG. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2018 AND EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. A YAG WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS CURRENT CONDITION WAS ANGLE OPENED AND ANTERIOR CHAMBER WELL FORMED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8359916&pc=MTA | 2023826-2019-00291 | 2019-01-08 | Injury | STAAR SURGICAL COMPANY | 2019-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -9.00/1.5/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, IOP AND IRIDO-CORNEAL ANGLES RETURNED TO NORMAL. THE DOCTOR HAS NO PLANS TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223685&pc=MTA | 2023826-2015-01409 | 2015-07-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.50/+0.5/X094 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. UNREACTIVE PUPIL AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -4.50/0.50/X094. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756735&pc=MTA | 2023826-2016-00792 | 2016-01-21 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -9/+4/167 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED IOP. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE BUT PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891300&pc=MTA | 2023826-2016-00455 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). TWENTY-FOUR REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21 AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891288&pc=MTA | 2023826-2016-00453 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077547&pc=MTA | 2023826-2016-01562 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -18.0/./+20 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND PUPIL BLOCK (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: PREEXISTING MEDICAL CONDITION: KERATOCONUS. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH KERATOCONUS IS CONTRAINDICATED. CORRECTED DATA: CORRECT LENS DIOPTER IS -18.0/+2.0/X074 (SPHERE/CYLINDER/AXIS). (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6222127&pc=MTA | 2023826-2017-00012 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/096 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6420639&pc=MTA | 2023826-2017-00446 | 2016-12-13 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIA. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6482395&pc=MTA | 2023826-2017-00601 | 2016-12-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIAE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS IS TWISTED DUE TO POSITION IN VIAL. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260641&pc=MTA | 2023826-2017-00101 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/95 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED TO THE MANUFACTURER BUT NOT YET BEEN EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653556&pc=MTA | 2023826-2017-00998 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/072 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A VICMO LENS WITH THE SAME SIZE AND DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE AND DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711929&pc=MTA | 2023826-2017-01185 | 2017-03-11 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/098 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED INFORMATION: THE LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS ADVERSE EVENTS WERE NOT INCLUDED IN INITIAL MDR: PATIENT HAD ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6554053&pc=MTA | 2023826-2017-00781 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXPLANTED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: STATED LENS IS NOT RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION REVEALED LENS WAS WITHIN SPECIFICATIONS. METHOD, RESULT CODES: UPDATED. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587670&pc=MTA | 2023826-2017-00872 | 2017-03-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+5/X126 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, THE PUPIL CAN NOT CONTRACT, AND ELEVATED IOP. IN THE SAME DAY, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637203&pc=MTA | 2023826-2017-00964 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/88 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RED AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: PREVIOUS CODE NOT APPLICABLE. AS PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR IMPLANTATIONS IN PATIENTS WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637164&pc=MTA | 2023826-2017-00961 | 2017-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637205&pc=MTA | 2023826-2017-00965 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/88 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740871&pc=MTA | 2023826-2017-01240 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5/+3.0/162 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED AS AN EMERGENCY IN THE HOSPITAL. THE LENS WAS EXPLANTED BECAUSE OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF PATIENT'S IRIDO-CORNEAL ANGLES (ICA), AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850323&pc=MTA | 2023826-2017-01394 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT DATE: (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114677&pc=MTA | 2023826-2017-01841 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15.0/+3.0/079 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS CURVED AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114680&pc=MTA | 2023826-2017-01859 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -14.0/+1.5/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF THE HAPTIC AND THE PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955546&pc=MTA | 2023826-2017-01576 | 2017-09-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS WAS EXPLANTED DUE TO IOP SPIKE, EXCESSIVE VAULT, AND NARROW ANGLES. REPORTER STATES LENS IS PLANNED TO BE EXCHANGED WITH A SMALLER LENS; ORIGINAL LENS HAS BEEN LOST BY PHYSICIAN. ADDITIONAL INFORMATION IS NOT KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, ON (B)(6) 2017, A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5/+2.5/108 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS DAMAGED DURING EXPLANT. THE DOCTOR IS "WAITING FOR PUPIL TO CONSTRICT" BEFORE DECIDING ON FURTHER ACTION. THE DOCTOR SUSPECTS URRETS-ZAVALIA. PREVIOUS REPORT INCLUDED ELEVATED IOP; CURRENT REPORT FROM COMPLAINANT DOES NOT. CURRENT SYMPTOMS REPORTED: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE (FIXED) PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (FIXED PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA). (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871042&pc=MTA | 2023826-2018-01367 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: DEVICE WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION TOLERANCES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871107&pc=MTA | 2023826-2018-01368 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND DEFORMED. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO BROKEN HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978046&pc=MTA | 2023826-2017-01624 | 2017-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/090 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356130&pc=MTA | 2023826-2018-00437 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC TEAR AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356135&pc=MTA | 2023826-2018-00435 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884850&pc=MTA | 2023826-2018-01266 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/+1.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ASSESMENT ON (B)(6) 2018, ANGLE CLOSURE WITH ELEVATED IOP (26MMHG) AND EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621051&pc=MTA | 2023826-2018-00923 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXCHANGE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621071&pc=MTA | 2023826-2018-00921 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA; DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8275983&pc=MTA | 2023826-2019-00170 | 2018-04-08 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -10.00/2.0/85 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO "FALSE W-TO-W, MEASUREMENT PROVIDED." THE REPORTER STATED THAT THE PATIENTS EYE IS STABLE AND IN GOOD CONDITION. IT WAS REPORTED THAT THE DOCTOR PLANS TO IMPLANT ANOTHER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589676&pc=MTA | 2023826-2018-00879 | 2018-05-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AND WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7772418&pc=MTA | 2023826-2018-01180 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -12.0/1.00/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PREVIOUSLY STATED, "VISUAL INSPECTION FOUND THE HAPTIC TORN." CORRECTED STATEMENT: "VISUAL INSPECTION FOUND THE HAPTICS BENT." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903849&pc=MTA | 2023826-2018-01373 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 EXCESSIVE VAULT WAS OBSERVED, PATIENT WAS ASSESSED WITH ANGLE CLOSURE WITH ELEVATED IOP (50 MMHG) ON (B)(6) 2018 THE LENS WAS EXPLANTED. SURGEON PLANS TO IMPLANT ANOTHER ICL LENS IN (B)(6) 2018. EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "ADVERSE EVENT" IN THE INITIAL MDR. CORRECTED TO: "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)" IN THE INITIAL MDR. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765108&pc=MTA | 2023826-2018-01169 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON REPORTS VAULT VALUES OF 800UM/1.2CT, ELEVATED IOP, AND REFRACTIVE TREATMENT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, PUPIL BLOCK). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884841&pc=MTA | 2023826-2018-01392 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -09.5/6/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: "LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, SHALLOWING ACD, AND HIGHER IOP" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090043&pc=MTA | 2023826-2018-01696 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090042&pc=MTA | 2023826-2018-01699 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8277800&pc=MTA | 2023826-2019-00177 | 2018-08-25 | Injury | STAAR SURGICAL COMPANY | 2019-01-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). ADDITIONAL INFORMATION REPORTED: PATIENT CANCELLED THE SURGERY FOR BOTH EYES AFTER THE EXPLANTED OD LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE: CORRECTED FROM "(B)(6) 2018" TO (B)(6) 2018" FOR INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). PATIENT CANCELLED THE SURGERY AFTER THE EXPLANTED OD LENS AND HAS NOT YET SCHEDULED FOR ANY SURGERIES. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CAS/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON PRODUCT AND A BENT HAPTIC. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183420&pc=MTA | 2023826-2018-00045 | NA | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.0/2/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2 IMPLANTABLE COLLAMER LENS, OF -15.0/2/110 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018, DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4) Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183423&pc=MTA | 2023826-2018-00046 | NA | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -14/1/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, OF -14/1/101 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM NUMBER OF THE FILE IN THE PREVIOUS SUPPLEMENTAL IS CORRECTED FROM "CLAIM# (B)(4)" TO " CLAIM# (B)(4)". Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7627647&pc=MTA | 2023826-2018-00932 | NA | Injury | STAAR SURGICAL COMPANY | 2018-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/+3/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT SIZE LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE REPORTED STATED THAT THE PATIENT'S CONDITION IS VERY GOOD NOW. Manufacturer Narrative: CORRECTED DATA: COMPLAINT "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)". OUTCOME SHOULD HAVE BEEN MENTIONED IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EXCHANGE TOOK PLACE ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210642&pc=MTA | 2023826-2015-01434 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/X107 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01433 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00/+2.00/X107 (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210651&pc=MTA | 2023826-2015-01433 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/X064 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01434 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: --11.50/+2.00/X064. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086238&pc=MTA | 2023826-2015-01237 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -23.0/+1.0/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086239&pc=MTA | 2023826-2015-01238 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -22.5/+1.5/070 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4) ANGLE CLOSURE, LENS EXCHANGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736049&pc=MTA | 2023826-2015-00478 | 2015-01-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AS WELL AS ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOW BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. ANGLE CLOSURE. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES. THE LENS COULD NOT BE RE-MEASURED. THE LENS WAS RETURNED DRY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, EVALUATION OF THE RETURNED PRODUCT AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759443&pc=MTA | 2023826-2015-00522 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/100 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743860&pc=MTA | 2023826-2015-00499 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00498 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. LENS EXPLANTED. VAULTING. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A HAPTIC WAS TORN AND THE LENS WAS RETURNED DRY. THE WORK ORDER SEARCH SHOWED NO OTHER COMPLAINT RECORDED RELATED TO THE SAME WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106976&pc=MTA | 2023826-2015-01186 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/X101 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON(B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01185 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.5/+2.0/X101. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106991&pc=MTA | 2023826-2015-01185 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/X091 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01186 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.0/+4.0/X091. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405797&pc=MTA | 2023826-2016-00045 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -11.5/-4.50/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR. #2023826-2016-00046 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-11.50/+4.50/X090. EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178915&pc=MTA | 2023826-2015-01356 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -13.00/+1.0/X133 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.0/X133 (B)(4). METHOD CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919323&pc=MTA | 2023826-2015-00835 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS OBSERVED. THE PATIENT'S LEFT PUPIL IS LARGER THAN THE RIGHT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219006&pc=MTA | 2023826-2015-01410 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/X094 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONALS INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50/+0.50/X094. DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208843&pc=MTA | 2023826-2015-01429 | 2015-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.00/+1.50/X089 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01428 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S.,BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.00/1.50/X089. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206816&pc=MTA | 2023826-2015-01475 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/X082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206819&pc=MTA | 2023826-2015-01476 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/X075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS FIBERS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208884&pc=MTA | 2023826-2015-01428 | 2015-07-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.50/+2.5/X084 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01429 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.5/+2.5/X084. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154950&pc=MTA | 2023826-2015-01354 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536900&pc=MTA | 2023826-2016-00314 | 2015-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.0/+3.5/91 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+3.5/91. (B)(4). METHOD CODE (EVALUATION METHOD): WORK ORDER SEARCH RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536860&pc=MTA | 2023826-2016-00321 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+3.00/X80 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.5/+3.00/X80. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308661&pc=MTA | 2023826-2015-01659 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/X100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535797&pc=MTA | 2023826-2016-00320 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+2.50/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525093&pc=MTA | 2023826-2016-00329 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891295&pc=MTA | 2023826-2016-01131 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21. ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632888&pc=MTA | 2023826-2016-00535 | 2016-03-06 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT ONE WEEK POST-OP THE LENS HAD A HIGH VAULT WITH CONTROLLED INTRAOCULAR PRESSURE ELEVATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670128&pc=MTA | 2023826-2016-00643 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS LENS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670132&pc=MTA | 2023826-2016-00642 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955105&pc=MTA | 2023826-2016-01258 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810832&pc=MTA | 2023826-2016-00942 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.0/+1.0./X180 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. Manufacturer Narrative: CORRECTED INFORMATION: REPORT SOURCE: "USER FACILITY" TO " DISTRIBUTOR". (B)(4) . ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 WITH A SHORTER LENS AND THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OPERATION BEST-CORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047723&pc=MTA | 2023826-2015-01113 | NA | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/106 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, IRIS ATROPHY AND UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218062&pc=MTA | 2023826-2015-01498 | NA | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT DATE: UNK. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064630&pc=MTA | 2023826-2016-01552 | 2016-10-04 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/040 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, ELEVATED INTRAOCULAR PRESSURE (IOP), AND LENS DISLOCATION/SUBLUXATION. THE PATIENT HAD TO UNDERGO AN ENLARGEMENT OF PERIPHERAL IRIDOTOMY (PI) OR ADDITION OF NEW PI USING YAG. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891293&pc=MTA | 2023826-2016-00454 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. ELEVATED IOP NOT REPORTED FOR PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308634&pc=MTA | 2023826-2015-01687 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.50/+2.50/X076 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE) WAS NOTED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSION AT INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US, BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.5/+2.5/X076. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824382&pc=MTA | 2023826-2018-01253 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/063 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021353&pc=MTA | 2023826-2017-01672 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -15.5/1.0/105 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT EXPERIENCED RIGHT EYE PAIN BEFORE THE REPLACEMENT SURGERY. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872195&pc=MTA | 2023826-2017-01440 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -18.00/+3.20/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, EYE PAIN, SHALLOW AC DEPTH, AND HIGH EYE PRESSURE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP ICL EXCHANGE IS STABLE. THE STATUS OF THE EYE IS QUITE. THE PATIENT HAS NO COMPLAINT OF PAIN OR OTHER DISCOMFORTS. THE IOP IS ALSO WITHIN NORMAL LIMITS. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/28 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/28. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6962654&pc=MTA | 2023826-2017-01609 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED HIGH EYE PRESSURE AND PAIN. THE LENS WAS DISCARDED BY THE FACILITY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017, FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING WELL, THE VAULT IS SUFFICIENT AND THE INTRAOCULAR PRESSURE IS WITHIN NORMAL LIMITS. (B)(4). REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636103&pc=MTA | 2023826-2018-00953 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Known Impact Or Consequence To Patient; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED INTRAOCULAR PRESSURE INCREASE (ASSOCIATED WITH PAIN) DUE TO A HYPERVAULT OF THE ICL LEADING TO A NARROW CHAMBER ANGLE AND A FLATTENING OF THE ANTERIOR CHAMBER. PIGMENT DISPERSION AND IRIS ATROPHY DUE TO A HYPERVAULT OF THE ICL. NARROW CHAMBER ANGLE, FLATTENING OF ANTERIOR CHAMBER. PIGMENT DISPERSION, IRIS ATROPHY. DEVICE HISTORY RECORD REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4721367&pc=MTA | 2023826-2015-00429 | 2015-02-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Vomiting; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED SEVERE HEADACHE, VOMITING AND PAIN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. ANGLE CLOSURE/ELEVATED IOP IN THE PRESENCE OF A VICMO/VTICMO MODEL MAY OCCUR DUE TO: REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), BLOCKAGE OF THE PORT BY REMAINING VISCOELASTIC, OR BY AN INFLAMMATORY MEMBRANE, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5766297&pc=MTA | 2023826-2016-00871 | 2016-05-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/123 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE PUPIL (FIXED), GLARE/HALOS AND BLURRED VISION. THE SURGEON ENLARGED/ADDITION THE PERIPHERAL IRIDOTOMIES (YAG) AND ADDED A SUTURE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -1 AND THE PATIENT HAS PHOTOPHOBIA. Manufacturer Narrative: (B)(4): IMPROPER OR INCORRECT PROCEDURE OR METHOD - THE SURGEON REALIZED THAT THE IMPLANTED ICL MODEL REQUIRED PREOPERATIVE PERIPHERAL IRIDOTOMIES BUT HE FAILED TO DO IT DUE TO CONFUSION WITH THE ICL MODEL ORDERED. CORRECTION: FROM "13.7MM VTICMO13.7" TO "13.7MM VTICM13.7". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734555&pc=MTA | 2023826-2017-01130 | 2016-08-22 | Injury | Unknown Manufacturer | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.50/+3.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP WAS NOTICED. THE LENS WAS THEN REPOSITIONED. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/+3.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP WAS NOTICED. THE LENS WAS THEN REPOSITIONED. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PRIOR TO LENS EXCHANGE, BCVA WAS DECREASED TO 20/30 COMPARED TO BASELINE (20/20). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6425385&pc=MTA | 2023826-2017-00441 | 2017-02-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1/82 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND SIGNIFICANT REDUCTION OF ICA. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE AS THE PUPIL REMAINS FIXED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS HAPTIC TORN. THERE WAS EVIDENCE OF THICK STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. OFF-LABEL, CONTRAINDICATED USE. ACD < 3.00MM (2.89MM). (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686682&pc=MTA | 2023826-2016-00778 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/064 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248647&pc=MTA | 2023826-2015-01460 | 2015-06-23 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/-1.50/X155 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. THE LENS WAS REMOVED AND NO OTHER LENS WAS IMPLANTED. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/30 AND UCVA WAS 20/800. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS ALSO CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.00/-1.50/X155. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923030&pc=MTA | 2023826-2015-00827 | NA | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PUPIL HAS REMAINED DILATED SINCE SURGERY AND THE INTRAOCULAR PRESSURE SPIKED UP TO 40MM HG (PUPIL BLOCK). THE PATIENT IS PHOTOPHOBIC. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. EVALUATION CODES: RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540552&pc=MTA | 2023826-2017-00732 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED IOP. ON THE SAME DAY, IN A SECOND SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CUSTOMER STATED THE DR. THOUGHT THE ISSUE WAS CAUSED BY THE LENS SIZE, THOUGH HE DID NOT MEASURE THE VAULT. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR IN THE HAPTIC. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7219339&pc=MTA | 2023826-2018-00142 | 2017-12-27 | Injury | STAAR SURGICAL COMPANY | 2018-01-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PUPIL BLOCK, PIGMENT DISPERSION AND UNREACTIVE PUPIL (FIXED). THE REPORTER INDICATED THERE IS SUSPECTED URRETS-ZAVALIA DUE TO TEMPORARY AQUAPORT BLOCK AND INTRAOCULAR PRESSURE SPIKE. AT LAST VISIT ON (B)(6) 2017, THE INTRAOCULAR PRESSURE WAS 10 MMHG, DILATED PUPIL NOT VERY REACTIVE AND SOME PIGMENT ON THE LENS. THE LENS REMAINS IMPLANTED. THE PATIENT WAS REPORTED AS BEING WELL AND WITHOUT PRESSURE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7236056&pc=MTA | 2023826-2018-00143 | NA | Injury | STAAR SURGICAL COMPANY | 2018-02-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED SIMILAR SYMPTOMS AS THE RIGHT EYE BUT NOT AS STRONG - ELEVATED INTRAOCULAR PRESSURE, PUPIL BLOCK, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED) AND IRIS ATROPHY. THE REPORTER INDICATED THERE IS SUSPECTED URRETS-ZAVALIA DUE TO TEMPORARY AQUAPORT BLOCK AND INTRAOCULAR PRESSURE SPIKE. THE LENS REMAINS IMPLANTED. THE PATIENT WAS REPORTED AS BEING WELL AND WITHOUT PRESSURE. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031008&pc=MTA | 2023826-2016-01435 | 2016-06-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13/1/93 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ELEVATED IOP, PUPIL BLOCK (WITH ELEVATED IOP) AND ANGLE CLOSURE (WITH ELEVATED IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/BROKEN/DEFORMED. UNIDENTIFIED FOREIGN MATERIAL WAS PRESENT ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453289&pc=MTA | 2023826-2017-00532 | 2016-07-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK, AND ANGLE CLOSURE (WITH ELEVATED IOP). ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498926&pc=MTA | 2023826-2017-00646 | 2016-07-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/1/81 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK, AND ANGLE CLOSURE (WITH ELEVATED IOP). ON (B)(6) 2016 THE LENS WAS EXPLANTED, AND THEN EXCHANGED FOR SHORTER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FOREIGN MATERIAL PRESENT ON THE LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ANGLE CLOSURE). RESULT CODE: WORK ORDER SEARCH WAS PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6327728&pc=MTA | 2023826-2017-00122 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/91 DIOPTER, IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, WHICH LEAD TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK (WITH ELEVATED IOP), AND ANGLE CLOSURE (WITH ELEVATED IOP) IN PATIENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016, AND THE PROBLEM WAS RESOLVED. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6582998&pc=MTA | 2023826-2017-00864 | 2017-04-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -8.50/+1.0/014 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE).(B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670126&pc=MTA | 2023826-2017-01042 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.0/164 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED IOP. POST-OP FIRST DAY, THE SURGEON NOTED THE LENS PUSHED THE IRIS BACKWARD TO THE CRYSTALLINE AND IRIS HAS TOUCHED THE CRYSTALLINE LENS. HE PERFORMED THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS TWICE TO SOLVED THE PROBLEM BUT NOTHING HAS CHANGED. ON (B)(6) 2017 THE LENS WAS EXPLANTED. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PUPIL FIXED DILATED POSTERIOR SYNECHIA, AND LESS PUPILLARY MOVEMENT. ON (B)(6) 2017, THE CUSTOMER STATED THE PATIENT'S EYE BACK TO NORMAL AS BEFORE, AND THE DR. IS PLANNING TO IMPLANT ANOTHER LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673332&pc=MTA | 2023826-2017-01061 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. THE SURGEON STATED THAT THE PATIENT ALSO HAS "CORTICAL PERIPHERIC ANTERIOR CATARACT. FREE VISUAL AXIS." THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. THE STATUS OF THE EYE IS PUPIL SPHINCTER PARALYSIS. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4). Manufacturer Narrative: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7076607&pc=MTA | 2023826-2017-01845 | 2017-11-13 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS REPOSITIONED AND THE PERIPHERAL IRIDECTOMIES WERE ENLARGED/ ADDITIONAL PERIPHERAL IRIDECTOMIES BY YAG. THE SURGEON ALSO DILATED THE PUPIL AND EVACUATED/IRRIGATED THE ANTERIOR CHAMBER OF ANY REMAINING VISCO/FLUIDS. THE LENS REMAINS IMPLANTED BUT THE SURGEON IS CONSIDERING EXPLANT AND EXCHANGE. AT LAST POST-OP VISIT ON (B)(6) 2017, THE PATIENT'S UCVA WAS 20/25++, INTRAOCULAR PRESSURE (IOP) WAS 12 AND THE ICL WAS COVERED WITH IRIS PIGMENT. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 AND WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP CONDITION BEFORE THE LENS WAS EXPLANTED WAS HIGH VAULT - 3 CCT AND INTRAOCULAR PRESSURE OF 19. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SMALL PLASTIC BAG. VISUAL INSPECTION FOUND A PIECE OF ONE HAPTIC BROKEN OFF. A DIMENSIONAL INSPECTION OF THE LENS WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786337&pc=MTA | 2023826-2015-00593 | 2014-05-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND NOTICED PUPILLARY BLOCK WITH ELEVATED IOP IN THE POST-OPERATIVE PERIOD. THIS WAS ASSOCIATED WITH A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE ICL VAULTING WAS CONSIDERED TO BE EXCESSIVE. THREE WEEKS LATER THE SURGEON DECIDED TO REMOVE AND REPLACE THE LENS WAS A SHORTER ONE AND THIS RESOLVED THE PROBLEM. AFTER THE EXCHANGE THE PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH NO VISIBLE DAMAGE. THE LENS WAS REHYDRATED AND BOTH THE LENGTH AND WIDTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - PUPIL BLOCK WITH ELEVATED IOP IN THE PRESENCE OF HIGH VAULT EARLY AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATIENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCO), REMAINING VISCOELASTIC IN THE ANTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL L ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MAL-POSITIONED FOOTPLATES, ECT. THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THIS EVENT, HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ECT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933090&pc=MTA | 2023826-2015-00912 | 2014-11-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2014. PUPILLARY BLOCK WITH ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXPLANTED AND THE SURGEON PLANS ON IMPLANTING A SHORTER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937377&pc=MTA | 2023826-2015-00902 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PI WAS EITHER ENLARGED OR THERE WAS AN ADDITION OF A NEW PI. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE MEDICAL REVIEW STATED THE SAME MODEL LENS WAS IMPLANTED IN BOTH EYES OF THIS PATIENT BY THE SAME DOCTOR WITH PREOPERATIVE ACD (ENDO) OF 2.64MM IN OD AND 2.71MM IN OS AND HISTORY OF KERATOCONUS. THE ACD IS OUTSIDE THE RECOMMENDED RANGE OF ICL IMPLANTATION WHICH IS GREATER OR EQUAL TO 3MM. AS REPORTED ON THE MEDICAL REVIEW THE MOST LIKELY CAUSE OF THE EVENT IS TOO SHALLOW EYES (OUTSIDE APPROVED INDICATION). ICL VAULTING MAY HAVE A CONTRIBUTING ROLE IN THIS CASE. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937426&pc=MTA | 2023826-2015-00903 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PI WAS EITHER ENLARGED OR THERE WAS AN ADDITION OF A NEW PI. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/120. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE MEDICAL REVIEW STATED THE SAME MODEL LENS WAS IMPLANTED IN BOTH EYES OF THIS PATIENT BY THE SAME DOCTOR WITH PREOPERATIVE ACD (ENDO) OF 2.64MM IN OD AND 2.71MM IN OS AND HISTORY OF KERATOCONUS. THE ACD IS OUTSIDE THE RECOMMENDED RANGE OF ICL IMPLANTATION WHICH IS GREATER OR EQUAL TO 3MM. AS REPORTED ON THE MEDICAL REVIEW THE MOST LIKELY CAUSE OF THE EVENT IS TOO SHALLOW EYES (OUTSIDE APPROVED INDICATION). ICL VAULTING MAY HAVE A CONTRIBUTING ROLE IN THIS CASE. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5323196&pc=MTA | 2023826-2015-01707 | 2015-07-10 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM TICM130V4 IMPLANTABLE COLLAMER LENS, -16.0/+4.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED IRIS ATROPHY. THE PERIPHERAL IRIDOTOMIES WERE SURGICAL ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223686&pc=MTA | 2023826-2015-01408 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -4.00/+2.00/X096 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. A YAG AND ENLARGEMENT OF PI WAS PERFORMED. THE LENS WAS EXPLANTED AND WAS EXCHANGED FOR A SHORTER LENS. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015; UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -4.00/+2.00/X096. EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632893&pc=MTA | 2023826-2016-00536 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER, ANGLE CLOSURE, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT 4 DAYS POST-OP THE LENS HAD A HIGH VAULT WITH NON-CONTROLLED INTRAOCULAR PRESSURE ELEVATION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (UNSPECIFIED). PATIENT CODE - (B)(4) ANGLE CLOSURE, SHALLOW ANTERIOR CHAMBER. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685192&pc=MTA | 2023826-2016-00669 | NA | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.50/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND REMAINS IMPLANTED. THE LENS IS PLANNED TO EXCHANGE WITH ANOTHER LENS. THE PATIENT EXPERIENCED PUPIL BLOCK (WITH ELEVATED IOP). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT AGE AND WEIGHT: UNK. COMMON DEVICE NAME: UNK. EVENT DATE: UNK. EXPLANT DATE: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487402&pc=MTA | 2023826-2017-00630 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+4/101 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, PUPIL BLOCK WITH ELEVATED IOP, AND AN IRIS THAT WILL NOT DECREASE IN SIZE. THE LENS WAS REPOSITIONED. THE SURGEON NOTES THAT AN EXCHANGE IS PLANNED, REPOSITIONING DID NOT WORK. TO DATE, UNABLE TO OBTAIN INFORMATION REGARDING IF/WHEN AN EXCHANGE OCCURRED. THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5750966&pc=MTA | 2023826-2016-00862 | 2016-05-03 | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE AND THE SURGERY WAS UNEVENTFUL. SOON AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED SEVERE PUPIL BLOCK DESPITE TWO ADEQUATE IRIDOTOMIES. INITIALLY, RETAINED VISCOELASTIC WAS THE CAUSE BECAUSE DILATING THE PUPIL DID NOT COMPLETELY RESOLVE THE PUPIL BLOCK. AFTER A FEW DAYS, THE DILATING DID RESOLVE IT, BUT EACH TIME THE PUPIL CAME DOWN, THE PATIENT HAD A PUPIL BLOCK. THE EVENT WAS RESOLVED BY GOING BACK TO THE OR AND PERFORMING A SURGICAL IRIDECTOMY. THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS REQUESTED THE LENS BE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+3.0/100 DIOPTER, WAS IMPLANTED IN THE LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DATE OF BIRTH - UNK. EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577252&pc=MTA | 2023826-2017-00848 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; No Known Impact Or Consequence To Patient; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED PROBLEMS WITH A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18/+2.0/X082 DIOPTER. HOWEVER, WE RECEIVED CONFLICTING INFORMATION REGARDING THIS LENS. REPORTER LISTS BOTH EYES AS BEING AFFECTED, STATES THAT THE LENS WAS NEITHER IMPLANTED NOR EXPLANTED DUE TO HIGH VAULT, PUPIL BLOCK, AND ELEVATED IOP. SURGERY DATE IS LISTED AS (B)(6) 2016, HOWEVER, LENS WAS NOT SHIP TO THE CUSTOMER UNTIL (B)(6) 2016. ATTEMPTS TO CLARIFY HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THE REPLACEMENT LENS WAS SUBMITTED IN THE INITIAL MDR, INCORRECTLY. THERE ARE NO ADVERSE EVENTS ASSOCIATED WITH THE REPLACEMENT LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037544&pc=MTA | 2023826-2016-01480 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS 16.50/25/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/2.5/74 DIOPER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. PATIENT ALSO EXPERIENCED DIPLOPIA IN LOW LIGHT. THE LENS WAS EXCHANGED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DATE OF EVENT LISTED ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2016. DATE RETURNED TO MANUFACTURER ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037545&pc=MTA | 2023826-2016-01481 | 2016-06-11 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS - -16.00/25/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). METHOD: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.00/2.5/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARES/HALOES. DIPLOPLIA WAS ALSO EXPERIENCED IN LOW LIGHT. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254868&pc=MTA | 2023826-2019-00062 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, GLARE/HALOS, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE OF LENS BECAUSE OF RIGHT EYE ANATOMY". Manufacturer Narrative: THE LENS WAS RETURNED MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5985749&pc=MTA | 2023826-2016-01347 | 2016-05-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7409133&pc=MTA | 2023826-2018-00523 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6015356&pc=MTA | 2023826-2016-01434 | 2016-08-30 | Injury | STAAR SURGICAL COMPANY | 2016-10-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11/+6/10 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Manufacturer Narrative: LENS DESCRIPTION IS -11.0/+6.0/140 (SPHERE/CYLINDER/AXIS). VAULT VALUE PRIOR TO EXPLANT WAS REPORTED AS 1500 UM. THE LENS WAS EXCHANGED FOR A SHORTER SAME MODEL LENS WITH POST-OP VAULT VALUE OF 1 MM AND NO REPORTED PATIENT ISSUES. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7629830&pc=MTA | 2023826-2018-00945 | NA | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/123 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED OCULAR PAIN IN DARKNESS, WITH IOP SPIKES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867141&pc=MTA | 2023826-2017-01464 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Nausea; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. INVESTIGATION CONFIRMED THE LABELLING ISSUE. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THAT THE PATIENT COULD SEE "NOTHING WITH OD AND CANNOT MOVE DUE TO NAUSEA WITH HIS ASTIGMATISM." THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721248&pc=MTA | 2023826-2018-01078 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Pain; Blurred Vision; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, BLURRED VISION, PAIN, AND "NO COMFORT FOR NEAR AND MIDDLE VISION." THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403798&pc=MTA | 2023826-2019-00367 | 2019-02-04 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2017. THE REPORTER INDICATED THAT AN ACCIDENT TO THE PATIENTS EYE HAS RESULTED IN BLUR VISION WHEN READING OR CONCENTRATING, AND PAIN WITH LIGHT CHANGES AND WHILE LOOKING AT A MOVING OBJECT. REPORTEDLY, THE PATIENT HAD A VISUAL CHECK WITH THE SURGEON AND THE VISUAL CHECK WAS "PERFECT." THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2017. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248831&pc=MTA | 2023826-2015-01539 | 2015-05-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -04.00/+4.50/X114 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT STARTED TO HAVE SEVERE PAIN 15 MINUTES POST-OP. THE LENS WAS EXPLANTED THAT SAME DAY. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PEN MARKS ON LENS SURFACE. (B)(4). CONCLUSION: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. PATIENT'S ACD IS 2.97 MM. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -04.00/+4.50/X114. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7261493&pc=MTA | 2023826-2018-00173 | 2018-01-10 | Injury | STAAR SURGICAL COMPANY | 2018-02-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Pain; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.11/+0.33/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE PATIENT BEGAN TO COMPLAIN OF OCULAR PAIN. REPORTEDLY, ALSO NOTED IS EXCESSIVE VAULT AND THAT THE LENS IS "COMPROMISING THE ANGLES." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: LENS WAS EXCHANGED AS THE VAULT WAS EXCESSIVE. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown