## Lens-related issue-1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code...36 | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241293&pc=QCB | 2023826-2021-02893 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS PREPARING TO IMPLANT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND THE LENS CAPSULE WAS NICKED DURING THE PRIMARY INCISION POINT. THE ICL LENS WAS IMPLANTED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292528&pc=QCB | 2023826-2023-02649 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/176 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE ANTERIOR CAPSULE RUPTURED AT THE TIME THE MAIN INCISION WAS MADE. PHACOEMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PHACOEMULSIFICATION (CATARACT SURGERY), IOL IMPLANTED. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10247584&pc=QCB | 2023826-2020-01427 | 2020-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/073 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). WHEN THE LENS WAS REMOVED, THE LENS CONTACTED THE CRYSTALLINE LENS AND A MILD ANTERIOR SUBCAPSULAR OPACITY DEVELOPED. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS EYE MOVED WHEN THE SURGEON INSERTED THE INITIAL ICL. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR/UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014526&pc=QCB | 2023826-2020-02985 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. ANOTHER ICL LENS HAS NOT BEEN IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10990261&pc=QCB | 2023826-2020-02963 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LOW VAULT AND LENS OPACITY ASC (OBSERVED: (B)(6) 2020). ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: B5- PER NEW INFORMATION NO INTERVENTION OR TREATMENT WAS PERFORMED FOR LENS OPACITY. IT WAS REPORTED THAT THE LENS OPACITY WAS NOT PRODUCT RELATED, THE REPORTER STATED THAT THE DOCTOR WAS NOT SURE HOW THE LENS OPACITY CAME ABOUT. IT MAYBE RELATED TO HIS MANIPULATION DURING SURGERY. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT, AND DEFORMED WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894548&pc=QCB | 2023826-2021-01709 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER STATES A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -12.0/+1.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO A POSTERIOR SUBCAPSULAR CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. MODERATE CORNEAL EDEMA AND MODERATE ANTERIOR CHAMBER CELL ARE REPORTED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "WHILE THE MECHANISM OF CATARACT FORMATION REMAINS SPECULATIVE, THE TIME COURSE FOR THE DEVELOPMENT OF THE CATARACT AND THE PRESENCE OF ANTERIOR CAPSULE FINDINGS IN THIS CASE STRONGLY SUGGEST AN ASSOCIATION BETWEEN THE LASER PERIPHERAL IRIDOTOMY AND THE POSTERIOR SUBCAPSULAR CATARACT." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: RESOLUTION DATE WAS 07-JUN-2021. ALL SYMPTOMS RETURNED TO NORMAL. (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986324&pc=QCB | 2023826-2021-01855 | 2021-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC. PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. THE SURGEON REPORTS THIS WAS AN ICL-INDUCED CATARACT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: LENS OPACITY-ASC (ANTERIOR SUBCAPSULAR), REPORTED IN INITIAL MDR IS CORRECTED TO LENS OPACITY-PSC (POSTERIOR SUBCAPSULAR). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021151&pc=QCB | 2023826-2021-01906 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/1.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC). IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4) OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064338&pc=QCB | 2023826-2021-02028 | 2021-05-26 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM, TMICL13.2, -10.0/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED THAT PATIENT IS HAPPY WITH VISION AND THE SURGEON IS NOT TREATING IT AT THE TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872840&pc=QCB | 2023826-2021-04560 | 2021-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.5/1.5/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO SUBCAPSULAR CATARACT (ASC) WHICH WAS OBSERVED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G3: INITIAL MDR DATE SHOULD BE CORRECTED TO 11-NOV-2021. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12957841&pc=QCB | 2023826-2021-04691 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.00/1.0/129 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ANTERIOR SUBCAPSULAR CATARACT GRADE 1 (ASC). LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082661&pc=QCB | 2023826-2021-04920 | NA | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED AT ONE WEEK POST-OP, A SMALL ANTERIOR CATARACT WAS OBSERVED. THE PATIENT IS 20/15 OU AND SO FAR THE CATARACT IS A NON ISSUE. THE LENS REMAINED IMPLANTED AND THE PATIENT WILL BE MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED AT ONE WEEK POST-OP, A SMALL ANTERIOR CATARACT WAS OBSERVED. THE PATIENT IS 20/15 OU AND SO FAR THE CATARACT IS A NON ISSUE. THE LENS REMAINED IMPLANTED AND THE PATIENT WILL BE MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK; WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; DATE OF EVENT: UNK; EXPLANT DATE: UNK. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: SEX: UNK; WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; DATE OF EVENT: UNK; EXPLANT DATE: UNK. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14197863&pc=QCB | 2023826-2022-01310 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM12.0; -10.0/1.5/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. THE SURGEON REPORTS OPACITY - NS CATARACT WHICH WAS OBSERVED ON (B)(6) 2021. ALSO ON THIS DATE THE LENS WAS EXPLANTED AND REPLACED WITH ANOTHER IOL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506297&pc=QCB | 2023826-2022-03329 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/016 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO A CATARACT DEVELOPED. THE LENS WAS REPLACED WITH AN IOL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PHACOEMULSIFICATION (CATARACT SURGERY), IOL IMPLANTED. B3: DATE OF EVENT (B)(6) 2022. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/016 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND THE DEVELOPMENT OF A CATARACT (NUCLEAR SCLEROTIC), OBSERVED ON (B)(6) 2022. PHACOEMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. D6B: (B)(6) 2022. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15679254&pc=QCB | 2023826-2022-03845 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -12.5/2.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2022 LENS OPACITY (ASC) WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2022 CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. PROBLEM RESOLVED. IN THE REPORTER OPINION THE LENS DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "PATIENT DEVELOPED EARLY CATARACT". Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15679288&pc=QCB | 2023826-2022-03846 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -12.0/3.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 LENS OPACITY (ASC) WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2022 CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. PROBLEM RESOLVED. IN THE REPORTER OPINION THE LENS DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED " PATIENT DEVELOPED EARLY CATARACT ". Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018250&pc=QCB | 2023826-2022-04428 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 . LOW VAULT WITH LENS OPACITY (CATARACT) ASC WAS REPORTED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY BUT DIFFERENT SURGERY A REPLACEMENT LENS OF DIFFERENT MODEL AND LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE; USER ERROR AND UNKNOWN. THE DEVICE FAILED TO PERFORM AS INTENDED. "EXCESSIVE LOW ICL VAULT MOST LIKELY CONTRIBUTED TO LENS OPACITY FORMATION." SEE MFR # 2023826-2022-04656 FOR EXCHANGE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITH LENS OPACITY (CATARACT) ASC WAS REPORTED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY BUT DIFFERENT SURGERY A REPLACEMENT LENS OF DIFFERENT MODEL AND LONGER LENGTH WAS IMPLANTED. REPORTEDLY THE EXCHANGE LENS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16345841&pc=QCB | 2023826-2023-00483 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00/3.0/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LENS OPACITY (CATARACT); ANTERIOR SUBCAPSULAR CATARACTS (ASC). PHACOEMULSIFICATION WAS DONE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "LENS TOUCH BETWEEN ICL AND ANTERIOR CAPSULE." Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16526528&pc=QCB | 2023826-2023-00873 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6 -11.0/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT LENS OPACITY (ASC) WAS OBSERVED ON THE SAME DATE OF SURGERY, THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY AND THE PROBLEM RESOLVED. IN THE REPORTER OPINION USER ERROR WAS THE CAUSE OF THE EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE DOCTOR MADE AN INCISION THAT DAMAGED THE PATIENT'S ANTERIOR CRYSTALLINE CAPSULE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16541830&pc=QCB | 2023826-2023-00855 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND AN ANTERIOR SUBCAPSULAR CATARACT (ASC) HAD DEVELOPED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545625&pc=QCB | 2023826-2023-00856 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND AN ANTERIOR SUBCAPSULAR CATARACT (ASC) HAD DEVELOPED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFORMATION: UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16569934&pc=QCB | 2023826-2023-00972 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/0.5/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING; AND LENS OPACITY (ASC) THAT WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF A DIFFERENT MODEL AND LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - (B)(4): OFF-LABEL USE (ACD < 3.0MM) (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694265&pc=QCB | 2023826-2023-01299 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -12.0/+1.5/17 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT EXPERIENCED SUBCAPSULAR OPACITY. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "HIPERLAXITY" AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16770068&pc=QCB | 2023826-2023-01385 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13,2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791232&pc=QCB | 2023826-2023-01444 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.1 -8.0/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791528&pc=QCB | 2023826-2023-01445 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6 -7.0/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON 13-FEB-2023 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989892&pc=QCB | 2023826-2023-01992 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1 -10.50/1.00/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY (CORTICAL) WAS OBSERVED ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT CATARACT SURGERY WAS ALSO PERFORMED. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: A1: (B)(6). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17226939&pc=QCB | 2023826-2023-02566 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/052 (SPHERE / CYLINDER / AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY (OBSERVED: ON (B)(6) 2023) AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT AN IOL WAS IMPLANTED AND PATIENT IS HAPPY. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17263291&pc=QCB | 2023826-2023-02664 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, ASC LENS OPACITY WAS OBSERVED AND THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17308923&pc=QCB | 2023826-2023-02762 | NA | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/0.5/97 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LENS OPACITY, ASC THAT WAS OBSERVED IN (B)(6) 2023. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "SURGEON TOLD: LENS OPACITY WAS NOT PRODUCT RELATED." THE PROBLEM WAS RESOLVED. THE REPORTER ALSO STATED, "THE SURGEON EXPLANTED THE ICL AND IMPLANTED A MONFOCAL IOL ALL FINE! PATIENT IS HAPPY." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342578&pc=QCB | 2023826-2023-02887 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/019 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO AN ANTERIOR SUBCAPSULAR (ASC) OPACITY (CATARACT) WAS NOTED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17595282&pc=QCB | 2023826-2023-03530 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT CATARACTS WAS OBSERVED IN PATIENT. LENS WAS IMPLANTED IN 2021. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS OPACITY (CATARACT) ASC WAS OBSERVED ON (B)(6) 2023. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS TRAILING A SCL FOR THE REFRACTIVE SHIFT AND WILL REFER CATARACT SURGERY FOR NOW. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTION: G2: REPORTED AS FOREIGN - CORRECT TO REMOVE FOREIGN. CLAIM #(B)(4). Manufacturer Narrative: A1 - UNK A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, H4 - UNK. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007848&pc=QCB | 2023826-2023-04808 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+2.50/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS OPACITY (CORTICAL) ASSESSED ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A CATARACT LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18268802&pc=QCB | 2023826-2023-05293 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/131 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS EXPLANTED DUE TO THE DEVELOPMENT OF AN ANTERIOR CHAMBER CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18419826&pc=QCB | 2023826-2023-05835 | 2023-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -0.00/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO LENS OPACITY. PROBLEM RESOLVED. IT WAS REPORTED THAT CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED AND THE PATIENT IS HAPPY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960156&pc=QCB | 2023826-2020-00852 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, THE PATIENT DEVELOPED PSC CATARACTS ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY (B)(6) 2019. THE CATARACT REAPPEARED IN (B)(6) 2019. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.6MM TMICL 12.6 OF -7.50/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT DEVELOPED PSC CATARACT ON (B)(6) 2019, THAT APPEARED TO RESOLVE BY SLIT LAMP AND IMAGING BY (B)(6) 2019. THE CATARACT RE-APPEARED IN (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020, AND THIS RESOLVED THE PROBLEM. STATUS OF EYE: S/P CATARACT EXTRACTION WITH PCIOL IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B3- DATE OF EVENT: (B)(6) 2020 B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT DEVELOPED PSC CATARACT ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY SLIT LAMP AND IMAGING BY (B)(6)2019. THE CATARACT RE-APPEARED IN (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 AND THIS RESOLVED THE PROBLEM. STATUS OF EYE: S/P CATARACT EXTRACTION WITH PCIOL IMPLANTATION. D6A- (B)(6) 2019, D6B- (B)(6) 2020. H6- CLINICAL CODE 4581- "TRANSIENT LOSS OF VISION". TYPE OF INVESTIGATION 4110- LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4) Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437079&pc=QCB | 2023826-2021-03475 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/044 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE AND THE PATIENT DEVELOPED A CATARACT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) -2021. THE PATIENT HAD BLURRED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS DEBRIS ON LENS. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM #:(B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840495&pc=QCB | 2023826-2021-04372 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS.THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN AND BENT AND FIBERS WERE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146115&pc=QCB | 2023826-2021-04373 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15578068&pc=QCB | 2023826-2022-03617 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -3.0/1.50/064 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT INDICATED THE PATIENT WILL BE MEASURED AGAIN IN A FEW MONTHS; THE INITIAL REFRACTIVE MUST BE RE-DETERMINED THEN A DECISION ABOUT A NEW IMPLANTATION WILL OCCUR. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. A CATARACT WAS OBSERVED. A IOL WAS IMPLANTED AND THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND A CATARACT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND AN IOL WAS IMPLANTED. THIS RESOLVED THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018251&pc=QCB | 2023826-2022-04429 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE; USER ERROR AND UNKNOWN. THE DEVICE FAILED TO PERFORM AS INTENDED. "EXCESSIVE LOW ICL VAULT MOST LIKELY CONTRIBUTED TO LENS OPACITY PROGRESSION." SEE MFR# 2023826-2022-04657. Manufacturer Narrative: ADDITIONAL INFORMATION: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. REPORTEDLY THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547301&pc=QCB | 2023826-2023-03523 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/051 9SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/051 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM #(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741313&pc=QCB | 2023826-2023-04041 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.00/+0.50/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN AND THE PATIENT HAD BLURRY VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPOSITIONED AND IMPLANTED THE CORRECT WAY. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND DEVICE CAUSALITY IS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.00/0.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A CORTICAL LENS OPACITY AND BLURRED VISION. ON (B)(6) 2023, WAS EXPLANTED AND REPOSITIONED THE CORRECT WAY. THE PROBLEM WAS NOT RESOLVED, "REFRACTIVE LENS EXCHANGE/CATARACT SURGERY IS REQUIRED AS PATIENT IS +6 AND HAS A LENS OPACITY" AND " I WILL INFORM YOU OF THE RLE SURGERY DATE AND OUTCOME". STATUS OF THE EYE: "VERY BLURRY REFRACTION". THE CAUSE OF THE EVENT IS USER ERROR. CAUSE DETAILS: "LENS TWISTED ON IMPLANTATION. AIR BUBBLE CONFUSED POSITIONING HOLE." CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019193&pc=QCB | 2023826-2022-04657 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS. IT IS REPORTED THAT "THE ICL HAS LOW VAULT WITH SIGNIFICANT INFERIOR DECENTRATION AND CCW ROTATION. DESPITE REPLACING ICL WITH ONE STEP LARGER SIZED ICL THE VAULT REMAINS SHALLOW AND INFERIOR DECENTRATION SUPPORTS UNDERSIZING OF REPLACEMENT ICL DESPITE INCREASING SIZE ONE STEP." THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2022-04429 FOR INITIAL LENS. Manufacturer Narrative: CLINICAL CODE: 4581: SIGNIFICANT INFERIOR DECENTRATION. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: RECEIVED INFORMATION THAT THIS EXCHANGE LENS DID RESOLVE THE ISSUE OF LOW VAULT. THIS CLAIM IS NOT A COMPLAINT DUE TO NO PRODUCT DEFICIENCY AGAINST THE LENS. PLEASE DISREGARD THE INITIAL MDR. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16688954&pc=QCB | 2023826-2023-01202 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/2.5/84(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASSESSED: (B)(6) 2022), PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT ON (B)(6) 2022 PILOCARPINE WAS USED FOR AROUND 1 MONTH, PATIENT PUPIL STILL UNREACTIVE AND STOP USING PILOCARPINE THEN. CATARACT SURGERY IS PLANNED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: UNREACTIVE/FIXED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678146&pc=QCB | 2023826-2023-01164 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EVENT DATE:UNK. EXPIRATION DATE: UNK. EXPLANT DATE: UNK. MFR DATE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16692052&pc=QCB | 2023826-2023-01165 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: A4: WEIGHT UNK. A5: ETHNICITY UNK. A6: RACE UNK. B3:EVENT DATE UNK. D4: EXPIRATION DATE UNK. D6B: EXPLANT UNK. H4: DEVICE MANUFACTURE DATE UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # 432841 | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17610615&pc=QCB | 2023826-2023-03705 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS OPACITY (CATARACT) ASC WAS OBSERVED ON (B)(6) 2023. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS TRAILING A SCL FOR THE REFRACTIVE SHIFT AND WILL REFER CATARACT SURGERY FOR NOW. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" AND "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED. "4614 " SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED. "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED AN ASC LENS OPACITY OBSERVED ON 17/MAY/2023, REFRACTIVE CHANGE OVER TIME, AND A SPHERICAL CONTACT LENS PRESCRIBED. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE:" PATIENT IS TRIALING SCL (SPHERICAL CONTACT LENS) OS FOR THE REFRACTIVE SHIFT; WILL DEFER CATARACT SURGERY FOR NOW." CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611928&pc=QCB | 2023826-2019-00824 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -5.50/2.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (ASC), SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS, AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. AS OF (B)(6) 2019 PATIENT'S "VISION IS GREAT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE: 114 SHOULD BE CORRECTED TO 213 IN INITIAL MEDWATCH REPORT. PATIENT CODE: 1766 SHOULD BE INCLUDED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: 04/17/2019 SHOULD BE CORRECTED TO 04/16/2019 IN INITIAL MEDWATCH REPORT. DEVICE CODE: 1494 OFF-LABEL USE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611958&pc=QCB | 2023826-2019-00791 | 2019-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: "REMOVED ON 10-MAR-2019" SHOULD BE CORRECTED TO "REMOVED ON (B)(6) 2019" IN INITIAL MDR. "10-MAR-2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936253&pc=QCB | 2023826-2020-00745 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -8.0/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY PSC CORTICAL. TRANSIENT LOSS OF BCVA IS ALSO REPORTED. REPORTEDLY, THE PATIENT DEVELOPED PSC CATARACTS ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY (B)(6) 2019. THE CATARACT REAPPEARED IN (B)(6) 2019 AND BCVA WAS LIMITED TO 20/50. Manufacturer Narrative: H6 - METHOD CODE 3331. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND A CIRCULAR MARK ON THE OPTIC. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE PROBLEM WAS RESOLVED WITH S/P CATATARCT EXTRACTION AND A PCIOL IMPLANTATION." THIS STATEMENT SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4-DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. D10- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 13-MAR-2020. G4- DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476134&pc=QCB | 2023826-2020-01935 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. SEE MFR REPORT # 2023826-2020-01934 FOR INITIAL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS INSPECTED UNDER 2X MAGNIFICATION AND THERE WAS NO VISIBLE DAMAGE OR RESIDUE NOTED. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. IT IS LIKELY THAT THESE LENSES WERE INADVERTENTLY SWAPPED DURING RETURN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904423&pc=QCB | 2023826-2023-01686 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Cataract; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS OPACITY (CATARACT, ANTERIOR SUBCAPSULAR) WAS OBSERVED ON (B)(6) 2022. THERE WAS PIGMENT DISPERSION AND THE PATIENT COMPLAINED OF BLURRED VISION AND GLARE/HALOES. THE PROBLEM WAS RESOLVED. CHAFING OF IRIS WITH RESULTING SIGNIFICANT IRIS TRANSILLUMINATION BETWEEN 1 WEEK AND 1 MONTH POST OP WITH SLOW PROGRESSION; DESPITE EXCELLENT VAULT. BELIEVED TO BE MECHANICAL ISSUE RELATED TO LENS/IRIS APPROXIMATION. THE POST-OP STATUS OF THE EYE AFTER LENS REMOVAL WAS STABLE IRIS TRANSILLUMINATION. ANISOCORIA IMPROVED BUT STILL WITH MYDRIASIS AND OVOID PUPIL. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, CHAFING OF IRIS, IRIS TRANSILLUMINATION, MYDRIASIS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706991&pc=QCB | 2023826-2020-02407 | 2020-09-15 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Capsular Bag Tear | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/094 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY, DUE TO AN ANTERIOR SUBCAPSULAR OPACITY AND A CAPSULAR TEAR HAD OCCURRED. CATARACT SURGERY WAS PERFORMED AND AN IOL LENS WAS IMPLANTED IN THE CILIARY SULCUS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO 2 PITS WERE NOTED IN THE LENS CAPSULE (ANTERIOR) DIRECTLY BELOW THE IRIDOTOMIES. WHEN THE LENS WAS INSERTED A CAPSULAR TEAR EXTENDED FROM THE PITS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTICS TORN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8970000&pc=QCB | 2023826-2019-01541 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (26 MMHG) AND LOW VAULT WERE ASSESSED ON (B)(6) 2019. EYE STATUS ON (B)(6) 2019, "AC QUIET WITH MINIMAL CORNEAL EDEMA. NEAR VA J8." LENS REMAINS IMPLANTED. Manufacturer Narrative: "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SAME LENGTH, SAME MODEL LENS. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED IN THE INITIAL MDR. PATIENT CODE 3191: SECONDARY SURGERY- EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON A FOLLOW UP VISIT ON (B)(6) 2019, IT WAS REPORTED THAT THE PATIENT HAS "SLIGHT ANTERIOR CAPSULE OPACITY, ANTERIOR CHAMBER QUIET AND OPEN ANGLES". SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. EXPLANTED DATE: (B)(6) 2019. PATIENT CODE: 1791, PATIENT CODE: 1766, PATIENT CODE 3191: NO CODE AVAILABLE (ANGLE CLOSURE), DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), DEVICE CODE 2017: IMPROPER OR INCORRECT PROCEDURE OR METHOD. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874226&pc=QCB | 2023826-2020-00634 | 2020-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN-HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUES EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK; ELEVATED IOP AND INFLAMMATION WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE ON (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN- HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUED EVALUATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- INCLUDED "ELEVATED IOP AND INFLAMMATION" TO STATEMENT. H6-PATIENT CODE 1937- ELEVATED IOP. 1932- INFLAMMATION. H6-DEVICE CODE: 1401-BLURRED VISION. H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107370&pc=QCB | 2023826-2019-01751 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Inflammation; Intraocular Pressure Increased; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -18.00/+1.0/177 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED. THE LENS WILL BE EXPLANTED BUT TO DATE REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1937: INFLAMMATION. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: LENS IMPLANTED UPSIDE DOWN. PATIENT CODE 1937 IN PREVIOUS MDR CORRECTED TO 1932. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16261376&pc=QCB | 2023826-2023-00326 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER (PATIENT CURRENT OBSERVING DOCTOR) STATES THAT A SURGEON AT A DIFFERENT CLINIC IMPLANTED A 13.7MM VTICMO 3.7 IMPLANTABLE COLLAMER LENS OF -7.00/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE CURRENT OBSERVING DOCTOR SAW PATIENT FOR THE FIRST TIME ON (B)(6) 2022, WHERE HE UNFORTUNATELY BECAME AWARE OF SEVERAL MEDICAL ERRORS. "TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT." COMPLAINT QUESTIONNAIRE PROVIDED REPORTS CATARACT (ASC); ANGLE CLOSURE WITH ELEVATED IOP; LENS DISLOCATION/ SUBLUXATION; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; BLURRED VISION; AND IRIS ATROPHY WAS REPORTED. CURRENT STATUS OF THE EYE: "EYE ALREADY CALM, CORNEA CLEAR, ANTERIOR CHAMBER SHALLOW, PUPIL WIDE DILATED, ICL IN SITUATION ROTATED COUNTER CLOCKWISE TO INTENDED POSITION, ANTERIOR SUBCAPS CAT. CATARACT FORMATION." CAUSE REPORTED AS USER ERROE; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE IMPLANTING SURGEON HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538715&pc=QCB | 2023826-2020-02049 | 2020-07-08 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Cataract; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; Vomiting; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/1.50/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE SURGEON ASSESSED AN ELEVATED IOP OF 22MMHG. THIS WAS CONTROLLED WITHIN THE FIRST 24 HOURS WITH MEDICATION. ALSO REPORTED IS PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, DILATED PUPIL, GLARE/HALOES, IRIS ATROPHY, PUPIL BLOCK, AND ANTERIOR SUBCAPSULAR OPACITIES. THE PATIENT ALSO EXPERIENCED NASEAU AND EXCESSIVE MOVEMENT DUE TO VOMITING CAUSED BY THE ANESTHESIA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE PUPILLARY DILATION WAS "SLOWLY CORRECTED WITHIN THE FIRST MONTH, LEAVING AN ANISOCORIA OF APPROX. 2.5MM AND PRESENTING SECTORAL ATROPHY OF THE IRIS." THE PATIENT ALSO REPORTS "DECREASE IN THE INTENSITY OF THE COLORS." THE PATIENT STILL EXPERIENCES GLAUKOMFLECKEN, IRIS PIGMENT DISPERSION, AND GLARES AT NIGHT, BUT IS OVERALL "REALLY SATISFIED." THE CAUSE OF THE EVENT IS REPORTED AS "REMNANT OF OVD BEHIND THE ICL." Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8807833&pc=QCB | 2023826-2019-01275 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -11.00/1.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE DAY AFTER OPERATION THE EYE HAD UNDERGONE MYDRIASIS, WHICH WAS IMPOSSIBLE TO NARROW. ON (B)(6) 2019 THE LENS WAS EXPLANTED. THE SURGEON STATED " THE EVENT WAS DUE TO AN UNKNOWN PROBLEM. I CAN ONLY SUPPOSE, THAT DIAMETER ICL WAS BIGGER. WE STILL DONT KNOW WHAT HAPPENED. CATARACT (BEHIND CAPSULA ANTERIOR) WAS OBSERVED AFTER IMPLANTATION. SHE STILL HAS PARALYTIC MYDRIASIS (5-6 MM).'' PATIENT CONDITION IS STILL THE SAME, ON (B)(6) 2019 THE DIAMETER OF THE PUPIL WAS 5-6 MM AND IT DIDN'T RESPOND TO THE LIGHT. PATIENT USED PILLOCARPIN 1% THE LAST 2 MONTHS. Manufacturer Narrative: DATE OF THIS REPORT: 20-JUN-2019. CONCOMITANT MEDICAL PRODUCTS: CARTRIDGE MODEL-SFC-45; LOT#-UNK SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT'S DATE OF BIRTH: (B)(6) 1996. PATIENT CODE 3191: NO CODE AVAILABLE (UNREACTIVE FIXED PUPIL) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 26 AUG 2019. OVD USED: EYEVISC PFS 2 ML (HYDROXYPROPYL METHYLCELLULOSE 2% W/V) OPHTALMIC SOLUTION USP FROM BIOTECH IOP WAS 50.0MM THE DAY AFTER SURGERY, WHEN MYDRIASIS WAS NOTICED. THE ANGLES WERE OPEN. MEDICATIONS USED TO TREAT PATIENT: DEXAMETASON-INIECTION 3 DAYS ONE TIME PER DAY, BETOPTIC AND AZOPT IN DROPS FOR 3 DAYS AND THEN THE IOP WAS OK -14 MM. PATIENT CODE 3191: MEDICAL INTERVENTION. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620830&pc=QCB | 2023826-2022-03493 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/084, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR (ASC) CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. POST-OP, THE PUPIL IN MYDRIASIS NON REACTIVE AND SLIGHT EPITHELIAL HAZE. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EPITHELIAL HAZE, MYDRIATIC PUPIL SHOULD HAVE BEEN INCLUDED. B5: POSTOP COMMENTS ADDITIONAL NOTED PUPIL IN MYDRISASI NON REACTIVE; SLIGHT EPITHELIAL HAZE. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366584&pc=QCB | 2023826-2021-00480 | 2021-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/010 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY, LENS OPACITY ASC WAS REPORTED; OBSERVED ON (B)(6) 2021, AFTER LENS WAS EXPLANTED. MIOL SURGERY PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO BROKEN HAPTIC. H6: 1766 SHOULD HAVE BEEN INCLUDED PREVIOUSLY. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368673&pc=QCB | 2023826-2021-00463 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -11.50/+2.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE ICL WAS TOO SMALL. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: B5: THE ICL LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. H6: HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY AND IOL IMPLANT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021862&pc=QCB | 2023826-2021-02006 | 2020-08-06 04:00 | Death | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Contamination; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+6.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS DUE TO LENS OPACITY-CORTICAL. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES THAT THE LENS HAS A DEPOSIT ON ITS SURFACE. Manufacturer Narrative: B5-FINAL BCVA 0.85, UCVA 0.7. H6-METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS IS A CASE OF FOREIGN DEBRIS ON THE LENS THAT CAUSED THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. AND THERE WAS RESIDUE AND DEBRIS ON THE LENS SURFACE. RED PARTICULATES WERE NOTED, ON LENS OPTIC. CLAIM# (B)(4). Manufacturer Narrative: 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS; -9.5/+6.0/107. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465183&pc=QCB | 2023826-2022-03166 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS FLIPPED IN THE EYE. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT: ADDITIONAL SURGERY: 4625: CATARACT SURGERY, IOL IMPLANTED. H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR/CORTICAL CATARACT. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED THE NEXT DAY AND TORIC IOL WAS IMPLANTED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (THE EYE WAS TOO SOFT AND PATIENT SQUEEZED WHEN THE INCISION FOR ICL WAS MADE, CUTTING THE ANTERIOR CAPSULE, WHICH INDUCED THE CATARACT). CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766105&pc=QCB | 2023826-2023-01343 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "PAT STILL WITH LENS OPACITY SUPRISINGLY, BUT BETTER VAULT." Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO ADVERSE EVENT. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE ADDED FOR CORRECTION. H1: SERIOUS INJURY SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986861&pc=QCB | 2023826-2023-04603 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. MILD LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A MILD ASC. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT VERY HAPPY." H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4614" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986877&pc=QCB | 2023826-2023-04604 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. MILD LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A MILD ASC. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT VERY HAPPY". H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4614" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307838&pc=QCB | 2023826-2019-00213 | 2018-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED, THE SURGEON NOTING LENS OPACITY CORTICAL. REPORTEDLY, THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A CAPSULAR TEAR OCCURRED. A CORTICAL CATARACT WAS NOTED. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND THE PATIENT UNDERWENT PHACOEMULSIFICATION AND IOL IMPLANTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN THE ORIGINAL MDR PATIENT CODE 3191: CAPSULAR TEAR, SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8426297&pc=QCB | 2023826-2019-00417 | 2018-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORT STATED "BEFORE THE SURGERY, DOCTORS HAD OBSERVED CATARACTS. AND THE DOCTOR HAD FULL COMMUNICATION WITH THE PATIENT". A VTICMO13.2, -8.5/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AND LATER REMOVED ON THE SAME DAY. LENS OPACITY (CORTICAL, DATE: (B)(6) 2018) WAS REPORTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. NO TREATMENT OR INTERVENTION HAS BEEN PERFORMED. NON-TRAUMATIC CATARACT IN THE FELLOW EYE WAS REPORTED AS A PRE-EXISTING CONDITION. Manufacturer Narrative: THE REPORTER INDICATED THAT THAT THIS OPACITY WAS NOT PRESENT BEFORE THE ICL SURGERY. TO AVOID UNNECESSARY COMPLICATIONS, THE DOCTOR DECIDED NOT TO IMPLANT THE LENS FOR THE TIME BEING. THE PATIENT IS UNDER OBSERVATION AND LENS IMPLANTATION MAY BE CONSIDERED FOR THE FUTURE. IT WAS ALSO REPORTED THAT THE OPACITY MIGHT HAVE OCCURRED DUE TO LENS TOUCH DURING ICL IMPLANTATION. CONCLUSION CODE- CORRECTED FROM "24" TO "4315" FOR INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601622&pc=QCB | 2023826-2019-00811 | 2018-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-10.50/1.0/096 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: OTHER RELEVANT HISTORY: CATARACT IN THE OPERATIVE EYE. DEVICE CODE- 1494: OFF-LABEL USE (CATARACT IN OPERATIVE EYE). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932863&pc=QCB | 2023826-2019-01545 | NA | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO A CATARACT HAD DEVELOPED. THE PATIENT HAD A LASIK PROCEDURE AFTER THE ICL SURGERY. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LASIK) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. METHOD CODE: 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. IMPLANT DATE: (B)(6) 2018. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10011427&pc=QCB | 2023826-2020-00996 | 2020-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH AN UNSPECIFIED LENS ROTATION AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. D7-EXPLANT DATE-(B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. PATIENT CODE 1766-LENS OPACITY(CATARACT). DEVICE CODE 2923- ROTATION. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2020 REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED, PATIENT IS IN VERY GOOD CONDITION. THE REPORTER STATED THAT LENS OPACITY ASC WAS NOTED BEFORE FIRST SURGERY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATON : LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285662&pc=QCB | 2023826-2020-01480 | NA | Injury | STAAR SURGICAL COMPANY | 2020-07-16 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. REPORTEDLY, THE PATIENT WAS OPERATED ON WITHOUT COMPLICATIONS. AT THE END OF THE SURGERY THE DOCTOR NOTICED AN OPACIFICATION IN THE CENTER OF THE LENS CORTEX BUT WASN'T SURE IF IT WAS THERE PREVIOUSLY. VAULTING INITIALLY HIGH BUT IS OK NOW. REFRACTION IS +0.75 INSTEAD OF THE EXPECTED -0.5D. ANTERIOR SUBCAPSULAR OPACITIES SEEN BUT VISION IS "TOO LOW." IT IS REPORTED THAT THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10337903&pc=QCB | 2023826-2020-01604 | 2020-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/2.00/95 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED ON THE SAME DAY DUE TO LENS OPACITY (ASC). THE REPORTER STATED THAT THE OPACITY WAS OBSERVED ON THE DAY OF SURGERY, THE PATIENTS CONDITION IS WELL NOW AND UCVA IS 20/32. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429096&pc=QCB | 2023826-2020-01830 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC. CATARACT SURGERY WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429117&pc=QCB | 2023826-2020-01846 | 2019-05-15 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9742235&pc=QCB | 2023826-2020-00373 | NA | Injury | STAAR SURGICAL COMPANY | 2020-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: TICL LENS ROTATION WAS REPORTED. EMAIL STATES "WE HAVE A PATIENT THAT HAS HAD HIS ICL ROTATE AND WE'LL NEED TO REPOSITION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: A1: 231999120 TK. B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/2.5/111 (SPHERE/CULINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. " DECREASED VISION RELATED TO CRYSTALLINE LENS CHANGED (PSC) SO NO TX UNTIL PT READY FOR CATARACT SURGERY. PRIMARY OPEN ANGLE OR NARROW GLAUCOMA WAS REPORTED AS A PRE-EXISTING CONDITION. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO OTHER SERIOUS IN INITIAL MDR. B4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. G4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR & SUPPLEMENTAL MDR#1. H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 1766 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689304&pc=QCB | 2023826-2020-00184 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.50/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE TO OCCUR. Manufacturer Narrative: CORRECTED DATA: B4- "16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. G4- "16-JAN-2020" SHOULD BE CORRECTED TO "15-JAN-2020" IN THE INITIAL MDR. ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 THE LENS WAS EXPLANTED WITH A PLAN FOR EXCHANGE. HOWEVER, THE PATIENT DEVELOPED AN INTRAOPERATIVE CATARACT DURING EXPLANTATION AND THE DECISION WAS MADE TO ABORT IMPLANTATION OF REPLACEMENT ICL." H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: PATIENT CODE 1766 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 AN ICL IMPLANTATION IN THE RIGHT EYE WITH A LARGER LENS WAS PERFORMED." CLAIM# (B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844162&pc=QCB | 2023826-2023-04346 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6) 1983, A3 - M, B3 - 13-SEP-2022. CORRECTION: B5 - REPORTED AS: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND GLARE/HALOES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023 AND IN AS SECOND SURGERY ON (B)(6) 2023 A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) WAS IMPLANTED. PROBLEM RESOLVED. REPORTEDLY, MILD ASC AND PATIENT IS VERY HAPPY. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844244&pc=QCB | 2023826-2023-04347 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6)1983, A3 - M, B3 - 13-SEP-2022. CORRECTION B5 - REPORTED AS THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXPLANTED AND A NEW LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) WAS LATER IMPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND GLARE/HALOES ON (B)(6) 2023. PROBLEM RESOLVED. REPORTEDLY MILD ASC, PATIENT IS VERY HAPPY. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941309&pc=QCB | 2023826-2019-01571 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -13.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 CATARACT SURGERY WAS PERFORMED DUE TO LENS OPACITY-CORTICAL, GLARE/HALOS, POOR VISION, AND CORTICAL CHANGES. THE SURGEON STATES, "VISION POOR WITH GLARE/HALOS, SOME DOUBLE." Manufacturer Narrative: "IT WAS LATER REPORTED AT A FOLLOW UP DATE THAT THE PATIENT STATUS IS GOOD AND STABLE WITH NO VISUAL COMPLAINTS." SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171666&pc=QCB | 2023826-2020-01273 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+4.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171670&pc=QCB | 2023826-2020-01272 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+4.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON 14-JUL-2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CORRECTIONS: B5 - "BLURRED VISION WAS ALSO REPORTED." SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 2140 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522593&pc=MTA | 2023826-2017-00710 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -11.0/1.0/091 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(6). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6534124&pc=MTA | 2023826-2017-00711 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -10.0/2.0/093 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.9MM)]. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873717&pc=MTA | 2023826-2016-01069 | 2015-07-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.50/+1.50/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR/POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/50 +1. Manufacturer Narrative: CONCLUSION: THE PRE-EXISTING "GRAINY PERIPHERAL CORTICAL CLOUDING" WAS LISTED ON THE FORM AND THAT CONDITION MIGHT HAVE CAUSED THE CATARACT DEVELOPMENT SINCE IT IS WELL KNOWN THAT OPACITIES START AS VACUOLES AND CLEFTS BETWEEN THE LENS FIBERS AND PROGRESS TOWARDS ANTERIOR, POSTERIOR OR EQUATORIAL CORTEX. IT SHOULD BE NOTED THAT AT THE TIME OF SURGERY THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. GIVEN ALL THIS INFORMATION, OBSERVED LENS OPACITY AND SUBSEQUENT CATARACT DEVELOPMENT WAS AGE RELATED (PATIENT FACTOR) AND NOT CAUSED BY THE DEVICE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC BROKEN OFF. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873711&pc=MTA | 2023826-2016-01070 | 2004-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.75/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/30 -2. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955552&pc=MTA | 2023826-2017-01577 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Activation Failure | Cataract, Induced; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5/+2.0/097 INTO THE PATIENT'S LEFT EYE (OS), THE LENS WOULD NOT PROPERLY UNFOLD. THE SURGEON STATES THAT, "AFTER ALL POSSIBLE MANEUVERS" THE LENS DIDN'T UNFOLD PROPERLY AND THE LENS WAS EXPLANTED. DURING EXPLANT THE LENS WAS SCRATCHED. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THE SURGEON THEN PEFORMED PHACO-EMULSIFICATION ON THE EYE AND IMPLANTED AN EXCHANGE IOL ON (B)(6) 2017 "WITHOUT COMPLICATIONS." Manufacturer Narrative: BEST CORRECTED VISUAL ACUITY (BCVA), ON (B)(6) 2017, WAS 20/30 WHICH WAS A DECREASE FROM BASELINE BCVA (20/25). (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PHACO-EMULSIFICATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176008&pc=MTA | 2023826-2016-01763 | 2014-09-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract, Induced; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/042 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2014. THE PATIENT EXPERIENCED BLURRED VISION AND CLOUDY VISUAL ACUITY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25 +2. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. THE LENS WAS REMEASURED AND WAS FOUND TO BE IN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608173&pc=MTA | 2023826-2017-00900 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Intraocular Pressure Increased; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/043 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE, AND IRIS ATROPHY. ON (B)(6) 2017 THE PATIENT ALSO EXPERIENCED LENS OPACITY ASCO. THE SURGEON PLANS TO EXCHANGE TO A SHORTER LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. (B)(4). Manufacturer Narrative: (B)(6). THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). H6_WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684194&pc=MTA | 2023826-2017-01086 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ACCORDING TO THE DR. THAT BECAUSE THE SURGERY OPERATING TIME WAS TOO LONG, THE SUSPENSORY LIGAMENT NEAR THE LENS CAPSULE MEMBRANE DOT LIGHT LENS OPACITY. THE LENS WAS EXPLANTED IN THE SAME SURGERY, AND THE PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED, NO IOL LENS WAS IMPLANTED. THE CUSTOMER STATED THAT THE SITUATION IS STABLE, AND OBSERVATION ON REGULAR BASIC. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE : OFF LABEL USE (ACD<3.0MM) (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5711853&pc=MTA | 2023826-2016-00720 | 2016-05-08 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+0.5/151 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY - ASCO) ON (B)(6) 2016 AND UNREACTIVE PUPIL (FIXED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4) LENS REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16474967&pc=MTA | 2023826-2023-00725 | 2022-07-20 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.0/1.0/99 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT FOR A PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THIS WAS DUE TO A EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 12.6, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/99 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE PATIENT EXPERIENCED EARLY LENS OPACITY. REPORTEDLY "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 13.2, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS BECAME CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721163&pc=MTA | 2023826-2017-01196 | 2017-06-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/4.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT, BLURRED VISION, AND LENS REPOSITIONING. LENS WAS "REMOVED AND CATARACT PROCEDURE HAVE BEEN DONE DUE TO CHANGES ON NATURAL LENS". THE PATIENT'S POST-OP CONDITION IS "WITH APHAKIC LENS AND HIS VISION IS 6/12". Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS MISSING PIECE OF HAPTIC. THE PATIENT'S POST-OP VISION IS 6/12. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874039&pc=MTA | 2023826-2017-01447 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 13.2 MM VTICMO13.2, -7.0/+4.5/107 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2015, THE LENS WAS EXPLANTED DUE TO POST TRAUMATIC DISLOCATION IN A/C AND TRAUMATIC CATARACT DUE TO FIREWORKS MUCH LATER THAN THE IMPLANTATION. IT HAD NO CONNECTION TO THE SURGERY. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND PRESENCE OF FIBERS ON LENS SURFACE. THE PATIENT ALSO HAD BLURRED VISION. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE:TRAUMA CAUSED BY FIRWORK. NO CONNECTION TO SURGERY. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6681259&pc=MTA | 2023826-2017-01089 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/40 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND THE DEVELOPMENT OF A CATARACT. A REPOSITIONING OF THE LENS, ENLARGEMENT OF THE INCISION (IN ORDER TO PERFORM CATARACT SURGERY), PHACO-EMULSIFICATION, AND IOL IMPLANTATION WAS PERFORMED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND THE PATIENT'S POST OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25 WITH STATUS OF THE EYE BEING PSEUDOPHAKIC. Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. Manufacturer Narrative: "HOSPITALIZATION-INITIAL OR PROLONGED" WAS CHECKED BY MISTAKE. THERE WAS NO LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5762131&pc=MTA | 2023826-2016-00817 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -11.00/ +6.0/100 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL) BLURRED VISION, THE CATARACT WAS REMOVED, AND A TORIC IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513549&pc=MTA | 2023826-2015-00149 | 2014-11-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LOW VAULT AND LENS OPACITY WAS OBSERVED AND THE PATIENT EXPERIENCED BLURRED VISON. THE LENS WAS EXCHANGED FOR A LONGER LENGTH WHICH RESOLVED THE VAULTING ISSUE, HOWEVER, THE CATARACT PROGRESED. THE SECOND LENS WAS REMOVED, THE CATARACT WAS EXTRACTED AND AN IOL WAS IMPLANTED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: .(B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. THERE WAS EVIDENCE OF A CLEAR SURGICAL RESIDUE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). IN CASES OF POSTOPERATIVE LOW VAULTING STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS OR ROTATION ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPUSLAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE MOST PROBABLE CAUSE OF THIS EVENT IS LOW VAULT SECONDARY TO A SHORT LENS. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247963&pc=MTA | 2023826-2017-00071 | 2011-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+1.50/094 DIOPTER, IN THE PATIENT'S LEFT EYE. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, REFRACTIVE SURPRISE, LENS OPACITY, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS, GLARE/HALOES, AND BLURRED VISION. THE PATIENT UNDERWENT PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6129538&pc=MTA | 2023826-2016-01682 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract; Blurred Vision; Visual Impairment; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND SURGEON REPOSITIONED THE NEW LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE SURGEON EXCHANGED THE LENS FOR A LONGER LENS AND REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS OPACITY, LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. REPOSITIONING OF THE LENS WAS PERFORMED AND THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. (B)(4). Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LENS ROTATION, DISLOCATION, REFRACTIVE CHANGE, AND BLURRED VISION. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE SURGEON REPORTED LENS OPACITY (ASCO) AND LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694861&pc=MTA | 2023826-2016-00799 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Loss of Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED ASC (ANTERIOR SUBCAPSULAR CATARACT) ON (B)(6) 2016 AND SUBSEQUENTLY LOST BCVA (BEST-CORRECTED VISUAL ACUITY). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572298&pc=MTA | 2023826-2016-00432 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/X010 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015, AND THE PATIENT EXPERIENCED LOSS OF BCVA. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/200. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378618&pc=MTA | 2023826-2015-01669 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2015) UNCORRECTED VISUAL ACUITY WAS 20/15. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671142&pc=MTA | 2023826-2016-00741 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5./X160 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015. THE PATIENT HAD CATARACT SURGERY AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. LENS RETURNED BENT WAS ALSO NOTED. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5814335&pc=MTA | 2023826-2016-00916 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR)/PERIPHERAL TOUCH. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND SPOTS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987375&pc=MTA | 2023826-2018-01533 | 2018-08-08 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Corneal Edema; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/060 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), THE LAST WEEK OF (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A FIXED DILATED PUPIL POST-OP, SEVERE CORNEAL EDEMA, WHICH HAS BEEN PERSISTENT UNTIL NOW AND AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, BLURRED VISION AND PUPIL/IRIS ATROPHY. THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED ONE MONTH AFTER THE SURGERY AND AN IOL WAS IMPLANTED. THE PUPIL REMAINED DILATED AND THE CORNEA WAS EDEMATOUS AND WHITE. THE PATIENT HAS HAD A CORNEAL TRANSPLANT. THE REPORTER INDICATED THE INTRAOCULAR PRESSURE HAS BEEN STABLE AND NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENTS VISION IS STILL HAND MOTION POST OPERATIVELY AND INTRAOCULAR PRESSURE IS 12MMHG. THE FIRST CHECKUP AFTER KERATOPLASTY INDICATED ALL IS GOOD. CORRECTED DATA: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6220084&pc=MTA | 2023826-2017-00009 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Glaucoma; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DESCRIPTION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/13 AS OF 04/21/2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937371&pc=MTA | 2023826-2015-00893 | 2015-03-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY, WHICH WAS FIRST NOTED ON (B)(6) 2015. THE REPORTER INDICATED IT WAS AN ANTERIOR SUBCAPSULAR OPACITY. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARE/HALOS AND BLURRED VISION. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926670&pc=MTA | 2023826-2016-00374 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED AT FOLLOW-UP VISITS, THE LENS HAPTIC WAS FOLDED AND THE PUPIL WAS OVALISED. THE LENS WAS TILTED ON ONE SIDE AND WAS RE-POSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP #4 FOR MDR # 2023826-2017-00374 WAS SUBMITTED IN ERROR. THE REPORTED INFORMATION WAS MEANT TO BE FOLLOW-UP #4 FOR MDR # 2023826-2016-00374. AS A RESULT, ONLY THE INITIAL MDR # 2023826-2017-00374 PERTAINS TO CLAIM# 710106. FOLLOW-UP #4 FOR MDR # 2023826-2016-00374 WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794667&pc=MTA | 2023826-2018-01196 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. ANOTHER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), DURING THE SAME SURGERY. THE LENS WAS REMOVED DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE AND THE SURGEON DAMAGED THE CAPSULE. THE SURGEON PERFORMED PHACOEMULSIFICATION, WITH IOL IMPLANTATION. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED IN A SECOND SURGERY AND THE PROBLEM WAS RESOLVED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: PREVIOUS DATE NOT APPLICABLE, LENS WAS NOT RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300934&pc=MTA | 2023826-2018-00292 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+1.0/150 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS LENS OPACITY NS. REPORTEDLY, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. THE SURGEON STATES "CATARACT SURGERY PERFORMED. ICL REMOVED DUE TO CATARACT. NOW PERFECT VISION." Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721184&pc=MTA | 2023826-2017-01190 | 2016-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4/87 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (POSTERIOR SUBCAPSULAR CATARACT). ON (B)(6) 2017 THE LENS WAS EXPLANTED FOR PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. THE PRODUCT HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL AND HAS SOME MOISTURE. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6805946&pc=MTA | 2023826-2017-01283 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO LENS OPACITY (ASCO). PHACO- EMULSIFACTION WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS FOUND TO BE 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: IN THE PROCESS OF THE IOL IMPLANT, DUE TO THE LACK OF PATIENT'S COOPERATION (TURN THE HEAD) LEAD TO THE OPERATION CAUSED BY NATURAL CRYSTALS BECOME CLOUDY. THE IOL IMPLANTATION SOLVED THE PROBLEM AND THE PATIENT'S CONDITION IS GOOD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077381&pc=MTA | 2023826-2016-01531 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PMA/510(K) : NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074345&pc=MTA | 2023826-2016-00912 | 2016-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -22.50/+3.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6566884&pc=MTA | 2023826-2017-00810 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/82 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED LOW VAULTING AND LENS OPACITY (ASCO) EARLY POSTOPERATIVELY. THE LENS WAS EXCHANGED FOR A LONGER LENS OF A DIFFERENT MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/14 (BETTER THAN PREOPERATIVE). Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE THE LENS WAS RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: (B)(4). AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER (B)(6)YEARS OF AGE AND WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6590465&pc=MTA | 2023826-2017-00825 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/88 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT AND LENS OPACITY (ASCO). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED WITH A LONGER LENS WITH A DIFFERENT MODEL AND DIOPTER, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE OF RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE- AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) UNDER 3.0MM AND FOR PATIENTS OVER THE AGE OF (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059292&pc=MTA | 2023826-2016-01530 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+1.5/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/13 AND UNCORRECTED VISUAL ACUITY WAS 20/13. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850306&pc=MTA | 2023826-2017-01413 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS OBSERVED ON (B)(6) 2017. POST CATARACT OPERATION PATIENT IS HAPPY WITH THE VISION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (ACD<3.0MM). (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704662&pc=MTA | 2023826-2015-00424 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED DISCOMFORT, LIGHT SENSITIVITY AND ANTERIOR UVEITIS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704665&pc=MTA | 2023826-2015-00425 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, INFLAMMATION, ANTERIOR UVEITIS, DISCOMFORT, LIGHT SENSITIVITY, PAIN AND TEARING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved