## Vision Issue-1
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code...35 | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Exemption Number | Number of Events | Event Text | Management | Outcomes
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372525&pc=QCB | 2023826-2019-00315 | 2019-01-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.5/146 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS DEVICE CODE 1494: PREVIOUS CODE NOT APPLICABLE CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8498014&pc=QCB | 2023826-2019-00583 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-04-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS THE VAULT IS A "BIT HIGH" AND THAT THE PATIENT COMPLAINS OF BLURRED VISION IN SLIGHTLY DIM LIGHTING "LIKE A LAYER OF OIL COVERED ON YOUR GLASSES." LENS REMAINS IMPLANTED. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511547&pc=QCB | 2023826-2019-00600 | 2019-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE REPORTER INDICATED THERE WAS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8511615&pc=QCB | 2023826-2019-00601 | 2019-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE REPORTER INDICATED THERE WAS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519002&pc=QCB | 2023826-2019-00613 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS AN UNEXPECTED RESIDUAL REFRACTION OF -1.25 CYL POSTOPERATIVELY. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: IMPLANT DATE: PREVIOUSLY MARKED UNK IN ERROR. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8524100&pc=QCB | 2023826-2019-00619 | 2018-09-05 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -11.0/5.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 AND REFRACTIVE SURPRISE HAS BEEN OBSERVED. THE REPORTER STATED THAT THE LENS REMIANS IMPLANTED WITH POOR VA (VISUAL ACUITY). PATIENT BCVA 1.0 WITH OVER CORRECTION +0.75-1.25X63. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611967&pc=QCB | 2023826-2019-00675 | 2019-04-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. THE DOCTOR STATES "THE PATIENT IS HAPPY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692892&pc=QCB | 2023826-2019-01070 | 2017-10-19 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+0.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708611&pc=QCB | 2023826-2019-01090 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1MM, -1.00/1.0/032 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE REPORTER STATED " THE PATIENT HAD A SMALL REFRACTIVE ERROR ON THE RIGHT EYE PRE-OP, BUT IT HAS BEEN WORSE AFTER ICL IMPLANT ". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT THE PATIENT HAS NOT BEEN TO THE CLININC SINCE MAR 20TH, BCVA WAS REPORTED AS 1.2 , PATIENT WANTS A LENS EXCHANGE. PATIENT WAS TRAINED WITH A COVER AS A CHILD, RIGHT EYE DOMINANT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8708784&pc=QCB | 2023826-2019-01091 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1MM, -2.00/2.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER STATED THAT THE PATIENT HAS NOT BEEN TO THE CLININC SINCE MAR 20TH, BCVA WAS REPORTED AS 1.2 , PATIENT WANTS A LENS EXCHANGE. PATIENT WAS TRAINED WITH A COVER AS A CHILD, RIGHT EYE DOMINANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8773867&pc=QCB | 2023826-2019-01187 | 2019-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/1.0/144 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2019 AND REFRACTIVE SURPRISE WAS OBSERVED. PATIENT IS UNHAPPY WITH NEAR VISION, AWAITING ICL EXCHANGE FOR HYPEROPIC SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE IS (B)(6) 2019; THE EXACT DAY IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE FOR ACD <3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795444&pc=QCB | 2023826-2019-01242 | 2019-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798875&pc=QCB | 2023826-2019-01267 | 2019-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -0.50/6.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798976&pc=QCB | 2023826-2019-01266 | 2019-06-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -0.50/6.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813862&pc=QCB | 2023826-2019-01290 | 2019-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/020 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813865&pc=QCB | 2023826-2019-01289 | 2019-05-04 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/139 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852208&pc=QCB | 2023826-2019-01323 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-08-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/141 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852214&pc=QCB | 2023826-2019-01324 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-08-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/079 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871962&pc=QCB | 2023826-2019-01395 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894318&pc=QCB | 2023826-2019-01433 | NA | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. LOSS OF BCVA IS ALSO REPORTED, HOWEVER THE PRE-OP BCVA IS LISTED AS "?" AND POST-OP IS "20/25 GLASSES." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ONLY YEAR OF BIRTH REPORTED; EXACT DATE NOT KNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966498&pc=QCB | 2023826-2019-01648 | 2019-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -0.5/+6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND THAT UCVA HAS WORSENED FROM THE 1ST WEEK OF SURGERY. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD CODE 4110: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9088369&pc=QCB | 2023826-2019-01670 | 2019-07-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -9.50/3.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136940&pc=QCB | 2023826-2019-01809 | 2019-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-30 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -3.5/+6.0/118 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND NO SURGICAL INTERVENTION WILL BE PERFORMED AS "PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9148352&pc=QCB | 2023826-2019-01853 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-10-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WITH POST LASIK ECTASIA WHO HE CROSS LINKED AND IMPLANTED 13.7MM, VTICMO13.7, -2.00/4.0/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE RIGHT EYE (OD) HAD AN EXCELLENT OUTCOME, BUT OVER THE PAST 2 YEARS HIS CORNEAL SHAPE HAS CHANGED SLIGHTLY AND HAS BECOME MORE MYOPIC. PATIENT HAS NO CATARACT AND HIS CORNEAL ENDOTHELIAL CELL COUNT IS PERFECT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454032&pc=QCB | 2023826-2019-02344 | 2019-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -8.50/0.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1- "PRODUCT PROBLEM" SHOULD BE MARKED IN THE INITIAL MDR. B2- "REQUIRED INTERVENTION" IS NOT APPLICABLE IN THE INITIAL MDR. H6- WORK ORDER SEARCH SHOULD SAY: "NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9469760&pc=QCB | 2023826-2019-02383 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-12-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/065 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476270&pc=QCB | 2023826-2019-02391 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+3.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - POSTOPERATIVE REPORT INDICATED PATIENT'S IOP IS STABLE AND REPORTED BLURRED VISION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476271&pc=QCB | 2023826-2019-02390 | 2019-11-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+4.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - POSTOPERATIVE REPORT STATED PATIENT'S IOP IS STABLE AND REPORTS BLURRED VISION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476283&pc=QCB | 2023826-2019-02389 | 2017-01-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS "HYPEROPIC SHIFT 2D." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE "PATIENT HAS REFRACTIVE CHANGES 'SIGNIFICANT' OVER TIME AFTER 3 YEARS OF ICL IMPLANTATION." ALSO STATED, "PATIENT IS DOING FINE NOW WITH ACCOMMODATION AND BRAIN ADAPTATION." REFRACTIVE SURPRISE IS ALSO REPORTED. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9491791&pc=QCB | 2023826-2019-02421 | 2020-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/3.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS IMPLANTED ON (B)(6) 2019. REFRACTIVE SURPRISE WAS ASSESSED ON (B)(6) 2020. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516576&pc=QCB | 2023826-2019-02450 | 2019-12-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.00/6.0/156 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9536347&pc=QCB | 2023826-2019-02508 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-12-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50/+6.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588359&pc=QCB | 2023826-2020-00120 | 2019-11-04 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/5.00/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 REFRACTIVE SURPRISE WAS NOTED. LENS REMAINS IMPLANTED, PATIENTS CURRENT CONDITION IS GOOD, BCVA 20/20. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614033&pc=QCB | 2023826-2020-00138 | 2019-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.5/+6.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. AT POST-OP VISIT ON (B)(6) 2020, THE REFRACTION WAS STABLE. THE PATIENT WAS PRESCRIBED SPECTACLES TO CORRECT RESIDUAL ASTIGMATISM WHILE DRIVING. THE PATIENT IS SATISFIED WITH THE FINAL OUTCOME. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9614063&pc=QCB | 2023826-2020-00139 | 2019-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.5/+6.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. AT POST-OP VISIT ON (B)(6) 2020, THE REFRACTION WAS STABLE. THE PATIENT WAS PRESCRIBED SPECTACLES TO CORRECT RESIDUAL ASTIGMATISM WHILE DRIVING. THE PATIENT IS SATISFIED WITH THE FINAL OUTCOME. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689160&pc=QCB | 2023826-2020-00282 | 2018-02-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/2.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT "ENDED UP WITH SMALL MINUS ON ONE EYE AND SMALL PLUS ON OTHER EYE. WE WILL CHANGE RIGHT LENS AND HOPE THAT THE BALANCE OF THE EYES WILL RELIEVE SYMPTOMS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706694&pc=QCB | 2023826-2020-00306 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.00/2.0/084 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. "VERY LOW VAULT AND REFRACTIVE SURPRISE" WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706704&pc=QCB | 2023826-2020-00305 | 2020-01-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.50/2.5/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. "VERY LOW VAULT AND REFRACTIVE SURPRISE" WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721915&pc=QCB | 2023826-2020-00308 | 2020-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+4.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721913&pc=QCB | 2023826-2020-00309 | 2020-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Appropriate Term/Code Not Available | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION:: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6; -11.5/4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL/SIZE LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE:(B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- IN PREVIOUS MDR RESULT CODE 114 IS CORRECTED TO INVESTIGATION FINDINGS 213. H6- IN PREVIOUS MDR INVESTIGATION CONCLUSION CODE 4310 NEEDED TO BE REPORTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9759568&pc=QCB | 2023826-2020-00325 | 2019-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+6.0/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. PRE-EXISTING: PREVIOUS CORNEAL OR REFRACTIVE SURGERY; KERATOCONUS; POST PK SURGERY. IMPLANT DATE: (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+6.0/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME AND TRANSIENT LOSS OF BCVA. THE LENS REMAINS IMPLANTED. AT THE TIME OF IMPLANTATION WITH HISTORY OF KERATOCONUS AND POST PK SURGERY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9850132&pc=QCB | 2023826-2020-00571 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE AND DIFFERENT POWER LENS. ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. D6A: (B)(6) 2019. D6B: (B)(6) 2020. H1- SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE LENS AND DIFFERENT POWER LENS. THE EXCHANGED RESOLVED THE PROBLEM. H6-IMPACT CODE 4629 IS NOT APPLICABLE; THE EXCHANGE LENS WAS NOT REPOSITONED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960153&pc=QCB | 2023826-2020-00853 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.5/+3.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND POSSIBLE LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960163&pc=QCB | 2023826-2020-00854 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-04-14 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND POSSIBLE LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025646&pc=QCB | 2023826-2020-01007 | 2020-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025656&pc=QCB | 2023826-2020-01006 | 2020-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2020-05-05 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10130964&pc=QCB | 2023826-2020-01217 | 2020-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -1700/3.50/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED HIS OR IS (B)(6). HE HAS NOT RECEIVED ANY TREATMENT AND INTERVENTION AT THIS TIME. THE REFRACTIVE SURPRISE WAS FIRST NOTED ON (B)(6) 2020, 10 DAYS POST-OP. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - CORRECT LENS (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -17.00/+3.50/89. B5 - DATE OF IMPLANT SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10131012&pc=QCB | 2023826-2020-01218 | 2020-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/3.00/77 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED HIS OR IS (B)(6). HE HAS NOT RECEIVED ANY TREATMENT AND INTERVENTION AT THIS TIME. THE REFRACTIVE SURPRISE WAS FIRST NOTED ON (B)(6) 2020, 10 DAYS POST-OP. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10162715&pc=QCB | 2023826-2020-01294 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER REPORTED REFRACTIVE CHANGE OVERTIME IN THE LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/0.50/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE, ASSESSED ON (B)(6) 2020 WAS REPORTED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE PLANNED A TARGET REFRACTION THE LEFT EYE FOR PROXIMITY, THE RESULT WAS EMMETROPIA. WE ASK FOR A LENS EXCHANGE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D2: CORRECTED FROM PHAKIC INTRAOCULAR LENS,MTA TO PHAKIC TORIC INTRAOCULAR LENS,QCB. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE CHECKED IN PREVIOUS MDR. B2- REQUIRED SURGICAL INTERVENTION SHOULD BE CHECKED IN PREVIOUS MDR. B5- "LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS ON (B)(6) 2020, AND THE PROBLEM WAS RESOLVED." SHOULD BE IN THE PREVIOUS MDR. D6B- (B)(6) 2020, SHOULD BE IN THE PREVIOUS MDR. H1- "SERIOUS INJURY" SHOULD BE CHECKED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE NAME: UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222177&pc=QCB | 2023826-2020-01420 | 2020-06-08 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.0/+5.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222170&pc=QCB | 2023826-2020-01418 | 2020-06-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254180&pc=QCB | 2023826-2020-01444 | 2020-03-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -12.5/+4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B1- CORRECTED TO PRODUCT PROBLEM-MALFUNCTION. H1- CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332955&pc=QCB | 2023826-2020-01568 | NA | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION : B5- THE LENS WAS ROTATED ON (B)(6) 2020 8 DEGREES. IT IS REPORTED THAT THE PATIENT IS STILL NOT SATISFIED WITH RESULT. H1- SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356745&pc=QCB | 2023826-2020-01608 | 2020-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+4.0/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED A 20 DEGREE MISALIGNMENT HAD LED TO THE SURPRISE REFRACTION. DOCTOR DECIDED TO RE-ALIGN LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403862&pc=QCB | 2023826-2020-01730 | 2020-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-13 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+4.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508720&pc=QCB | 2023826-2020-02061 | 2020-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.00/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE RECEIVED BY MFR: THIS PRODUCT IS NOT MARKETED IN THE U.S.. HEALTH EFFECT IMPACT : LENS REMAINS IMPLANTED, NO SECONDARY INTERVENTION HAS BEEN PERFORMED. UNABLE TO ENTER FIELD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508814&pc=QCB | 2023826-2020-02063 | 2020-07-23 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/2.00/61 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10508931&pc=QCB | 2023826-2020-02065 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -12.50/4.5/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PROBLEM RESOLVED, PATIENT IS HAPPY. USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524412&pc=QCB | 2023826-2020-02039 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | N/A#S-P#N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE CHANGE OVER TIME WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER/AXIS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551357&pc=QCB | 2023826-2020-02122 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. THE SURGEON STATES THE LENS WAS REPOSITIONED 3 TIMES AND THE LENS ROTATED AT THE SAME ANGLE AND DIRECTIN AFTERWARD. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- LENS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS. THIS RESOLVED THE PROBLEM. - SHOUDL BE IN THE PREVIOUS MDR. D6B- (B)(6) 2020 - SHOULD BE IN PREVIOUS MDR. H6- HEALTH IMPACT CODE 4629 - SHOULD BE IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10556954&pc=QCB | 2023826-2020-02154 | 2020-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.50/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED, NO FURTHER INTERVENTION HAS BEEN PERFORMED, VISION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY, THE "PATIENT VISION IS 6/6 AND IS HAPPY". NO INTERVENTION WAS DONE. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10560526&pc=QCB | 2023826-2020-02112 | 2020-08-26 | Injury | STAAR SURGICAL COMPANY | 2020-09-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM TMICL12.6, -10.50/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. BLURRED VISION WAS OBSERVED AND LENS WAS REPOSITIONED ON (B)(6) 2020, THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, SPHERICAL LENS. THIS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10600979&pc=QCB | 2023826-2020-02232 | 2020-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/1.5/021 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, SURGEON "NOT YET DECIDED WHETHER TO TRY ROTATION, REPLACEMENT, OR LASER CORRECTION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: LENS ROTATION ALSO REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635460&pc=QCB | 2023826-2020-02247 | 2020-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -17.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL. PATIENT CONTINUES TO BE UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H10 "(B)(4)" CORRECTED TO "CLAIM#: (B)(4)" IN INITAL MDR. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT 13.2MM VTICMO_13.2 -17.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, THE SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL." PATIENT CONTINUES TO BE UNDER OBSERVATION. THE PATIENT IS UNHAPPY WITH VISION BUT REPORTEDLY THE PATIENT'S VA IS GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635463&pc=QCB | 2023826-2020-02246 | 2020-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -17.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL. "PATIENT CONTINUES TO BE UNDER OBSERVATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT 13.2MM VTICMO_13.2 -17.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER REPORTER ON (B)(6) 2020, THE SURGEON "HAS GIVEN PATIENT CONTACT LENSES TO CORRECT THE TINY RESIDUAL." PATIENT CONTINUES TO BE UNDER OBSERVATION. THE PATIENT IS UNHAPPY WITH VISION BUT REPORTEDLY THE PATIENT'S VA IS GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H10 - "CLAIM# (B)(4)" CORRECTED TO "CLAIM#: (B)(4)" IN INITAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659329&pc=QCB | 2023826-2020-02317 | 2020-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION IS UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659337&pc=QCB | 2023826-2020-02313 | 2020-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -11.5/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659339&pc=QCB | 2023826-2020-02310 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ASTIGMATISM, AND BLURRY VISION. THE LENS WAS REPOSITIONED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.5/+2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTED RESIDUAL ASTIGMATISM AND BLURRY VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2020, AND THIS RESOLVED THE PROBLEM. LENS REMAINS IMPLANTED. SECONDARY SURGICAL INTERVENTION OCCURRED ON (B)(6) 2023, FOR LASIK ENHANCEMENT FOR RESIDUAL ASTIGMATISM AFFECTING VISUAL OUTCOME. REPORTEDLY THIS WAS NOT RELATED TO THE LENS OR PROCEDURE. THIS RESOLVED THE PROBLEM. H6: 4625-LASIK ENHANCEMENT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10678723&pc=QCB | 2023826-2020-02357 | 2020-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.00/2.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. BLURRED VISION AND POSSIBLE REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "A WHITE FIBROUS MEMBRANE ADHERENT TO POSTERIOR ICL." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 OF -7.0/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS A WHITE FIBROUS MEMBRANE ADHERENT TO POSTERIOR ICL. THE PATIENT EXPERIENCED BLURRED VISION AND DUREZOL DROPS WERE PRESCRIBED. THE SURGEON TRIED TO REMOVE MEMBRANE WITH LOW POWER LASER ENERGY ANTERIOR TO ICL WITH NO SUCCESS. THE LENS WAS EXPLANTED ON (B)(6) 2020. D6A- (B)(6) 2020. H3- DEVICE EVALUATION: LENS RETURNED IN A SPECIMEN CUP WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. H6- HEALTH EFFECT: 4624- LOW POWER LASER ENERGY ANTERIOR TO ICL. CORRECTED DATA: G4- IN PREVIOUS MDR SHOULD BE CORRECTED TO 14-OCT-2020. H6-TYPE OF INVESTIGATION CODE 4110: UPDATED ERIAL NUMBER LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: REPORTED AS (B)(6). UPDATED / CORRECTED TO (B)(6). D9: REPORTED AS NO. UPDATED / CORRECTED TO YES, 04/NOV/2020. H4: REPORTED AS 11/OCT/2018. UPDATED / CORRECTED TO 29/OCT/2018. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687742&pc=QCB | 2023826-2020-02382 | 2020-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.5/+6.0/175 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXPLANTED WITH A PLANNED REPLACEMENT IMPLANTATION IN 2 MONTHS. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749959&pc=QCB | 2023826-2020-02527 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/+5.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT , ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749979&pc=QCB | 2023826-2020-02528 | 2020-09-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.5/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -8.5/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, THE LENS WAS REPOSITIONED. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098630&pc=QCB | 2023826-2020-03245 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTPRY RECORD: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BENN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 TYPE OF INVESTIGATION- "4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE IN THE PREVIOUS MDR. H6- TYPE OF INVESTIGATION: 3331 SHOULD BE IN THE PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065615&pc=QCB | 2023826-2020-03110 | 2020-11-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.50/6.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024114&pc=QCB | 2023826-2020-03041 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -6.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ERROR AND LENS ROTATION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "-6.5/+2.5/095" FOR IMPLANTED LENS SPHERE/CYLINDER/AXIS SHOULD BE CORRECTED TO "-8.50/2.5/095" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024133&pc=QCB | 2023826-2020-03040 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -6.0/+2.5/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS RESIDUAL ERROR AND LENS ROTATION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: BIRTH YEAR: (B)(6). DATE OF EVENT- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11024129&pc=QCB | 2023826-2020-03049 | 2020-06-29 | Injury | STAAR SURGICAL COMPANY | 2020-12-16 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+3.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS POSTOP UCVA OF 20/50. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SPHERE, AND AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR, THE WRONG POWER LENS WAS CHOSEN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010445&pc=QCB | 2023826-2020-02979 | 2020-11-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.50/0.5/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, THE REPORTER STATED "THE DOCTOR DECIDED TO WAIT 3 MONTHS FOR THE PATIENT TO GET USED TO THE CONDITION. THEN A DECISION IS MADE ABOUT FURTHER MEASURES. THE PATIENT IS DOING WELL SO FAR." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.50/0.5/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, THE REPORTER STATED "THE DOCTOR DECIDED TO WAIT 3 MONTHS FOR THE PATIENT TO GET USED TO THE CONDITION. THEN A DECISION IS MADE ABOUT FURTHER MEASURES. THE PATIENT IS DOING WELL SO FAR." PER LATEST UPDATED MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY. IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR PATIENT CONDITION (LAST VISIT) BCVA: 20/22, VAULT 740 AND IOP 12MMHG WAS REPORTED. THE REPORTER STATED " A BETTER RESULT IS NOT POSSIBLE DUE TO THE ORIGINAL CONDITIONS." CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012259&pc=QCB | 2023826-2020-03019 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, "PATIENT IN ACCIDENT. EYE SURGERY IS NOT POSSIBLE FOR THE NEXT TWO MONTHS." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, "PATIENT IN ACCIDENT. EYE SURGERY IS NOT POSSIBLE FOR THE NEXT TWO MONTHS. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021 WHICH RESOLVED THE PROBLEM. REPORTEDLY, THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012186&pc=QCB | 2023826-2020-03008 | 2020-11-20 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984844&pc=QCB | 2023826-2020-02879 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED AND REFRACTIVE CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2018 AND (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/1.0/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WAS OBSERVED AND REFRACTIVE CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2018 AND (B)(6) 2020. THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6- HEALTH EFFECT IMPACT CODE REPORTED AS 2199 - UPDATE TO 4629. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980485&pc=QCB | 2023826-2020-02934 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/020 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH AND MODEL, DIFFERENT SPHERE AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITIH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980505&pc=QCB | 2023826-2020-02920 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -1.0/+2.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WASEXPLANTED AND REPLACED WITH A LONGER LENGTH LENS DUE TO BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5-CLAIM IS FOR LEFT (OS) AND NOT RIGHT (OD) EYE AS PREVIOUSLY REPORTED. ADDITIONAL SYMPTOMS: LENS ROTATION, VISION LOSS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INPSECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: H1- IN PREVIOUS MDR SHOULD BE SERIOUS INJURY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980516&pc=QCB | 2023826-2020-02910 | 2020-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN, LEADING TO REFRACTIVE SURPRISE. REPORTEDLY, THE LENS WAS EXCHANGED. Manufacturer Narrative: B5 - "REPORTEDLY OD EYE WAS DONE." REPLACES "REPORTEDLY, THE LENS WAS EXCHANGED." H10 - 4629 CORRECTED TO SURGICAL CODE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980478&pc=QCB | 2023826-2020-02911 | 2020-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN, LEADING TO REFRACTIVE SURPRISE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942061&pc=QCB | 2023826-2020-02845 | 2020-11-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/+5.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - "REFRACTIVE TREATMENT PERFORMED" SHOULD BE REMOVED AS IT IS NOT APPLICABLE. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT IMPACT CODE: CORRECTED TO 2199, AS 4625 IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942069&pc=QCB | 2023826-2020-02846 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -14.0/+2.00/119. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS TILTED WAS LISTED AS CAUSE OF EVENT, THE REPORTER STATED " OBSERVED DURING SURGERY AND THIS TILT CAN ALSO EXPLAIN THE REFRACTIVE SURPRISE. TILTED CERTAINLY BY BIG VAULTING." IT WAS REPORTED THAT THE PATIENT WEARS GLASSES AND THAT NO ADDITIONAL INTERVENTION IS BEING PERFORMED FOR THE MOMENT. "PUPIL CLOSED AND NOT VERY REACTIVE" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS), NOT PRODUCT RELATED "ALWAYS NOT VERY REACTIVE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ETHNICITY & RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10942215&pc=QCB | 2023826-2020-02847 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -15.50/+3.0/060. (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS TILTED WAS LISTED AS CAUSE OF EVENT, THE REPORTER STATED " OBSERVED DURING SURGERY AND THIS TILT CAN ALSO EXPLAIN THE REFRACTIVE SURPRISE. TILTED CERTAINLY BY BIG VAULTING." IT WAS REPORTED THAT THE PATIENT WEARS GLASSES AND THAT NO ADDITIONAL INTERVENTION IS BEING PERFORMED FOR THE MOMENT. "PUPIL CLOSED AND NOT VERY REACTIVE" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS), NOT PRODUCT RELATED "ALWAYS NOT VERY REACTIVE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ETHNICITY & RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937922&pc=QCB | 2023826-2020-02787 | 2020-09-30 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -5.50/2.5/161 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES, AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR (USER TRANSPOSE ERROR, PLUS TO MINUS CYLINDER RESULTED IN THE INCORRECT DATA ENTRY.). REPORTEDLY, PATIENT SATISFIED WITH VISION. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. IMPLANT DATE - UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930495&pc=QCB | 2023826-2020-02836 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, SIMILAR SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899949&pc=QCB | 2023826-2020-02740 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED POOR VISUAL ACUITY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CORRECTED DATA: B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899956&pc=QCB | 2023826-2020-02741 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.50/+4.0/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED POOR VISUAL ACUITY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CORRECTED DATA: B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905489&pc=QCB | 2023826-2020-02738 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT REFRACTIVE SURPRISE HAS BEEN OBSERVED IN PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.00/0.5/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10: "EXCESSIVE VAULT" SHOULD BE CORRECTED TO "LOW VAULT" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN. SEX: UNKNOWN. WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. COMMON DEVICE NAME-PRODUCT CODE: UNKNOWN. MODEL #: UNKNOWN. IMPLANT DATE: UNKNOWN. PMA/510K-THIS PRODUCT IS NOT MARKETED IN THE US, DEVICE MANUFACTURE DATE: UNKNOWN. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905491&pc=QCB | 2023826-2020-02726 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905564&pc=QCB | 2023826-2020-02725 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: D6A - CORRECTED TO 07-MAY-2020 IN INITIAL MDR. G4 - COMBINATION PRODUCTION IS N/A IN INITIAL MDR. CLAIM#: (B)(4) . Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851802&pc=QCB | 2023826-2020-02723 | 2020-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/4.0/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXCHANGED FOR SAME LENGTH LENS ON (B)(6) 2020 RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, "SURGEON FORGOT TO ENTER THE NEGATIVE SIGN BEFORE ASTIGMATISM DIOPTER WHEN ENTERING DATA." REPORTEDLY, "THE CORNEA IS CLEAR, THE ANTERIOR CHAMBER IS DEEP, AND THE LENS IS CENTERED." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852506&pc=QCB | 2023826-2020-02673 | 2020-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856344&pc=QCB | 2023826-2020-02700 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851799&pc=QCB | 2023826-2020-02664 | 2020-10-13 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE OVERTIME WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2020, SURGEON "HAS DECIDED TO MONITOR AND ON HOLD FOR THE LENS EXCHANGE AT THE MOMENT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10851807&pc=QCB | 2023826-2020-02724 | 2020-10-21 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7 -16.00/3.5/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "CONSIDERING ROTATION IF NEXT EXAM REVEALS ROTATION OFF AXIS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823677&pc=QCB | 2023826-2020-02598 | 2020-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- BS THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 OF -7.00/3.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE DID NOT SOLVE THE PROBLEM. D6B. (B)(6) 2021. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. SEE MFR# 2023826-2021-01504 (FOR EXCHANGE LENS). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10824095&pc=QCB | 2023826-2020-02661 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10795929&pc=QCB | 2023826-2020-02572 | 2020-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+6.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, LOSS OF BCVA, BLURRED VISION AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN (WRONG SIZE). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED. CORRECTED DATA: D10: INJECTOR MODEL MSI-PF. CLAIM # (B)(4). Manufacturer Narrative: B5: A LONGER LENS WAS IMPLANTED ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. THE PATIENT HAS NO PROBLEMS WITH THE NEW LENS AND IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11122217&pc=QCB | 2023826-2021-00072 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129308&pc=QCB | 2023826-2021-00078 | 2020-11-19 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.0/1.5/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT ,REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159195&pc=QCB | 2023826-2021-00128 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE WRONG CALCULATED LENS WAS IMPLANTED. Manufacturer Narrative: B1-UPDATED. B5-UPDATED INFORMATION: ON (B)(6) 2021, THE LENS WAS ECHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS AND PATIENT IS HAPPY. H6-HEALTH IMPACT CODE UPDATED . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159208&pc=QCB | 2023826-2021-00121 | 2020-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS; DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159253&pc=QCB | 2023826-2021-00123 | 2020-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+6.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS RESPOSITIONED DUE TO REFRACTIVE SURPRISE WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED; REFERENCE MFR REPORT# 2023826-2021-00124 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - "ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED." SHOULD BE CORRECTED TO "ON (B)(6) 2020 THE LENS WAS EXPLANTED AND LATER THE SAME DATE WAS REPLACED". H6 - HEALTH EFFECT IMPACT CODE: 4627. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159266&pc=QCB | 2023826-2021-00124 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G6/H10 - PREVIOUS SUPPLEMENTAL #1 IS TO BE DISREGARDED. CONTENTS OF THAT REPORT ARE NOT ASSOCIATED WITH THIS CLAIM. H1 - CHANGED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION (B)(4)-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161084&pc=QCB | 2023826-2021-00130 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/1.0/090(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS DIOPTER -5.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161342&pc=QCB | 2023826-2021-00131 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/1.0/090(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS DIOPTER -5.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168375&pc=QCB | 2023826-2021-00142 | 2017-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+3.0/152 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE 1494 THIS CODE WAS USED IN ERROR. THE PATIENTS AGE WAS WITHIN LABELED PARAMETERS. CLAIM (B)(4). Manufacturer Narrative: B5: A POSSIBLE CAUSE OF THE EVENT WOULD BE ICL ROTATION CAUSED THE REFRACTIVE CHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11196998&pc=QCB | 2023826-2021-00180 | 2020-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11197000&pc=QCB | 2023826-2021-00179 | 2020-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231160&pc=QCB | 2023826-2021-00208 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -05.50/+4.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VAIL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238059&pc=QCB | 2023826-2021-00174 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -13.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -13.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238071&pc=QCB | 2023826-2021-00166 | 2020-03-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/2.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2021, "PATIENT WILL GET LASIK ENHANCEMENT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC'D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239144&pc=QCB | 2023826-2021-00215 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.5/+3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THAT THE PATIENT ISN'T COMING IN FOR A FOLLOW-UP. THE PATIENT IS HAPPY WITH THEIR VISION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11239149&pc=QCB | 2023826-2021-00214 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.5/+3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ACCORDING TO THE REPORTER, THE PATIENT ISN'T COMING IN FOR A FOLLOW-UP. THE PATIENT IS HAPPY WITH THEIR VISION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11273779&pc=QCB | 2023826-2021-00228 | 2020-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.50/1.5/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, ROTATION SUSPECT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11277279&pc=QCB | 2023826-2021-00298 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/73, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. " WOUND STABLE, ICL ON AXIS, EYE COMFORTABLE PER PT. AC CLEAR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/73, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. " WOUND STABLE, ICL ON AXIS, EYE COMFORTABLE PER PT., AC CLEAR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF EVENT IS UNKNOWN. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11280319&pc=QCB | 2023826-2021-00252 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHTT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DO TO THE LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED. B5 STATEMENT CORRECTED TO: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS; -14.00/+1.0/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTER HAVING LOW VAULTING WITH LENS ROTATION AND BLURRED VISION. ON (B)(6) 2020, THE LENS WAS REPOSITIONED; AND THE ISSUE WAS NOT RESOLVED. ON (B)(6) 2021; THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS STATED TO BE UNKNOWN: ICL 80 DEGREE ROTATION FROM HORIZONTAL TO VERTICAL. CLAIM: (B)(4). Manufacturer Narrative: D9: LENS RETURNED (B)(6) 2021. B5: LENS EXCHANGED ON (B)(6) 2021 FOR A LONGER LENTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11299007&pc=QCB | 2023826-2021-00349 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: THE SURGEON SWAPPED THE LENSES FOR THE LEFT AND RIGHT EYE AND IMPLANTED THE INCORRECT LENS. THE LENS WAS CHANGED FOR THE OD EYE, WHICH WAS ASTIGMATIC AND HAD POOR QUALITY OF VISION. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11310525&pc=QCB | 2023826-2021-00253 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/176 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT COMPLAINED OF BLURRED VISION, THE PATIENT IS AMBLYOPIA AND FELT UNSATISFACTION ABOUT FAR. THE PROBLEM WAS RESOLVED. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11342672&pc=QCB | 2023826-2021-00436 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.0/1.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THIS EXPLANT RESOLVED THE PROBLEM, THE REPORTER STATED " THE PATIENT HAS RECOVERED HIS PREOPERATIVE VISION AND NO DECISION HAS BEEN MADE TO OPERATE AGAIN." CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11352786&pc=QCB | 2023826-2021-00493 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT, AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11358972&pc=QCB | 2023826-2021-00498 | 2021-01-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.00/5.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366596&pc=QCB | 2023826-2021-00486 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.0/+1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTSBLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366580&pc=QCB | 2023826-2021-00482 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372200&pc=QCB | 2023826-2021-00441 | 2020-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.00/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. "USED WRONG REFRACTION FOR THE CALCULATION." .REPORTEDLY, EXCHANGE IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400355&pc=QCB | 2023826-2021-00526 | 2020-12-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.5/+1.0/018 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400354&pc=QCB | 2023826-2021-00527 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11423056&pc=QCB | 2023826-2021-00573 | 2020-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/2.0/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -14.00/2.0/031 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN LENS VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT/DEFORMED AND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453117&pc=QCB | 2023826-2021-00601 | 2020-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "NO REPOSITION: THE TICL ROTATES UNARTIFICIALLY COUNTERCLOCKWISE 92°AFTER THE SURGERY. THE PATIENT COMPLAINED OF BLURRED VISION AND THE AUTOMATIC REFRACTOR RESULTS SHOWS A CHANGE IN THE ASTIGMATISM AXIS.(THE POSITION ARE STABLE DURING FOLLOW-UP.) Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-RESULT CODE 213 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. RESULT CODE 114 SHOULD HAVE BEEN REPORTED IN SUPPLEMENTAL MDR #1. CONCLUSION CODE 4315 SHOULD HAVE BEEN REPORTED IN INITIAL MDR CLAIM# (B)(4) Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11467987&pc=QCB | 2023826-2021-00632 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. "BLURRINESS" WAS REPORTER FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PER UPDATED INFORMATION THE PUPIL IS REACTING, IOP 24MMHG. BLURINESS HAS NOT RESOLVED. PATIENT DOES HAVE A LEFT OVER RX OF : -0.25 -0.75 X 023. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND PROBLEM RESOLVED. FINAL POST-OP ON (B)(6) 2021 REPORTS: UCVA 20/20 AND IOP 18MMHG AND FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTED STATED " EXCELLENT OUTCOME. PUPIL ROUND AND REACTIVE WITH NORMAL IOP." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507463&pc=QCB | 2023826-2021-00677 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507471&pc=QCB | 2023826-2021-00675 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: INITIALLY THE PATIENT'S OBJECTIVE REFRACTION VALUE WAS DIFFERENT FROM THE SUBJECTIVE REFRACTION VALUE. THE DOCTOR PERFORMED MANY TESTS, BUT THE SUBJECTIVE REFRACTION VALUE SHOWED THE SAME VALUE. THEREFORE, THE DOCTOR DECIDED THE LENS POWER WITH AN EMPHASIS ON THE SUBJECTIVE REFRACTION VALUE. HOWEVER, THE CYLINDRICAL POWER WAS INSUFFICIENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507474&pc=QCB | 2023826-2021-00662 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+6.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5- THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+6.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. ON (B)(6) 2021 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A SAME MODEL AND SIZE BUT DIFFERENT POWER REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D6B-(B)(6) 2021. H1- SERIOUS INJURY. H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/ DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507487&pc=QCB | 2023826-2021-00676 | 2021-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11515292&pc=QCB | 2023826-2021-00586 | 2020-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED (B)(6) 2021, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC SPLIT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED (B)(6) 2021. THE LENS WAS REPLACED WITH A SAME LENGTH, HIGHT POWER (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "PATIENT HEALED OFF TARGET." REPORTEDLY, "HEALING WELL, APPEARS UNTARGET AT INTIAL POST OP." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558783&pc=QCB | 2023826-2021-00723 | 2021-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE 2 MONTHS POSTOPERATIVELY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1-UPDATED. B5-UPDATED: ON (B)(6) 2021 THE LENS WS REPOSITIONED. THERE IS RESIDUAL REFRACTION AND THE LENS "APPEARS TO BE SITTING OFF-AXIS A BIT." H6-HEALTH IMPACT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+0.5/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2021. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. DATE REC¿D BY MFR-UNK. DATE REC¿D BY MFR-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE -UNK. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563257&pc=QCB | 2023826-2021-00763 | 2020-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -04.50/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH, LENS WITH A -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) POWER ON (B)(6) 2020 TO RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 -04.50/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SUPRISE REPORTED. THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) POWER ON (B)(6) 2020 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563279&pc=QCB | 2023826-2021-00753 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -8.00/2.0/054 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (10.00/2.0/050(SPHERE/CYLINDER/AXIS)) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "STILL LETTLE EDEMA OF THE CORNEA, IMPROVING.". PER THE REPORTER ON (B)(6) 2021, "ACTUALLY THE PATIENT CAN SEE 20/20 WITHOUT CORRECTION AND VAULT IS CORRECT. HE PREVIOUSLY HAD A LITTLE BIT OF A HIPEROPIA AND A VERY UNSATISFACTORY CLOSE VISION. MUCH BETTER NOW." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND WITH SPOTS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620609&pc=QCB | 2023826-2021-00880 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -3.0/+4.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SUPRISE, RESIDUAL ASTIGMATISM, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE: PREOPERATIVE AND POSTOPERATIVE COURSE WAS SMOOTH WITHOUT ANY COMPLICATION. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: "HIGHER TOTAL ASTIGMATISM MORE DUE TO LENTAL ASTIGMATISM; VOD=0.05CC+5.50X90=0.1." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11632409&pc=QCB | 2023826-2021-00940 | 2021-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN 2012. ZERO VAULTING AND REFRACTIVE CHANGE OVERTIME WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THA A 12.6MM, VTICM12.6, -5.5/6.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. LOW VAULT WAS OBSERVED, PATIENT IS FINE AND THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654115&pc=QCB | 2023826-2021-00960 | 2021-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -13.5/+1.5/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11676163&pc=QCB | 2023826-2021-01056 | NA | Injury | STAAR SURGICAL COMPANY | 2021-04-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/4.0/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS IMPLANTED AS A REPLACEMENT LENS TO CORRECT REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS REPOSITIONED, IT WAS REPORTED THAT AFTER REPOSITIONING THE LENS UCVA WAS 20/25. THE REPORTER STATED " THE CURRENT VISUAL OUTCOME IS GOOD AND THE REFRACTIVE SURPRISE PROBLEM IS RESOLVED WITH REPOSITIONING AS OF NOW. PROBLEM WAS RESOLVED AND HE WILL CONTINUE TO MONITOR". Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11676122&pc=QCB | 2023826-2021-01055 | 2020-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION WAS OBSERVED. LENS REPOSITIOING WAS PERFORMED BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME DIOPTER, LONGER LENGTH LENS BUT IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE. THE REPORTER STATED " THE SURGEON DECIDED THAT THE VISUAL ACUITY WAS NOT GOOD DUE TO THE ROTATION AT THAT TIME. HE PERFORMED REPOSITIONING SURGERY, BUT THE OUTCOME WAS NOT BETTER. SO, HE EXCHANGED THE LENS WITH SAME DIOPTER. THE DOCTOR MEASURED THE DATA AGAIN BEFORE THE EXCHANGE SURGERY, AND VAULT WAS LOWER THAN BEFORE. SO, HE DECIDED TO REPLACE WITH A LARGER LENGTH LENS. AFTER EXCHANGING THE LENS, THE VAULT WAS 0.94CT ((B)(6) 2021)." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697123&pc=QCB | 2023826-2021-01081 | NA | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.5/+5.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B1-UPDATED. B5-UPDATED INFO: THE REPORTER STATES THE "DOCTOR SOLVED THIS CASE ROTATING THE LENS." H6-HEALTH IMPACT CODE UPDATED. INVESTIGATION CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11714600&pc=QCB | 2023826-2021-01131 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/2.0/079 9SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721900&pc=QCB | 2023826-2021-01017 | 2019-05-25 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REPORTEDLY, BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND , THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS POSITION IS STABLE DURING FOLLOW-UP. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721910&pc=QCB | 2023826-2021-01026 | 2018-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS ROTATED CLOCKWISE 90 DEGREES AFTER SURGERY. BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WERE OBSERVED. ON (B)(6) 2018 THE LENS WAS REPLACED FOR A SAME LENGTH LENS POSITIONED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721911&pc=QCB | 2023826-2021-01027 | 2018-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/3.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS ROTATED CLOCKWISE 90 DEGREES AFTER SURGERY. BLURRED VISION AND CHANGE IN THE ASTIGMATISM AXIS WERE OBSERVED. ON (B)(6) 2018 THE LENS WAS REPLACED FOR A SAME LENGTH LENS POSITIONED VERTICALLY; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721923&pc=QCB | 2023826-2021-01045 | 2021-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/0.5/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721928&pc=QCB | 2023826-2021-01046 | 2021-03-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/0.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721885&pc=QCB | 2023826-2021-01004 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.50/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA; H3-DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721890&pc=QCB | 2023826-2021-01008 | 2021-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/6.0/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, "PATIENT IS NOT HAPPY WITH VISION." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED "SURPRISE REFRACTION OR REFRACTIVE CHANGE OVER TIME AFTER ALMOST 2 YEARS." PATIENT NOT HAPPY WITH HIS VISION. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721908&pc=QCB | 2023826-2021-01024 | 2018-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2017. LEN OPACITY (NS) WAS OBSERVED ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSEVERED. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US . CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721919&pc=QCB | 2023826-2021-01033 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER LENS REPLACEMENT, PATIENT HAD 1.0 VISION AND FELT VERY SATISFIED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11729939&pc=QCB | 2023826-2021-01238 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: XIAOJUN821229. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # ((B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734843&pc=QCB | 2023826-2021-01180 | 2020-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+5.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: LOW VAULT WAS ALSO OBSERVED. THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734852&pc=QCB | 2023826-2021-01179 | 2018-01-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS RESIDUAL ASTIGMATISM AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5- ADDITIONAL INFO: A 12.5MM TICM12.5, -7.5/+2.0/164 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2012. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.5MM TICM 12.5 OF -7.5/2.0/164 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2012. TICL ROTATED 1 DEGREE COUNTERCLOCKWISE AFTER IMPLANTATION. REFRACTIVE SURPRISE WAS NOTED. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750600&pc=QCB | 2023826-2021-01297 | 2019-02-21 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED . THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS (DIFFERENT AXIS) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750604&pc=QCB | 2023826-2021-01298 | 2019-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751162&pc=QCB | 2023826-2021-01282 | 2018-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779017&pc=QCB | 2023826-2021-01329 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OD -8.50/2.5/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE AND DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2021. H6- HEALTH IMPACT CODE 4629- LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779026&pc=QCB | 2023826-2021-01330 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OD -8.0/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE AND DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2021. H6- HEALTH IMPACT CODE 4629- LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11787301&pc=QCB | 2023826-2021-01387 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/0.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. SECONDARY INTERVENTION/ ADDITIONAL TREATMENT HAS NOT BEEN PERFORMED OR DECIDED UPON. FOR STATUS OF THE EYE " PATIENT CANNOT READ ANYMORE " WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER REPORTED " REFRACTIVE SURPRISE. DOCTOR WANTS TO DO A LASER TO SOLVE THE REFRACTIVE SURPRISE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781988&pc=QCB | 2023826-2021-01384 | 2021-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/4.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTD INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. "SIGNIFICANT REFRACTIVE SURPRISE AFTER RE ICL IMPLANTATION" WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITRIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/4.0/155 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE (INCREASED ASTIGMATISM). THE LENS WAS EXPLANTED AND A REPLACEMENT LENS WAS IMPLANTED. REPORTEDLY "THE ORIGINAL LENS WAS EXPLANTED AND REPLACED WITH A NON-TORIC - VA WAS AROUND (B)(6) UNAIDED AND VAULT OK". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788224&pc=QCB | 2023826-2021-01349 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -4.0/+1.0/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) -2020. ON (B)(6) -2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788229&pc=QCB | 2023826-2021-01348 | 2021-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4) : LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820908&pc=QCB | 2023826-2021-01406 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.50/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820910&pc=QCB | 2023826-2021-01407 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823183&pc=QCB | 2023826-2021-01480 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS -13.0/+1.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS RESIDUAL ASTIGMATISM AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823341&pc=QCB | 2023826-2021-01455 | 2020-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+0.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823528&pc=QCB | 2023826-2021-01440 | 2021-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.00/3.00/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WA REPORTED TO BE UNKNOWN, THE REPORTER STATED "IN DR. OPINION MUST HAVE A MISTAKE ABOUT ASTIGMATISM. THE MANIFEST REFRACTION AND THE ORDERED IOL WAS -3.00 D OF ASTIGMATISM AND THAT IS ALSO THE ASTIGMATISM THE PATIENT HAVE IN HIS GLASSES. HOWEVER TOPOGRAPHIC ASTIGMATISM WAS NEAR 4.00D. ALSO IN THE CONTRALATERAL EYE ( RIGHT ) BOTH MANIFEST AND TOPOGRAPHIC ASTIGMATISM WERE 4.25D AND WITH 4.50D ORDERED IN THE ICL LENSES THE PATIENT IS FINE. DR. (B)(6) BELIEVES THAT A REPLACEMENT OF THE IOL FOR A NEW ONE WITH 4.00 D OR 4.25 D OF ASTIGMATISM WILL BE OK." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021 AND LATER REPLACED WITH A LONGER LENGTH LENS OF DIFFERENT POWER AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823603&pc=QCB | 2023826-2021-01441 | 2020-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6 MM, TMICL12.6, -9.00/3.00/067(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION (TORIC ICL CONTINUED TO ROTATE) AND BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CONTAINER, IN A STERILE POUCH. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS REPORTED THAT A 12.6MM TMICL12.6 -9.00/3.00/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS SIZE TOO SMALL; BLURRED VISION AND LENS ROTATION WERE REPORTED. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11829659&pc=QCB | 2023826-2021-01538 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. PER UPDATED INFORMATION PATIENT HAS LEFT OVER RX OF -0.25 -0.50X026. PUPIL IS REACTING IOP 23MMHG. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER LATEST UPDATE, A MEDICAL/SECONDARY SURGICAL INTERVENTION WAS PERFORMED AND THE PROBLEM RESOLVED; HOWEVER NO DETAILS WERE PROVIDED. FINAL POST-OP ON (B)(6) 2021 REPORTS: UCVA 20/25 AND IOP 15MMHG AND FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTED STATED " EXCELLENT OUTCOME. PUPIL ROUND AND REACTIVE WITH NORMAL IOP. PATIENT SATISFIED. PRN." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Unknown | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839592&pc=QCB | 2023826-2021-01558 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6MM, -4.50/+4/71 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839593&pc=QCB | 2023826-2021-01557 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6MM, -4.50/+4.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855234&pc=QCB | 2023826-2021-01516 | 2020-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL OBSERVATIONS FOUND HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: DATE RECD BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11886266&pc=QCB | 2023826-2021-01614 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WILL BE EXPLANTED ON (B)(6) 2021. DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IS HAPPY. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891207&pc=QCB | 2023826-2021-01699 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -10.0/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891209&pc=QCB | 2023826-2021-01698 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -9.5/1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11899111&pc=QCB | 2023826-2021-01682 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -15.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE LENS HAPTIC, BENT DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION WAS PERFORMED, AND THE DIAMETER VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION PERFORMED FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6- WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11904673&pc=QCB | 2023826-2021-01689 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -7.0/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED DRY, INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC STRETCHED OUT, HAPTIC DEFORMED, AND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE OUT OF SPECIFICATION (ONLY LENS SPHERE, CYLINDER AND AXIS WITHIN SPECIFICATION).. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: A2: DOB (B)(6) 1992. CLAIM (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11904725&pc=QCB | 2023826-2021-01690 | 2020-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL WIDTH AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. UNABLE TO PERFORM LENGTH VERIFICATION, DUE TO HAPTIC FOOTPLATES DAMAGE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938949&pc=QCB | 2023826-2021-01648 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/5.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938956&pc=QCB | 2023826-2021-01649 | 2021-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/4.0/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939030&pc=QCB | 2023826-2021-01673 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938963&pc=QCB | 2023826-2021-01658 | 2020-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED . B3: DATE CHANGED TO 11/10/2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938964&pc=QCB | 2023826-2021-01657 | 2021-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.00/0.05/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH LIQUID ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942673&pc=QCB | 2023826-2021-01780 | 2020-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942675&pc=QCB | 2023826-2021-01778 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942687&pc=QCB | 2023826-2021-01772 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. H6-METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT; UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966802&pc=QCB | 2023826-2021-01813 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020 . REFRACTIVE SURPRISE WAS OBSERVED. REFRACTIVE TREATMENT BY CORNEAL T-CUT WAS PERFORMED. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, AN EXCHANGE IS PLANNED AND WILL UPDATE WHEN EXCHANGE IS DONE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT. H1: SERIOUS INJURY. CLAIM : (B)(4). Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, DATE OF EVENT- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986307&pc=QCB | 2023826-2021-01875 | 2021-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED NIGHT VISION DUE TO LARGE PUPIL SIZE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005350&pc=QCB | 2023826-2021-01939 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-06-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+4.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT WAS REPORTED AS COMPLAINING OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005489&pc=QCB | 2023826-2021-01940 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-06-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT WAS REPORTED AS COMPLAINING OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009118&pc=QCB | 2023826-2021-01896 | 2021-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/.05/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, THERE IS STILL A RESIDUAL REFRACTION, AN EXCHANGE IS PLANNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/.05/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, THERE IS STILL A RESIDUAL REFRACTION, AN EXCHANGE IS PLANNED." ON (B)(6) 2021 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. REPORTEDLY, ALL FINE. CLINIC TOLD THERE IS GOOD VAULTING. PATIENT IS HAPPY!. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021861&pc=QCB | 2023826-2021-02001 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/176 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. DIMENSION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029743&pc=QCB | 2023826-2021-01916 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.0/1.0/16 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM, VTICM5_ 12.6, -5.0/1.0/16 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THIS EXCHANGED RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060456&pc=QCB | 2023826-2021-01953 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.00/2.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BROKEN, BENT, DEFORMED, AND STRETCHED OUT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060470&pc=QCB | 2023826-2021-01954 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060480&pc=QCB | 2023826-2021-01966 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060483&pc=QCB | 2023826-2021-01967 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.00/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE, SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064274&pc=QCB | 2023826-2021-02062 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+5.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING, WITH LENS ROTATION. THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED, THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENS. AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND LENS IS IN POSITION. PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H3: DEVICE EVALUATION IS NOT REQUIRED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072254&pc=QCB | 2023826-2021-02088 | 2020-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072262&pc=QCB | 2023826-2021-02075 | 2021-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.5/+3.0/049 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE SURGEON STATES, "THE LENS IS WELL ALIGNED IN PLACE AND THE EYE IS QUITE AND STABLE BUT THE SURPRISE REFRACTION IS UNEXPLAINABLE TO US." Manufacturer Narrative: B5- SURGEON PLANS TO REMOVE THE LENS. RECEIVED UPDATE ON 24-MAR-2022 "PATIENT JUST SHOW UP TODAY SAME POST OP REFRACTION 3 D ASTIGMATISM. SURGEON LOOKING FOR ADVISE, HE IS PLANNING TO REMOVE THE LENS AND ACCESS AGAIN". CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072269&pc=QCB | 2023826-2021-02097 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. PER EMAIL, "ALL IS WELL WITH THE PATIENT." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072274&pc=QCB | 2023826-2021-02072 | 2021-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/+1.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B4: IN FOLLOW-UP 1 ENTERED IN ERROR. SHOULD BE BLANK AND DATE ENTERED IN D6B. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -9.5/+1.0/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS REPORTED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH/DIFFERENT DIOPTER LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - CORRECTION TO INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102507&pc=QCB | 2023826-2021-02132 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND POOR UNAIDED VISION WERE OBSERVED. LENS REMAINS IMPLANTED AND AN EXCHANGE IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: 1 SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102535&pc=QCB | 2023826-2021-02131 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "CHANGED POSITION AND NOW OFF AXIS AND POOR UNAIDED VISION. " ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE CHANGE OVERTIME AND POOR UNAIDED VISION WERE OBSERVED. LENS REMAINS IMPLANTED AND A LENS EXCHANGE IS PLANNED. TH REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN AND STATED "BASICALLY VISION AND TICL POSITION CHANGED AFTER 6/12 AND NOW OFF AXIS AND POOR UNAIDED VISION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED CQ: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.50/+4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2019. LOW VAULT WITH ROTATION, REFRACTIVE CHANGE OVERTIME AND POOR UNAIDED VISION WERE OBSERVED. LENS WAS EXPLANTED IN (B)(6) 2021 AND REPLACED WITH A LONGER LENGTH LENS WITH SIMILAR POWER (VTICMO 13.2, -15.50/+4.0/095) ON (B)(6) 2021. THE PROBLEM WAS RESOLVED AND "VAULT HAS IMPROVED AND PATIENT IS HAPPY". CAUSE OF EVENTS WAS REPORTED AS UNKNOWN. D6B- (B)(6) 2021. H6 DEVICE CODE: 1583 ADDED; LOW VAULT WITH LENS ROTATION REPORTED. H6 DEVICE CODE: 1494: OFF LABEL USE- AGE<21 YEARS OLD. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102798&pc=QCB | 2023826-2021-02145 | 2021-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: HE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D4: WRONG SERIAL NUMBER PROVIDED IN INITIAL MDR B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -2.50/+6.0/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). REFRACTIVE SURPRISE AND BLURRY VISION REPORTED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 WHICH DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. REPORTEDLY MISTAKE WAS MADE WITH MIXING UP THE SERIAL NUMBERS. THEY ORDERED THE SAME LENS FOR THE SAME PATIENT UNDER 2 DIFFERENT DOCTOR NAMES. AN ANALYSIS WAS MADE BASED ON THE WRONG LENS. THE INFORMATION CORRECT LENS IMPLANTED HAS A DIFFERENT IOD AND NOW PROBABLY A NEW ANALYSIS IS MADE WITH THE CORRECT IOD. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112717&pc=QCB | 2023826-2021-02226 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/1.0/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AFTER IMPLANT WAS OBSERVED. " THE PATIENT HAS GOT AN HYPEROPIA AND STILL SMOOTH LOSS OF UCVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATEINT IS HAPPY WITH THE RESULTS AND WEARS GLASSES IF NECESSARY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112726&pc=QCB | 2023826-2021-02225 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.0/86, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AFTER IMPLANT WAS OBSERVED. " THE PATIENT HAS GOT AN HYPEROPIA AND STILL SMOOTH LOSS OF UCVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATIENT IS HAPPY WITH THE RESULTS AND WEARS GLASSES IF NECESSARY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147536&pc=QCB | 2023826-2021-02268 | 2021-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, "PENTACAM HAS NOT BEEN PERFORMED YET." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147853&pc=QCB | 2023826-2021-02481 | 2021-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.00/4.00/092, IMPLANTABLE COLLAMER LENS WAS IMLPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. . IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT; B2: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE; B5: THE REPORTER INDICATED, THAT A 12.6MM, VTICM5 12.6, -11.00/4.00/092, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME SIZE SPHERICAL LENS, DUE TO REFRACTIVE SURPRISE. RESIDUAL ASTIGMATISM WAS CORRECTED BY EXCIMER LASER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS DID NOT MEET THE ORIGINAL VALUES MEASURED AT TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147855&pc=QCB | 2023826-2021-02497 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.0/100, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H10: IT WAS LATER CLARIFIED THAT THIS CLAIM WAS REPORTED IN ERROR WITH INCORRECT INFORMATION AND SEE MFR#: 2023826-2021-04168 WHICH HAS THE CORRECT DETAILS (CLAIM#: (B)(4). DISREGARD INITIAL MDR SUBMISSION. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148279&pc=QCB | 2023826-2021-02476 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -13.5/+6.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND POOR NEAR VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148267&pc=QCB | 2023826-2021-02470 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+6.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. REPORTEDLY AN ADDITIONAL SUTURE WAS ADDED AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE LENS WAS REPOSITIONED AGAIN AND THE PATIENT IS "FINE FOR NOW." CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148274&pc=QCB | 2023826-2021-02471 | 2021-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12160678&pc=QCB | 2023826-2021-02556 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -10.0/4.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH AND MODEL LENS BUT DIFFERENT SPHERE/AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND DEPOSIT ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167123&pc=QCB | 2023826-2021-02530 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS TOO SHORT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167124&pc=QCB | 2023826-2021-02531 | NA | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0 12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -15.00/4.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS TOO SHORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12208007&pc=QCB | 2023826-2021-02760 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -14.5/2.0/162, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT AXIS LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, THE REPORTER STATED " THE SURGEON INPUT THE WRONG ASTIGMATISM AXIS IN OCOS SYSTEM". Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12211083&pc=QCB | 2023826-2021-02758 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN /STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214331&pc=QCB | 2023826-2021-02694 | 2021-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND FIBER ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214338&pc=QCB | 2023826-2021-02696 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.50/3.50/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: ON (B)(6) 2021, THE LENS WAS EXPLANTED AND AT A UNKNOWN DATE, PHACO-EMULSIFICATION AND IOL IMPLANTATON WAS PERFORMED. THE PROBLEM RESOLVED. REPORTEDLY, "PHACO WITH MIOL. PATIENT IS VERY HAPPY!" CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214343&pc=QCB | 2023826-2021-02695 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/3.00/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - ON (B)(6) 2021, THE LENS WAS EXPLANTED. CLAIM #: (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2021, THE LENS WAS EXPLANTED AND AT A UNKNOWN DATE, PHACO-EMULSIFICATION AND IOL IMPLANTATON WAS PERFORMED. THE PROBLEM RESOLVED. REPORTEDLY, "PHACO WITH MIOL. PATIENT IS VERY HAPPY!" CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222277&pc=QCB | 2023826-2021-02815 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222317&pc=QCB | 2023826-2021-02805 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS -14.5/+2.0/064 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS A SMALL REFRACTIVE ERROR AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12255959&pc=QCB | 2023826-2021-02976 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/015 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR BUT THE DEVICE DID FAIL TO PERFORM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257216&pc=QCB | 2023826-2021-02971 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -17.5/+1.0/015 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257224&pc=QCB | 2023826-2021-02964 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+5.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT-UNK IMPLANT DATE-UNK PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257227&pc=QCB | 2023826-2021-02970 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -16.5/+0.5/178 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12263420&pc=QCB | 2023826-2021-02555 | 2021-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.00/+2.50/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT "THE RESULT IS OK FOR THE PATIENT AND WISH NO FURTHER SURGICAL INTERVENTION". PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278288&pc=QCB | 2023826-2021-03053 | 2021-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -4.0/+4.0/140 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. MANIFEST REFRACTION IS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-ADDITIONAL AND CORRECTED INFORMATION: THE LENS WAS IMPLANTED ON (B)(6) 2014. REPORTEDLY, "THE RESULT WAS OK AND THE PATIENT WAS HAPPY" AFTER THE INITIAL LENS WAS IMPLANTED. REFRACTIVE CHANGE OVER TIME IS REPORTED DUE TO MYOPIA GROWTH. (B)(4). Manufacturer Narrative: SEX WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281699&pc=QCB | 2023826-2021-03066 | 2021-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.MM, VTICH13.2, -5.00/5.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281714&pc=QCB | 2023826-2021-03067 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.MM, VTICH13.2, -5.00/5.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321646&pc=QCB | 2023826-2021-03148 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.50/1.0/171, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12326006&pc=QCB | 2023826-2021-03182 | NA | Injury | STAAR SURGICAL COMPANY | 2021-08-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2MM; -1.0/+6.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE; LENS ROTATION WHICH THE SURGEON PERFORMED REPOSITIONING ON (B)(6) 2021 WHICH DID IMPROVE VISION AND REFRACTIVE BUT AGAIN THE VISION HAS BLURRED AND LENS HAD ROTATED AGAIN. THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378554&pc=QCB | 2023826-2021-03370 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.5/+0.5/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SUPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -10.5/+0.5/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378674&pc=QCB | 2023826-2021-03369 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -9.0/+0.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SUPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -9.0/+0.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND EXCESSIVE VAULT ARE REPORTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426115&pc=QCB | 2023826-2021-03452 | NA | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS "REDIALED" DUE TO LENS ROTATION AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS ROTATED AGAIN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443369&pc=QCB | 2023826-2021-03537 | 2021-08-03 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457493&pc=QCB | 2023826-2021-03557 | 2021-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0/+6.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. AT A LATER DATE A SAME SIZE DIFFERENT SPHERE/AXIS LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON REPORTS, "WE IMPLANTED THE TICL AND WE GOT A REFRACTIVE SURPRISE AND LOW VA OF 0.1...SO WE REMOVED THE LENS AND DID SUBJECTIVE REFRACTION, THERE WAS A CHANGE IN REFRACTION, SO I REQUESTED A TRANSPOSED BCVA TEST, WE ACHIEVED THE 20/40 WHICH IS THE TARGET." Manufacturer Narrative: B5-THE REPORTER NOW REPORTS PERMANENT LOSS OF BCVA AS THE REASON FOR THE LENS CHANGE. ALSO REPORTED, "ICL WAS IMPLANTED AND HAD A REFRACTIVE ERROR. I REQUESTED A NEW MANIFEST REFRACTION AND BCVA WITH A TRANSPOSED REFRACTION AND CONCLUDED THAT THE OLD LENS REFRACTION WAS NOT ACHIEVING THE TARGET SO WE HAD TO REDO THE REFRACTION AGAIN." CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES, "AFTER THE PATIENT IMPLANTED THE FIRST ICL HE HAD REDUCTION IN THE BCVA. AFTER EXCHANGING IT THE NEW LENS RETRIEVED THE BCVA." AFTER EXCHANGE THE BCVA IS 0.5. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION CODE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457494&pc=QCB | 2023826-2021-03558 | 2021-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/+2.0/059 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER DIFFERENT MODEL LENS DUE TO LENS ROTATION AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. THE REPORTER STATES, "THERE WERE BIG DIFFERENCES IN WTW MEASUREMENT BY OPD, IOL MASTER AND CALIPER." Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE LENGTH STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479353&pc=QCB | 2023826-2021-03598 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/2/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AS REPLACEMENT LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. THE LENS ROTATED AND THE PATIENTS VISION AT EARLY POSTOPERATIVE EXAMINATION WAS ONLY 20/40. THE REPORTER STATED "THE DOCTOR DECIDED TO CONTINUE THE OBSERVATION. THE PATIENT WAS LAST REVIEWED ON (B)(6) 2021 AND HIS VISUAL ACUITY REACHED 20/20. THE SURGEON CONSIDERED THAT THE REASON FOR 20/20 VISION WAS DUE TO THE EFFECT OF THE CORNEAL INCISION AND THE POSITION OF THE LENS BUT DID NOT DILATE THE PUPIL." THE CORNEAL INCISION WAS PERFORMED WITHIN THE SAME SURGERY AS IMPLANTATION OF THE LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479078&pc=QCB | 2023826-2021-03590 | 2021-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM512.1, -8.00/2.50/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479327&pc=QCB | 2023826-2021-03657 | 2021-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS LATER REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS BUT THE PROBLEM DID NOT RESOLVE. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12492856&pc=QCB | 2023826-2021-03651 | 2021-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND OPTIC AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL VERIFICATION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524398&pc=QCB | 2023826-2021-03718 | 2021-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-11.5/1.0/68, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND REFRACTIVE SURPRISE. ON (B)(6) 2021 A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT DIOPTER WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12537688&pc=QCB | 2023826-2021-03697 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE PLAN IS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH ANOTHER ICL LENS (SPHERICAL) AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - LENS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE : NA. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12550693&pc=QCB | 2023826-2021-03794 | 2021-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.50/5.0/02 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2021 DUE TO A MINOR ROTATION. IT WAS REPORTED THAT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED AND ON (B)(6) 2021 A REPLACEMENT LENS OF DIFFERENT POWER WAS IMPLANTED, PATIENT IS IN EXCELLENT CONDITION AND SATISFIED. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR, LENS POWER WAS NOT CALCULATED CORRECTLY. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558537&pc=QCB | 2023826-2021-03808 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR, "MAY BE CAUSE OF THE OCULAR ANATOMY." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599770&pc=QCB | 2023826-2021-03908 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/+2.0/020 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "PATIENT IS UNHAPPY POST OPERATIVELY AS HE ACCEPTS +1/-1.75@145. WITH THIS CORRECTION PATIENT IMPROVES TO 20/20." Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY THE LENS WAS REPOSITIONED. AS OF 1-DAY POSTOP THE PATIENT IS "DOING WELL.". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643335&pc=QCB | 2023826-2021-04012 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/2.0/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. IT WAS REPORTED THAT IT WAS UNKNOWN WHEN THE DAMAGE OCCUR AND THAT THE DAMAGE WAS NOTED WHEN WITHDRAWING THE LENS. LOSS OF BCVA WAS ALSO NOTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-PER UPDATED INFORMATION THE REPORTER STATED " THE DOCTOR NOTICED THAT THE LENS WAS BROKEN, WHEN THE NEXT DAY, WHEN CHECKING THE EYE. THE PATIENT'S EYE IS FINE. HE BELIEVES THAT THE FAILURE MUST HAVE BEEN WHEN INSERTING THE LENS INTO THE EYE WAS BROKEN IN THE INJECTION." CLAIM # (B)(4). Manufacturer Narrative: G3- CORRECTED TO (B)(6) 2021, IN INITIAL MDR. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643795&pc=QCB | 2023826-2021-04018 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.00/+3.0/083, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 03 SEP 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. " PATIENT IS ACCEPTING +1/-2.5@12 AND IMPROVING TO 6/5" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "REPORTEDLY PER EMAIL ON 06/06/2022, 'SURGEON RE-ROTATED THE LENS ON (B)(6) 2022 TO POSITION. THE REFRACTION WAS REPORTED AS +1.5/-2.5/30. THE PATIENT IS HAPPY WITH 6/6 VISION.'" SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B5 - "REPORTEDLY, 'LAST MONTH THIS PATIENT VISITED THE SURGEON ON (B)(6) 2022 AND THAT TIME THE REFRACTION WAS +1.5/-2.5/30. SURGEON RE-ROTATED THE LENS ON (B)(6) 2022 TO POSITION AND THE PATIENT IS HAPPY WITH 6/6 VISION.'" SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12678194&pc=QCB | 2023826-2021-04106 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/3.0/122, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS NOTED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/ MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: .UNK PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724193&pc=QCB | 2023826-2021-04156 | NA | Injury | STAAR SURGICAL COMPANY | 2021-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -7.5/+2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED AND THE LENS ROTATED AGAIN. REPORTEDLY, "THE PATIENT COMPLAINED OF REDUCTION OF VISION. AFTER THE FIRST SURGERY THE IOL WAS AT 4 DEGREES. THE TARGET WAS 178 DEGREES. REFRACTION WAS +0.50/-1.00X160." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725216&pc=QCB | 2023826-2021-04179 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.50/018, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PER UPDATED INFORMATION AN EXCHANGED WAS CARRIED OUT ON (B)(6) 2021. THE REPORTER STATED THAT THE SURGEON HAS CONFIRMED THAT THE PATIENT IS NOW DOING WELL. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725051&pc=QCB | 2023826-2021-04168 | 2021-03-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.0/102, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE PRODUCT WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725068&pc=QCB | 2023826-2021-04167 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/172, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR, THE REPORTER STATED "THE DOCTOR DID NOT ENTER WHEN CALCULATING THE CYLINDER." Manufacturer Narrative: B3- CORRECTED TO (B)(6) 2021 IN INITIAL MDR. D4 MODEL #: CORRECTED TO VTICMO12.6 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE PRODUCT WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED LENS FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING; H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749569&pc=QCB | 2023826-2021-04228 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/1.5/101, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS CYLINDER TO BE OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755752&pc=QCB | 2023826-2021-04209 | 2021-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.5/+0.5/116 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: H6: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY THE SYMPTOMS HAVE RESOLVED AND NO INTERVENTION IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755753&pc=QCB | 2023826-2021-04208 | 2021-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.5/+0.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "THE REFRACTION WAS REMINDED UNCORRECTED (OR WORSE) AT 1D POST-OPE AND THE PATIENT GOT VERY ANGRY ABOUT THE RESULT, BUT AT 1W, THE REFRACTION BECAME NORMAL AS EXPECTED. WE DON'T HAVE ANY ESTIMATIONS OF REASON WHY THE TRANSIENT REFRACTIVE SURPRISE WAS HAPPENED." Manufacturer Narrative: B5 ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT'S SYMPTOMS HAVE RESOLVED AND NO INTERVENTION IS PLANNED. (B)(4). Manufacturer Narrative: H6: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763784&pc=QCB | 2023826-2021-04231 | 2021-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.5/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807798&pc=QCB | 2023826-2021-04263 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/174 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808809&pc=QCB | 2023826-2021-04333 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/+0.5/138, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION ON 27-DEC-2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND WIDTH TO BE WITHIN SPECIFICATION; HOWEVER, FUNCTIONAL INSPECTION FOUND THE LENS AXIS OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849444&pc=QCB | 2023826-2021-04442 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+1.50/97, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872781&pc=QCB | 2023826-2021-04487 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING, LENS ROTATION AND BLURRED VISION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872828&pc=QCB | 2023826-2021-04488 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+4.0/079 (SPHERE/ CYLINDER/ AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING, LENS ROTATION AND BLURRED VISION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879840&pc=QCB | 2023826-2021-04526 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.5/5.0/100, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN/BROKEN . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12911285&pc=QCB | 2023826-2021-04623 | 2021-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/3.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH SIMILAR POWER DIFFERENT CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973012&pc=QCB | 2023826-2021-04653 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+5.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973346&pc=QCB | 2023826-2021-04707 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/000 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISIBLE INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973347&pc=QCB | 2023826-2021-04711 | 2020-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-REPORTEDLY THE SURGEON REPOSITIONED THE LENS AND AS OF (B)(6) 2021 THE UCVA IS 20/20. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973353&pc=QCB | 2023826-2021-04697 | 2020-03-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023235&pc=QCB | 2023826-2021-04887 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -12.0/+0.5/119 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13040123&pc=QCB | 2023826-2021-04906 | 2021-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.5/3.0/95 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND IT WAS INDICATED THAT THE 13.2MM LENS CAUSED A OVERCORRECTION; WHICH MADE IT DIFFICULT TO READ UP CLOSE FOR THE PATIENT. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM AND NOTED "MYOPIA UNDER CORRECTION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MDR # 2023826-2021-04907 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082112&pc=QCB | 2023826-2021-04987 | 2021-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -5.5/5.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATE RECEIVED ON 09-MARCH-2022 "NO REPLACEMENT WILL BE PLANNED. VAULTING WILL BE MONITORED. RESIDUAL REFRACTION WILL BE TREATED BY LASER TOUCH UP." THE PROBLEM WAS MARKED AS RESOLVED. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082432&pc=QCB | 2023826-2021-04962 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AGE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AGE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082443&pc=QCB | 2023826-2021-04968 | 2021-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "CILIARY BODY CYSTS ROTATING ICL TO 165 DEGREES AND ROTATION ICL ON (B)(6) 2021 DIDN'T RESOLVED THE PROBLEM. ON 1ST DAY AFTER SURGERY ICL WAS AGAIN ROTATED TO 165 DEGREES." Manufacturer Narrative: (B)(6). PATIENT EIGHT, ETHNICITY AND RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090705&pc=QCB | 2023826-2021-04997 | 2021-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/3.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115042&pc=QCB | 2023826-2021-05018 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.5/1.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT LENGTH/POWER LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115047&pc=QCB | 2023826-2021-05021 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.50/+0.50/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN, BENT, DEFORMED, AND STRETCHED OUT. A BROKEN LENS HAPTIC PIECE WAS FOUND ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122738&pc=QCB | 2023826-2021-05007 | 2021-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6-INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13144019&pc=QCB | 2023826-2022-00016 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. LENS REMAINS IMPLANTED. PATIENT IS HAPPY. CAUSE OF THE EVENT WAS REPORTED AS UKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER CQ RECEIVED ON 02-FEB-2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/6.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. LENS REMAINS IMPLANTED. PATIENT IS HAPPY. CAUSE OF THE EVENT WAS REPORTED AS UKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182567&pc=QCB | 2023826-2022-00026 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.5/+3.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. BLURRED VISION IN THE IMMEDIATE POST-OPERATIVE CONTROL IS REPORTED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- ADDITIONAL INFORMATION: PER THE REPORTER BLURRED VISION IMMEDIATELY POST-OP IS AN EXPECTED OCCURRENCE. THE PATIENT IS "HAPPY AND ADAPTED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233320&pc=QCB | 2023826-2022-00154 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.5/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233423&pc=QCB | 2023826-2022-00153 | 2021-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.00/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256317&pc=QCB | 2023826-2022-00140 | 2021-11-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256345&pc=QCB | 2023826-2022-00139 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS AXIAL MISALIGNMENT. IT WAS "CORRECTED" AND THE LENS ROTATED AGAIN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -4.5/+5.0/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AND (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO AXIAL MISALIGNMENT AND ROTATION BUT PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH LONGER LENS AND PROBLEM WAS RESOLVED. "LENS SIZE IS NOT A PROBLEM. A LITTLE ASTIGMATISM REMAINS. CONSIDERING WHETHER TO RECORRECT THE AXIS OR GET LASIK." THE CAUSE OF THE EVENT WAS NOT REPORTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256368&pc=QCB | 2023826-2022-00152 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -7.0/+4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A DIFFERENT SPHERE/CYLINDER LENS. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286530&pc=QCB | 2023826-2022-00179 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE (VA WAS 20/400) AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H3) DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. CLAIM # (B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286581&pc=QCB | 2023826-2022-00194 | 2021-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/4.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THE LENS HAD ROTATED 30 DEGREES COUNTER CLOCKWISE AND THE PATIENT COMPLAINED OF BLURRY VISION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS AND DIFFERENT AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13296036&pc=QCB | 2023826-2022-00241 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.00/+1.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED BLURRED VISION AND WAS DISSATISFIED WITH THE QUALITY OF VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH BUT DIFFERENT POWER/AXIS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. PATIENT IS MORE SATISFIED AND EYE IS QUITE. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "IN MY OPINION MAYBE INTRAOCULAR REACTION LED TO THIS BAD QUALITY OF VISION OR LENS DEPOSITS". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: D4: CORRECTED TO VTICM5_13.2. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389436&pc=QCB | 2023826-2022-00356 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389437&pc=QCB | 2023826-2022-00355 | 2020-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE SIZE STATED IN THE CLAIM. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459226&pc=QCB | 2023826-2022-00393 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/+5.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - "THE SURGEON REPOSITIONED THE LENS TWO DAYS LATER." SHOULD BE ADDED TO THE INITIAL MDR. H1- CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: D4- MODEL# CORRECTED TO VTICMO12.6 IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488089&pc=QCB | 2023826-2022-00426 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -4.00/-3.00/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS REPORTED A LENS EXCHANGED IS PLANNED TO RESOLVE THE ISSUE. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13542049&pc=QCB | 2023826-2022-00439 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO 12.1, -17.50/1.0/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SAME SIZE SPHERICAL LENS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588892&pc=QCB | 2023826-2022-00599 | 2017-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND NOTED "OVER CORRECTION +1.75/0.25/45-0.7. IN A SEPARATE SURGERY THAT SAME DAY, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. POSTOP STATUS INDICATED "ECD 2283/MM2". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT THE LENS WAS NOT EXCHANGED ON THE SAME DAY OF (B)(6) 2017, BUT EXPLANTED ON 05-JUL-2017 AND LATER ON THIS SAME DATE REPLACED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13608356&pc=QCB | 2023826-2022-00618 | 2021-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAS REFRACTIVE SURPRISE AND THERE ARE "NO ISSUES. VAULT WAS NOT RECORDED ON MOST RECENT VISION; BUT AT A PREVIOUS VISIT IT WAS 3+ ICL THICKNESS; VECTOR ANALYSIS WAS DONE BUT NO DECISION ON FURTHER TREATMENT". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS INDICATED THAT THE PATIENT HAD LENS ROTATION. THIS SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13608364&pc=QCB | 2023826-2022-00619 | 2021-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAS REFRACTIVE SURPRISE AND THERE ARE "NO ISSUES. VAULT WAS NOT RECORDED ON MOST RECENT VISION; BUT AT A PREVIOUS VISIT IT WAS 3+ ICL THICKNESS; VECTOR ANALYSIS WAS DONE BUT NO DECISION ON FURTHER TREATMENT". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS INDICATED THAT THE PATIENT HAD LENS ROTATION. THIS SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13712176&pc=QCB | 2023826-2022-00760 | 2021-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/+4.0/146 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725581&pc=QCB | 2023826-2022-00772 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2; 5.50/3.0/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION AND LOSS OF VISUAL ACUITY WHICH THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2022; BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. IT WAS ALSO INDICATED THAT THE RESIDUAL CYLINDER WILL BE TREATED BY LASER TOUCH UP AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831254&pc=QCB | 2023826-2022-00882 | NA | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.00/1.5/92 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO CORRECT REFRACTIVE SURPRISE BUT IT DID NOT RESOLVE THE PROBLEM, UNAIDED VA RE 6/9+1 WAS REPORTED. LENS ROTATION WAS OBSERVED AT 20 DAYS POST-OP. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED, UCVA 20/20-2 WAS REPORTED AT PATIENTS LAS VISIT ON (B)(6) 2022. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-00435 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13872480&pc=QCB | 2023826-2022-00903 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE BUT DIFFERENT POWER -10.5/2.0/90 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND HAPATIC BROKEN. UNABLE TO IDENTIFY LENS MODEL OR PROCEED WITH DIMENSIONAL AND REFRACTIVE MEASUREMENTS DUE TO LENS RETURNED SPLIT IN TWO. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878813&pc=QCB | 2023826-2022-00943 | 2022-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+3.50/77 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2022. THE PATIENT EXPERIENCED BLURRED VISION. THEN ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS NOTED "THERE IS NO PLAN TO IMPLANT OTHER LENS". THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911552&pc=QCB | 2023826-2022-00964 | 2022-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 -7.0/+0.5/051 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911603&pc=QCB | 2023826-2022-00963 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7, -6.0/+1.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SUPRISE IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS REPLACED WITH A DIFFERENT MODEL/SAME LENGTH LENS. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS NO LONER REMAINS IMPLANTED. PATIENT IS HAPPY. THE PROBLEM IS RESOVED." CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911619&pc=QCB | 2023826-2022-00965 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -15.5/+2.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13932582&pc=QCB | 2023826-2022-00948 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.0/5.0/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. REPORTER STATES " IN THE FIRST POSTOPERATIVE DAY FOLLOWING PHAKIC LENS EXPLANTATION CORNEA AND CRYSTALLINE WERE CLEAR. IT IS EXPECTED THAT THE CORNEA DEFLATES AFTER THE EXTRACTION OF THE LENS TO TAKE MEASUREMENTS AGAIN AND REQUEST THE NEW REPLACEMENT LENS." CAUSE WAS REPORTED AS THE PATIENT RELATED FACTOR AND USER ERROR. "DUE TO PREVIOUS KERATOCONUS IT WAS POSSIBLE THAT THE PRE-OPERATIVE MANIFEST REFRACTION WAS NOT ACCURATE." Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13997876&pc=QCB | 2023826-2022-00836 | 2021-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME SIZE DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. SEE MFR. REP#2023826-2022-00838 FOR EXCHANGE LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998048&pc=QCB | 2023826-2022-00838 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS. THIS DID NOT RESOLVE REFRACTIVE SURPRISE AND IS UNDER FURTHER OBSERVATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14006177&pc=QCB | 2023826-2022-01047 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/2.5/060 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO REFRACTIVE SURPRISE AND ROTATION. IT WAS REPORTED THAT AT 1-DAY POST-REPOSITIONING THE PATIENT WAS FINE AND FELT BETTER. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061348&pc=QCB | 2023826-2022-01129 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL ANF FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061353&pc=QCB | 2023826-2022-01135 | 2022-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -14.5/+2.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SUPRISE, BLURRED VISION, AND LENS ROTATION ARE REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED. THE "PATIENT'S UCVA IS 20/20. OS MR IS PLANO/-0.75X180 FOR THE FIRST REVIEW." CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, "PATIENT CAME BACK ON (B)(6) 2022 AND REPORTED BLURRED VISION IN THE LEFT EYE. AFTER DILATION, DR LIM HAS NOTICED THAT THE LENS HAS YET AGAIN ROTATED OUT OF POSITION." CLAIM# (B)(4). Manufacturer Narrative: B5 - ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED LATER ON THE SAME DAY WITH A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. "ICL ON AXIS AT 170". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138859&pc=QCB | 2023826-2022-01232 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -12.0/6.0/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145816&pc=QCB | 2023826-2022-01204 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -12.5/+4.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING LENS ROTATION, REFRACTIVE SURPRISE, AND DROP OF UCVA. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "1ST ROTATION OF THE IOL (20DEGREES) AND RE-ROTATION AFTER REPOSITIONING, NO CORRECTION OF THE ASTIGMATISM; POSTOP REFRACTION=-4X40DEGREES." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14164530&pc=QCB | 2023826-2022-01271 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2022 LASIK ENHANCEMENT WAS PERFORMED. ON (B)(6) 2022 THE LENS ROTATED AGAIN, IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT MODEL AND POWER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION PROBLEM HAS BEEN RESOLVED . CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190875&pc=QCB | 2023826-2022-01295 | 2021-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B3- DATE OF EVENT- (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3-DATE OF EVENT: 03-DEC-2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICTY:UNKNOWN, RACE:UNKNOWN. DATE OF EVENT:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190880&pc=QCB | 2023826-2022-01291 | 2021-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) IN (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION AND LENS ROTATION. REPORTEDLY THE LENS WAS REPOSITIONED AND IT ROTATED AGAIN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER: UNKNOWN, WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED THAT A 13.7MM, VTICMO13.7, -14.5/+3.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION AND INDUCED ASTIGMATISM. THE REPORTER STATED THAT THE PROBLEM RESOLVED BUT AFTER THREE MONTHS THE LENS ROTATED AGAIN. ON (B)(6) 2022 LASIK SURGERY WAS PERFORMED TO CORRECT THE ASTIGMATISM. IT WAS REPORTED THAT THE PROBLEM RESOLVED, EYE IS STABLE POST LASIK ENHANCEMENT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14197862&pc=QCB | 2023826-2022-01309 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -12.0/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - UPDATED DEVICE EVALUATION: LENS WAS RETURNED IN AN EYE DROP CONTRAINER IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A EYE DROP CONTAINER IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274820&pc=QCB | 2023826-2022-01393 | 2022-03-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/+1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE AND PATIENT-RELATED FACTOR WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14274990&pc=QCB | 2023826-2022-01392 | 2022-03-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/+1.5/157 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE AND PATIENT-RELATED FACTOR WAS REPORTED AS CAUSE OF EVENT. THE REPORTER STATED "REGARDING THE REFRACTIVE ERROR, HER RIGHT EYE REFRACTIONS WERE A BIT UNRELIABLE PRE-OP AS MY OPTOMETRIST NOTED HER RIGHT EYE WAS ACCOMMODATING VERY STRONGLY." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14288001&pc=QCB | 2023826-2022-01302 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE LENS ROTATED. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED ON (B)(6) 2022 WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311898&pc=QCB | 2023826-2022-01242 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -17.5/2.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISABLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342556&pc=QCB | 2023826-2022-01475 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -3.5/+1.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT NOW HAS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS DID NOT RESOLVE THE ISSUE AS THE SURGEON NOTED "UCVA IS VERY GOOD; VAULTING STILL VERY LOW. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342603&pc=QCB | 2023826-2022-01474 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -3.50/+1.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT NOW HAS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS DID NOT RESOLVE THE ISSUE AS THE SURGEON NOTED "UCVA IS VERY GOOD; VAULTING STILL VERY LOW. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439265&pc=QCB | 2023826-2022-01369 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -6.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440972&pc=QCB | 2023826-2022-01603 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: BS THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 31-DEC-2021. ON 26-MAR-2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466440&pc=QCB | 2023826-2022-01379 | 2022-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -8.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VIISBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493479&pc=QCB | 2023826-2022-01685 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -2.00/4.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493491&pc=QCB | 2023826-2022-01686 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1.50/4.00/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14523289&pc=QCB | 2023826-2022-01699 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -15.0/2.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND BLURRED VISION. IT WAS REPORTED THAT BLURRY VISION WAS STILL PRESENT AFTER IMPLANTATION OF THIS REPLACEMENT LENS. THE REPORTER STATED, "WILL CONTINUE TO MONITOR". SEE MFR # 2023826-2022-01316 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14570417&pc=QCB | 2023826-2022-01761 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ON (B)(6) 2022; THE SURGEON EXCHANGED THE LENS WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. IT WAS NOTED "DUE TO UNDERCORRECTED 100 DEGREE MYOPIA; HIS VISION HAS BEEN POOR AFTER THE OPERATION; AND HE HAS NOT IMPROVED AFTER VISUAL FUNCTION TRAINING. HE REPLACED THE LENS ON (B)(6) 2022 AND HIS VISION REACHED 1.0". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 -DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH"FIBER" ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14605631&pc=QCB | 2023826-2022-01800 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1.00/+2.50/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "APPEARANCE IS THAT THE IMPLANTED ICL DID NOT CORRESPOND TO THE PATIENT'S RX". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC BROKEN,BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATION. WIDTH AND FUNCTIONAL INSPECTION COULD NOT BE PERFORMED DUE TO HAPTIC/OPTIC DAMAGE. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650054&pc=QCB | 2023826-2022-01838 | 2021-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.0/1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650083&pc=QCB | 2023826-2022-01837 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.50/+5.0/127 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706977&pc=QCB | 2023826-2022-01750 | 2004-03-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.5MM TICM12.5 -13.0/1.2/064 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. REFRACTIVE CHANGE OVERTIME WAS REPORTED. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH, WEIGHT-RACE:UNK. DATE OF EVENT: UNK CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707086&pc=QCB | 2023826-2022-01888 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -5.50/+5.50/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022, THE SURGEON REPOSITIONED THE LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: . CLAIM#: (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707407&pc=QCB | 2023826-2022-01863 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS THE DEVICE. FOR CAUSE DETAILS THE REPORTER STATED " THE TORIC LENS IS CONTINUOUSLY ROTATING WITHIN 5 MONTH FOLLOW-UP INTERVAL, OVER 330 DEGREES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862313&pc=QCB | 2023826-2022-02092 | 2022-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2.5/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862676&pc=QCB | 2023826-2022-01908 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 -5.5/1.0/066 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A BACKUP LENS OF DIFFERENT AXIS WAS IMPLANTED. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: PER ADDITIONAL INFORMATION, "UNDESIRABLE REFRACTIVE OUTCOME" AFTER THE LENS WAS EXCHANGED. REFRACTIVE SURPRISE WAS ALSO REPORTED. PROBLEM WAS NOT RESOLVED. CLAIM #(B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14885725&pc=QCB | 2023826-2022-02095 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/54 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1: SHOULD BE CORRECTED TO: ADVERSE EVENT. B2: SHOULD BE CORRECTED TO: REQUIRED INTERVENTION PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SHOULD BE CORRECTED TO: SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS REMOVED AND INTRAOPERATIVELY REPLACED ON (B)(6) 2022. D6B: ON (B)(6) 2022. CORRECTED DATA: H6: 4627 SHOULD BE OMITTED AND REPLACED WITH 4629 FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE:23-MAR-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN DEVICE SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010429&pc=QCB | 2023826-2022-02233 | 2020-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/0.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED, HAPTIC BENT AND STRETCHED OUT TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15017540&pc=QCB | 2023826-2022-02026 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM TICM12.0 -15.0/2.0/122 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. REFRACTIVE CHANGE OVERTIME AND PROGRESSION OF MYOPIA WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15022376&pc=QCB | 2023826-2022-02270 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7 VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER. THE EXCHANGE LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA; H10 IN PREVIOUS MDR CLAIM NUMBER (B)(4), SHOULD BE CORRECTED TO CLAIM NUMBER (B)(4). DEVICE EVALUATION: H6-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISYURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15097257&pc=QCB | 2023826-2022-02279 | NA | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE INITIAL LENS WAS EXPLANTED AND A REPLACEMENT LENS OF SAME MODEL/SIZE AND DIFFERENT POWER WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE CAUSE IS REPORTED AS DEVICE "MISMARKED LENS RESULTED IN REFRACTIVE SURPRISE." Manufacturer Narrative: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15102877&pc=QCB | 2023826-2022-02436 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -17.0/1.5/119 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15103623&pc=QCB | 2023826-2022-02448 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -10.5/1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT POWER LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATIION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIMN# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118601&pc=QCB | 2023826-2022-02443 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2M VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.0/+1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CASUSE OF THE EVENT WAS USER ERROR. FOR CAUSE DETAILS THE REPORTER STATED " SUMM UP OF MANY FACTORES: ERROR IN REFRACTION (CYCLOPLEGIC REFRACTION WAS NOT NOTICED), INTRAOPERATIVELY INDUCED CYLINDER AND CYLINDRIC LENSPOWER GAIND A RESIDUAL REFRACTIVE ERROR." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2022. THE LENS WAS REPLACED LATER IN ANOTHER SURGERY ON (B)(6) 2022 WITH A SAME LENGTH DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY BUT DISTANT VISION HAS YET TO STABILIZE." CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118602&pc=QCB | 2023826-2022-02445 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6M VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT CYLINDER AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN (FIBERS)ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118603&pc=QCB | 2023826-2022-02444 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6M VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT CYLINDER AND AXIS LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND THE HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H3: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153360&pc=QCB | 2023826-2022-02581 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS OF LONGER LENGTH DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. BLOOD STAINS WERE OBSERVED ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180105&pc=QCB | 2023826-2022-02630 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/98 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE LENS BUT DIFFERENT POWER LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180107&pc=QCB | 2023826-2022-02629 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/88 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY A DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE LENS BUT DIFFERENT POWER LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 -DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204034&pc=QCB | 2023826-2022-02701 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/4.00/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND UNDESIRABLE REFRACTIVE OUTCOME WAS REPORTER, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5: LENS WAS REPOSITIONED ON (B)(6) 2022 AND PROBLEM WAS RESOLVED. STATUS OF THE EYE IS "QUIET". DEVICE DID NOT FAIL TO PERFORM. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204622&pc=QCB | 2023826-2022-02627 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.00/2.50/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022; AS A REPLACEMENT LENS TO A PREVIOUS LOW VAULT WITH ROTATION ISSUE. THE PATIENT NOW HAS REFRACTIVE SURPRISE; AND REFRACTIVE CHANGE OVER TIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: "THE LENS WAS EXPLANTED ON AN UNKNOWN DATE" SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 12-APR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15217420&pc=QCB | 2023826-2022-02695 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/+2.0/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR DIFFERENT POWER LENS DUE TO THE DOCTOR SELECTING THE WRONG LENS POWER. THE EXCHANGE RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS USER ERROR AND THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED . Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15227610&pc=QCB | 2023826-2022-02783 | 2022-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/3.86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND OPTIC TORN ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.50/3.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022." CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15276388&pc=QCB | 2023826-2022-02826 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4. MODEL # VTICMO12.6, EXPIRATION DATE:(B)(6) 2023. Manufacturer Narrative: CORRECTED DATA: H4: DEVICE MANUFACTURE DATE: 02-SEP-2020. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299051&pc=QCB | 2023826-2022-02919 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -7.5/1.5/168 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 21-OCT-2022 . H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN DEVICE SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: EXPIRATION DATE: 12-31-2024. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299255&pc=QCB | 2023826-2022-02530 | 2022-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH LENS ROTATION AND BLURRED VISION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299617&pc=QCB | 2023826-2022-02925 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/161 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT . A REPLACEMENT LENS OF SAME LENGTH, DIFFERENT MODEL AND POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. THE REPORTER STATED " TOO MUCH VAULT AND ASTIGMATISM CORRECTION IN THE LENS. WITH THE SECOND LENS IMPLANTED WITHOUT ASTIGMATISM CORRECTION THE FINAL REFRACTION WAS CLOSE TO EMETROPY. VAULT ALSO WAS BETTER WITH THE SECOND IMPLANT." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 02/28/2025. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299656&pc=QCB | 2023826-2022-02927 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 12/31/2024. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325099&pc=QCB | 2023826-2022-02977 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/1.5/065 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 03-31-2025. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325103&pc=QCB | 2023826-2022-02976 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 01-31-2025. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342016&pc=QCB | 2023826-2022-02969 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -10.50/0.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WITH LOW VAULT WAS RECORDED. A LENS EXHANGE IS PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM (E.G. DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B1 CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE CHANGE OVERTIME AND LOW VAULT. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342195&pc=QCB | 2023826-2022-02970 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WITH LOW VAULT WAS RECORDED. A LENS EXHANGE IS PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM (E.G. DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR . B5 "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ LONGER LENS DUE TO REFRACTIVE CHANGE OVERTIME AND LOW VAULT. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 HEALTH EFFECT IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15385938&pc=QCB | 2023826-2022-03048 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/5.0/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 02-28-2025. (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5: "ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE." SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE: ON (B)(6) 2022, SHOULD BE ADDED TO INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15407358&pc=QCB | 2023826-2022-03069 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -1.5/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE ON THE LENS SPECIFICALLY FIBRE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423219&pc=QCB | 2023826-2022-03072 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION.. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF 11.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION.. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D4: CORRECTED TO MODEL#: VTICM5 12.6 IN THE INITIAL MDR. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423220&pc=QCB | 2023826-2022-03071 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/4.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND BLURRED VISION. THE VAULT WAS GOOD BUT THE LENS ROTATED OFF AXIS AND AFFECTED THE REFRACTION. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER- LENS ROTATED DUE TO SMALL LENS. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423292&pc=QCB | 2023826-2022-03053 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER THE FIRST OPERATON, THE VAULT OF THE LENS WAS LOW , AND THE SIZE OF THE REPLACEMENT LENS WAS CHANGED FROM SMALL (12.1MM) TO LARGE (12.6MM). CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423498&pc=QCB | 2023826-2022-03054 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER THE FIRST OPERATON, THE VAULT OF THE LENS WAS LOW , AND THE SIZE OF THE REPLACEMENT LENS WAS CHANGED FROM SMALL (12.1MM) TO LARGE (12.6MM). CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBER ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15441785&pc=QCB | 2023826-2022-03151 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/2.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A REPLACEMENT LENS OF DIFFERENT POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE (ACD<3.0MM) AND (AGE<21). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15459789&pc=QCB | 2023826-2022-03239 | 2022-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/+3.0/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS DUE TO REFRACTIVE SURPRISE." SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE: ON (B)(6) 2022, SHOULD BE ADDED TO INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO INITIAL MDR. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498967&pc=QCB | 2023826-2022-03371 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.0/1.5/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498988&pc=QCB | 2023826-2022-03369 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -2.50/+2.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499015&pc=QCB | 2023826-2022-03370 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.00/+4.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15508658&pc=QCB | 2023826-2022-03344 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.5/3.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS HAPTIC BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15524739&pc=QCB | 2023826-2022-03366 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.0/5.0/59 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "THE DOCTOR PROBABLY DID NOT POSITION LENS CORRECTLY." CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON 12 DEC 2022 AND PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM#(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534242&pc=QCB | 2023826-2022-03428 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED, BUT WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "TWO WEEKS AFTER THE OPERATION, THE PATIENT FOUND THAT THE LENS OF HIS RIGHT EYE WAS ROTATED, AND THE VISUAL ACUITY DECREASED TO 0.6. AFTER CONSULTATION BY THE OPERATING DOCTOR, HE DECIDED TO PERFORM THE ADJUSTMENT OPERATION. HOWEVER, IT WAS FOUND THAT THE LENS ROTATED GAIN 3 MONTHS AFTER THE ALIGNMENT OPERATION, AND THE VISUAL ACUITY DECREASED TO 0.6.0. AFTER CONSULTATION AND COMMUNICATION WITH THE PATIENT, THE DOCTOR DECIDED TO CUSTOMIZE THE LENS WITH SPECIAL AXIAL POSITION. HOWEVER, DUE TO THE LONG TIME OF CRYSTAL CUSTOMIZATION AND THE EPIDEMIC SITUATION, THE PATIENT UNDERWENT THE RIGHT EYE TICL LENS REPLACMENT ON AUGEST 19." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS, CLAIM# (B)(4) 4B LOW VAULT, CLAIM#(B)(4)4A EXCESSIVE VAULT. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED, BUT WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "TWO WEEKS AFTER THE OPERATION, THE PATIENT FOUND THAT THE LENS OF HIS RIGHT EYE WAS ROTATED, AND THE VISUAL ACUITY DECREASED TO 0.6. AFTER CONSULTATION BY THE OPERATING DOCTOR, HE DECIDED TO PERFORM THE ADJUSTMENT OPERATION. HOWEVER, IT WAS FOUND THAT THE LENS ROTATED GAIN 3 MONTHS AFTER THE ALIGNMENT OPERATION, AND THE VISUAL ACUITY DECREASED TO 0.6.0. AFTER CONSULTATION AND COMMUNICATION WITH THE PATIENT, THE DOCTOR DECIDED TO CUSTOMIZE THE LENS WITH SPECIAL AXIAL POSITION. HOWEVER, DUE TO THE LONG TIME OF CRYSTAL CUSTOMIZATION AND THE EPIDEMIC SITUATION, THE PATIENT UNDERWENT THE RIGHT EYE TICL LENS REPLACEMENT ON (B)(6)." CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534408&pc=QCB | 2023826-2022-03427 | 2022-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITION AND WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL, SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ON (B)(6), THE PATIENT IMPLANTED TICL CRYSTAL IN HIS LEFT EYE. AFTER THE OPERATION, IT WAS FOUND THAT THE CRYSTAL WAS ROTATED 36 DEGREES ANTICLOCKWISE. AFTER THE FACE-TO-FACE CONSULTATION OF THE OPERATIVE DOCTOR, THE PATIENT DECIDED TO PERFORM TICL CRYSTAL ADJUSTMENT. AFTER FACE-TO-FACE REVIEW, IT WAS FOUND THAT THE CRYSTAL WAS ROTATED AGAIN. AFTER THE FACE-TO-FACE CONSULTATION, THE OPERATING DOCTOR COMMUNICATED WITH THE PATIENT AND FINALLY DECIDED TO REPLACE THE ICL CRYSTAL. ASTIGMATISM WAS CORRECTED BY LASER SURGERY." Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15557486&pc=QCB | 2023826-2022-03452 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OBSERVED REFRACTIVE CHANGE OVER TIME IN THE PATIENT'S LEFT EYE (OS). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK CLAIM# (B)(4). Manufacturer Narrative: B1: PRODUCT PROBLEM. B2: ENTRY SHOULD BE LEFT BLANK . CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A -3.0/05/082 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT ON (B)(6) 2022. ON (B)(6) 2022, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "MONOVISION ON THE LEFT EYE WAS PLANNED. THE RESULT IS NOT ACHIEVED WITH THE VALUE -9.25-9.50/88. WE ASK FOR EXCHANGE. TARGET 1.25." H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND WITH ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH A -3.0/0.5/082 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. RETURN DATE- 12-APR-2023 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL/RESIDUE/DEBRIS ON LENS SURFACE, SPECIFICALLY FIBERS. NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15609430&pc=QCB | 2023826-2022-03642 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -18.0/6.0/80 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE FIRST TIME, AN INCORRECT LENS WAS IMPLANTED, THE REASON WAS THAT THE CALCULATION WERE NOT DONE CORRECTLY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629399&pc=QCB | 2023826-2022-03705 | 2022-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/+3.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H1: CORRECTED TO - MALFUNCTION IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15650328&pc=QCB | 2023826-2022-03704 | 2022-09-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.00/+2.5/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D4 - CORRECTED TO : MODEL # VTICM5 13.2 IN THE INTIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651860&pc=QCB | 2023826-2022-03636 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H2: DEVICE EVALUATION SHOULD HAVE BEEN SELECTED IN PREVIOUS SUPPLEMENTAL. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND RESIDUE/DEBRIS ON THE LENS - SHOULD HAVE ADDITIONALLY BEEN INCLUDED IN PREVIOUSLY SUPPLEMENTAL. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697248&pc=QCB | 2023826-2022-03860 | 2022-02-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 2.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. A MEDICAL CONSULT IS BEING PROVIDED TO GIVE GUIDANCE ON HOW TO CORRECT THE PATIENT'S VISION. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15727569&pc=QCB | 2023826-2022-03976 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -6.50/2.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. . Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15728310&pc=QCB | 2023826-2022-03975 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/2.0/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15782580&pc=QCB | 2023826-2022-03959 | 2022-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -1.00/4.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE IS REPORTED AS "USER ERROR." REPORTEDLY, "AWAITING ADVICE FROM MEDICAL AFFERS". Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. PATIENT RELATED FACTOR (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" SHOULD BE CORRECTED TO "PRODUCT PROBLEM". H6- MEDICAL DEVICE PROBLEM CODE- REMOVE CODE: 2682. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.00/+4.00/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "AWAITING ADVICE FROM MEDICAL AFFAIRS" AND "AN 10º ERROR WAS MADE WHEN PUTTING THE SUBJECTIVE REFRACTION IN OCOS." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15842667&pc=QCB | 2023826-2022-04287 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.00/1.50/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 13-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15843892&pc=QCB | 2023826-2022-04289 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850084&pc=QCB | 2023826-2022-04265 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D4- UNIQUE IDENTIFIER (UDI) #: (B)(4). ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850228&pc=QCB | 2023826-2022-04267 | 2022-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/0.50/80 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY PROBLEM RESOLVED. THE PATIENT DOES NOT WANT A TREATMENT AND THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854790&pc=QCB | 2023826-2022-04293 | NA | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/+3.00/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/3.00/003 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE REPLACEMENT LENS HAD "LOSS OF STERILIZATION", AND WAS NOT IMPLANTED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. D6B. "12/13/2022" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15874950&pc=QCB | 2023826-2022-04304 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/1.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5 - "ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/ MODEL LENS DUE TO REFRACTIVE CHANGE OVERTIME. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2023 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15875543&pc=QCB | 2023826-2022-04305 | 2021-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.5/1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021 . THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - CORRECTED TO: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15891930&pc=QCB | 2023826-2022-04367 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.50/0.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF 4.50/0.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL SAME LENGTH LENS. THE PROBLEM RESOLVED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15912687&pc=QCB | 2023826-2022-04282 | 2022-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/069 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975198&pc=QCB | 2023826-2022-04416 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975200&pc=QCB | 2023826-2022-04414 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.00/4.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5: CORRECTED TO : "THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.00/4.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022. THE LENS REMAINS IMPLANTED." SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990827&pc=QCB | 2023826-2022-04591 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM SO THE LENS WAS EXPLANTED. A SPHERICAL LENS OF LONGER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT LASIK SURGERY WAS NECESSARY DUE TO EXISTING RESIDUAL ASTIGMATISM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15991920&pc=QCB | 2023826-2022-04417 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/4.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000596&pc=QCB | 2023826-2022-04599 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -10.5/3.0/038 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2022-02093 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019046&pc=QCB | 2023826-2022-04708 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON(B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "USER ERROR" AND NOTED "THERE IS A SUBJECTIVE OVERCORRECTION IN LENS POWER". Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SIMILAR LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16024296&pc=QCB | 2023826-2022-04693 | 2020-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH THE REPLACEMENT LENS DUE TO A LENS TEAR/BREAK DURING INJECTION DELIVERY INTO THE EYE AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT/ ETHINICITY/ RACE:UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: BASED ON CORRECTED ADDITIONAL INFORMATION FROM THE REPORTER, REFRACTIVE SURPRISE WAS NOT OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: 26-NOV-2022 SHOULD BE CORRECTED TO 26-NOV-2020. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030437&pc=QCB | 2023826-2022-04435 | 2020-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -7.5/1.5/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2020 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF SHORTER LENGTH AND DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030860&pc=QCB | 2023826-2022-04694 | 2021-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTER AS, OTHER. Manufacturer Narrative: B5: THE CAUSE OF THE EVENT IS UNCERTAIN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030861&pc=QCB | 2023826-2022-04695 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: B5: THE CAUSE OF THE EVENT IS UNCERTAIN. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16034661&pc=QCB | 2023826-2022-04644 | 2021-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/72 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 13-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057495&pc=QCB | 2023826-2022-04733 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064724&pc=QCB | 2023826-2022-04691 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075117&pc=QCB | 2023826-2022-04773 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/1.0/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075678&pc=QCB | 2023826-2022-04747 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS WITH A -6.0/1.0/089 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED AS A REPLACEMENT INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: B3 - DATE OF EVENT: 20-MAY-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO "SULCUS TOO SMALL FOR 12.1." AND BLURRY VISION. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075681&pc=QCB | 2023826-2022-04746 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A -6.0/1.0/107 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: DEVICE EVALUATED BY MFR: OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16077729&pc=QCB | 2023826-2022-04792 | 2022-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -6.5/1.5/11 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A LENS OF DIFFERENT POWER WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-FEB-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099355&pc=QCB | 2023826-2023-00083 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13/1.5/173 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "ERROR IN ASTIGMATISM DATA". Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16112609&pc=QCB | 2023826-2023-00079 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -0.50/6.00/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE EYE IS QUIET. IT WAS REPORTED THAT VISION IS 20/60 UNDILATED WITH A VAULT OF 660 AND 20/30 DILATED WITH A VAULT OF 442. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16113739&pc=QCB | 2023826-2023-00081 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16117774&pc=QCB | 2023826-2023-00054 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATAS FOR CALCULATION IN OCOS, DATA FOR OD USED FOR OS. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD HAVE BEEN PRODUCT PROBLEM IN THE ORIGINAL MDR SUBMISSION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN OMITTED FROM THE ORIGINAL MDR SUBMISSION. B5: REFRACTIVE CHANGE OVER TIME SHOULD BE CORRECTED TO STATE REFRACTIVE SURPRISE. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS. REPORTEDLY, "SURGEON USED WRONG DATA FOR CALCULATION IN OCOS, DATA FOR OD USED FOR OS. D6B: EXPLANT DATE: ON (B)(6) 2023. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16117832&pc=QCB | 2023826-2023-00053 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATAS FOR CALCULATION IN OCOS, DATA FOR OS USED FOR OD. Manufacturer Narrative: B1: ADVERSE EVENT SHOULD HAVE BEEN PRODUCT PROBLEM IN THE ORIGINAL MDR SUBMISSION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN OMITTED FROM THE ORIGINAL MDR SUBMISSION. B5: REFRACTIVE CHANGE OVER TIME SHOULD BE CORRECTED TO STATE REFRACTIVE SURPRISE. H6: OMIT 1494 FOR CORRECTION. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SAME MODEL AND LENGTH BUT DIFFERENT DIOPTER LENS. ON (B)(6) 2022, THE LENS WAS REPLACED WITH A SAME MODEL AND LENGTH BUT DIFFERENT DIOPTER LENS. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "SURGEON USED WRONG DATA FOR CALCULATION IN OCOS, DATA FOR OS USED FOR OD. D6B: EXPLANT DATE: ON (B)(6) 2023. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16141698&pc=QCB | 2023826-2023-00120 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM513.2, -7.50/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITONAL INFORMATION: B5- ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142821&pc=QCB | 2023826-2023-00123 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -3.50/1.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY SIZE BUT A DIFFERENT MODEL AND POWER LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: D4 SERIAL # - CORRECTED TO: (B)(6) IN THE INITIAL MDR. CLAIM (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16142822&pc=QCB | 2023826-2023-00124 | 2022-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.50/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED. ON A LATER DATE A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9:29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153694&pc=QCB | 2023826-2023-00107 | NA | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS USER UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154820&pc=QCB | 2023826-2023-00130 | 2022-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.00/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9:29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189407&pc=QCB | 2023826-2023-00173 | 2022-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/+3.50/97 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT CALCULATION IS BEING DONE IN PLAN TO REPOSITION THE LENS. LENS REMAINS IMPLANTED. THE PROBLEM IS REPORTED AS NOT RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189410&pc=QCB | 2023826-2023-00199 | 2022-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/3.50/64 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE LEFT RIGHT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT CALCULATION IS BEING DONE IN PLAN TO REPOSITION THE LENS. LENS REMAINS IMPLANTED. THE PROBLEM IS REPORTED AS NOT RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199991&pc=QCB | 2023826-2023-00202 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.50/1.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE 4628. CLAIM (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199992&pc=QCB | 2023826-2023-00203 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS OF -13.50/4.00/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. PER UPDATED INFORMATION THE LENS WAS ROTATED ON (B)(6) 2022 AND THE PATIENT NOW HAS 1.2 VA. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER THE LENS ROTATION. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210693&pc=QCB | 2023826-2023-00244 | 2022-11-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -1/+1/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT HAS REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED "UDVA IS NOT SUFFICIENT DUE TO RESIDUAL MYOPIA AND ASTIGMATISM. WE SHOULD REPLACE THE ICL". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16214291&pc=QCB | 2023826-2023-00204 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -1.50/2.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON(B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. D6B. "(B)(6) 2023" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_ 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/2.0/174(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "ALL FINE!" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219689&pc=QCB | 2023826-2023-00251 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED.THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS OTHER. DEVICE DID NOT FAULT O PERFORM AS INTENDED. "THERE WAS A MISUNDERSTANDING ABOUT ALIGNMENT." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248972&pc=QCB | 2023826-2023-00337 | 2022-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE DID NOT RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FIBRE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16249317&pc=QCB | 2023826-2023-00338 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS REFRACTIVE SURPRISE BUT IT DID NOT RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2023-00337 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE; BUT THAT DID NOT RESOLVE THE PROBLEM. THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16297884&pc=QCB | 2023826-2023-00435 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -2.50/5.5/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION AND BAD VISION. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: EXPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16300819&pc=QCB | 2023826-2023-00440 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.00/1.50/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301296&pc=QCB | 2023826-2023-00442 | 2021-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16328105&pc=QCB | 2023826-2023-00380 | 2022-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.0/3.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343807&pc=QCB | 2023826-2023-00390 | 2023-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/4.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343985&pc=QCB | 2023826-2023-00391 | 2023-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.50/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WITH REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICM 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.50/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. LOW VAULT WAS OBSERVED.THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344319&pc=QCB | 2023826-2023-00402 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5/2.0/83 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: UNK. OFF-LABEL - AGE<21. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359096&pc=QCB | 2023826-2023-00487 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/87(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD)ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/87(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD)ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SIMILAR MAKE AND MODEL BUT DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE-12-APR-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376165&pc=QCB | 2023826-2023-00563 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DIFFERENT POWER; THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385392&pc=QCB | 2023826-2023-00561 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADVERSE EVENT PROBLEM - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390126&pc=QCB | 2023826-2023-00591 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN/BROKEN TO THE LENS. REHYDRATION OF A LENS HAS CONFIRMED INADEQUACY OF PERFORMING DIM/REF INSPECTION DUE TO SIGNIFICANT DAMAGE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390268&pc=QCB | 2023826-2023-00592 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: THE DHR REVIEW INDICATES THAT THE PRODUCT HAS BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS NO INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406313&pc=QCB | 2023826-2023-00599 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 AS AND EXCHANGE LENS TO RESOLVE LOW VAULT. TRANSIENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED. "A LASIK TOUCH UP IS PLANNED. SEE MDR # (B)(4) FOR INITIAL LENS. Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS AN EXCHANGE LENS TO RESOLVE LOW VAULT. TRANSIENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED. "A LASIK TOUCH UP IS PLANNED." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFOMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/153 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS. TRANSIENT LOSS OF BCVA WAS OBSERVED. ON (B)(6) 2023 LASIK WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT IS HAPPY". THE REPORTER MARKED THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411082&pc=QCB | 2023826-2023-00569 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. POST-OP THE PATIENTS VA WAS BLURRY AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00568. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411155&pc=QCB | 2023826-2023-00567 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. POST-OP THE PATIENTS VA WAS BLURRY. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00566. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16417982&pc=QCB | 2023826-2023-00608 | 2023-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/064 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT POWER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16425177&pc=QCB | 2023826-2023-00620 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF -3.50/2.5/034 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. VERY HIGH RESIDUAL ASTIGMATISM WAS PRESENT. A BAROMETRY DEVICE WAS USED WHICH SHOWED THAT THERE WAS NO INTERNAL CYLINDER POWER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433883&pc=QCB | 2023826-2023-00693 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50/6.0/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433993&pc=QCB | 2023826-2023-00694 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.50/5.0/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16471700&pc=QCB | 2023826-2023-00724 | 2021-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.0/4.0/58 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH ROTATION AND REFRACTIVE SUPRISE WERE OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PATIENT IS FINE WITH THE RESULT. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16471706&pc=QCB | 2023826-2023-00723 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -5.5/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WERE OBSERVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITIAL MDR. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.5/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021.LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WERE OBSERVED. THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THE PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 11-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: SERIAL NUMBER SHOULD BE CORRECTED TO (B)(6). EXPIRATION DATE: SHOULD BE CORRECTED TO 31-MAR-2022. UDI: SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492175&pc=QCB | 2023826-2023-00779 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -11.5/2.5/170 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023, AND IN A SUBSEQUENT SURGERY LATER THAT DAY, WAS REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "NO PROBLEM" IS STATED FOR STATUS OF THE EYE. Manufacturer Narrative: CORRECTED DATE: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAM#: (B)(4). Manufacturer Narrative: PATIENT INFO:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492446&pc=QCB | 2023826-2023-00780 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.50/2.5/178 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "POSTOPERATIVE DAY 1 IOP IS HIGH AND I/A IS PERFORMED. CURRENTLY, DOCTOR IS TAKING A WAIT-AND-SEE APPROACH." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATE: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE- 27-MAR-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492703&pc=QCB | 2023826-2023-00781 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -16.5/2.0/179 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY STATUS OF THE EYE IS, "STABLE." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/ DEFORMED WITH FIBRE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507974&pc=QCB | 2023826-2023-00835 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.50/29 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS A REFRACTIVE ERROR DUE TO WRONG LENS CALCULATION; AXIS 117 INSTEAD OF 177 WAS USED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16508352&pc=QCB | 2023826-2023-00788 | 2022-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5/3/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT DISSATISFACTION/STORE AND BLURRY SYMPTOM. PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519835&pc=QCB | 2023826-2023-00813 | 2023-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -8.50/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT MODEL AND POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION-SHOULD BE CORRECTED TO STATE "FOUND THE OPTIC AND THE HAPTIC TORN WITH RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS". CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16521994&pc=QCB | 2023826-2023-00789 | 2022-09-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITHOUT ROTATION AND ON (B)(6) 2022, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. THE LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: PATIENT INFORMATION: UNK. EVENT PROBLEM AND EVALUATION CODES: OFF-LABEL AGE<21 CLAIM #: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16538767&pc=QCB | 2023826-2023-00792 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: OFF-LABEL - ACD < 3.0 AGE > 45. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 03/23/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: 728275 2023 00792 SUP2. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16549892&pc=QCB | 2023826-2023-00905 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/4.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "GOOD HOWEVER ICL IN @ 158 POST OPERATIVELY AND WAS SUPPOSED TO BE AT 167. I WANT TO MOVE LENS BUT WHICH ONE WOULD BE BEST?" AND "ICL HAS NOT BEEN EXPLANTED. DR (B)(6) FIRST WANTS TO SEE, IF BY ROTATING THE ICL HE CAN IMPROVE THE REFRACTIVE OUTCOME." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: G2: SOUTH AFRICA. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16567677&pc=QCB | 2023826-2023-00940 | 2023-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT COMPLAINED OF BLURRY VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/+3.50/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, BLURRED VISION, AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY " PATIENT HAS A HIGH VAULT OU BUT THERE HAVE BEEN NO ISSUES WITH IOP. HIS PATIENT ALSO HAS RESIDUAL REFRACTIVE ERROR OD GREATER THAN OS. HOWEVER, THE PATIENT WAS AT UPPER EXTREME OF MYOPIA OD VS OS. HE IS NOT COMPLAINING BUT WANTED TO RECEIVE SOME ADVICE ON FOLLOWING AND OPTIONS FOR THE PATIENT." THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: " BLURRY VISION". THE CAUSE OF THE EVENT IS UNKNOWN. CAUSE DETAILS- "HIGH VAULT". CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584143&pc=QCB | 2023826-2023-00987 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/+0.5/48 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 LENS REPLACED ON (B)(6) 2023, WITH A DIFFERENT POWER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584242&pc=QCB | 2023826-2023-00986 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+0.5/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - LENS REPLACED ON (B)(6) 2023 WITH A DIFFERENT POWER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16606958&pc=QCB | 2023826-2023-01003 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5/3.5/016 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO DOUBLE VISION. Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO 07-NOV-2022. D9: RETURN DATE: 15-APR-2023 G2: SHOULD BE OMITTED FOR CORRECTION. ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A ZIPLOCK BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT PROBLEM WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610082&pc=QCB | 2023826-2023-00980 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS OF DIFFERENT POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610212&pc=QCB | 2023826-2023-00981 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED AS CAUSE OF EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS OF DIFFERENT POWER ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633210&pc=QCB | 2023826-2023-01010 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/3.0/061 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "SO FAR PATIENT NOT COME FOR FOLLOW UP AND MORE INVESTIGATION." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633211&pc=QCB | 2023826-2023-01009 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/4.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "SO FAR PATIENT NOT COME FOR FOLLOW UP AND MORE INVESTIGATION." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16640306&pc=QCB | 2023826-2023-01122 | NA | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.0/+3.0/002 (SPHERE,CYLINDER,AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649748&pc=QCB | 2023826-2023-01078 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/013 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16668634&pc=QCB | 2023826-2023-01193 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A TICM12.0; -13.5/+3.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON NOTES THE PATIENT HAS REFRACTIVE CHANGE OVER TIME WHICH WAS ASSESSED ON (B)(6) 2023. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVEN WAS REPORTED AS UNKNOWN. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16670694&pc=QCB | 2023826-2023-01198 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -7.00/1.0/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. A PLANNED LENS REMOVAL IS REPORTED DUE TO POSTOPERATIVE REFRACTION (INTENDED PLANO), LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678777&pc=QCB | 2023826-2023-01181 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE CHANGE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "REPOSITIONING OF THE LENS BY PROF. PARASTA WAS NOT SUCCESSFUL. PATIENT IS WEARING GLASSES AGAIN. COLE MONITORING. IF THE PATIENT WANT A LENS EXCHANGE THE SURGEON WOULD LIKE TO EXPLANT THE CURRENT LENS. WAIT FOR 3 MONTHS AND THEN REMEASURE THE EYE AND RECALCULATE THE LENS FOR EXCHANGE. WE WILL KEEP YOU UPDATED". THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-8.5/+1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678779&pc=QCB | 2023826-2023-01184 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -6.0/1.0/76 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 08-FEB-2023. ON 17-FEB-2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON (B)(4) 2023 DUE TO REFRACTIVE SURPRISE. PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678780&pc=QCB | 2023826-2023-01180 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE CHANGE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "REPOSITIONING OF THE LENS BY PROF. PARASTA WAS NOT SUCCESSFUL. PATIENT IS WEARING GLASSES AGAIN. COLE MONITORING. IF THE PATIENT WANT A LENS EXCHANGE THE SURGEON WOULD LIKE TO EXPLANT THE CURRENT LENS. WAIT FOR 3 MONTHS AND THEN REMEASURE THE EYE AND RECALCULATE THE LENS FOR EXCHANGE. WE WILL KEEP YOU UPDATED". THE CAUSE OF THE EVENT AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678782&pc=QCB | 2023826-2023-01182 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION, WHICH DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REPLACED FOR A SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER IMPLANTATION OF TICL LENS IN THE RIGHT EYE, THE PATIENT'S VISION WAS DECREASED AND THE LENS POSITION WAS CHECKED FOR DEVIATION. AFTER THE SURGEON'S INTERVIEW, HE DECIDED TO PERFORM THE REPOSITIONING OPERATION ON (B)(6), AFTER THE REPOSITIONING OPERATION, THE PATIENT'S EYE CONDITION WAS RECHECKED AGAIN. THE INSPECTION RESULTS SHOWED THAT THE VISION WAS STILL NOT IMPROVED, THE LENS POSITION WAS NOT ROTATED, AND THE ARCH HEIGHT WAS NORMAL. AFTER CONSULTATION WITH THE SURGEON AND EXPERTS, THE CRYSTAL WAS INITIALLY SUSPECTED. AFTER COMMUNICATION WITH THE PATIENT, THE NEW TICL LENS WAS IMPLANTED AGAIN ON SEPTEMBER 30. THE VISION WAS STILL NOT IMPROVED ONE DAY AFTER SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-APR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: PATIENT DEMOGRAPHICS:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678785&pc=QCB | 2023826-2023-01177 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/0.5/140 (SPHERE,CYLINDER,AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERIC MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678786&pc=QCB | 2023826-2023-01183 | 2022-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE EXCHANGE WAS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER IMPLANTATION OF TICL LENS IN THE RIGHT EYE, THE PATIENT'S VISION WAS DECREASED AND THE LENS POSITION WAS CHECKED FOR DEVIATION. AFTER THE SURGEON'S INTERVIEW, HE DECIDED TO PERFORM THE REPOSITIONING OPERATION ON AUGUST 5. AFTER THE REPOSITIONING OPERATION, THE PATIENT'S EYE CONDITION WAS RECHECKED AGAIN. THE INSPECTION RESULTS SHOWED THAT THE VISION WAS STILL NOT IMPROVED, THE LENS POSITION WAS NOT ROTATED, AND THE ARCH HEIGHT WAS NORMAL. AFTER CONSULTATION WITH THE SURGEON AND EXPERTS, THE CRYSTAL WAS INITIALLY SUSPECTED. AFTER COMMUNICATION WITH THE PATIENT, THE NEW TICL LENS WAS IMPLANTED AGAIN ON SEPTEMBER 30. THE VISION WAS STILL NOT IMPROVED ONE DAY AFTER SURGERY." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728332&pc=QCB | 2023826-2023-01362 | 2021-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/2.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE SURGEON REPORTS LENS ROTATION; REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773435&pc=QCB | 2023826-2023-01464 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -10.5/+0.50/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791810&pc=QCB | 2023826-2023-01446 | 2023-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.5/159 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REPORTEDLY, MEDICAL OR SURGICAL INTERVENTION WERE NOT NECESSARY. PER THE COMPLAINT QUESTIONNAIRE THE ISSUE WAS REPORTED AS "RESOLVED". THE STATUS OF THE EYE IS "ALL FINE". NO INJURY REPORTED. ADDITIONAL INFORMATION PROVIDED STATES "THE CLINIC PLANNED AN OUTCOME OF -0.75DPT BUT THE FINAL OUTCOME WAS ABOUT 0.00DPT. IT'S OK FOR THE PATIENT AND CURRENTLY WISH NO FURTHER TREATMENT". CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16791940&pc=QCB | 2023826-2023-01447 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.00/1.0/29 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2023. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. OVERALL LENGTH (DIAMETER) VERIFICATION AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WIDTH VERIFICATION COULD NOT BE PERFORMED DUE TO HAPTIC DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792761&pc=QCB | 2023826-2023-01428 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811280&pc=QCB | 2023826-2023-01517 | 2022-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -2.5/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811756&pc=QCB | 2023826-2023-01558 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.00/+4.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022.THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!" AND "PATIENT IS HAPPY". THE THE CAUSE OF THE EVENT IS UNKNOWN. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16828371&pc=QCB | 2023826-2023-01531 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16839441&pc=QCB | 2023826-2023-01568 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.50/+3.0/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING; BLURRY VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ORIGINAL CALCULATION & RECOMMENDED 12.6 SIZING TOO SMALL; = +1- 90 DEGREES ROTATION CAUSING BLURRY VISION. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT - IMPACT CODE: "4628" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.50/+3.00/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS 'OTHER'. CAUSE DETAILS: "ORIGINAL CALCULATION AND RECOMMENDED (12.6) SIZING TOO SMALL= +/- 90 DEGREES. ROTATION CAUSING BLURRY VISION". DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION WILL NOT BE PERFORMED DUE TO SIGNIFICANT HAPTIC DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16839442&pc=QCB | 2023826-2023-01569 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/+4.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING; BLURRY VISION. IT WAS NOTED THAT AFTER IMPLANT; THE PATIENT WENT TO THE EXAM ROOM TO BE EVALUATED AND THE SURGEON NOTICED THE LEFT ICL HAD ROTATED. THE SURGEON HAD REPOSITIONED THE ICL TO IT'S ORIGINAL POSITION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ORIGINAL CALCULATION 12.6 SIZING TOO SMALL; = +1- 90 DEGREES ROTATION CAUSING BLURRY VISION". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850168&pc=QCB | 2023826-2023-01598 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON THIS SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. IT WAS ADDITIONALLY NOTED "AFTER SURGERY THE PATIENT COMPLAINED OF VISUAL PROBLEMS IN THE OPTICAL ZONE AND STRONGLY REQUESTED TO REMOVE THE LENS. AFTER REMOVAL A TRACE WAS FOUND NEAR THE OPTICAL ZONE MARKER LINE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16864790&pc=QCB | 2023826-2023-01663 | 2023-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -14.00/4.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE AND BLURRY VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16888245&pc=QCB | 2023826-2023-01638 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A -15.0/3.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2023, DUE TO THE INITIAL LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. BLURRED VISON AND MILD CORNEAL EDEMA WAS OBERSRVED. THE STATUS OF THE EYE REPORTED, "MILD CORNEA EDEMA." Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT: SHOULD BE CORRECTED TO 08-04-2023. B5: OD SHOULD BE CORRECTED TO OS. D6B: EXPLANT DATE SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16890526&pc=QCB | 2023826-2023-01746 | 2023-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM013.2, -15.50/1.5/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO CORRECT LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM, " LENS HAS ROTATED/ INDUCED ASTIGMATISM" WAS REPORTED. LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTOR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892246&pc=QCB | 2023826-2023-01747 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894692&pc=QCB | 2023826-2023-01556 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT REFRACTIVE SURPRISE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM12.6 -3.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2023 INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6) 2023 AND THIS RESOLVED THE PROBLEM. STATUS OF THE EYE IS "EXPECTANT". REPORTEDLY, "IN FUTURE VISITS WE WILL SEE THE EVOLUTION." CLAIM# (B)(4). Manufacturer Narrative: H4 - CORRECTED TO: 09-AUG-2022 IN THE INTIAL MDR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT SHOULD BE ADDED TO THE INITIAL MDR. (B)(4) Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894694&pc=QCB | 2023826-2023-01555 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/1.5/159 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - REPORTEDLY "THE PATIENT IS ALREADY OPERATED ON AND SEEING 20/20. IT IS CONFIRMED THAT THE INTIAL LENSES WERE SWAPPED." CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894700&pc=QCB | 2023826-2023-01554 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/1.5/041 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - REPORTEDLY "THE PATIENT IS ALREADY OPERATED ON AND SEEING 20/20. IT IS CONFIRMED THAT THE INTIAL LENSES WERE SWAPPED." CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16897246&pc=QCB | 2023826-2023-01750 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -9.5/1.0/019 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT POWER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908166&pc=QCB | 2023826-2023-01776 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/048 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING AND THE PATIENT EXPERIENCED BLURRED VISION AND A REFRACTIVE SURPRISE. THE LENS VAULT WAS LOW AND SAT TOO HIGH IN THE SULCUS CAUSING A "SMUDGE" IN VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16920640&pc=QCB | 2023826-2023-01802 | 2023-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50/0.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16920684&pc=QCB | 2023826-2023-01801 | 2023-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50/0.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922176&pc=QCB | 2023826-2023-01781 | 2023-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD A 1 + CELL AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01687. Manufacturer Narrative: B5: THE PATIENT WAS SEEN ON (B)(6) 2023 AT POST-OP VISIT. THE PATIENT IS HAPPY WITH IMPROVED VISION, VA UNCORRECTED 20/30-2, BCVA 20/30. VISUAL ACUITY IS AT POTENTIAL DUE TO AMBLYOPIA. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16934647&pc=QCB | 2023826-2023-01876 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+1.0/142 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16962318&pc=QCB | 2023826-2023-01924 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -5.5/1.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991212&pc=QCB | 2023826-2023-01771 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -14.5/2.5/055 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.5/55 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022 .THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16994612&pc=QCB | 2023826-2023-01993 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998681&pc=QCB | 2023826-2023-01998 | NA | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.00/3.0/178 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH,DIFFERENT MODEL AND POWER LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT ASTIGMATISM WAS TREATED WITH RELAXING INISION. PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007679&pc=QCB | 2023826-2023-02045 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.5/1.5/066 (SPHERE/CYLINDER/AXIS), INTO SON'S LEFT EYE (OS) ON (B)(6) 2023 FOR CORRECTION OF -12 DIOPTERS OF MYOPIA AND -2 DIOPTERS OF ASTIGMATISM. THE REPORTER STATES " AFTER THE SECOND POST-OPERATIVE VISIT MY SON COMPLAINS OF BLURRY VISION AND THE SURGEON REALIZES THAT THE 2 DIOPTERS OF ASTIGMATISM HAVE NOT BEEN CORRECTED ". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REFRACTIVE SURPRISE WAS REPORTED. CAUSE OF EVENT IS UNKNOWN WITH LENS ROTATION NOTED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039566&pc=QCB | 2023826-2023-01885 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Optical Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.00/4.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- MEDICAL DEVICE PROBLEM CODE: "3001" SHOULD BE REMOVED AND "1401" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21) . CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039571&pc=QCB | 2023826-2023-01888 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/3.0/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - A6: UNK. B3: UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D6A - UNK. D6B - UNK. H4 - UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.5/3.0/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE (B)(6) 2023. REFRACTION WAS OBSERVED. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039736&pc=QCB | 2023826-2023-02153 | 2023-04-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.00/+1.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17054422&pc=QCB | 2023826-2023-02173 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -10.0/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH, MODEL, AND POWER WAS IMPLANTED BUT THIS DID NOT RESOLVE THE PROBLEM. IN WAS ADDITIONALLY NOTED AFTER THE REPLACEMENT LENS WAS IMPLANTED THE PATIENT HAS CORNEAL EDEMA, AND THE POST OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2023-02174 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4627" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION: "PATIENT IS HAVING RESIDUAL OF +1.75/3@175. CHECKED POST OP ON (B)(6) WITH CORNEAL EDEMA, AND THE POST OP REFRACTION WAS +0.75/-1.75/165. SURGEON ADVISED THE NEXT REVIEW AFTER 6 WEEKS." IT WAS ALSO REPORTED THAT "SUGGEST REPOSITIONING OF THIS SAME LENS TO CORRECT THE RESIDUAL ERROR". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17061058&pc=QCB | 2023826-2023-02039 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -5.0/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SAME LENGTH DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17062290&pc=QCB | 2023826-2023-02038 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066749&pc=QCB | 2023826-2023-02221 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM). (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076192&pc=QCB | 2023826-2023-02226 | 2023-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.00/1.0/070 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS: DEVICE, A PATIENT RELATED FACTOR, AND UNKNOWN. THE DEVICE DID NOT FAIL AS INTENDED. REPORTEDLY, "NO PROBLEM WITH THE PROCEDURE AND TEST DATA, BUT NO IMPROVEMENT IN VISION." THE PROBLEM HAS NOT RESOLVED. STATUS OF THE EYE SAID, "THERE IS NOTHING WRONG WITH THE EYE, BUT THE VISION IS NOT GOOD," AND "DOCTOR SUSPECTS A PARALYSIS OF THE CILIARY MUSCLE OR A PROBLEM WITH THE LENS ITSELF." Manufacturer Narrative: B5 - LENS WAS EXPLANTED. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID WITHIN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS REPLACED WITH A DIFFERENT POWER LENS ON (B)(6) 2023. STATUS OF THE EYE IS "GOOD". CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE DHR REVIEW INDICATES THAT THE PRODUCT HAS NOT BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. ROOT CAUSE WAS NOT DETERMINED. PROBABLE CAUSE IS LIKELY ASSOCIATED TO THE LENS LABEL REWORK PROCESS. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084116&pc=QCB | 2023826-2023-02230 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -.05/6/140 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON 18-MAY-2023, REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT WAS REPORTED AS BOTH THE DEVICE AND A PATIENT RELATED FACTOR. REPORTEDLY, "LENS WAS NOT AVAILABLE. HAD TO WAIT FOR STAAR TO MANUFACTURE (NOT AN OFF SHELF LENS). OUTCOME IS OFF TARGET." STATUS OF THE EYE IS STATED, "UNHAPPY PATIENT BECAUSE OF THE RESULTANT REFRACTIVE SURPRISE ANISOMETROPIA MAKES IT DIFFICULT TO CORRECT WITH GLASSES CL INTOLERANT." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5: DIOPTER SHOULD BE CORRECTED TO: -.50/6/140. CLAIM#(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17089408&pc=QCB | 2023826-2023-02232 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/4.5/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "BOTH LENS ARE PLACED VERTICALLY." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093105&pc=QCB | 2023826-2023-02236 | 2023-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/035 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO BLURRED VISION. REPORTEDLY, "THE PATIENT WAS NOT SATISFIED WITH THE POSTOPERATIVE VISUAL QUALITY, AND THE CLINICIAN FOUND THAT THE PATIENT LACKED CORRECTION FOR 0.5D THROUGH MYDRIASIS AND OPTOMETRY, AND FULLY COMMUNICATED THAT THE PATIENT WAS EXTREMELY REQUIRED TO REPLACE THE LENS." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK.". Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD <3.0. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102366&pc=QCB | 2023826-2023-02186 | 2023-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS. THE PROBLEM WAS RESOLVED AND THE EYE WAS QUIET. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 729514 | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17105578&pc=QCB | 2023826-2023-02268 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -09.50/4.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 24-FEB-2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17105579&pc=QCB | 2023826-2023-02269 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50/3.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119691&pc=QCB | 2023826-2023-02163 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0 12.6, -9.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. LENS WAS RETURNED IN ONE HALF. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL/ FUNCTIONAL INSPECTION DUE TO LENS DAMAGE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140261&pc=QCB | 2023826-2023-02405 | 2019-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -09.00/+20/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. THE PATIENT NOW HAS REFRACTIVE SURPRISE. LEN REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159843&pc=QCB | 2023826-2023-02439 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+2.5/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+2.5/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INSERTED UPSIDE DOWN. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENGTH AND MODEL LENS. THIS DID NOT RESOLVE THE PROBLEM. LASER TOUCH UP SCHEDULED. D6B: EXPLANT DATE: (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17185665&pc=QCB | 2023826-2023-02524 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/2.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WHICH WAS NOTED ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210598&pc=QCB | 2023826-2023-02570 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.0/3.0/7 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, WITH FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216367&pc=QCB | 2023826-2023-02428 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/1.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. REFRACTIVE SURPRISE WAS OBSERVED. LASER TOUCH UP WAS PERFORMED ON (B)(6)2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS FINE AND ALL IS FINE. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/78 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LASER VISION CORRECTION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "CORRECTED WITH TOUCH UP LVC". THE STATUS OF THE EYE AS "PATIENT IS HAPPY & ALL IS FINE!". ADDITIONAL INFORMATION PROVIDED: "AFTER THE SURGERY THERE WAS RESIDUAL REFRACTION OF -0.50/-0.25/35 UCVA 20/30" THE CAUSE OF THE EVENT IS UNKNOWN, AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4625: REFRACTIVE TREATMENT(LASER TOUCH UP). H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17217896&pc=QCB | 2023826-2023-02504 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.0/1.0/82 (SPHERE/CYLIDNER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY "DARE TO CORRECT LOW MONOVISION." Manufacturer Narrative: H6: OFF-LABEL AGE<45. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17222268&pc=QCB | 2023826-2023-02601 | 2021-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.5/3/009 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REFRACTIVE TREATMENT WAS PERFORMED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "ALL IS GOOD PATIENT IS HAPPY," AND "THE RESIDUAL RX: +1.25/-1.25/189 (UCVA 20/40) WAS CORRECTED BY LVC TOUCH UP ON (B)(6) 2022." Manufacturer Narrative: A4-A6:UNK. H3: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253208&pc=QCB | 2023826-2023-02693 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT LENGTH AND POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17261853&pc=QCB | 2023826-2023-02737 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2022. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273273&pc=QCB | 2023826-2023-02722 | NA | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/+4.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE SURGEON PERFORMED PRK SURGERY ON THE PATIENT. LENS REMAINS IMPLANTED. THE PROBLEM WAS REPORTED AS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278996&pc=QCB | 2023826-2023-02610 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM0_12.6, -10.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DUE TO SIGNIFICANT LENS HAPTIC DAMAGE, DIMENSIONAL INSPECTION NOT POSSIBLE TO PERFORM. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283187&pc=QCB | 2023826-2023-02724 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283882&pc=QCB | 2023826-2023-02676 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -5.00/2.0/092(SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "DATA STATUS IS FINE. THE PATIENT IS UNDER OBSERVATION BECAUSE OF LACK OF VISUAL ACUITY". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4) | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283884&pc=QCB | 2023826-2023-02675 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5/1.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS LATER IMPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "DATA STATUS IS FINE. THE PATIENT IS UNDER OBSERVATION BECAUSE OF LACK OF VISUAL ACUITY." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297503&pc=QCB | 2023826-2023-02785 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER, SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297504&pc=QCB | 2023826-2023-02786 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A DIFFERENT POWER, SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313215&pc=QCB | 2023826-2023-02878 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -3.50/2.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE A REPLACEMENT LENS OF A LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322358&pc=QCB | 2023826-2023-02846 | 2023-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/2.0/163 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND RESIDUAL MANIFEST REFRACTION. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. " ERROR IN MRX VS POOR ICL SELECTION" WAS REPORTED FOR CAUSE DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION THE PATIENT IS CURRENTLY STABLE REFRACTIVELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322375&pc=QCB | 2023826-2023-02847 | 2023-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/032 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON \(B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND RESIDUAL MANIFEST REFRACTION. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. " ERROR IN MRX VS POOR ICL SELECTION" WAS REPORTED FOR CAUSE DETAILS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PATIENT IS CURRENTLY STABLE REFRACTIVELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330143&pc=QCB | 2023826-2023-02929 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: RETURNED DEVICE EVALUATION, LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330144&pc=QCB | 2023826-2023-02927 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17338891&pc=QCB | 2023826-2023-02980 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.50/6.0/138 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN 2016. THE SURGEON REPORTS THE ICL AXIS IS AT 5 DEGREES BUT REFRACTIVE IS +2.50-6.25@157 AND INDICATED THE PATIENT HAS ASTIGMATISM. POSSIBLE LENS ROTATION PLANNED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (KERATOCONUS). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17380973&pc=QCB | 2023826-2023-03138 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.00/3.5/162 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.0/+3.5/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR WANTS TO REMOVE THE ICL AND IMPLANT ANOTHER ONE". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17383221&pc=QCB | 2023826-2023-03139 | NA | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.50/3.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS REPORTED THAT ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17385329&pc=QCB | 2023826-2023-03154 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/+5.0/58 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391261&pc=QCB | 2023826-2023-03039 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6) 2023; A3 - M; B3 - 10-MAY-2023. B5 - REPORTED AS: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. - UPDATE TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/2.0/128 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING D4 - VTICM5_13.2, T1347711, 31-JAN-2026; D6A - (B)(6) 2023; D10 - INJECTOR MODEL#MSI-PF,LOT#UNK, CARTRIDGE MODEL#SFC-45,LOT#UNK, FOAM TIP PLUNGER MODEL#FTP,LOT#UNK; H4 - 07-FEB-2023; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: A2 - UNK A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. D6B - UNK. H4 - UNK. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417150&pc=QCB | 2023826-2023-03133 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.0/3.0/002 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED A SAME LENGTH SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMETRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMATISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS CORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CORRECTED DATA: 4627-LENS EXPLANT ADDED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417153&pc=QCB | 2023826-2023-03132 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/2.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED A SAME LENGTH SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMETRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMATISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS CORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING CORRECTED DATA: 4627 EXPLANT ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440469&pc=QCB | 2023826-2023-03207 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.00/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SUPRISE HAS BEEN OBSERVED. REPORTEDLY, "PATIENT UNDERWENT SURGERY ON (B)(6) 2023, THE NEXT DAY SHE HAD VA 0.4. IN THE REVISION OF (B)(6) 2023, SHE STILL HAD VA 0.4 WITH RESIDUAL REFRACTION OF +1.00-0.75X165 VA 1.0. AT THE LAST EXAMINATION ON (B)(6) 2023, SHE HAS A +1.25-1.00X165 AV 1.0. LENS IS AT 163." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446725&pc=QCB | 2023826-2023-03291 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.50/1.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WHICH WAS NOTED ON (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023, THE LENS WAS EXPLANTED AND REPLACED WITH A SIMILAR LENS. PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17447261&pc=QCB | 2023826-2023-03213 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0/0.5/010 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, "BOTH TIMES IN HORIZONTAL POSITION." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 "THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE" SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED FOR A SAME SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE". CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17476696&pc=QCB | 2023826-2023-03334 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.; -2.50/+6.0/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477169&pc=QCB | 2023826-2023-03333 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -0.50/+2.0/143 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17483721&pc=QCB | 2023826-2023-03348 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484473&pc=QCB | 2023826-2023-03355 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. PATIENTS VISION AFTER SURGERY WAS 20/50 +1 WITH VAULT AROUND 1000. AFTER 2 WEEKS WITH NO VISUAL IMPROVEMENT, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.00/2.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY "NO VISUAL IMPROVEMENT. (VA 20/50). HIGHER VAULTING WAS ALSO NOTED". ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENS LENGTH AND A DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486827&pc=QCB | 2023826-2023-03266 | 2023-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.50/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505812&pc=QCB | 2023826-2023-03324 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17545344&pc=QCB | 2023826-2023-03436 | 2023-06-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/6.00/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546728&pc=QCB | 2023826-2023-03519 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/157 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED DUE TO REFRACTIVE SURPRISE. PER SURGEON ROTATION WENT WELL, NO PROBLEMS ENCOUNTERED, PATIENT NOTICED INSTANTANEOUS IMPROVEMENT POST OP. PER UPDATED INFORMATION TORIC MARKINGS ARE @ 98 RATHER THAN 103, THE SURGEON STATES " 5 DEGREE OUT IS MINIMAL. IT SHOULDN'T BE SIGNIFICANT. I'D WAIT AND REVIEW MONTHLY." THE CAUSE OF THIS EVENT WAS DUE TO USER ERROR (TYPING ERROR), THE AXIS ENTERED WAS INCORRECT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549362&pc=QCB | 2023826-2023-03447 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -4.0/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# 730787. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549364&pc=QCB | 2023826-2023-03446 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/5.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL AND CORRECTIONS: ADDITIONAL INFORMATION: B1 - REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B3 - DATE OF EVENT (B)(6) 2023. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/5.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, EYE IS QUIET AND LENS IN POSITION AT DESIRED AXIS. PATIENT AND DOCTOR ARE HAPPY. D6A - (B)(6) 2023. H6 - HEALTH EFFECT - IMPACTT CODE REPORTED AS 2199 - UPDATE TO 4627. CORRECTION: A1 - REPORTED AS (B)(6). CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE REPORTED AS 4627 IN SUPP1 - UPDATE TO 4628. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549367&pc=QCB | 2023826-2023-03452 | 2023-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549368&pc=QCB | 2023826-2023-03451 | 2023-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -11.50/3.5/048 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549381&pc=QCB | 2023826-2023-03453 | 2023-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE AND BLURRED VISION WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17552477&pc=QCB | 2023826-2023-03524 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -5.0/1.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS MARKED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS EXPLANTED ON (B)(6) 2023 AND REPLACED THE SAME DAY WITH SAME LENGTH DIFFERENT DIOPTER LENS. PROBLEM RESOLVED. STATUS OF EYE "ALL FINE! PATIENT IS HAPPY! SUBJ REFRACTION: -1.50DPT." H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17568859&pc=QCB | 2023826-2023-03512 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A REPLACEMENT 13.2MM VTICMO13.2 WITH A -15.0/1.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "PT NOTES LESS ABBERRATIONS, SORE, STILL BLURE BUT BETTER." Manufacturer Narrative: A4-A6:UNK. H6: OFF LABEL: ACD<3.0. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581263&pc=QCB | 2023826-2023-03595 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50/0.50/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17595281&pc=QCB | 2023826-2023-03531 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2023 TO RESOLVE THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. LENS MODEL SIZE RETURNED IS NOT LENS MODEL SIZE STATED IN CLAIM. DIMENSIONAL INSPECTION FOUND THE OVERALL WIDTH WITHIN SPECIFICATIONS BUT OVERALL LENGTH FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599751&pc=QCB | 2023826-2023-03624 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A SMALL REFRACTIVE ERROR WAS OBSERVED. REPORTEDLY, "THE TORIC ICL REMAINS IN PLACE AND THE PATIENT DOES NOT WANT TO DO ANYTHING AS SHE IS SO MUCH BETTER THAN SHE WAS BEFORE SURGERY." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -10.0/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627778&pc=QCB | 2023826-2023-03720 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/1.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT POWER/BRAND ICL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627779&pc=QCB | 2023826-2023-03721 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/1.0/058 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT POWER/BRAND ICL AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17628351&pc=QCB | 2023826-2023-03605 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637154&pc=QCB | 2023826-2023-03753 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.0/6.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE SURGEON REPOSITIONED THE LENS AND THE PATIENT IS PLANO (NO VISUAL PROBLEMS). THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17644938&pc=QCB | 2023826-2023-03693 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/94 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SLIGHTLY WEAKER POWER; SAME MODEL AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17646185&pc=QCB | 2023826-2023-03794 | 2022-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655764&pc=QCB | 2023826-2023-03800 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED AND NOTED " DUE TO PSYCHOLOGICAL REASONS, THE PATIENT COMPLAINT OF SUBNORMAL VISION AND SORE EYES MULTIPLE TIMES WITHIN A MONTH AFTER SURGERY. SHE ALSO REFUSED CONSERVATIVE TREATMENT AND REQUESTED THE REMOVAL OF THE ICL FROM BOTH EYES". Manufacturer Narrative: (B)(4) Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H4 - 30-JUN-2025 CORRECTED TO 22-JUL-2022. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655766&pc=QCB | 2023826-2023-03802 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/0.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED AND NOTED "DUE TO PSYCHOLOGICAL REASONS, THE PATIENT COMPLAINT OF SUBNORMAL VISION AND SORE EYES MULTIPLE TIMES WITHIN A MONTH AFTER SURGERY. SHE ALSO REFUSED CONSERVATIVE TREATMENT AND REQUESTED THE REMOVAL OF THE ICL FROM BOTH EYES". Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND HAPTIC BROKEN. ADDITIONAL COMMENTS STATE, "UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL VERIFICATION DUE TO SIGNIFICANT LENS DAMAGE". CLAIM #(B)(4). Manufacturer Narrative: (B)(4) | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669071&pc=QCB | 2023826-2023-03737 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/5.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS ROTATED DUE TO DUE TO LARGE AMOUNT OF ASTIGMATISM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, AFTER ROTATION THE ASTIGMATISM WAS HIGHLY IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17671454&pc=QCB | 2023826-2023-03839 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, REFRACTIVE SURPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695113&pc=QCB | 2023826-2023-03946 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.0/+2.5/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6). 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT STILL HAS THE SAME CHECK UP DATA TO A PRIOR UPSIDE-DOWN LENS IMPLANT AND THE PATIENT CURRENTLY HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IT WAS REPORTED THE SURGEON HAS PLANS TO PERFORM A LASER TOUCH UP ON THE PATIENT. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17696270&pc=QCB | 2023826-2023-03773 | 2023-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: PRODUCT WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5: LENS ROTATION AND REPOSITIONING SHOULD BE INCLUDED. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716004&pc=QCB | 2023826-2023-03955 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.50/3.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS VAULT AND 20/50 +1 VISION. THE REPORTER STATES THAT THIS REPLACEMENT RESOLVED THE PROBLEM BUT THERE IS SOME RESIDUAL REFRACTIVE ERROR. SURGEON WILL MONITOR FOR STABILITY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716019&pc=QCB | 2023826-2023-03956 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.50/3.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS VAULT AND 20/50 +1 VISION. THE REPORTER STATES THAT THIS REPLACEMENT RESOLVED THE PROBLEM BUT THERE IS SOME RESIDUAL REFRACTIVE ERROR. SURGEON WILL MONITOR FOR STABILITY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716167&pc=QCB | 2023826-2023-03957 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.00/1.0/17 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722426&pc=QCB | 2023826-2023-03740 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO BLURRED VISION AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725920&pc=QCB | 2023826-2023-04007 | 2023-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.5/+0.5/158 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME WHICH WAS ASSESSED ON (B)(6) 2023. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733557&pc=QCB | 2023826-2023-04035 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE SURGEON INDICATED THAT A 13.2MM IMPLANTABLE COLLAMER LENS OF -7.5/1.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON STATES THAT THE LENS ROTATED 15 MONTHS AFTER SURGERY, BLURRED VISION WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT HAD REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THIS WAS ADDITIONALLY REPORTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS LATER ADDITIONALLY REPORTED THAT THE SURGEON PERFORMED LENS REPOSITIONING ON (B)(6) 2023 AND "THE PATIENT HAD A POSTOP ON (B)(6) 2023 AND IS NOW SEEING 20/20 AGAIN. IF ANYTHING ELSE CHANGES AFTER HIS POSTOP ON (B)(6) 2023, DR. (B)(6) WILL HAVE US REACH OUT." CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733896&pc=QCB | 2023826-2023-04023 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+1.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE SAME MODEL AND LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6A: "ON (B)(6) 2023" SHOULD BE ADDED. D6B: "ON (B)(6) 2023" SHOULD BE ADDED. D3: "08/18/2023" SHOULD BE REMOVED AND "07/26/2023" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733897&pc=QCB | 2023826-2023-04024 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/+2.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17751507&pc=QCB | 2023826-2023-04000 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL, SAME LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "LRI WAS PERFORMED FOR ASTIGMATISM." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH FIBER ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17769660&pc=QCB | 2023826-2023-04103 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- SHOULD BE CORRECTED TO STATE: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IN A SEPARATE SURGERY THAT SAME DATE REPLACED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS UNKNOWN. H6: 4627-SHOULD BE ADDED AND 4629 SHOULD BE OMITTED FOR CORRECTION. H6: 1494-OFF-LABEL : IT SHOULD BE NOTED THAT THE PATIENT'S AGE FELL OUTSIDE THE INDICATED AGE RANGE AT THE TIME OF IMPLANTATION. ADDITIONAL INFORMATION: D9: 15-SEP-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DAMAGED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND SPLIT AND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE UNABLE TO BE PERFORMED DUE TO THE DAMAGED STATE OF THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778161&pc=QCB | 2023826-2023-04110 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.0/4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17799577&pc=QCB | 2023826-2023-04172 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.0/+6.0/23 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED THERE WAS DEVICE CAUSALITY AND "REFRACTIVE IN OCOS WAS -0.25/-5.0X108 BUT SHOULD HAVE BEEN -0.25/-5.0X168 (60 DEGREES DIFFERENT FROM WHAT IT SHOULD HAVE BEEN)". Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON 11-SEP-2023 THE LENS WAS REPOSTIONED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4) | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814792&pc=QCB | 2023826-2023-04276 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 31-JUL-2023. THE REPORTER REPORTS EXCESSIVE VAULT AND STATES THAT THERE IS STILL RESIDUAL ASTIGMATISM, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/85 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814804&pc=QCB | 2023826-2023-04277 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE REPORTER REPORTS EXCESSIVE VAULT AND STATES THAT THERE IS STILL RESIDUAL ASTIGMATISM, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/92 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814901&pc=QCB | 2023826-2023-04230 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. CAUSE OF THE EVENT IS UNKNOWN. THE DEVICE DID NOT PERFORM AS INTENDED. REPORTEDLY,"OS ICL HAS ROTATED AT 8 WEEKS POST-OP. CAUSE UNKNOWN." STATUS OF EYE IS STATED, "DECREASED VISION, BURNING SENSATION," AND "PATIENT HAD ICL INSERTED ON (B)(6) 2023. SURGERY WAS ROUTINE, AND ICL WAS INSERTED INTO THE SULCUS AT THE TARGET AXIS OF 175 DEGREES (5 DEGREE ROTATION). VISION WAS 6/6 (20/20) POST OP, INCLUDING UPON EXAMINATION ON (B)(6) 2023. ON (B)(6) 2023 THE PATIENT PRESENTED TO THE CLINIC WITH BLURRED VISION WHICH HAD STARTED THAT MORNING AND WAS NOT PRESENT BEFORE GOING TO BE THE NIGHT BEFORE. NO PAIN WAS NOTED BUT PATIENT MENTIONED THE EYE FELT A BURNING SENSATION 4 DAYS EARLIER, PRIOR TO VISION CHANGE. NO TRAUMA WAS NOTED BY THE PATIENT. UPON EXAMINATION THE ICL HAD ROTATED TO AN AXIS OF 70 DEGREES SINCE THE LAST POST-OP VISIT." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822039&pc=QCB | 2023826-2023-04304 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/+2.0/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822446&pc=QCB | 2023826-2023-04281 | 2022-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -3.50/1.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PER REPORTER THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED "PATIENT DIDNT REACH EXPECTED VA" WAS REPORTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840454&pc=QCB | 2023826-2023-04286 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.50/+3.50/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LENS REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT AFTER TRAUMA, "POSSIBLE SMALL LENS" REPORTED. PER REPORTER PATIENT WAS RECENTLY SEEN WITH DECREASED VA (NOTICED 1 MONTH AGO), EXAMINTATION SHOWED A DISLOCATED IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840487&pc=QCB | 2023826-2023-04287 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.50/+2.50/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. ON (B)(6) 2018 LENS REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT AFTER TRAUMA, "POSSIBLE SMALL LENS" REPORTED. PER REPORTER PATIENT WAS RECENTLY SEEN WITH DECREASED VA (NOTICED 1 MONTH AGO), EXAMINATION SHOWED A DISLOCATED IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846529&pc=QCB | 2023826-2023-04241 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: A4-A6:UNK. H6: USER ERROR . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 4627-EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE ON THE LENS, SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848603&pc=QCB | 2023826-2023-04371 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/4.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857611&pc=QCB | 2023826-2023-04373 | 2023-09-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -7.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857612&pc=QCB | 2023826-2023-04374 | 2023-09-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -8.50/1.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857763&pc=QCB | 2023826-2023-04405 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -2.00/1.00/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE PATIENT WAS UNHAPPY AND THERE WERE SOME VAGUE VISUAL COMPLAINTS. THE SURGEON INDICATED THAT THERE ARE GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE, AND A PATIENT RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862021&pc=QCB | 2023826-2023-04064 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -12.0/3.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS WAS REPOSITIONED ON (B)(6) 2023, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2022). THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE-14-NOV-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#731573. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874041&pc=QCB | 2023826-2023-04416 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882808&pc=QCB | 2023826-2023-04422 | 2023-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17890610&pc=QCB | 2023826-2023-04449 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.0/1.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING, GLARE/HALOES, BLURRED VISION, ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND A LONGER LENGTH LENS WAS IMPLANTED ON THIS SAME DATE AND RESOLVED THE PROBLEM. AN ADDITIONAL SUTURE WAS ALSO LATER USED ON (B)(6) 2023. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "ICL TOO SMALL". Manufacturer Narrative: B5 - NO REPOSITIONING WAS DONE FOR RIGHT EYE (OD). TRACE CATARACT WAS NOTED AS WELL. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17892475&pc=QCB | 2023826-2023-04454 | 2023-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.00/1.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. EXCESSIVE VAULT WAS REPORTED AND BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936453&pc=QCB | 2023826-2023-04364 | 2023-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.3 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.00/4.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE AND ROTATION WAS OBSERVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.00/4.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. CLAIM #(B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4) | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936654&pc=QCB | 2023826-2023-04366 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.50/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936660&pc=QCB | 2023826-2023-04367 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.50/2.50/091(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936727&pc=QCB | 2023826-2023-04544 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND HAPTIC TORN. THE RETURNED LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954380&pc=QCB | 2023826-2023-04626 | 2023-09-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.50/4.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17966720&pc=QCB | 2023826-2023-04648 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.0/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT), ACD <3.00MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970392&pc=QCB | 2023826-2023-04668 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/2.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH, SPHERICAL LENS WITH DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (AGE < 21). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981364&pc=QCB | 2023826-2023-04658 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/4.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE CODE 1494: OFF-LABEL USE: PRE-EXISTING KERATOCONUS; PREVIOUS CORNEAL OR REFRACTIVE SURGERY. CLAIM#(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981402&pc=QCB | 2023826-2023-04659 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/4.0/024 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE CODE 1494: OFF-LABEL USE: PRE-EXISTING KERATOCONUS; PREVIOUS CORNEAL OR REFRACTIVE SURGERY. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986850&pc=QCB | 2023826-2023-04595 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH DIFFERENT MODEL AND LONGER LENGTH LENS DUE TO ROTATION ASSOCIATED WITH LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6: PATIENT CODE: 2137 - REFRACTIVE SURPRISE CORRECTED TO 2923 LENS ROTATION MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17998684&pc=QCB | 2023826-2023-04774 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.50/2.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999785&pc=QCB | 2023826-2023-04730 | 2023-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN WITH THE DEVICE NOT PERFORMING AS INTENDED. Manufacturer Narrative: WEIGHT-RACE:UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18008643&pc=QCB | 2023826-2023-04811 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18030398&pc=QCB | 2023826-2023-04843 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/2.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036112&pc=QCB | 2023826-2023-04812 | 2023-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/+3.5/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, LIGHT ANISOCORIA/BLURRY VISION. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052034&pc=QCB | 2023826-2023-04847 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO REFRACTIVE SUPRISE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18058847&pc=QCB | 2023826-2023-04795 | 2023-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.50/4.5/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED.THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18066237&pc=QCB | 2023826-2023-04856 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/007 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE FIRST TICL IS IN A VERTICAL POSITION." Manufacturer Narrative: H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075201&pc=QCB | 2023826-2023-04920 | 2023-09-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED. LENS REMIANS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-12.0/+1.0/80 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075202&pc=QCB | 2023826-2023-04919 | 2023-09-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED. LENS REMIANS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075217&pc=QCB | 2023826-2023-04927 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18077635&pc=QCB | 2023826-2023-04797 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2019. REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18077667&pc=QCB | 2023826-2023-04798 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)V2019. REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080760&pc=QCB | 2023826-2023-04994 | 2023-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/009 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18101041&pc=QCB | 2023826-2023-05056 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.00/5.5/020 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. THE REPORTER STATES " ASTIGMATISM AXIS WAS WRONG". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18125522&pc=QCB | 2023826-2023-05097 | 2023-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.50/+1.0/125 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT HAD LOW VAULTING; REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED THAT "CONSULTING WITH MA WHAT TO DO WITH THE VAULT AND REFRACTIVE SURPRISE". LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129704&pc=QCB | 2023826-2023-05001 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -13.0/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129705&pc=QCB | 2023826-2023-05004 | NA | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -3/0.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE/CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2020 AND (B)(6) 2023. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -3/0.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE/CHANGE OVERTIME WAS ASSESSED ON (B)(6) 2020 AND (B)(6) 2023. LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+0.5/96 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129709&pc=QCB | 2023826-2023-05000 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -13.0/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136692&pc=QCB | 2023826-2023-05059 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.50/4.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS REPORTED. PER UPDATED INFORMATION LENS ROTATION WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18151978&pc=QCB | 2023826-2023-05087 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/3.0/132 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE IS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166188&pc=QCB | 2023826-2023-05169 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE IN A PATIENT WHO HAD SURGERY IN LATE MARCH/EARLY (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.5/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED LASIK SURGERY TO REDUCE REMAINING CYLINDER POST ICL AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170503&pc=QCB | 2023826-2023-05148 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.00/+4.0/141 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE, LENS CONTINUES TO ROTATE. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATES THAT THE LENS MAY BE A SMALL SIZE FOR THE PATIENTS EYE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170905&pc=QCB | 2023826-2023-05149 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50/+2.5/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE, LENS CONTINUES TO ROTATE. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATES THAT THE LENS MAY BE A SMALL SIZE FOR THE PATIENTS EYE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18173823&pc=QCB | 2023826-2023-05218 | 2023-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/+5.5/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6 2023. THE SURGEON REPORTS THERE IS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING; REFRACTIVE SURPRISE AND ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "CALCULATION ERROR CAUSED BY USE OF THE WRONG AXIS; AFTER REPOSITION INTO THE INTENDED AXIS; THE LENS ROTATES AGAIN WITHOUT LOW VAULT". Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18173829&pc=QCB | 2023826-2023-05217 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179664&pc=QCB | 2023826-2023-05220 | 2023-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -2.00/1.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188684&pc=QCB | 2023826-2023-05254 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.50/-1.50/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER AND SPHERICAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188687&pc=QCB | 2023826-2023-05255 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -2.00/-0.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER AND SPHERICAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18230975&pc=QCB | 2023826-2023-05364 | 2023-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.60/4.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 05-JUL-2023. REFRACTIVE SURPRISE IS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258190&pc=QCB | 2023826-2023-05310 | 2023-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.0/2.5/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18268549&pc=QCB | 2023826-2023-05386 | 2023-11-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2023 AS AN EXCHANGE LENS. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MDR # 2023826-2023-01948 FOR INITIAL LENS. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271595&pc=QCB | 2023826-2023-05503 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.00/0.5/021 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279460&pc=QCB | 2023826-2023-05477 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A LONGER LEGNTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298063&pc=QCB | 2023826-2023-05433 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MMVTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18298064&pc=QCB | 2023826-2023-05434 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MMVTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18314824&pc=QCB | 2023826-2023-05633 | 2023-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/1.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND AXIS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18314840&pc=QCB | 2023826-2023-05632 | 2023-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.5/2.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON 26-OCT-2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18332220&pc=QCB | 2023826-2023-05584 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18340519&pc=QCB | 2023826-2023-05680 | 2023-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -6.0/1.0/050 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18352639&pc=QCB | 2023826-2023-05699 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM513.7, -8.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18352702&pc=QCB | 2023826-2023-05700 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM513.7, -9.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369992&pc=QCB | 2023826-2023-05623 | 2023-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.00/3.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378510&pc=QCB | 2023826-2023-05717 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.5/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION AND REFRACTIVE SURPRISE BUT IT DID NOT RESOLVE THE PROBLEM, LENS ROTATED AGAIN. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387275&pc=QCB | 2023826-2023-05783 | 2023-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.0/1.0/56 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432640&pc=QCB | 2023826-2024-00106 | NA | Malfunction | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/3.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LENS ROTATION AND REFRACTIVE SURPRISE. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BLUR VISION FOR DISTANCE AND NEAR". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18438447&pc=QCB | 2023826-2024-00120 | 2023-10-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -6.0/+6.0/19 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES THERE IS LENS ROTATION AND THE PATIENTS VISION IS BLURRY. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18438448&pc=QCB | 2023826-2024-00121 | 2023-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.00/+1.5/118 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAS LOW VAULTING, REFRACTIVE SURPRISE AND BLURRED VISION. LENS REMAINS IMPLANTED. THE SURGEON ADDITIONALLY NOTES THE MAIN CONCERN IS THE LOW VAULTING. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18452220&pc=QCB | 2023826-2024-00108 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -14.50/2.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 12-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE.THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT THE FIRST LENS WAS IN VERTICAL POSITION. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18454578&pc=QCB | 2023826-2024-00109 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.50/1.5/007 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 12-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT BOTH LENSES WAS IN VERTICAL POSITION. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473804&pc=QCB | 2023826-2024-00135 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/4.0/082 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2015. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18499009&pc=QCB | 2023826-2024-00211 | 2023-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM513.2, -5.5/0.5/042 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18505077&pc=QCB | 2023826-2024-00118 | 2023-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.00/2.00/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PROBLEM RESOLVED. PATIENT-RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THE EVENT. HOWEVER, IT WAS REPORTED THAT THE LENS DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18521107&pc=QCB | 2023826-2024-00229 | NA | Malfunction | STAAR SURGICAL COMPANY | 2024-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VTICM512.6 -1.00/5.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525784&pc=QCB | 2023826-2024-00174 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -6.0/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 05/DEC/2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11014078&pc=QCB | 2023826-2020-02992 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.50/+6.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11270315&pc=QCB | 2023826-2021-00283 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS TEMPORAL IRIS ATROPHY-MYDRIASIS ASPECT AND IRIS PIGMENT DISPERSION. THE PATIENT COMPLAINS ABOUT THE QUALITY OF VISION. Manufacturer Narrative: DATE OF EVENT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788215&pc=QCB | 2023826-2021-01350 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+6.0/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE A SURGICAL PI WAS ADDED, THE SURGEON REPORTING EXCESSIVE VAULT, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF ICA, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED IS DEVICE: "THE ICL PHAKIC LENS IS LONGER THAN THE WTW MEASURE IN HER EYE, THE EXCESSIVE VAULT IS CAUSING SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES MERITING A PERIPHERAL IRIDECTOMY. ALSO THERE WAS A REFRACTIVE SURPRISE." THE SURGEON HAS DECIDED TO LEAVE THE LENS IMPLANTED, KEEPING THE PATIENT UNDER SURVEILLANCE. Manufacturer Narrative: B5-ADDITIONAL INFO: THE DOCTOR REPORTS THE PATIENT IS IMPROVING REFRACTION AND LEAVES THE PATIENT WITH A SPHERE OF +0.50 WHICH IS BETTER THAN THE PUBLICATION RESULT. IRIDOTOMIES WERE PERFORMED AND THE PATIENT IS STABLE. THE DOCTOR CONSIDERS THE CASE CLOSED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147111&pc=QCB | 2023826-2021-02410 | 2021-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/161 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME AND TRANSIENT LOSS OF BCVA. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE, RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS FIBER ON LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4) Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. REFRACTIVE CHANGE OVERTIME. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12182327&pc=QCB | 2023826-2021-02729 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM, VTICM5_13.2, -11.0/+6.0/148, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH KERATOCONUS. IT WAS REPORTED THAT EXCESSIVE VAULT, REFRACTIVE SURPRISE, NARROWING OF ANGLES DUE TO VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: ON (B)(6) 2021;. LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874528&pc=QCB | 2023826-2021-04516 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, TMICL12.1, -15.0/2.0/88, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS. ON PATIENTS POST-OP VISIT ON (B)(6) 2021 FOR STATUS OF THE EYE THE REPORTER REPORTED "BLURRY VISION". ON (B)(6) 2021 VISIT "DECOMPRESSION RETINOPATHY AFTER ICL EXCHANGE" WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED "THE INITIAL ICL WAS TOO LARGE. AFTER REPLACEMENT EMERGENTLY, THE RETINOPATHY OCCURRED." STEROIDS AND NSAIDS WERE PRESCRIBED, IT WAS REPORTED THAT APPROXIMATELY 90% OF THE BLURRY VISION AND RETINOPATHY HAS RESOLVED. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879841&pc=QCB | 2023826-2021-04537 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.50/1.5/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND UNREACTIVE PUPIL. THIS EXCHANGE RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. "OTHER" WAS REPORTED TO BE THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTER STATED " ICL TOO LARGE." Manufacturer Narrative: B5 - CORRECTED TO :"THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -6.50/1.5/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND UNREACTIVE PUPIL. THIS EXCHANGE RESOLVED THE PROBLEM. NO PATIENT INJURY WAS REPORTED. "OTHER" WAS REPORTED TO BE THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTER STATED " ICL TOO LARGE." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286578&pc=QCB | 2023826-2022-00175 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -6.50/+2.50/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; BLURRY VISION AND THE LENS WAS EXCHANGED ON (B)(6) 2022 BUT THIS DID NOT RESOLVE THE PROBLEM. IT WAS LATER CLARIFIED THAT THE PATIENT STILL HAS FUZZY VISION; PUPILS ARE DILATED AND PRK SURGERY WAS PERFORMED. PATIENT STILL HAS BLURRY VISION. CAUSE OF THE EVENT WAS REPORTED AS "OTHER- UNEXPECTED SHORT CILLARY SULCUS SPACE". SEE MDR# 2023826-2022-00177 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311391&pc=QCB | 2023826-2022-00177 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -6.00/+2.50/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. IT WAS REPORTED THE PATIENT FELT IMPROVEMENT BUT VISION IS STILL FUZZY; THE PUPIL WAS STILL DILATED AFTER SURGERY. AT A LATER POSTOP VISIT THE PATIENT HAD -1.00/+1.00/107 VISION;THE SURGEON HAD PERFORMED PRK SURGERY. PATIENT HAS REPORTED VISION IS STILL BLURRY. LENS REMAINS IMPLANTED. SEE MDR# 2023826-2022-00175 FOR INITIAL LENS IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459220&pc=QCB | 2023826-2022-00387 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE AND DIZZINESS. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, "AS OVERCORRECTION OF ASTIGMATISM, THE PATIENT FELT DIZZY AND COULDN'T ADAPT TO IT, SO A NON-TORIC ICL WAS REPLACED." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869343&pc=QCB | 2023826-2022-00886 | NA | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+6.0/159 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT FELT HIS VISION WAS DISTORTED AND BLURRED. HIS FINAL ASTIGMATISM WAS EXTREMELY HIGH AT -6.50D. THE SURGEON PRESCRIBED PILOCARPINE TO SEE IF THIS WOULD IMPROVE HIS VISION. THE LENS REMAINED IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02729. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061336&pc=QCB | 2023826-2022-01131 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.5/+1.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION, IRIS SPHINCTER LOSS AT 5 O'CLOCK, HEAVILY PIGMENTS TM, KIV EXPLANT, AND IRREGULAR PUPIL. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5- PER EMAIL ON 23-JUN-2022 THE REPORTER STATED THAT "THERE IS NO ADVERSE EVENT FOR OS". CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14863456&pc=QCB | 2023826-2022-02037 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF -9.0/4.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS DISLOCATION/SUBLUXATION, AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A NEW MODEL/SAME LENGTH LENS OF DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: 0000001876. UNK. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15386494&pc=QCB | 2023826-2022-03049 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/+0.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "THE SURGEON SUSPECTED THAT THE LENS WAS WITH THE WRONG LABEL". Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN, AND THE RETURNED LENS IS NOT THE SIZE STATED IN THE CLAIM. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465942&pc=QCB | 2023826-2022-03251 | 2022-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+1.0/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN 2019. IT WAS REPORTED THE PATIENT HAS REFRACTIVE SURPRISE AND "PATIENT HAS BEEN PUNCHED IN THE EYE". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801908&pc=QCB | 2023826-2022-04084 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED O (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE INTRAOCULAR PRESSURE (IOP) WAS ELEVATED. INITIATED COMBIGAN BID AND PRED FORTE BID. THE ELEVATED IOP WAS RESOLVED AND WAS MEASURED WITHIN NORMAL RANGE ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS DUE TO LENS WAS TOO SMALL. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15813337&pc=QCB | 2023826-2022-03932 | 2022-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -8.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION IN IRIDO-CORNEAL ANGELS, AND BLURRED VISION.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTION DATA D4: REPORTED EXPIRATION DATE WAS 31-MAY-2022 - CORRECT TO 31-MAY-2025. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856726&pc=QCB | 2023826-2022-04213 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: 1494- OFF LABEL USE: PRE-EXISTING CATARACTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856830&pc=QCB | 2023826-2022-04214 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: 1494- OFF LABEL USE: PRE-EXISTING CATARACTS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882930&pc=QCB | 2023826-2022-04339 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "MOST PROBABLY THE POSTERIOR SULCUS IS NARROWER THAN AVERAGE STATISTICALLY EXPECTED"; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882947&pc=QCB | 2023826-2022-04338 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; GLARE/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "MOST PROBABLY THE POSTERIOR SULCUS IS NARROWER THEN AVERAGE STATISTICALLY EXPECTED"; THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15898502&pc=QCB | 2023826-2022-04377 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND COMPLAINED OF DOUBLE VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309073&pc=QCB | 2023826-2023-00443 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE WAS A HYPEROPIC OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VTICM5_13.2; -6.00/+1.00/157 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME/REFRACTIVE SURPRISE; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT AND HISTORY OF REFRACTIVE SURGERY). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16358874&pc=QCB | 2023826-2023-00497 | 2022-12-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: H3: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC SHOULD BE CORRECTION TO H6 -HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC. ADDITONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VTICM5_13.2; -6.00/+1.00/157 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THERE IS A HYPEROPIC OUTCOME/REFRACTIVE SURPRISE; IT WAS ADDITIONALLY INDICATED LATER THAT LENS HAD ALSO ROTATED. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: 13.2MM VTICM5_13.2; -6.00/+1.00/157 SHOULD BE UPDATED TO 13.2MM VTICM5_13.2; -6.00/+1.00/162. CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIC. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083434&pc=QCB | 2023826-2023-02276 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.5/3.5/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT EXPERIENCED EXCESSIVE VAULTING, HALO/GLARES AND IRIS TRANSILLUMINATION. IT WAS INDICATED THAT IT IS PLANNED TO MONITOR THE PATIENT OVER 6 MONTHS. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - IRIS TRANSILLUMINATION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295356&pc=QCB | 2023826-2023-02833 | 2023-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -12.0/+1.0/135 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTES THERE IS LENS GLISTENING'S ON THE ICL, THE LENS WAS OBSERVED WITH GLISTENING'S/MILD OPACIFICATION MATERIAL ON THE LENS. PATIENT COMPLAINTS OF SLIGHT GLARE/DISTURBANCE DURING DIM LIGHTS CONDITION. LENS REMAINS IMPLANTED. THE SURGEON REPORTS THE PATIENT WILL BE MONITORED WITH NO PLANNED INTERVENTION AT THIS TIME. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIPPER -12.0/+1.0/135 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARES/HALOES, BLURRED VISION, AND 'LENS GLISTENINGS'. REPORTEDLY "BLURRING OF VISION DURING DIM LIGHT CONDITION AND DR YONG FINDS THERE IS A LAYER OF GLISTENING/MILD OPACIFICATION ON THE LENS SURFACE." THE LENS REMAINS IMPLANTED. REPORTEDLY "THERE IS NO PLANNED SURGICAL INTERVENTION AT THIS TIME". THE CAUSE OF THE EVENT IS UNKNOWN. CAUSE DETAILS PROVIDED: "THE ICL LENS IS OBSERVED WITH GLISTENINGS/MILD OPACIFICATION WHERE PT COMPLAINTS SLIGHT MORE GLARE AND DISTURBANCE DURING DIM LIGHTS CONDITION." HEALTH EFFECT CLINICAL CODE: "4581 LENS GLISTENING" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080354&pc=QCB | 2023826-2023-04850 | 2023-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/122 (SPHERE/CYLINDER/AXIS) LENS TORE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA, ON (B)(6) 2023. LENS DISLOCATION/SUBLUXATION AND DOUBLE VISION ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "HAPTICS CUT." THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6: UNK. H6: OFF-LABEL: PRE-EXISTING MYOPIA. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136617&pc=QCB | 2023826-2023-05053 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/100, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: "4581-DOUBLE VISION" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-10.0/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE PATIENT EXPERIENCED BLURRED VISION AND DOUBLE VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH, AND POWER LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437834&pc=QCB | 2023826-2024-00029 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/101(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A SMALLER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH - ONE SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND: 731987. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18446384&pc=QCB | 2023826-2024-00050 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTEDLY IMPLANTED UPSIDE DOWN; SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE IS USER ERROR. . Manufacturer Narrative: H6-CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082445&pc=QCB | 2023826-2021-04976 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON STATES THERE IS A HIGH VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION - UNK. DATE OF THE EVENT -UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. (B)(4). Manufacturer Narrative: PATIENT INFORMATION - UNK. DATE OF THE EVENT -UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). Manufacturer Narrative: B5- ADDITIONAL INFORMATION: THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+3.5/083 (SPHERE/CYLINDER/AXIS) LENS INTO THE PATIENT'S EYE ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE CAUSE IS REPORTED AS DEVICE: "OVERALL LENS DIAMETER TOO BIG." H6-CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER AND THE PROBLEM RESOLVED. FOR STATUS OF THE EYE (SIGN/SYMPTOMS) "MILD IRIS ATROPHY SUBINCISIONAL" WAS REPORTED. HOWEVER, PER UPDATED INFORMATION THE REPORTER STATED "IT'S JUST PAR FOR THE COURSE" CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627814&pc=QCB | 2023826-2022-01824 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.50/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO HIGH VAULT. Manufacturer Narrative: ADDITO NAL INFORMATION: B5 - IT WAS ADDITIONALLY LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED BLURRED VISION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SIZING". THE PROBLEM WAS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627817&pc=QCB | 2023826-2022-01825 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.00/1.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO HIGH VAULT. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS ADDITIONALLY LATER REPORTED THAT THE PATIENT ALSO EXPERIENCED BLURRED VISION AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251274&pc=QCB | 2023826-2023-00242 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/018 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). RESIDUAL MYOPIA WAS OBSERVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50/1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "RESIDUAL MYOPIA". REPORTEDLY " RESIDUAL MYOPIA AND HE WOULD LIKE TO DO AN ICL EXCHANGE". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS "WELL PLACED EVO ICL; EYE IS QUIET". ADDITIONAL INFORMATION PROVIDED WAS "PT'S COMPLAINTS RESOLVED WITH ICL EXCHANGE". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS "OTHER" AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY "OVER ACCOMMODATION DURING PRE-OP MR". CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- "RESIDUAL MYOPIA." CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411020&pc=QCB | 2023826-2023-00568 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT'S VA WAS BLURRY, SEE MFR. REPORT # 2023826-2023-00569 FOR REPLACEMENT LENS. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. EVENT DATE UNKNOWN. ADVERSE EVENT PROBLEM WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) DESPITE NORMAL VAULT. PATIENT'S VA WAS MUCH BLURRIER AT INITIAL POST-OP, WHICH IS WHY THE LENS WAS REMOVED AND REPLACED WITH LONGER LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/2.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT, REPOSITIONING, BLURRED VISION AND INTACS INTRACORNEAL RINGS. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE; "CLEAR VA AND HAD ADJUSTED TO HAZE/SHADOW USING ARTIFICIAL TEARS PRN." ADDITIONAL INFORMATION PROVIDED WAS "CLEAR VA AT NIGHT, DILATED EXAM SHOWS INTACS INTACT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN- " ROTATED VISIAN ICL DESPITE NORMAL VAULT. EXPLANT AND INSERTION OF LONGER ICL TO HELP PREVENT ROTATION. VA MUCH BLURRIER DURING POST OP". H6- HEALTH EFFECT- CLINICAL CODE: 4581-INTACS INTRACORNEAL RINGS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411273&pc=QCB | 2023826-2023-00566 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT'S VA WAS BLURRY, SEE MFR. REPORT # 2023826-2023-00567 FOR REPLACEMENT LENS. THE PATIENT HAD INTACS IMPLANTED AND WAS USING ARTIFICIAL TEARS. Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN; DATE OF EVENT: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) DESPITE NORMAL VAULT. PATIENT'S VA WAS MUCH BLURRIER AT INITIAL POST-OP, WHICH IS WHY THE LENS WAS REMOVED AND REPLACED WITH LONGER LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT, LENS REPOSITIONING, BLURRED VISION AND INTACS (INTRACORNEAL RINGS). ON (B)(6) 2022, THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE; "CLEAR VA AND HAD ADJUSTED TO INTACS USING ARTIFICIAL TEARS PRN." ADDITIONAL INFORMATION PROVIDED WAS "CLEAR VA AT NIGHT, DILATED EXAM SHOWS INTACS INTACT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN- " ROTATED VISIAN ICL DESPITE NORMAL VAULT. EXPLANT AND INSERTION OF LONGER ICL TO HELP PREVENT ROTATION. VA MUCH BLURRIER DURING POST OP". H6- HEALTH EFFECT- CLINICAL CODE: 4581-INTACS INTRACORNEAL RINGS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093103&pc=QCB | 2023826-2023-02235 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.5/3.5/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- KERATOCONUS" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/3.5/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD).THE PATIENT EXPERIENCED KERATOCONUS AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093104&pc=QCB | 2023826-2023-02234 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -14.50/3.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: "4581- KERATOCONUS" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/3.5/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED KERATOCONUS AND RESIDUAL REFRACTIVE ERROR. REPORTEDLY "1 WEEK POST OP KERATOCONUS PATIENT'S OUTCOME". THE LENS REMAINS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18008540&pc=QCB | 2023826-2023-04810 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL "KERATOCONUS" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.00/2.0/165 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY "POST-OPERATIVELY "HIGH VAULTS & REFRACTIVE SURPRISE" AND "2 PATIENTS I RECENTLY DID TORIC ICLS AND UNFORTUNATELY THEY BOTH SEEM TO HAVE HIGH VAULT. THEY ARE BOTH KERATOCONIC WITH HIGH PRESCRIPTIONS SO THEY HAD HIGH POWER LENSES, AND ALTHOUGH THEIR VAULT IS >1000MICRONS THEY HAVE GOOD IOPS AND OPEN ANGLES."THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00/2.0/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540747&pc=QCB | 2023826-2022-03403 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/1.5/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO OVERCORRECTION AND DISCOMFORT IN LOOKING CLOSE. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500060&pc=QCB | 2023826-2023-00803 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.50/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 TO ADDRESS EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. IT WAS REPORTED THAT THE VAULT ISSUE RESOLVED BUT REPORTER STATES THAT THE PATIENT IS VERY UNCOMFORTABLE AND DISAPPOINTED WITH HER REFRACTION RESULTS, LENS REMAINS IMPLANTED. ADDITIONAL COMMENTS FROM REPORTER- "GOOD JUMP BUT REFRACTIVE ERROR. GOOD LENGTH OF THE IMPLANT BUT NOT HYPERMETROPIA POWER. PATIENT WITH KERATOCONUS OD/OS, STRABISM - RETINA LASER." Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND DISCOMFORT. REPORTEDLY THE PATIENT IS "VERY UNCOMFORTABLE AND DISAPPOINTED WITH REFRACTION RESULTS". ADDITIONAL COMMENTS FROM REPORTER- "GOOD JUMP BUT REFRACTIVE ERROR. GOOD LENGTH OF THE IMPLANT BUT NOT HYPERMETROPIA POWER. PATIENT WITH KERATOCONUS OD/OS, STRABISM - RETINA LASER." THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541911&pc=QCB | 2023826-2020-00022 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/6.00/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT RIGHT AFTER SURGERY. THE REPORTER STATED " THE PATIENT CONDITION IS GOOD AND HAVE SAME CONDITION BEFORE THE SURGERY, BCVA 20/25." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATENT CODE 2330 SHOULD BE ADDED TO THE INITIAL MDR. H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541936&pc=QCB | 2023826-2020-00023 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/6.00/78 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT RIGHT AFTER SURGERY. THE REPORTER STATED " THE PATIENT CONDITION IS GOOD AND HAVE SAME CONDITION BEFORE THE SURGERY, BCVA 20/25." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2330 SHOULD BE ADDED TO THE INITIAL MDR. H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563252&pc=QCB | 2023826-2021-00692 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/5.5/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME WERE OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS DEVICE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -5.50/5.5/122 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PATIENT HAD PRE-EXISTING KERATOCONUS AND HISTORY OF PREVIOUS NON-REFRACTIVE CORNEAL SURGERY PRIOR TO IMPLANT. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVERTIME; BLURRED VISION (HAZE); HALOES WERE OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567718&pc=QCB | 2023826-2022-03573 | NA | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -16.0/2.0/124 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. PATIENT EXPERIENCED HALOS AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15847565&pc=QCB | 2023826-2022-04169 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A SMALLER SIZE LENS. THE PROBLEM WAS RESOLVED. THE POST-OP RESULTS WERE VAULT IS 75%, UCVA 20/100, BCVA 20/80, NO EDEMA AND AC IS CLEAR. AT THE (B)(6) 2022 POST-OP VISIT, THE PATIENT WAS UCVA 20/70 -2 @ DISTANCE AND NEAR WAS CF. HIS REFRACTION WAS +1.25 -0.50 X 077, BCVA 20/70 AND THE VAULT WAS AT 100%. PATIENT WILL RETURN FOR A YAG PC ON (B)(6) 2023 DUE TO PIGMENT DUSTING ACROSS THE ICL. THE PATIENT IS DOING WELL. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376164&pc=QCB | 2023826-2023-00550 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED "REMOVE AND NO LONGER IMPLANT". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376167&pc=QCB | 2023826-2023-00551 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS NOTED "REMOVE AND NO LONGER IMPLANT". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002296&pc=QCB | 2023826-2020-00994 | 2017-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.5/1.5/49 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT; B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES); B5- BS THE REPORTER INDICATED THAT A VTICM5_12.6, -10.5/1.5/49 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2020, THE LENS WAS EXCHANGED DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- (B)(6) 2020. H1- SERIOUS . CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO- CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002391&pc=QCB | 2023826-2020-00995 | 2017-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -11.5/1.0/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-BS THE REPORTER INDICATED THAT A VTICM5_12.6, -11.5/1.0/18 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXCHANGED DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION- LENS RETURNED IN A MICRO- CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708891&pc=QCB | 2023826-2020-02442 | 2019-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. REPORTEDLY, THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. THE EXCHANGED RESOLVED THE PROBLEM. D6B- (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10708899&pc=QCB | 2023826-2020-02441 | 2019-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. REPORTEDLY, THE PATIENT MAY HAVE "OVERLY RUBBED HIS EYES" AFTER SCUBA DIVING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 OF -12/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. D6B- (B)(6) 2020. H6-IMPACT CODE 4629; LENS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED MICROCENTRIFUGE VIAL; RETURN DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT WITH DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11033257&pc=QCB | 2023826-2020-03079 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.00/+3.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED, THE REPORTER STATED "VAULT IS 0.75CT, TO OBSERVE IN 3 MONTHS". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REFRACTIVE SURPRISE WAS REPORTED. APPLYING FOR EXCHANGE, REPLACEMENT SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852537&pc=QCB | 2023826-2020-02679 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED DUE TO BLURRED VISION. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158612&pc=QCB | 2023826-2021-00094 | 2020-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/162 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020 DUE TO BEING TOO SHORT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED AND EXCHANGED DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: A2: DATE OF BIRTH - (B)(6) 1979 B3: DATE OF EVENT -(B)(6) 2020 CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238061&pc=QCB | 2023826-2021-00175 | 2020-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -12.00/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -12.00/2.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651612&pc=QCB | 2023826-2021-00914 | 2021-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: G3: DATE ON FOLLOW-UP #1 CHANGED FROM (B)(6) 2021 TO (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: B5 - INCLUDE BLURRED VISION WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651657&pc=QCB | 2023826-2021-00915 | 2021-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - INCLUDE BLURRED VISION WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683204&pc=QCB | 2023826-2021-01077 | 2020-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.7MM, TMICL13.7, -14.0/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. FOR CAUSE OF EVENT OTHER: THE LENS WAS TOO LARGE FOR THE PATIENT. WAS REPORTED. Manufacturer Narrative: B5- IT WAS REPORTED THAT A 13. 7MM, TMICL 13. 7, -14.0/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND BLURRED VISION. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. FOR CAUSE OF EVENT OTHER: THE LENS WAS TOO LARGE FOR THE PATIENT. WAS REPORTED. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11750617&pc=QCB | 2023826-2021-01295 | 2017-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT IS CORRECTED TO (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED, THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. "THE PATIENT COMPLAINED OF BLURRED VISION. AND THE AUTOMATIC REFRACTOR RESULTS SHOWS -1.00DC*10 -> 1.0. THE PATIENT WHO STRONGLY DEMANDED, FOR TICL REPLACEMENT SURGERY". ON (B)(6) 2017, THE LENS WAS REMOVED. AND ON THE SAME DAY, BUT DIFFERENT SURGERY, A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K : THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781966&pc=QCB | 2023826-2021-01397 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: A TICL ROTATION CASE WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THE PROBLEM RESOLVED. CAUSE FOR THIS EVENT WAS PATIENT RELATED FACTOR AND IT WAS ALSO REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: - RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE (2021 REPORTS): UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891166&pc=QCB | 2023826-2021-01675 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P990013 | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -13/2/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (0S) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION (UNKNOWN WHETHER VAULT WAS LOW), IT WAS REPORTED THAT REPOSITIONING RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGN/SYMPTOMS): "RESOLVED" WAS REPORTED. PER STUDY INVESTIGATOR, THE CAUSE OF THIS EVENT WAS RELATED TO THE DEVICE. FOR ADDITIONAL INFORMATION THE REPORTER STATED "PER MEDICAL MONITOR OF NOTE, LENS POWER CALCULATION RECOMMENDED BY STAAR WAS TMICL 13.7, -13.00/2.0/080. HOWEVER, THE DOCTOR CHOSE TO GO WITH THE TMICL 13.2." Manufacturer Narrative: B5- PER UPDATED INFORMATION "AS OF (B)(6) 2021, LENS ROTATED AGAIN AND SUBJECT HAD INCREASED CYLINDER TO BASELINE (2.75 VS. BASELINE 2.25). SLE: ROTATED ICL @ 087 (83 DEGREES OF ROTATION). CLAIM # (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12104944&pc=QCB | 2023826-2021-02220 | 2021-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "BLURRED VISION, RESIDUAL REFRACTION, AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED." SHOULD BE ADDED TO INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE: 1401 - MISFOCUSING ADDED TO INITIAL MDR. H6- CLINICAL CODE: 2137 - BLURRED VISION ADDED TO INITIAL MDR CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171281&pc=QCB | 2023826-2021-02615 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS ROTATED AT ONE WEEK POST-OP. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/3.0/094, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGED RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND THAT PATIENT RELATED-FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER STATED " WTW LESS THAN ACTUAL SULCUS SIZE". WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171287&pc=QCB | 2023826-2021-02614 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS ROTATED AT ONE WEEK POST-OP. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/2.0/086, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THIS EXCHANGED RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED AND THAT PATIENT RELATED-FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER STATED " WTW LESS THAN ACTUAL SULCUS SIZE". H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478395&pc=QCB | 2023826-2021-03564 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON IS PLANNING AN EXPLANT AND EXCHANGE OF THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/035 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DOB: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EVENT DATE: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13515929&pc=QCB | 2023826-2022-00417 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+5.50/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS OF THE SAME LENGTH AND DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546061&pc=QCB | 2023826-2022-00435 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/1.5/99 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IN THE REPORTERS OPINION "PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/1.5/99 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. IN THE REPORTER'S OPINION "PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. H6- MEDICAL DEVICE PROBLEM CODE: 2923 TO 1401- PREVIOUS CODE NOT APPLICABLE, THERE WAS NO LENS ROTATION PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657019&pc=QCB | 2023826-2022-00667 | 2021-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS WHICH WAS ALSO IMPLANTED VERTICALLY WHICH HAD RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CORRECTION: B5 - PATIENT ALSO EXPERIENCED BLURRED VISION - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664194&pc=QCB | 2023826-2022-00659 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: REPORTEDLY, AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE AND THE LENS ROTATED 32 DEGREES 5 DAYS POSTOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 12.6MM VTICMO12.6 -14.5/+2.5/066 (SPHERE/CYLINDER/AXIS) ICL WAS IMPLANTED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE SURGEON REPORTS ADDITIONAL SYMPTOMS OF LOW VAULT, UNSTABLE REFRACTIVE CHANGE OVER TIME, BLURRED VISION, AND "REFRACTIVE ERROR KEEPS CHANGING, ICL IS PROPELLERING." THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY EXPIRATION DATE, IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14657149&pc=QCB | 2023826-2022-01841 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1; -10/4.0/123 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THERE WAS LENS ROTATION AND ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT HAD BLURRED VISION. AC IRRIGATION WAS ALSO PERFORMED. -THIS SHOULD HAVE BEEN INCLUDED IN PREVIOUS MFR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15022388&pc=QCB | 2023826-2022-02269 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.50/2.0/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "PATIENT ALSO EXPERIENCED REFRACTIVE CHANGE OVER TIME." CLAIM: (B)(4). Manufacturer Narrative: CORRECTED DATA: H4- "03/NOV/2013 SHOULD BE CORRECTED TO 11/MAR/2013." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/2.0/89 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT WITHOUT ROTATION AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- HEALTH EFFECT- IMPACT CODE: 2199 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. INVESTIGATION TYPE- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071012&pc=QCB | 2023826-2022-02325 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Insufficient Information | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND IT WAS REPORTED THAT THERE IS A PLANNED ICL EXCHANGE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.0/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. THE REPORTER ALSO STATED THAT THE LENS DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15103186&pc=QCB | 2023826-2022-02446 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -13.0/3.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -13.0/3.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A VERTICALLY IMPLANTED; SHORTER LENGTH AND DIFFERENT POWER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. H6- DEVICE CODE: 1494- OFF LABEL USE (AGE<21) . CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118180&pc=QCB | 2023826-2022-02283 | 2021-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D6A (B)(6) 2017. D6B- UNK. B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON AN UNKNOWN DATE THE LENS WAS REPLACED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT DIFFERENT AXIS LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN." CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136758&pc=QCB | 2023826-2022-02345 | 2021-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -3.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233578&pc=QCB | 2023826-2022-02726 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT (UNDESIRABLE REFRACTIVE OUTCOME). THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE AC IS DEEP AND QUIET. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233848&pc=QCB | 2023826-2022-02738 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -11.0/+4.50/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 16-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15246687&pc=QCB | 2023826-2022-02795 | 2022-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN THE YEAR 2014. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15246688&pc=QCB | 2023826-2022-02796 | 2022-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -8.0/3.50/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN THE YEAR 2014. REFRACTIVE SURPRISE; REFRACTIVE CHANGE OVER TIME WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: 1991. IMPLANT DATE: 2014. (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285627&pc=QCB | 2023826-2022-02521 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -1.0/2.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15294643&pc=QCB | 2023826-2022-02798 | 2022-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -4.0/1.5/63 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299118&pc=QCB | 2023826-2022-02801 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/6.0/025 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE IS NORMAL AC/ NORMAL PUPIL SIZE/ V. HIGH ASTIGMATISM. " IT IS A 5Y POST PKP CASE A PRE-OP CORRECTED VISION REACHES 20/20, FOLLOW UP DONE OVER 2 MONTHS WITH STABLE REF. OF 0.0/-8.0/AX 105." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL/CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/6.0/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND PRK WAS PERFORMED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "NORMAL AC/ NORMAL PUPIL SIZE/ V. HIGH ASTIGMATISM". ADDITIONAL INFORMATION REPORTED WAS "IT IS 5Y POST PKP CASE WITH A PRE OP CORRECTED VISION REACHED 20/20. FOLLOW UP DONE OVER 2 MONTHS WITH A STABLE REF OF 0.0/-18.0/AX 105 ALTHOUGH THE LEFT EYE DONE WITH AN EXCELLENT RESULT." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480682&pc=QCB | 2023826-2022-03282 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: REPORTEDLY, "UNSTERILE." Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS WITH A, -14.0/2.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555029&pc=QCB | 2023826-2022-03516 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS DID NOT RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS 45 YEARS OLD, PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." SEE MFR# 2023826-2022-03312 FOR EXCHANGE LENS. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS 45 YEARS OLD, PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." MFR# 2023826-2022-03312 FOR EXCHANGE LENS IS NOT A COMPLAINT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15566370&pc=QCB | 2023826-2022-03332 | 2022-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/155 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER INDICATED THE LENS HAD ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: UDI# - (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/155 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED AND REPLACED FOR A NEW LENS ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: B5 UPDATE - THE LENS WAS EXCHANGED FOR A FOR A NEW LENS OF DIFFERENT POWER ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCING A REFRACTIVE SURPRISE AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK , RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624407&pc=QCB | 2023826-2022-03547 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -08.50/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCAHNGED WITH THE SAME MODEL; STRONGER POWER; SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15795261&pc=QCB | 2023826-2022-04002 | 2022-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON AN UNKNOWN DATE. ON 01-SEPT-2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/ SIZE LENS BUT DIFFERENT POWER OF -10.50/3.0/061 (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS IMPLANTED; REMOVED AND REPLACED INTRAOPERATIVELY FOR A SAME MODEL/ SIZE LENS BUT DIFFERENT POWER OF -10.50/3.0/061 (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. " A WRONG LENS WAS ORDERED RESULTING IN A REFRACTIVE SURPRISE." H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849968&pc=QCB | 2023826-2022-04260 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/5.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: D6A- IMPLANT DATE IS (B)(6) 2022. CORRECTED DATA: H10 IN PREVIOUS MDR SHOULD CONTAIN THE FOLLOWING WORK ORDER SEARCH RESULTS. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+5.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY "HE DID NOT HAVE CYCLOREFRACTION. THE PREOP SPECTACLE FOR THE LEFT EYE WAS -8.00/-6.50 X 80 (BACK VERTEX DISTANCE 13MM). THE BEST CORRECTED LEFT VISION WAS 6/7.5 HIS REFRACTION WAS CHECKED AND THE CYL CONFIRMED AS STABLE. YES THERE IS A SIGNIFICANT RESIDUAL OCULAR ASTIGMATISM (ROA) BUT WE DECIDED TO TREAT ON REFRACTION. THERE WAS NO CAUSE FOUND TO ACCOUNT FOR THE DISPARITY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4- MODEL#: "VTICM5 12.6" SHOULD BE CORRECTED TO "VTICMO 12.6" CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035423&pc=QCB | 2023826-2022-04548 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, "THE PATIENT'S POSTOPERATIVE UCVA WAS NOT GOOD AND COULDN'T REACH THE PREOPERATIVE CDVA. THERE WAS HIGH VAULT AND REFRACTIVE RESIDUAL, SO THE REFRACTIVE POWER WAS ADJUSTED AT THE SAME TIME WHEN THE LENS WAS REPLACED'. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." IN THE INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065467&pc=QCB | 2023826-2022-04741 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: D9-RETURN DATE: 12-JAN-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC DEFORMED AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383938&pc=QCB | 2023826-2023-00488 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 12-JUL-2022. LOW VAULT WAS OBSERVED. ON 15-JAN-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT AND REFRACTIVE SUPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16397843&pc=QCB | 2023826-2023-00564 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/149 (SPHERE/CYLINDER/AXIS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED THAT VISION WAS NOT GOOD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND SIGNIFICANT BLURRING OF VA. POST-OP, THE BCVA WAS MUCH IMPROVED. THE CAUSE OF THE EVENT WAS THE DEVICE AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CORRECTED DATA: B3: DATE OF EVENT: (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742891&pc=QCB | 2023826-2023-01252 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON(B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: UNK. UNK. UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"2993" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16809012&pc=QCB | 2023826-2023-01477 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON 02-MAR-2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LENS ROTATION. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS EXCHANGED DUE TO REFRACTIVE CHANGE OVERTIME. REPORTEDLY, "LOW VAULT RESULTED IN ROTATION." LENS WAS REMOVED AND REPLACED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS "DEVICE" AND "OTHER." THE REPORTER NOTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. STATUS OF THE EYE IS REPORTED AS "DEEP AND QUIET." CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H3: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A VIAL. RETURNED PRODUCT EVALUATION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16871559&pc=QCB | 2023826-2023-01595 | 2023-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50/+2.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BENT. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16945974&pc=QCB | 2023826-2023-01839 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: A4-A6: UNK. B3: UNK. D4 (EXPERIATION DATE); UNK. NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.00/2.0/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "ASYMPTOMATIC" AND "IN FUTURE VISITS WE WILL SEE THE EVOLUTION." H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - BLURRED VISION WAS ALSO NOTED. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950559&pc=QCB | 2023826-2023-01857 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+5.0/168 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019319&pc=QCB | 2023826-2023-01948 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H5- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. ADDITIONAL DATA: B5- REPORTEDLY "EXCHANGE IS PLANNED END OF OCTOBER." CLAIM# (B)(4). Manufacturer Narrative: ADDITIOAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". D6B.- (B)(6) 2023 SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE:"2993" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056313&pc=QCB | 2023826-2023-02124 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.0/4.0/24 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS REPOSITIONING WAS PERFORMED AFTER A LENS ROTATION OCCURED. THE LENS ROTATED AGAIN. REPORTEDLY, "THIS LENS ICL SUFFERED A ROTATION OF APPROXIMATELY 16 DEGREES. MORE INFORM THAT THE LENS IN QUESTION WAS ALREADY REPOSITIONED.". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. ADDITIONAL DATA: B3- "(B)(6) 2023" SHOULD BE ADDED. B7- "PROGRESSIVE MYOPIA"SHOULD BE ADDED D6B. "(B)(6) 2023" SHOULD BE ADDED H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+4.0/024 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS ROTATION AND REPOSITIONING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. REPORTEDLY "THIS LENS ICL SUFFERED A ROTATION OF APPROXIMATELY 16 DEGREES. MORE INFORM THAT THE LENS IN QUESTION WAS ALREADY REPOSITIONED IN THE AXIS INDICATED IN THE DIAGRAM AND POSITIONED IN THE SAME PLACE". ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE DOCTOR WANT TO IMPLANT ANOTHER ONE TO SOLVE THE PROBLEM". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FIBRE ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119692&pc=QCB | 2023826-2023-02161 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -9.0/4.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS WAS DIRECTLY ON THE ANTERIOR LENS SURFACE. Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119698&pc=QCB | 2023826-2023-02168 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TRANSLENT LOSS OF BCVA AND LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: 2137 ADDED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119699&pc=QCB | 2023826-2023-02167 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.0/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. TRANSLENT LOSS OF BCVA AND LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: 2137 ADDED FOR CORRECTION. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17145853&pc=QCB | 2023826-2023-02393 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.50/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES A POSSIBLE LENS EXCHANGE DUE TO LENS ROTATION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE ON (B)(6) 2023. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17151996&pc=QCB | 2023826-2023-02323 | 2023-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AMD RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) /2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17167504&pc=QCB | 2023826-2023-02462 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.0/3.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. IT WAS REPORTED THAT THIS LENS WAS INCORRECTLY IMPLANTED INTO THE PATIENT'S LEFT EYE, AS IT WAS INTENDED FOR THE PATIENT'S RIGHT EYE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A SAME MODEL/LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213812&pc=QCB | 2023826-2023-02508 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | UNK | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -5.0/4.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO REFRACTIVE SURPRISE. REPOSITIONING DID NOT RESOLVE THE ISSUE. CAUSE OF THE EVENT IS UNKNOWN. H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17385337&pc=QCB | 2023826-2023-03155 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE WAS LENS ROTATION, THE LENS HAD ROTATED TO 161 DEGREES, 16 DEGREES OFF THE DESIRED LOCATION OF 177. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE WAS LENS ROTATION, THE LENS HAD ROTATED TO 161 DEGREES, 16 DEGREES OFF THE DESIRED LOCATION OF 177 AND REFRACTIVE CHANGE OVER TIME WAS OBSERVED. LENS REMAINS IMPLANTED. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669070&pc=QCB | 2023826-2023-03736 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICH12.6, -4.5/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION INFERIOR WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS ROTATED. THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, LENS TOO SMALL, KEEPS ROTATING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE ADDED. D6B. "ON (B)(6) 2023" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER +4.50/5.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. PRODUCT NONCONFORMITY WAS REPORTED. REPORTEDLY "THE ICL ROTATED INFERIORLY TWICE NOW" AND "ICL KEEPS ON ROTATING AND PREFERS TO SETTLE VERTICALLY". THE PATIENT EXPERIENCED BLURRED VISION AND INFERIOR ROTATION OF ICL X2. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. REPORTEDLY " ICL HAS ROTATED TWICE NOW. ONCE CLOCKWISE AND THE SECOND TIME COUNTER CLOCKWISE. ANTERIOR SEGMENT SCANS WILL BE PROVIDED". ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "THE ICL IS TOO SMALL AND HAS NOW ROTATED INFERIORLY TWICE. IT WAS ROTATED BACK ONCE BUT HAS AGAIN ROTATED INFERIORLY". CLAIM#: (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17927921&pc=QCB | 2023826-2023-04478 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -14.50/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO THE LENS BEING TOO SMALL. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, THE LENS ROTATED 90 DEGREES WITH THE SIZE OF THE LENS BEING PROBABLE ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT MUCH HAPPIER WITH IMPROVED VISION. ICL IN PROPER POSITION." CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17927922&pc=QCB | 2023826-2023-04479 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO THE LENS BEING TOO SMALL. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, THE LENS ROTATED 90 DEGREES WITH THE SIZE OF THE LENS BEING PROBABLE ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT MUCH HAPPIER WITH IMPROVED VISION. ICL IN PROPER POSITION." CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088953&pc=QCB | 2023826-2023-04764 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18275560&pc=QCB | 2023826-2023-05475 | NA | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.0/1.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT AND REFRACTIVE MISS WERE REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT: CLAIM# (B)(4) (FELLOW EYE). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18296805&pc=QCB | 2023826-2023-05487 | 2023-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.5/3.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18296806&pc=QCB | 2023826-2023-05486 | 2023-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/50 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE OBSERVED. THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4803722&pc=QCB | 2023826-2015-00169 | 2015-01-16 | Injury | STAAR SURGICAL COMPANY | 2015-05-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.0/+2.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2015 WITH A SAME MODEL AND SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH, DIOPTER POWER AND SPHERE WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYES MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). MDR RESUBMITTED PER FDA REQUEST. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794355&pc=QCB | 2023826-2018-01193 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS ROTATED NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT KEPT ROTATING AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: UPDATED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER ORIGINALLY NOTED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT THE LENS KEPT ROTATING. THE PROBLEM WAS NOT RESOLVED. ON 21-JAN-2019, ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE LENS ROTATED SEVERAL TIMES DUE TO LOW VAULT (100 MICRONS). IT WAS ALSO LATER REPORTED THAT THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY. ADDITIONAL INFORMATION RECEIVED ON 22-SEP-2019 STATED THAT THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS LATER REPLACED ON (B)(6) 2019 FOR A 13.7MM VTIMC5_13.7 LENS OF DIOPTER -10.0/4.5/082 (SPHERE/CYLINDER/AXIS) AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS A PATIENT-RELATED FACTOR. EXPLANT DATE UPDATED TO (B)(6) 2018 PER NEW INFORMATION RECEIVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY. CORRECTED DATA: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.50/+5.5/068 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER ORIGINALLY NOTED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES BUT THE LENS KEPT ROTATING. THE PROBLEM WAS NOT RESOLVED. ON 21-JAN-2019, ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE LENS ROTATED SEVERAL TIMES DUE TO LOW VAULT (100 MICRONS). THE REPORTER STATED THAT THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8067158&pc=QCB | 2023826-2018-01672 | 2018-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -2.5/5.0/41 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT REPOSITIONING OF THE LENS IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -2.5/5.0/41 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2018 REPOSITIONING OF THE LENS WAS PERFORMED. IT WAS REPORTED THAT ALL IS FINE AND THE PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REPOSITIONING. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182336&pc=QCB | 2023826-2018-01886 | 2018-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAD DECIDED NOT TO HAVE A LVC PROCEDURE TO CORRECT CURRENT REFRACTIVE ERROR AFTER TICL IMPLANTATION. THE PLAN IS TO EXPLANT THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5-THE REPORTER INDICATED THE "LENS WAS NOT STABLE" ON (B)(6) 2019. ON (B)(6) 2019 THE REPORTER STATED, "THE SURGEON TRIED TO CORRECT THE REFRACTIVE SURPRISE WITH A REPOSITIONING OF BOTH LENSES. FOR THIS HE ROTATED THE MISALIGNED TICLS TO HE AXIS OF THE IODS. AFTER THAT (1DAY POSTOP) FIRST THE REFRACTIONS WERE GOOD AND EVERYTHING WAS FINE. AT THE NEXT EXAMINATION MID OF JANUARY THE TICLS HAVE ROTATED AGAIN TO THE SAME MISALIGNED AXIS AS BEFORE. IN THIS POSITION THE LENSES WERE NOW STABLE AND SO THE SURGEON DECIDED TO MAKE A LASER TOUCH UP TO CORRECT THE RESIDUAL REFRACTION ON BOTH EYES". PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE: "2923" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2019. BECAUSE OF STRONG BLEEDING AFTER REMOVING THE LENS OUT OF THE SULCUS, THE SURGEON DECIDED TO WAIT TO IMPLANT THE EXCHANGE LENS FOR A COUPLE OF MONTHS. THE PATIENT AND EYE ARE FINE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, STRONG BLEEDING CORRECTED DATA: D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298862&pc=QCB | 2023826-2019-00192 | 2018-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.00/+1.5/038 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2019 AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307775&pc=QCB | 2023826-2019-00228 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -0.50/6.0/052 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT AND REFRACTIVE SURPRISE HAS BEEN REPORTED, IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437900&pc=QCB | 2023826-2019-00452 | NA | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE, BLURRED VISION, AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437914&pc=QCB | 2023826-2019-00456 | NA | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, BLURRED VISION, AND LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR -THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. H6-METHOD CODE ADDED CLAIM #(B)(4). Manufacturer Narrative: THE STATEMENT THAT THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN IS INCORRECT. PLEASE DISREGARD. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438505&pc=QCB | 2023826-2019-00435 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.00/+3.5/023 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE UNCORRECTED VISUAL ACUITY (UCVA) WAS LOWER THAN EXPECTED. THE CLINICAL REFRACTION AND UCVA DID NOT MATCH THE EXPECTED REFRACTION OF TICL CALCULATION. THE PATIENT AND SURGEONS ARE EXTREMELY DISSATISFIED WITH THE LOW UCVA THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8442017&pc=QCB | 2023826-2019-00439 | 2018-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, --11.5/+3.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8451630&pc=QCB | 2023826-2019-00504 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH12.6, -0.50/2.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LEN INTO PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAIN IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION: "THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND BLURRED VISION". DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTION INSPECTION FOUND THE LENS SPHERE AND CYLINDER TO BE WITHIN SPECIFICATIONS BUT FOUND THE LENS AXIS TO BE 3° OUT OF SPECFICIATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DATE OF EVENT IS CORRECTED TO (B)(6)2019 EVENT - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS OF -0.50/2.0/170 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019 . ON (B)(6)2019 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE REPORTER INDICATES THE CAUSE OF THE EVENT IS DEVICE. PATIENT CODE (B)(4). CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM #(B)(4). Manufacturer Narrative: "VTICH12.6" SHOULD BE CORRECTED TO "VTICMO12.6" IN ORIGINAL MDR AND MDR SUPPLEMENTAL #1. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452986&pc=QCB | 2023826-2019-00488 | 2017-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.5/+0.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE, THE PATIENT STATING HE HAS "BAD SIGHT". REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: EVENT DESCRIPTION: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. THE LENS WAS EXCHANGED AND PATIENT IS "VERY SATISFIED." SHE SEES 1.2 UNCORRECTED WITH REF. +0.25. EXPLANT DATE: UNK. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8466158&pc=QCB | 2023826-2019-00512 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.50/+1.0/028 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8466179&pc=QCB | 2023826-2019-00513 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+4.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -4.5/4.0/12 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2019 THE LENS WAS EXCHANGED INTRAOPERATIVELY AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS WAS EXCHANGED ON (B)(6) 2019 WITHIN SAME SURGERY AS REMOVAL OF INITIAL LENS. THE EXCHANGE RESOLVED THE PROBLEM. EXPLANTED DATE IN SUPPLEMENTAL MDR (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: THE STATEMENT " THE LENS REMAINS IMPLANTED" IN ORIGINAL MDR IS NOT APPLICABLE. THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. DEVICE CODE 4117: IS NOT APPLICABLE IN ORIGINAL MDR; LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485090&pc=QCB | 2023826-2019-00552 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (CALCULATION ERROR). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EVENT: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019. AT A FOLLOW-UP VISIT, THE PATIENT'S VISION HAD IMPROVED AND PATIENT WAS VERY SATISFIED. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485259&pc=QCB | 2023826-2019-00538 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VTICM5_13.2, -5.50/1.0/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. A CHANGE IN REFRACTION WAS OBSERVED, THE LENS REMAINS IMPLANTED. THE REPORTER STATED " THE CORNEA IS SIMPLY THIN, THERE ARE NO KC FFKC, ETC. DIAGNOSIS. THE SURGEON THOUGHT THE REFRACTIVE SURPRISE +1,5/-1,75/176 (B)(6) 2018) POST OPERATIVELY CAUSED BY THE 11 DEGREE MISALIGNMENT OF THE LENS. THEN HE REPOSITIONED THE LENS END OF NOVEMBER AND THOUGHT EVERYTHING WOULD BE GOOD. THE CURRENT REFRACTION OF +1,00/-1,75/172 WITH A CORRECT POSITIONED TICL LIE IOD OF 2° CCW HE CANNOT EXPLAIN AND DO NOT HAVE ANY IDEA ABOUT IT. HE'S ASKING TO RECALCULATE A NEW LENS BASED ON THE CURRENT RX AND TICL POSITION". PATIENT IS WELL AND THE SURGEON PLANS TO REPOSITION THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: SHOULD BE CORRECTED TO "NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT." CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492262&pc=QCB | 2023826-2019-00580 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANT DATE: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. PATIENT CODE 2137 SHOULD BE ADDED TO INITIAL MDR. DEVICE CODE 1401 SHOULD BE ADDED TO INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552784&pc=QCB | 2023826-2019-00666 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, TMICL12.1, -9.0/3.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINION THE CAUSE OF THE EVEN WAS DUE TO LENS BEING TOO SMALL. REPORTEDLY, REPOSITIONING WAS NOT PERFORMED. HOWEVER, IT WAS REPORTED THAT THE REPOSITIONING RESOLVED THE PROBLEM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE: 2137 - REFRACTIVE CHANGE OVER TIME. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED DATA: IF FOLLOW-UP, WHAT TYPE: "ADDITIONAL INFORMATION" IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "CORRECTION" IN PREVIOUS MDR. EVENT: "ON 04 APR 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND REFRACTIVE CHANGE OVER TIME." CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS NOT REPOSITIONED. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. THE VIAL WAS RETURNED TAPED TO PAPER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8579900&pc=QCB | 2023826-2019-00760 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND HAS RESIDUAL ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REFRACTIVE SURPRISE" IN INITIAL MDR NOT APPLICABLE. H6 - PATIENT CODE: 3191 - RESIDUAL ASTIGMATISM. H6 - PATIENT CODE: 2137 - PREVIOUS CODE NOT APPLICABLE. H6 - DEVICE CODE: 1401 - PREVIOUS CODE NOT APPLICABLE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611957&pc=QCB | 2023826-2019-00783 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS REMAINS IMPLANTED. REPORTEDLY "PATIENT IS FINE AND WAITING FOR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADVERSE EVENT: PRODUCT PROBLEM IN ORIGINAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2019, PATIENT STABLE AFTER REPOSITIOING. PER REPORTER ON (B)(6) 2019, "PATIENT IS HAPPY AFTER REPOSITIONING". PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621876&pc=QCB | 2023826-2019-00803 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE DOCTOR SUSPECTED THE REFRACTION OF THE FIRST LENS WAS INCORRECT. Manufacturer Narrative: RESULT CODE: 213, SHOULD BE CORRECTED TO RESULT CODE: 170. CONCLUSION CODE: 4315, SHOULD BE CORRECTED TO CONCLUSION CODE: 25. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH WAS IN SPECIFICATION BUT LENGTH AND REFRACTION (FUNCTIONAL) VERIFICATION FAILED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640314&pc=QCB | 2023826-2019-00905 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND ASTIGMATISM OVERCORRECTION. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/ VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8640353&pc=QCB | 2023826-2019-00910 | 2019-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 6.5/+3.0/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE.THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019 AND WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652062&pc=QCB | 2023826-2019-00914 | 2019-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+4.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. PATIENT CODE: 3191- SECONDARY SURGERY INTERVENTION, EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: "PROBLEM PRODUCT" SHOULD BE CORRECTED TO "ADVERSE EVENT" IN THE INITIAL MDR. "06-MAY-2019" SHOULD BE CORRECTED TO "02-MAY-2019" IN THE INTIAL MDR. - "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS INJURY" IN THE INITIAL MDR. ADDITIONAL INFORMATION: "REQUIRED INTERVENTION" SHOULD BE ADDED IN THE INTIAL MDR. "LENS ROTATION" SHOULD BE ADDED IN THE INITIAL MDR. DEVICE CODE: 2923- DEVICE DISLODGED/DISLOCATED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652063&pc=QCB | 2023826-2019-00913 | 2019-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. PATIENT CODE: 3191- SECONDARY SURGICAL INTERVENTION, EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652121&pc=QCB | 2023826-2019-00912 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON REPORTS REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, LENS REMAINS IMPLANTED. THE SURGEON STATES THE PATIENT'S VISION DECREASED GRADUALLY WITHIN THE LAST 2 MONTHS; NO PAIN, NO HISTORY OF EYE INJURY. Manufacturer Narrative: CORRECTED DATA: ALL PREVIOUS MDRS SHOULD BE CORRECTED TO SERIOUS DUE TO SECONDARY SURGERY FOR LENS REPOSITIONING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADD "LENS ROTATION WAS REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. DEVICE CODE 2923: DISLODGED OR DISLOCATED; LENS ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652191&pc=QCB | 2023826-2019-00902 | 2019-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, - 17.0/+1.0/083 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS AN ADDITION OF SUTURE. THE EYE IS QUIET AND VISION IS GOOD. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE LENS WIDTH WAS IN SPECIFICATION BUT LENGTH AND REFRACTION (FUNCTIONAL) VERIFICATION FAILED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION THE PROBABLE CAUSE OF THIS EVENT IS RELATED TO MANUFACTURING AND PROCESSING OF THIS LENS. CORRECTIVE ACTIONS ARE BEING DEVELOPED. RESULT CODE: 213 & 114 - SHOULD BE CORRECTED TO "RESULT CODE: 170." CONCLUSION CODE: 4315 - SHOULD BE CORRECTED TO "CONCLUSION CODE: 25." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8662649&pc=QCB | 2023826-2019-00865 | 2019-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA AND REFRACTIVE SURPRISE WAS OBSERVED. UPON FOLLOW-UP ON (B)(6) 2019, THE LENS WAS ROTATED/DISPLACED BY 25 DEGREES NOT RESOLVING THE PROBLEM. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "SMALL SIZE LENS." AS OF PATIENT'S LAST VISIT ON (B)(6) 2019 "LENS STABLE AFTER EXCHANGE AND VAULT INCREASED TO 500 MICRON." Manufacturer Narrative: DEVICE CODE 1494: OFF LABEL USE: ACD<3.0MM. "UPON FOLLOW-UP ON 06/MAY? 2019" SHOULD BE CORRECTED TO "UPON FOLLOW-UP ON 06/MAY/2019" CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671657&pc=QCB | 2023826-2019-01018 | 2019-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.5/+4.5/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS REFRACTIVE SURPRISE, WHICH WAS ATTEMPTED TO BE CORRECTED ON (B)(6) 2018 VIA REFRACTIVE TREATMENT ("LASER TOUCH UP"). THIS WAS NOT SUCCESSFUL AS CURRENTLY THERE IS AGAIN A REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: "REPORTEDLY, THE LENS REMAINS IMPLANTED." SHOULD BE CORRECTED TO "ON 08/08/2019 THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS. THE EXCHANGE RESOLVED THE PROBLEM." IN THE INITIAL MDR. PATIENT CODE 3191: SECONDARY SURGERY, EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: RESULT CODE 213 SHOULD BE ADDED TO THE PREVIOUS MDR. CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671659&pc=QCB | 2023826-2019-01015 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/103 (SPHERE/CYLINDAR/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE REFRACTIVE SURPRISE. THE REPORTER STATES "LENS WAS REMOVED AND WE'LL WAIT FOR ENSURE STABILITY." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8686760&pc=QCB | 2023826-2019-00869 | 2017-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692896&pc=QCB | 2023826-2019-01063 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.0/+5.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. SEE MFR# 2023826-2019-00057 FOR PREVIOUSLY IMPLANTED LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE . THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 OF -11.0/5.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS AN EXCHANGE LENS. THE SURGEON NOTES LOW VAULT; LENS ROTATION; AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED AGAIN ON (B)(6) 2019 FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTER STATED; "PATIENT IS HAPPY." EXPLANT DATE: (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710343&pc=QCB | 2023826-2019-01099 | 2018-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8710345&pc=QCB | 2023826-2019-01100 | 2018-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2018 A LONGER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8712615&pc=QCB | 2023826-2019-01096 | NA | Injury | STAAR SURGICAL COMPANY | 2019-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.50/2.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. IT WAS REPORTED THAT ON (B)(6) 2018 THE PATIENT NOTICED A SHARP REDUCTION OF VISUAL ACUITY, THE LENS HAD ROTATED 90 DEGREES. IN (B)(6) 2018 THE LENS WAS REPOSITIONED AND THE PATIENT RECOVERED VISION OF 20/20 WITHOUT GLASSES. A CONSULT ONE YEAR LATER (B)(6) 2019) AGAIN NOTICED DETERIORATION OF VISION AND THE LENS ROTATED AGAIN. THE SURGEON PLANS TO REMOVE AND REPLACE THE LENS WITH A LONGER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8741157&pc=QCB | 2023826-2019-01132 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; Material Split, Cut or Torn; Physical Resistance/Sticking | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO13.7, -15.50/4.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS "CUT AT THE END OF THE INSERTION" INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER INDICATED THE CAUSE OF THE CUT WAS DUE TO THE LENS BEING "STUCK TO FOAM TIP." THE LENS REMAINS IMPLANTED, "AS PART OF IT WAS CUT HENCE THE DIAMETER OF LENS WAS SMALLER AND THE LENS DISLOCATED." REPORTER INDICATED THE "ICL IS IN A VERTICAL POSITION" AND PATIENT'S "VISION IS POOR DUE TO ASTIGMATISM." THE REPORTER INDICATED NO PATIENT INJURY AND PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: "THE LENS REMAINS IMPLANTED" SHOULD BE CORRECTED TO "THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019" IN THE INITIAL MDR. "PLANS TO EXCHANGE THE LENS" SHOULD BE REMOVED IN THE INITIAL MDR. PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION, EXPLANT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8767927&pc=QCB | 2023826-2019-01204 | 2019-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.50/3.50/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, AT PATIENTS LAST VISIT, BLURRED VISION WAS REPORTED AS STATUS OF THE EYE. NO TREATRMENT OR INTERVENTION HAS BEEN PERFORMED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (PIGMENT DISPERSION, ANISOCORIA, LENS REPOSITIONING). LENS REPOSITIONING WAS PERFORMED APPROXIMATELY FIVE WEEKS AFTER IMPLANTATION, POST OP STATUS WAS RECORDED AS " OS ANISOCORIA LIKELY DUE TO PIGMENT DISPERSION/ MILD SPHINCTER DAMAGE". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774160&pc=QCB | 2023826-2019-01193 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/050 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: EVENT UPDATED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/050 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTED LOW VAULT AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2019FOR A LONGER LENGTH LENS, VTICM5 13.2MM -13.0/2.5/043 (SPHERE/CYLINDER/AXIS), AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTED TO BE "HAPPY"." (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774161&pc=QCB | 2023826-2019-01191 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. REPORTEDLY, THE DOCTOR CHOSE THE INITIAL LENS TO TARGET LOWER CORRECTION, BUT THE ACTUAL CORRECTION WAS EVEN LOWER THAN WHAT HE'D EXPECTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: HE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774162&pc=QCB | 2023826-2019-01194 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTED LOW VAULT AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS, VTICMO 13.2MM -15.0/3.5/138 (SPHERE/CYLINDER/AXIS), AND THE PROBLEM WAS RESOLVED. THE PATIENT IS REPORTED TO BE "HAPPY"." PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774166&pc=QCB | 2023826-2019-01192 | 2019-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, SIMILAR SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: ACD < 3.0MM NOT APPLICABLE IN INITIAL MDR (JAPAN APPROVED LENS USE FOR ACD > 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780383&pc=QCB | 2023826-2019-01181 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/3.0/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PANNS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. SURGEON NOTES THAT PATIENT'S "REFRACTION CHANGES AFTER THE ROTATION HAPPENED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019. REPORTEDLY A LONGER LENGTH LENS WAS IMPLANTED AND 'THE PATIENT IS DOING VERY GOOD.' PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION, LENS EXPLANT). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780390&pc=QCB | 2023826-2019-01182 | 2019-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.50/3.0/075 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION AND LOW VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY THE SURGEON PANNS TO EXCHANGE THE LENS WITH A LONGER LENGTH LENS. SURGEON NOTES THAT PATIENT'S "REFRACTION CHANGES AFTER THE ROTATION HAPPENED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT CODE 2137 NOT APPLICABLE IN INITIAL MDR. DEVICE CODE 1401 NOT APPLICABLE IN SUPPLEMENTAL MDR #1. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - THE LENS WAS EXPLANTED ON (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8790563&pc=QCB | 2023826-2019-01149 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VTICM5_12.6, -6.5/1.0/53 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. THE DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809463&pc=QCB | 2023826-2019-01260 | 2019-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826240&pc=QCB | 2023826-2019-01295 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM RESOLVED. POSTOPERATIVE REPORT STATED "PATIENT IS HAPPY" H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826253&pc=QCB | 2023826-2019-01215 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: 'A REPOSITIONING BY 0 DEGREES WAS ALSO REPORTED POST-OPERATIVELY - THIS IS NOT CLEAR.' PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK; ETHNICITY- UNK; RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851480&pc=QCB | 2023826-2019-01292 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/3.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871965&pc=QCB | 2023826-2019-01394 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/063 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED 'THE SURGEON REPORTED A LENS ROTATION.' (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS ROTATED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL WITH 6/6 VISION AND IS HAPPY. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED (ROTATED). CLAIM#: (B)(4). Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876530&pc=QCB | 2023826-2019-01401 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2019-08-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE DUE TO THE WRONG LENS (LENS CALCULATED FOR LEFT EYE(OS)) WAS IMPLANTED IN THE RIGHT EYE. THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ALL IS FINE AND THE PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889723&pc=QCB | 2023826-2019-01363 | 2018-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE CHANGE OVER TIME AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTER INDICATED THAT "REMAINING ASTIGMATISM WILL BE TREATED WITH LASER LATER." Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899033&pc=QCB | 2023826-2019-01421 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 12.50/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN A BIOHAZARD BAG SHOULD ALSO HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE CODE: 1401 SHOULD HAVE ALSO BEEN INCLUDED IN INITIAL MEDWATCH REPORT. PATIENT'S AGE: SHOULD BE CORRECTED TO (B)(6) 1984 IN INITIAL MEDWATCH REPORT.(B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941312&pc=QCB | 2023826-2019-01584 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND LOSS OF BCVA. REPORTEDLY, THE LENS ROTATED AFTER ONE WEEK, WAS REPOSITIONED, AND THEN ROTATED AGAIN AFTER ONE WEEK. THE LENS CURRENTLY REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. ADD TO PREVIOUSLY REPORTED PATIENT CODE 3191: (LENS EXCHANGE). DEVICE CODE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962475&pc=QCB | 2023826-2019-01590 | 2019-08-03 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+4.50/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS WAS REPOSITIONED ON (B)(6) 2019. THE PATIENTS VISION WAS IMPROVED TO 20/20 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966506&pc=QCB | 2023826-2019-01631 | 2019-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A LONGER AND SIMILAR POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL BUT DIFFERENT POWER LENS. DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966515&pc=QCB | 2023826-2019-01644 | 2018-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+3.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: RESULT CODE 213. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984516&pc=QCB | 2023826-2019-01674 | 2019-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.0/+2.0/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. EVENT: PER UPDATED INFORMATION, THE LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. METHOD CODE 4117 NO LONGER APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984800&pc=QCB | 2023826-2019-01653 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TMICL12.1 LENS, -5.00/+2.5/063 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT ,LENS ROTATION AND BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT CODE:3191- SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT, OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) THE INITIAL LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. REFERENCE CLAIM#: (B)(4).FOR REPLACEMENT LENS. EXPLANT DATE: (B)(6) 2019. DEVICE EVALUATION: LENS RETURNED IN LIQUID IN EYE DROP BOTTLE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8984823&pc=QCB | 2023826-2019-01652 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM TMICL12.1 LENS, -6.5/+4.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT ,LENS ROTATION AND BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE INITIAL LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019. THE REPLACEMENT LENS DID NOT RESOLVE THE PROBLEM. REFERENCE CLAIM# 430733 FOR REPLACEMENT LENS. EXPLANT DATE: (B)(6) 2019. LENS RETURNED IN LIQUID IN A EYE DROP BOTTLE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS TYPE OF REPORTABLE EVENT: SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE:3191- SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. DEVICE CODE:1401- BLURRED VISON. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9010339&pc=QCB | 2023826-2019-01656 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WITH LOSS OF BCVA AND LENS ROTATION WAS REPORTED. LENS WAS REPOSITIONED ON (B)(6) 2019 BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS SMALL LENS SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND AN OPTIC TEAR, THE HAPTIC TORN, AND DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE -UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9072751&pc=QCB | 2023826-2019-01724 | NA | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.50/1.5/082( SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE REPORTER STATED "PATIENT IS RESTLESS AND WORRIED ABOUT NOT SEEING SATISFACTORILY." LENS REMAINS IMPLANTED. SURGEON BELIEVES THE LENS WAS BADLY LABELED AT THE FACTORY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A VTICM5_12.6, -10.50/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE REPORTER STATED, "PATIENT IS RESTLESS AND WORRIED ABOUT NOT SEEING SATISFACTORILY". ON (B)(6) 2019 THE LENS WAS EXPLANTED, AND A REPLACEMENT LENS WAS IMPLANTED ON THE SAME DATE. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. DOCTOR IS SATISFIED WITH THE NEW RESULT. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT ON (B)(6) 2019 A NEW LENS WAS IMPLANTED AND THE PATIENTS CONDITION IS SATISFACTORILY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK ,DATES OF REPORT: UNK, UTENSILS: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9100740&pc=QCB | 2023826-2019-01734 | 2019-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: TTHE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/013 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115524&pc=QCB | 2023826-2019-01731 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SAME LENGTH , -10.00/2.0/091 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115526&pc=QCB | 2023826-2019-01732 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SAME LENGTH , -10.50/3.0/088 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175742&pc=QCB | 2023826-2019-01897 | 2019-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -6.00/1.5/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO A CALCULATION MISTAKE BY THE SURGEON AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE POST OPERATIVE DATA ARE ALL FINE BUT PATIENT IS NOT 100% HAPPY, THE REPORTER STATED " MAYBE PSYCHOLOGICAL REASONS". Manufacturer Narrative: THE LENS EXCHANGE DATE (B)(6) 2019 IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE OF EVENT IS INITIAL MDR IS CORRECTED TO (B)(6) 2019. DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 10-SEPT-2019. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188695&pc=QCB | 2023826-2019-01950 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE EVENT DESCRIPTION PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT POWER LENS. THE PROBLEM IS RESOLVED. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC BROKEN AND LENS STICKY/GUMMY. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: -UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199834&pc=QCB | 2023826-2019-01922 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+2.5/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO INCORRECT INPUT DATA. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND THE HAPTIC TORN AND BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL INSPECTION DUE TO DAMAGE OPTIC LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246796&pc=QCB | 2023826-2019-02028 | 2019-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+4.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, LOSS OF BCVA, AND PARTIAL PUPIL DILATION. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "EXCESSIVE SIZE." Manufacturer Narrative: REQUIRED SURGICAL INTERVENTION. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/4.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; LOSS OF BCVA AND PARTIAL PUPIL DILATION. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTER INDICATED THAT THE "AV IS OK." PATIENT CODE:3191 SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276012&pc=QCB | 2023826-2019-02076 | 2019-01-11 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS HAD ROTATED AND WAS RE-POSITIONED ON (B)(6) 2019. THE LENS AGAIN ROTATED ON (B)(6) 2019 AND WAS RE-POSITIONED. TO DATE THE LENS REMAINS IMPLANTED AND IS IN A STABLE POSITION BUT THERE IS STILL A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276004&pc=QCB | 2023826-2019-02073 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279650&pc=QCB | 2023826-2019-02060 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS REPOSITIONED. THE PROBLEM IS RESOLVED. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279651&pc=QCB | 2023826-2019-02064 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/016 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: IT WAS REPORTED THAT ON(B)(6) 2020 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND SIZE BUT DIFFERENT POWER LENS AND THAT THE PATIENT WAS EXTREMELY HAPPY WITH 6/6 VISION. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279666&pc=QCB | 2023826-2019-02061 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019 THE LENS WAS REPOSITIONED AND THE PROBLEM IS RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306655&pc=QCB | 2023826-2019-02095 | 2019-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -10.50/2.00/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED DRY, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND A BENT HAPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION WAS PERFORMED AND LENS WAS FOUND TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: RESULT CODE 213 SHOULD BE IN THE PREVIOUS MDR. H6- CONCLUSION CODE 4310 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(6). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336756&pc=QCB | 2023826-2019-02176 | 2019-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, BLURRED VISION, RESIDUAL CYLINER, A "LINE" IN THE VISUAL FIELD, AND A WHITE LIGHT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE: 1401 SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9344421&pc=QCB | 2023826-2019-02185 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.50/1.00/100 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE(OD) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE ARE ASKING FOR AN ICL EXCHANGE BECAUSE OF TOUCHDOWN AND MYOPIZATION". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: SEX:UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9349221&pc=QCB | 2023826-2019-02186 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.50/0.5/90 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON (B)(6) 2015. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED " WE ARE ASKING FOR AN ICL EXCHANGE BECAUSE OF TOUCHDOWN AND MYOPIZATION". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON 28/JAN/2020 WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: SEX:UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364603&pc=QCB | 2023826-2019-02211 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.50/3.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 12/JUL/2019. REFRACTIVE SURPRISE WAS OBSERVED ON (B)(6) 2019, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "A REPOSITIONING IS PLANNED, PATIENT'S CURRENT CONDITION 20/32 AND BCVA IS 20/20." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE). B5- THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND THE REPORTER STATES "PATIENT IS HAPPY AFTER REPOSITIONING." H6-PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. THE REPOSITIONING WAS PERFORMED ON (B)(6) 2020, THE REPOSITIONING SOLVED THE PROBLEM AND PATIENT IS HAPPY. DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9406617&pc=QCB | 2023826-2019-02302 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS NOTED THAT THE AXIS USED IN CALCULATION WAS NOT ALIGNED WITH REFRACTION. ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED ACHIEVING GOOD OUTCOMES. CAUSE OF THIS EVENT WAS REPORTED AS A USER ERROR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED DUE TO USER ERROR (AXIS USED IN CALCULATION WAS NOT ALIGNED WITH REFRACTION). ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED AND THE PROBLEM WAS RESOLVED. PATIENT IS IN GOOD CONDITION AND HAS NORMAL VISION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIAL MDR. B4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIALL MDR. G4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE INITIALL MDR. G4- (B)(6) 2019" SHOULD BE CORRECTED TO (B)(6) 2019" IN THE PREVIOUS MDR. B5- "ROTATION WAS ALSO OBSERVED." SHOULD BE ADDED TO THE PREVIOUS MDR. H6- DEVICE CODE 2993 IS NOT APPICABLE CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417044&pc=QCB | 2023826-2019-02209 | 2019-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9461812&pc=QCB | 2023826-2019-02380 | 2019-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH THE SAME MODEL/LENGTH LENS BUT DIFFERENT DIOPTER AND THE PROBLEM RESOLVED, "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476277&pc=QCB | 2023826-2019-02392 | 2019-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/+3.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE REPORTER STATES THE SURGEON POSITIONED THE LENS IN THE WRONG AXIS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2019 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #:(B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512599&pc=QCB | 2023826-2019-02440 | 2015-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATINET'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2020 THE LENS WAS EXCHNAGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- ADVERSE EVENT SHOULD BE MARKED IN MDR SUPPLEMENTAL #1. H10- "WITH MOISTURE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512630&pc=QCB | 2023826-2019-02441 | 2015-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.50/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATINET'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2020 THE LENS WAS EXCHNAGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516211&pc=QCB | 2023826-2019-02483 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -6.5/3/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- WORK ORDER SEARCH: "NO" SHOULD BE CORRECTED TO "ONE" IN THE INITIAL MDR. H6- RESULT CODE 213 SHOULD BE CORRECTED TO 114 IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516575&pc=QCB | 2023826-2019-02451 | 2019-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -3.50/4.5/168 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020. D7: EXPLANT DATE:(B)(6) 2020. H6: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LEN WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4 IS NOT APPLICABLE IN MDR SUPP1 D7 SHOULD BE 19-JAN-2020 IN MDR SUPP1 H6- INVESTIGATING FINDINGS CODE 213 SHOULD BE IN THE PREVIOUS MDR H6- INVESTIGATING CONCLUSION SODE 4310 SHOULD BE IN THE PREVIOUS MDR CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532308&pc=QCB | 2023826-2019-02470 | 2019-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -1.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION-LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND STICKY SUBSTANCE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE) B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 ; -1.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON 24-SEP-2019. ON 17-MAR-2020 THE LENS WAS EXCHANED WITH ASAME MODEL AND LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES PATIENT IS HAPPY. CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. D7- EXPLANT DATE: (B)(6)-2020 H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532323&pc=QCB | 2023826-2019-02457 | 2016-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532331&pc=QCB | 2023826-2019-02456 | 2016-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON 18/03/2020 THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT MODEL LENS AND THE PROBLEM RESOLVED. PATIENT WAS REPORTED TO BE HAPPY. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538844&pc=QCB | 2023826-2020-00003 | 2019-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/1.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. REPORTEDLY, "DUE TO DATA INPUT ERROR DURING CALCULATIONS, THE DEGREE WAS NOT SUITABLE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2137 SHOULD BE ADDED TO THE INITIAL MDR. H6- DEVICE CODE 2923 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO CODE 1401 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541982&pc=QCB | 2023826-2020-00024 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AFTER MISALIGMENT WAS OBSERVED. IT WAS REPORTED THAT THE REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2019. PATIENTS CURRENT UCVA : 20/25, BCVA 20/20. LENS REMAINS IMPLANTED AND LENS REPOSITIONING IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS REPOSITIONED ON (B)(6) 2019. THE REPORTER INDICATED THE REPOSITIONING WENT WELL." H6- PATIENT CODE 3191- SECONDARY SURGICAL INTERVENTION- LENS REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT" SHOULD BE MARKED IN THE PREVIOUS MDR. B3- "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE PREVIOUS MDR. D11- "INJECTOR SYSTEM MODEL: FTP, LOT #: UNK" SHOULD BE ADDED IN THE INITIAL MDR. H1- "SERIOUS INJURY" SHOULD BE MARKED IN THE PREVIOUS MDR. H6- DEVICE CODE 2923 SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9604383&pc=QCB | 2023826-2020-00166 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 14.00/+1.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO SURGEON WAS NOT HAPPY WITH LENS POSITION. ANOTHER SAME MODEL LENS WAS IMPLANTED AS THE EXCHANGED LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE LENS EXCHANGE RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTICS TORN. THE LENS WAS RETURNED IN THREE PIECES. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9663126&pc=QCB | 2023826-2020-00221 | 2019-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+1.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721916&pc=QCB | 2023826-2020-00313 | 2019-11-22 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.0/+3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY THE LENS FIRST ROTATED BEGINNING (B)(6) 2019 BY 1 DEGREE. ON (B)(6) 2019 THE LENS HAD ROTATED 90 DEGREES AND WAS REPOSITIONED. ON (B)(6) 2019 THE LENS AGAIN ROTATED 10 DEGREES. THE REPORTER STATES, "THE VISION HAS DECREASED AFTER 1 WEEK POSTOPERATION. WE REPOSITIONED THE LENS 30 DEGREE[S]. THE LENS WAS MORE FLOATING DURING I/A, COMPARING TO OS." REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9720695&pc=QCB | 2023826-2020-00337 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5/5.5/158 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED "FOLLOWING EXPLANTATION ON (B)(6) 2020- NO OTHER PROCEDURE PERFORMED. PT. VISITED ON (B)(6) AND THE EYE HAS SETTLED WELL. VISUAL ACUITY 0.2 LOGMAR IN HER OLD SPECTACLES." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781532&pc=QCB | 2023826-2020-00341 | 2019-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS INDUCED ASTIGMATISM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT WITH BLOOD STAINS AND DEBRIS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781544&pc=QCB | 2023826-2020-00423 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. REPORTEDLY, PATIENT IS NOT HAPPY. PATIENT AND SURGEON ARE LOOKING TO REPLACE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION B5: ON (B)(6) 2020, THE LENS WAS REMOVED AND THE NEXT MONTH ON (B)(6) 2020 A REPLACEMENT LENS OF THE SAME SIZE DIFFERENT MODEL; SIMILAR POWER WAS IMPLANTED AND PROBLEM RESOLVED. D7: EXPLANT DATE: ON (B)(6) 2020. H6: PATIENT CODE 3191, SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9795627&pc=QCB | 2023826-2020-00418 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: CORRECTED DATA: H6-RESULT CODE 114 IN PREVIOUS MDR IS CORRECTED TO 213. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814234&pc=QCB | 2023826-2020-00472 | 2019-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819168&pc=QCB | 2023826-2020-00456 | 2019-03-21 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT SPHERE/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824536&pc=QCB | 2023826-2020-00359 | 2018-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.50/2.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THE PATIENT'S POSTOPERATIVE FEEDBACK WAS POOR VISUAL QUALITY, WITH DOUBLE SHADOW, AND BLURRED VISION." AFTER PATIENT AND SURGEON CONSULTATION, THEY DECIDED TO TAKE THE LENS OUT. PER THE REPORTER ON (B)(6) 2020, "PATIENT'S CURRENT CONDITION IS GOOD. AND THE POOR VISION, DOUBLE SHADOW, AND BLURRED VISION HAS CLEARED. THE REMOVAL OF LENS RESOLVED THE PROBLEM." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- "VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS" NEEDS TO BE INCLUDED IN PREVIOUS MDR. H10-PATIENT CODE 3191-POOR VISUAL QUALITY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842752&pc=QCB | 2023826-2020-00519 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE REPORTER STATES THAT THERE WAS PLANNED MONOVISION FOR THE LEFT EYE HOWEVER THE RESULT WAS EMMETROPIA. Manufacturer Narrative: LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTER STATES "PATIENT IS HAPPY". PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910962&pc=QCB | 2023826-2020-00687 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SAME MODEL/LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910963&pc=QCB | 2023826-2020-00695 | 2019-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/132 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS OUT OF SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910968&pc=QCB | 2023826-2020-00688 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+4.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913390&pc=QCB | 2023826-2020-00741 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928366&pc=QCB | 2023826-2020-00716 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.5/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE PATIENT EXPERIENCED BLURRED VISION." SHOULD BE IN THE INITIAL MDR. H6- PATIENT CODE 2137 SHOULD BE IN THE INITIAL MDR. H6- DEVICE CODE 1401 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936282&pc=QCB | 2023826-2020-00768 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.5/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND PERMANENT LOSS OF BCVA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR AND SAYS, "PROBABLY THE WTW WAS UNDER MEASURED THE 1ST TIME." Manufacturer Narrative: ADDITIONAL NARRATIVE: H6 - WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: B2 - CORRECTED TO 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE' IN INITIAL MDR. B5 - 'A REPOSITIONING WAS PERFORMED ON (B)(6) 2020, BUT THIS DID NOT RESOLVE THE PROBLEM.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS REPOSITIONING) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021, AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936302&pc=QCB | 2023826-2020-00769 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND PERMANENT LOSS OF BCVA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR AND SAYS, "PROBABLY THE WTW WAS UNDER MEASURED THE 1ST TIME." Manufacturer Narrative: ADDITIONAL NARRATIVE: H6 - WORK ORDER SEARCH - NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: B2 - CORRECTED TO 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE' IN INITIAL MDR. B5 - 'A REPOSITIONING WAS PERFORMED ON (B)(6) 2020, BUT THIS DID NOT RESOLVE THE PROBLEM.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - LENS REPOSITIONING) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# : (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6)2021 AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A SCRATCH ON THE LENS OPTIC AND DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941318&pc=QCB | 2023826-2020-00836 | 2020-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/6.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: ADVERSE EVENT OR PRODUCT PROBLEM: ADVERSE EVENT. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. EVENT: THE REPORTER INDICATED THAT MEDICAL INTERVENTION WAS REQUIRED. REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. PATIENT CODE 3191: REFRACIVE TREATMENT (PRK). METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941412&pc=QCB | 2023826-2020-00837 | 2020-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/6.0/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT MEDICAL INTERVENTION WAS REQUIRED. REFRACTIVE TREATMENT (PRK) WAS PERFORMED AND THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. H6-PATIENT CODE 3191: REFRACTIVE TREATMENT (PRK). METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965157&pc=QCB | 2023826-2020-00894 | 2020-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979190&pc=QCB | 2023826-2020-00925 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -07.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- "DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983361&pc=QCB | 2023826-2020-00935 | 2020-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B1 - ADVERSE EVENT. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5 - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+3.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS; DUE TO LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. POST EXCHANGE UCVA 20/20. "PATIENT IS HAPPY.". D7- EXPLANT DATE- (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATON H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043595&pc=QCB | 2023826-2020-01033 | 2018-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.50/4.0/094 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/ CYLINDER/ AXIS) LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043598&pc=QCB | 2023826-2020-01034 | 2018-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/2.0/075 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/ CYLINDER/ AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049158&pc=QCB | 2023826-2020-01060 | 2020-02-17 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 12.50/+1.0/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10053461&pc=QCB | 2023826-2020-01097 | NA | Injury | STAAR SURGICAL COMPANY | 2020-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/2.5/104 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH MYOPIA ASTIGMATISM SINCE AGE 10YRS ON (B)(6) 2016 AND VISION WAS CORRECTED, VA 20/16. ON (B)(6) 2019 THE PATIENT COMPLAINED OF BLURRED VISION "AFTER PLAYING FOOTBALL (BOUNCE A BALL WITH HEAD)." THE DOCTOR FOUND VA TO BE 20/25+1 PH20/20-1, REFRACTION +0.62 = 2.25 X 162 AND ICL (IMPLANTABLE COLLAMER LENS) HAD ROTATED. ON (B)(6) 2019 THE DOCTOR REPOSITIONED THE LENS AND VISION RETURNED TO 20/15-2. ON (B)(6) 2020 THE PATIENT RETURNED WITH COMPLAINT OF BLURRY VISION AFTER WAKING UP. HE DENIED ANY TRAUMA. HIS VA WAS 20/20-2 PH 20/16-2, REFRACTION +0.75 = -2.12 X 162 AND ICL (IMPLANTABLE COLLAMER LENS) WAS ROTATED COUNTERCLOCKWISE TO 14 DEGREE POSITION. THE DOCTOR STATED THAT " A WEEK LATER THE ICL IS STILL IN WORSEN ROTATED POSITION." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE SURGEON REPOSITIONED THE LENS AGAIN CLOCKWISE 9 DEGREES ON (B)(6) 2020 BUT DURING A FOLLOW UP ON (B)(6) 2020 THE SURGEON FOUND THAT THE LENS WAS ALIGNED CLOCKWISE 60 DEGREES. THE LENS REMAINED IMPLANTED. THE SURGEON WOULD LIKE TO EXCHANGE THIS LENS TO 13.2." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE SURGEON REMOVED THE INITIAL LENS ON (B)(6) 2021 AND IMPLANTED THE REPLACEMENT LENS VTICMO13.2-12.50/+2.5/095 ((B)(6)) IN THE SAME OPERATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075062&pc=QCB | 2023826-2020-01094 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PAITENT IS HAPPY. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATON - LENS RETURNED IN A CASE/VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075970&pc=QCB | 2023826-2020-01084 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: THE LENS WAS EXCHANGED FOR A SAME SIZE AND MODEL, DIFFERENT SPHERE/AXIS LENS. THE PROBLEM WAS RESOLVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. H6-PATIENT CODE 3191 (SECONDARY SURGERY, LENS EXCHANGE). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094011&pc=QCB | 2023826-2020-01124 | 2018-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140609&pc=QCB | 2023826-2020-01195 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTER THAT ON 10/06/2020 THE LENS WAS EXPLANTED. LATER ON THE SAME DAY A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY WITH VISION". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158148&pc=QCB | 2023826-2020-01242 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158172&pc=QCB | 2023826-2020-01244 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/017 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBER. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163423&pc=QCB | 2023826-2020-01299 | 2018-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.0/040 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163480&pc=QCB | 2023826-2020-01233 | 2018-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.5/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WIT MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10185146&pc=QCB | 2023826-2020-01344 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: IT WAS REPORTED THAT A TORIC ICL WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND ROTATED ABOUT 10 DEGREES. THE REPORTER STATED " LENS REPOSITIONING WAS PERFORMED, LENS HAS CURRENTLY REMAINED IN THE CORRECT POSITION AND THE PATIENT IS 20/20 UNCORRECTED AT THIS TIME." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.5/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM, THE LENS ROTATED BACK TO THE SAME POSITION ONE WEEK LATER. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189013&pc=QCB | 2023826-2020-01379 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -16.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTION DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189480&pc=QCB | 2023826-2020-01289 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE DIMENSIONAL VERIFICATION WAS IN SPECIFICATION BUT REFRACTION (FUNCTIONAL) VERIFICATION FAILED. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A NON-TORIC ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190922&pc=QCB | 2023826-2020-01389 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.50/+1.0/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE SURGEON IMPLANTED ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222184&pc=QCB | 2023826-2020-01419 | 2020-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/125 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SHORTER, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242474&pc=QCB | 2023826-2020-01425 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.50/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, NON TORIC LENS BUT SAME LENGTH LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271198&pc=QCB | 2023826-2020-01493 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -1/2.50/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BENT HAPTIC AND DEBRIS ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - "THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -1/2.50/95 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM." CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 1401 SHOULD BE ADDED TO INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10272062&pc=QCB | 2023826-2020-01471 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/1.5/10 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS IN A MICEO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10317988&pc=QCB | 2023826-2020-01533 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.0/+2.5/044 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353747&pc=QCB | 2023826-2020-01635 | 2020-06-29 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REDUCED VISION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401675&pc=QCB | 2023826-2020-01700 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432548&pc=QCB | 2023826-2020-01786 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHAGED FOR A DIFFERENT DIOPTER LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN / HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D10-DEVICE NOT AVAILABLE FOR EVALUATION. H1-THIS TYPE OF REPORT IS CORRECTED TO SERIOUS INJURY. H2- DEVICE EVALUATION IN PREVIOUS MDR IS NOT APPLICABLE. H3-DEVICE EVALUATED BY MANUFACTURER IN PREVIOUS MDR IS CORRECTED TO NO. H6- METHOD CODE 10 IN PREVIOUS MDR IS NOT APPLICABLE. RESULT CODE 114 IN PREVIOUS MDR IS NOT APPLICABLE. SUPPLEMENTAL #1 MEDWATCH REPORT SUBMITTED IN ERROR WITH WRONG MDR NUMBER YEAR. THE DEVICE EVALUATION IS FOR MDR# (2023826-2019-01786) CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN AMICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432698&pc=QCB | 2023826-2020-01791 | 2018-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-7.5/1.0/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHAGED FOR A DIFFERENT DIOPTER LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476366&pc=QCB | 2023826-2020-01926 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -5.50/+2.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H3-DEVICE EVALUATION IN PREVIOUS MDR: FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476403&pc=QCB | 2023826-2020-01928 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/158 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LEN. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476428&pc=QCB | 2023826-2020-01927 | 2020-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672562&pc=QCB | 2023826-2019-01031 | 2013-07-02 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/133 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. POST-OP WAS REPORTED AS PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672587&pc=QCB | 2023826-2019-01032 | 2013-07-02 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -5.0/+0.5/047 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL BUT SAME LENGTH SIZE LENS AND THE PROBLEM RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819696&pc=QCB | 2023826-2019-01303 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINS IMPLANTED BUT SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE LENS WAS REPORTED AS HAVING LOW VAULTING' CORRECTED TO 'THE LENS WAS REPORTED AS HAVING LOW VAULTING, LOSS OF VISUAL ACUITY, AND LENS DISLOCATION. THE SURGEON ALSO NOTED MICROSPHEROPHAKIA WHEN THE PUPIL WAS DILATED AS THE INFERIOR ZONULES WERE VISIBLE'. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. LENS INCIDENTS WERE REPORTED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819702&pc=QCB | 2023826-2019-01304 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINS IMPLANTED BUT SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE LENS WAS REPORTED AS HAVING LOW VAULTING' CORRECTED TO 'THE LENS WAS REPORTED AS HAVING LOW VAULTING, LOSS OF VISUAL ACUITY, AND LENS DISLOCATION. THE SURGEON ALSO NOTED MICROSPHEROPHAKIA WHEN THE PUPIL WAS DILATED AS THE INFERIOR ZONULES WERE VISIBLE'. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. LENS INCIDENTS WERE REPORTED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876501&pc=QCB | 2023826-2019-01393 | NA | Injury | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS SEVERAL MONTHS AGO IN THE PATIENTS RIGHT EYE (OD), AND NOW THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/168 (SPHERE/CYLINDER/AXIS) ON 19-FEB-2019 IN THE PATIENTS RIGHT EYE (OD). THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM, THEREFORE CREATING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.'PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT FOUND WITHIN ASSOCIATED LOTS. (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876502&pc=QCB | 2023826-2019-01392 | NA | Injury | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS SEVERAL MONTHS AGO IN THE PATIENTS LEFT EYE (OS), AND NOW THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: CORRECTED TO OTHER SERIOUS IN INITIAL MDR. 'THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/11 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2019 IN THE PATIENTS LEFT EYE (OS). THE SURGEON HAS REPORTED THE LENS HAS ROTATED CLOCKWISE AND HAS INDUCED SOME ASTIGMATISM, THEREFORE CREATING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.' TYPE OF REPORTABLE EVENT: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. PATIENT CODE 2692 NOT APPLICABLE IN INITIAL MDR. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANTED: UNK. EXPLANTED: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9272155&pc=QCB | 2023826-2019-02042 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL | Misfocusing | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - PATIENT CODE 2692 CORRECTED TO PATIENT CODE 2137. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113436&pc=QCB | 2023826-2018-01755 | 2018-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+6.0/002 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A CORNEAL ASTIGMATISM OF 8 DIOPTER. AFTER IMPLANTATION OF THE ICL THE CORNEAL ASTIGMATISM REDUCED TO 2 DIOPTER. THEREFORE THE CURRENT OUTCOME IS PLANO CYLINDER -6.00 DIOPTER. THE SURGEON PLANS TO EXPLANT THE LENS FOR ANOTHER LENS. THE LENS TO DATE REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.00/6.0/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THE PATIENT EXPERIENCED REFRACTIVE SURPRISE DUE TO POST-OPERATIVE REDUCTION IN CORNEAL ASTIGMATISM FROM 8 TO 2 DIOPTRES. THE LENS CURRENTLY REMAINS IMPLANTED WITH EXPLANT AND NEW LENS IMPLANTATION PLANNED FOR THE FUTURE. PATIENT CODE 2692: PREVIOUS CODE NOT APPLICABLE. DEVICE CODE 3189: PREVIOUS CODE NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED SUCCESSFULLY ON (B)(6) 2018. THE DOCTOR WILL WAIT APPROX. 1 MONTH BEFORE IMPLANTING A NEW LENS. AFTER THE EXPLANTATION THE PATIENT IS FINE. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CORRECTED DATA: COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. CLAIM#: (B)(4). Manufacturer Narrative: DISTRIBUTOR CHECKED IN ERROR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10164916&pc=QCB | 2023826-2020-01277 | 2020-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.50/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(4) 2020. THE LENS WAS EXPLANTED ON (B)(4) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURER DATE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17541754&pc=QCB | 2023826-2023-03429 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Blurred Vision; Swelling/ Edema | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, EYE SWELLING AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THAT THE SURGEON HAD REPOSITIONED THE LENS ON (B)(6) 2023 BEFORE EXCHANGING THE LENS. H3 - DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15008783&pc=QCB | 2023826-2022-02203 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.00/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. TASS WAS REPORTED AND PATIENT COMPLAINED OF BLURRED VISION. DIAGNOSTICS WAS PERFORMED - NON CONTACT TONOMETRY MEDICAL EXAMINATION. TREATMENT WAS PERFORMED - CLEANING, DROPPING MEDICINE IN THE EYE. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039569&pc=QCB | 2023826-2023-01883 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.0/6.0/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME(TASS) AND BLURRED VISION WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITION OF ORAL STEROIDS AND OINTMENTS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WASREPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6: HEALTH IMPACT CODE: "4625 - ADDITIONAL SURGERY: ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391254&pc=QCB | 2023826-2023-03035 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE AND TOXIC ANTERIOR SEGMENT SYNDROME(TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR, ASTIGMATISM CORRECTION. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -14.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME SIZE SIMILAR POWER LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR. THE DOCTOR BELIEVES THAT TASS MAY HAVE BEEN INDUCED BY USING AN INJECTOR THAT WAS NOT STAAR'S INJECTOR. H6 HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391263&pc=QCB | 2023826-2023-03036 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE AND TOXIC ANTERIOR SEGMENT SYNDROME(TASS, EYE DROP TREATMENT PERFORMED). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR, ASTIGMATISM CORRECTION. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -12.0/2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME SIZE SIMILAR POWER LENS DUE TO TOXIC ANTERIOR SEGMENT SYNDROME (TASS, EYE DROP TREATMENT PERFORMED). CAUSE OF THE EVENT IS REPORTED AS OTHER, TASS CAUSED BY INJECTOR. THE DOCTOR BELIEVES THAT TASS MAY HAVE BEEN INDUCED BY USING AN INJECTOR THAT WAS NOT STAAR'S INJECTOR. H6 HEALTH EFFECT CLINICAL CODE: 2137 SHOULD BE REMOVED. CLAIM#: 730392 Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323359&pc=QCB | 2023826-2016-01547 | 2016-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2017-02-10 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing | Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -8.0/+2.0/62 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, BLURRED VISION AND DIPLOPIA (DOUBLE VISION). THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/60. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703193&pc=QCB | 2023826-2020-00266 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THEIR LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT REPORTED HAS BLURRED VISION, ESPECIALLY AROUND LUMINOUS SIGNS AND DIFFICULTY READING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H1- "SERIOUS" SHOULD BE CORRECTED TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852522&pc=QCB | 2023826-2020-02677 | 2020-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -01.00/+6.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND PATIENT EXPERIENCED BLURRED VISION AND GLARE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO PRODUCT PROBLEM. B2 - OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE REMOVED. H1 - CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11606813&pc=QCB | 2023826-2021-00835 | 2020-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/2.0/088 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE AND DOUBLE VISION WAS OBSERVED. LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS THE DEVICE " SMALL DIAMETER OF ICL". THE REPORTER STATED " SURGEON HAS TRIED TO REPOSITION THE TICL TWO TIMES AND AFTER A COUPLE DAYS THE LENS WAS DECENTRED AGAIN, IN THE SAME POSITION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE DEVICE WAS THE CAUSE OF THE EVENT BUT THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "SMALL DIAMETER OF ICL, SULCUS ANATOMY?" CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12061900&pc=QCB | 2023826-2021-02038 | NA | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: ON (B)(6) 2021 WE RECEIVED NOTIFICATION FROM A PATIENT STATING THAT SHE HAD A 12.6MM, VTICM5_12.6, -6.50/0.5/077, IMPLANTABLE COLLAMER LENS IMPLANTED IN HER RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER STATED " AFTER ONE DAY I GOT TRAUMA IN MY HEAD AND I WENT FOR DOCTOR FOR EXAMINATION AND I COMPLAIN THAT MY VISION NOT CLEAR SHE ADVICE ME TO DO ANOTHER OPERATION FOR MY RIGHT EYE TO FIX THE AXIS AS MAY BE THE LENS ALIGNMENT MOVE FROM THE AXIS AND AFTER 2 WEEKS FROM SECOND OPERATION STILL I HAVE ASTIGMATISM IN MY RIGHT EYE." PER UPDATED INFORMATION FROM THE PATIENT SHE IS CONSULTING WITH A SURGEON FOR SECOND OPINION. THE PATIENT STATED " SHE EXAMINED MY EYES AFTER DILATING THE PUPILS AND SHE SAID TO ME THE AXIS OF THE LENS NOT ON THE CORRECT ALIGNMENT. SO SHE SAID TO ME FIRST TO FIX THE AXIS , AND IF STILL THERE IS ASTIGMATISM THAT MEAN NEED TO CHANGE THE LENS AND SHE SAID MAY BE THERE IS A NEW PRESCRIPTION FROM THE LENS COMPANY TO FIX THE AXIS IN MY CASE." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - DATE OF (B)(6) 2021, SHOULD BE CORRECTED TO (B)(6) 2021 IN INTIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - HEAD TRAUMA - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874022&pc=QCB | 2023826-2023-04415 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION. THE EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS, "PATIENT CANNOT ADAPT TO VISUAL GHOSTING AFTER IMPLANTATION OF CRYSTALS". THE REPORTER ALSO STATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473794&pc=QCB | 2023826-2024-00122 | NA | Malfunction | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -2.0/4.0/100 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. VISUAL ACUITY NOT IMPROVED. LENS REMAINS IMPLANTED. REPORTEDLY, VAULT HAS IMPROVED WITH GOOD RESULTS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634467&pc=QCB | 2023826-2020-02280 | NA | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED MILD GLARE AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02067. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED AT THE LAST VISIT THE PATIENT APPEARED TO BE STILL SYMPTOMATIC FOR GLARE, HALOS AND BLURRY VISION. A 30 DEGREE ROTATION OF THE LENS WAS MEASURED AT LAST VISIT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED MILD GLARE/HALOS AT NIGHT AND BLURRY VISION. THE LENS REMAINED IMPLANTED. AT THE POST-OP VISIT ON (B)(6) 2020 THE PATIENT WAS DOING WELL, EXPRESSED MILD DISCOMFORT BUT DIDN'T EXPRESS ANY SYMPTOMS OF GLARE/HALOS. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02067. H6: MEDICAL DEVICE PROBLEM CODE: 2923 - THIS CODE WAS USED IN ERROR, THE LENS DID NOT ROTATE. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2020. THE PATIENT REPORTED MILD GLARE AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT#: 2023826-2020-02067. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634506&pc=QCB | 2023826-2020-02282 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTED GLARE/HALOS AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02066. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 : THE REPORTER INDICATED AT POST-OP VISIT ON (B)(6) 2020, IT SHOWED THE PATIENT FELT VISION SEEMED BETTER. AT LAST VISIT ON (B)(6) 2020, THERE WERE NO COMPLAINTS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CHANGED TO PRODUCT PROBLEM. B4 - "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS" SHOULD BE REMOVED AS THIS CLAIM SHOULD BE CORRECTED TO NON SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573285&pc=QCB | 2023826-2021-00770 | NA | Injury | STAAR SURGICAL COMPANY | 2021-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES AND BLURRED VISION. REFRACTIVE SURPRISE WAS ALSO REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE RESIDUAL REFRACTION WAS ADDRESSED WITH LASIK SURGERY, "OD BCVA 20/20-2, UCVA 20/25, VAULTING 1.0CCT, IOP-NOT MEASURED". THE LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573286&pc=QCB | 2023826-2021-00769 | NA | Injury | STAAR SURGICAL COMPANY | 2021-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES AND BLURRED VISION. REFRACTIVE SURPRISE WAS ALSO REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE RESIDUAL REFRACTION WAS ADDRESSED WITH LASIK SURGERY, "OS BCVA 20/20-1, UCVA 20/20-1, VAULTING 1.0CCT, IOP-NOT MEASURED". THE LENS REMAINS IMPLANTED. CLAIM #(B)(4). Manufacturer Narrative: A THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11743352&pc=QCB | 2023826-2021-01264 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/3.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11743370&pc=QCB | 2023826-2021-01271 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-04-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/4.0/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION AND THE PROBLEM RESOLVED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14415838&pc=QCB | 2023826-2022-01589 | 2021-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.5/+5.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION; REFRACTIVE SURPRISE; GLARE/HALOES; LOW VISUAL ACUITY; BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER THE SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH; BUT DIFFERENT MODEL/AXIS/POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND "OTHER". Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. H3 - DEVICE ANTICIPATED BUT NOT YET BEGUN SHOULD, SHOULD HAVE BEEN SELECTED IN THE INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15219661&pc=QCB | 2023826-2022-02758 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS OF -2.00/1.5/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS OF +2.00/1.5/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE, GLARE/HALOES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. H6 - INVESTIGATION CONCLUSION CODE 4310 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#1. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937237&pc=QCB | 2023826-2022-04284 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER - 11.50/2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION, GLARES/HALOES, AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND POWER BUT LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "ALL WELL, LENS IS STABLE, REFRACTION OUTCOME IS SATISFACTORY". THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952662&pc=QCB | 2023826-2022-04354 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15953610&pc=QCB | 2023826-2022-04353 | 2020-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525763&pc=QCB | 2023826-2024-00168 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1, -6.0/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525764&pc=QCB | 2023826-2024-00167 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -6.0/1.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11382799&pc=QCB | 2023826-2021-00531 | 2020-12-18 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -11.50/2.0/67 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE. HALO AND GLARE BY LENS POSITION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-RACE- UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLE, PIGMENT DISPERSION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11382825&pc=QCB | 2023826-2021-00532 | 2020-12-18 05:00 | Injury | STAAR SUGICAL COMPANY | 2021-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. LENS REMAINS IMPLANTED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -7 .00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -7 .00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. "ICL TOO BIG" WAS REPORTED FOR CAUSE OF EVENT. "PIGMENT ON IOL. EXCESSIVE HIGH VAULT. IOP OK BUT NARROW ANGLE." WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). IT WAS REPORTED THAT THE LENS WAS REMOVED AND REPLACED AT AN UNKNOWN DATE AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT: CLINICAL CODE (B)(4): NARROW ANGLE, PIGMENT DISPERSION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555053&pc=QCB | 2023826-2023-00862 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555127&pc=QCB | 2023826-2023-00863 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+0.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16555168&pc=QCB | 2023826-2023-00861 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+2.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16559490&pc=QCB | 2023826-2023-00864 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNKNOWN ETHNICITY: UNKNOWN RACE: UNKNOWN (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636031&pc=QCB | 2023826-2023-01064 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). AT 3 WEEKS POST-OP, THE PATIENT COMPLAINED OF BLURRINESS (MILD) CENTRALLY, MONOCULAR DIPLOPIA, GLARE/HALOS AND GHOSTING/DOUBLING OF HER MIDPERIPHERAL VISION. THERE IS PERMANENT LOSS OF BCVA. THE CAUSE OF THE EVENT IS THE DEVICE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS ROTATED (REPOSITIONED) ON (B)(6) 2023. THE PATIENT IS UNCORRECTED 20/25- OD AT LAST VISIT ON (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. IMPLANTED/ EXPLANTED DATES UNKNOWN. ADVERSE EVENT PROBLEM: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE ERROR, PERMANENT LOSS OF BCVA; WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11158539&pc=QCB | 2023826-2021-00092 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2020 DUE TO BEING TOO LONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, BLURRED VISON AND GLARE/HALOS. THE LENS WAS WELL CENTERED, ANGLE CLOSURE, CLOSE TO SUTURE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THE PATIENT HAD AN ITCHY EYE BUT NO VISION COMPLAINTS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE; SUTURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8457907&pc=QCB | 2023826-2019-00477 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497526&pc=QCB | 2023826-2019-00568 | NA | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -11.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED A SAME SIZE, DIFFERENT MODEL LENS DUE TO LOW VAULT, LENS ROTATION, PIGMENT DISPERSION, REFRACTIVE CHANGE OVER TIME, GLARE/HALOS, BLURRED VISION, AND IRIS ATROPHY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736620&pc=QCB | 2023826-2019-01125 | 2019-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252685&pc=QCB | 2023826-2019-01959 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/071 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10077079&pc=QCB | 2023826-2020-01109 | 2018-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.00/1.00/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT WAS COMPLAINING OF EXCESSIVE GLARE/HALOES, MONOCULAR DIPLOPIA AND GHOSTING. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2018 TO CORRECT 2 DEGREES OF ERROR BUT SYMPTOMS PERSISTED DESPITE CORRECT IOL POSITION POST OP. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (GHOSTING, MONOCULAR DIPLOPIA). H6 - MEDICAL DEVICE PROBLEM CODE 2923. CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.00/+1.00/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, BLURRED VISION, GLARE AND HALOS, MONOCULAR DIPLOPIA, GHOSTING, ROTATION: REPOSITIONING ATTEMPTED. LENS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE: "BLURRED VISION IN RIGHT EYE AND HALOS AT NIGHT". ADDITIONAL INFORMATION PROVIDED: "WANTS REPLACEMENT OF ICL IN RIGHT EYE". IT WAS ALSO REPORTED, "THIS IS STILL UNDER CONSULT SO NO FURTHER UPDATES OR INTERVENTION DECIDED UPON". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9721918&pc=QCB | 2023826-2020-00310 | 2016-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2016. THE SURGEON REPORTS LENS ROTATION, GLARES/HALOS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), LENS DISLOCATION/SUBLUXATION, PIGMENT DISPERSION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. IN ADDITION REFRACTIVE TREATMENT WAS PERFORMED AFTER IMPLANT DUE TO LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR-THE SURGEON BELIEVES THE PATIENT'S SULCUS IS IRREGULAR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/SAME LENGTH LENS. THE PROBLEM WAS RESOLVED. H6- HEALTH EFFECT IMPACT CODE- 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: SYMPTOMS NOTED (B)(6) 2016. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9131800&pc=QCB | 2023826-2019-01761 | NA | Injury | STAAR SURGICAL COMPANY | 2019-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: RECEIVED AN ARTICLE INDICATING A SURGEON IMPLANTED A 13.7MM TORIC ICL IMPLANTABLE COLLAMER LENS, -7.50/+6.0/054 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) IN 2017. POSTOPERATIVELY, THE ICL WAS REPORTED AS BEING REPOSITIONED THREE TIMES BECAUSE OF ROTATION. THE PATIENT REPORTED BLURRY VISION WITH DIPLOPIA. THERE WAS SOME IRIS PIGMENT SEEN ON THE ANTERIOR SURFACE OF THE ICL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246802&pc=QCB | 2023826-2019-02026 | 2019-10-04 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -7.0/+4.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO BLURRED VISION, DOUBLE-VISION, AND NEAR VERTICAL ROTATION OF LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9246803&pc=QCB | 2023826-2019-02027 | 2019-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -5.5/+2.5/063 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO BLURRED VISION, DOUBLE-VISION, AND NEAR VERTICAL ROTATION OF LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - WOUND LEAKAGE AND ADDITION OF 3 SUTURES WERE REPORTED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE - WOUND LEAKAGE. H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION - ADDITION OF 3 SUTURES. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9406313&pc=QCB | 2023826-2019-02303 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC, DIFFERENT DIOPTER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BAD VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR G4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR H6 - PATIENT CODE 2137 CORRECTED TO PATIENT CODE 2140 IN INITIAL MDR H6 - DEVICE CODE 1401 CORRECTED TO DEVICE CODE 2923 IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938191&pc=QCB | 2023826-2020-00814 | 2020-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/121 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT HAS DIFFICULTY WITH SEEING NEAR AND READING. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER AND THE PROBLEM RESOLVED. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CORRECTED TO (B)(6) 2019. SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119235&pc=QCB | 2023826-2021-02320 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Impairment; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -14.0/+1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) HALOS, GLARE, SPARKLES, PIGMENT DISPERSION, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE. Manufacturer Narrative: H6: INCLUDE CLINICAL CODE: 2227. CLAIM #(B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -14.0/+1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH A HISTORY OF DIPLOPIA AND PROGRESSIVE MYOPIA ON (B)(6) 2021. EXCESSIVE VAULT; TRANSIENT LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA); HALOS; GLARE; SPARKLES; PIGMENT DISPERSION; AND BLURRED VISION WAS REPORTED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF A DIFFERENT MODEL. POSTOPERATIVE REPORTS THAT THE LENS WAS ROTATED 10 DEGREES AND THE PATIENT RE-ENTERED SURGERY TO ACCOMMODATE. AFTERWARDS REPORTS THAT THE PATIENT IS WELL, VISION WITHOUT HALOS OR FLASHES, VAULT OF 800, ANGLES DEGREE II, WILL CHECK UP EVERY 2 MONTHS TO SEE THE EVOLUTION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. IMPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119268&pc=QCB | 2023826-2021-02319 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Impairment; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/+0.5/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) HALOS, GLARE, SPARKLES, PIGMENT DISPERSION, AND BLURRED VISION WAS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS DEVICE. Manufacturer Narrative: H6 - INCLUDE CLINICAL CODE: 2227. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -12.0/+0.5/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH A HISTORY OF DIPLOPIA AND PROGRESSIVE MYOPIA ON (B)(6) 2021. EXCESSIVE VAULT; TRANSIENT LOSS OF BCVA; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA); HALOS; GLARE; SPARKLES; PIGMENT DISPERSION; AND BLURRED VISION WAS REPORTED. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF A DIFFERENT MODEL. POSTOPERATIVE REPORTS THAT THE PATIENT FEELS GOOD SEEING; VERY HAPPY; NORMAL IOP OF 12; VAULT OF 800; VISION WITHOUT HALOS OR FLASHES; ANGLES GRADE II-III BY GONIOSCOPY; WILL KEEP HIM IN CHECK-UPS EVERY 2 MONTHS TO SEE THE EVOLUTION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, SPARKLES). INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437079&pc=QCB | 2023826-2021-03475 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/044 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE AND THE PATIENT DEVELOPED A CATARACT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) -2021. THE PATIENT HAD BLURRED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS DEBRIS ON LENS. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM #:(B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840495&pc=QCB | 2023826-2021-04372 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS.THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN AND BENT AND FIBERS WERE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146115&pc=QCB | 2023826-2021-04373 | 2021-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS OPACITY-ASC, AND VISION LOSS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "SIZE ERROR, OCOS CALCULATION." REPORTEDLY, THE PROBLEM WAS NOT RESOLVED WITH EXPLANT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: CATARACT SURGERY IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15578068&pc=QCB | 2023826-2022-03617 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -3.0/1.50/064 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT INDICATED THE PATIENT WILL BE MEASURED AGAIN IN A FEW MONTHS; THE INITIAL REFRACTIVE MUST BE RE-DETERMINED THEN A DECISION ABOUT A NEW IMPLANTATION WILL OCCUR. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. A CATARACT WAS OBSERVED. A IOL WAS IMPLANTED AND THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+1.50/64 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND A CATARACT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND AN IOL WAS IMPLANTED. THIS RESOLVED THE PROBLEM. STATUS OF THE EYE: "THE PATIENT HAS ALSO DEVELOPED A CATARACT". ADDITIONAL INFORMATION: THE ICL WAS EXCHANGED FOR AN IOL. THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018251&pc=QCB | 2023826-2022-04429 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE; USER ERROR AND UNKNOWN. THE DEVICE FAILED TO PERFORM AS INTENDED. "EXCESSIVE LOW ICL VAULT MOST LIKELY CONTRIBUTED TO LENS OPACITY PROGRESSION." SEE MFR# 2023826-2022-04657. Manufacturer Narrative: ADDITIONAL INFORMATION: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT; PERMANENT LOSS OF BCVA AND LENS OPACITY (CATARACT) ASC WAS REPORTED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. REPORTEDLY THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547301&pc=QCB | 2023826-2023-03523 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/051 9SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/051 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM #(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741313&pc=QCB | 2023826-2023-04041 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.00/+0.50/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN AND THE PATIENT HAD BLURRY VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPOSITIONED AND IMPLANTED THE CORRECT WAY. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND DEVICE CAUSALITY IS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.00/0.5/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A CORTICAL LENS OPACITY AND BLURRED VISION. ON (B)(6) 2023, WAS EXPLANTED AND REPOSITIONED THE CORRECT WAY. THE PROBLEM WAS NOT RESOLVED, "REFRACTIVE LENS EXCHANGE/CATARACT SURGERY IS REQUIRED AS PATIENT IS +6 AND HAS A LENS OPACITY" AND " I WILL INFORM YOU OF THE RLE SURGERY DATE AND OUTCOME". STATUS OF THE EYE: "VERY BLURRY REFRACTION". THE CAUSE OF THE EVENT IS USER ERROR. CAUSE DETAILS: "LENS TWISTED ON IMPLANTATION. AIR BUBBLE CONFUSED POSITIONING HOLE." CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019193&pc=QCB | 2023826-2022-04657 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS. IT IS REPORTED THAT "THE ICL HAS LOW VAULT WITH SIGNIFICANT INFERIOR DECENTRATION AND CCW ROTATION. DESPITE REPLACING ICL WITH ONE STEP LARGER SIZED ICL THE VAULT REMAINS SHALLOW AND INFERIOR DECENTRATION SUPPORTS UNDERSIZING OF REPLACEMENT ICL DESPITE INCREASING SIZE ONE STEP." THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2022-04429 FOR INITIAL LENS. Manufacturer Narrative: CLINICAL CODE: 4581: SIGNIFICANT INFERIOR DECENTRATION. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: RECEIVED INFORMATION THAT THIS EXCHANGE LENS DID RESOLVE THE ISSUE OF LOW VAULT. THIS CLAIM IS NOT A COMPLAINT DUE TO NO PRODUCT DEFICIENCY AGAINST THE LENS. PLEASE DISREGARD THE INITIAL MDR. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16688954&pc=QCB | 2023826-2023-01202 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/2.5/84(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASSESSED: (B)(6) 2022), PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL AND BLURRED VISION AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT ON (B)(6) 2022 PILOCARPINE WAS USED FOR AROUND 1 MONTH, PATIENT PUPIL STILL UNREACTIVE AND STOP USING PILOCARPINE THEN. CATARACT SURGERY IS PLANNED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: UNREACTIVE/FIXED PUPIL, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678146&pc=QCB | 2023826-2023-01164 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EVENT DATE:UNK. EXPIRATION DATE: UNK. EXPLANT DATE: UNK. MFR DATE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16692052&pc=QCB | 2023826-2023-01165 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. A CATARACT WAS OBSERVED AND THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READING VA. THE LENS REMAINED IMPLANTED BUT THE PLAN WAS TO REMOVE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. AND LENS OPACITY (ANTERIOR SUBCAPSULAR) WAS OBSERVED. THE PATIENT COMPLAINED OF DISCOMFORT, BLURRED VISION AND DIFFICULTY READY VA. THE CAUSE WAS UNKNOWN. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: A4: WEIGHT UNK. A5: ETHNICITY UNK. A6: RACE UNK. B3:EVENT DATE UNK. D4: EXPIRATION DATE UNK. D6B: EXPLANT UNK. H4: DEVICE MANUFACTURE DATE UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY READING VA. CLAIM # 432841 | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17610615&pc=QCB | 2023826-2023-03705 | NA | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS OPACITY (CATARACT) ASC WAS OBSERVED ON (B)(6) 2023. LENS REMAINS IMPLANTED. REPORTEDLY, PATIENT IS TRAILING A SCL FOR THE REFRACTIVE SHIFT AND WILL REFER CATARACT SURGERY FOR NOW. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - CLINICAL CODE: "1766" AND "2137" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED. "4614 " SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED. "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED AN ASC LENS OPACITY OBSERVED ON 17/MAY/2023, REFRACTIVE CHANGE OVER TIME, AND A SPHERICAL CONTACT LENS PRESCRIBED. THE LENS REMAINS IMPLANTED. STATUS OF THE EYE:" PATIENT IS TRIALING SCL (SPHERICAL CONTACT LENS) OS FOR THE REFRACTIVE SHIFT; WILL DEFER CATARACT SURGERY FOR NOW." CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611928&pc=QCB | 2023826-2019-00824 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -5.50/2.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (ASC), SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS, AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. AS OF (B)(6) 2019 PATIENT'S "VISION IS GREAT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RESULT CODE: 114 SHOULD BE CORRECTED TO 213 IN INITIAL MEDWATCH REPORT. PATIENT CODE: 1766 SHOULD BE INCLUDED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: 04/17/2019 SHOULD BE CORRECTED TO 04/16/2019 IN INITIAL MEDWATCH REPORT. DEVICE CODE: 1494 OFF-LABEL USE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611958&pc=QCB | 2023826-2019-00791 | 2019-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: "REMOVED ON 10-MAR-2019" SHOULD BE CORRECTED TO "REMOVED ON (B)(6) 2019" IN INITIAL MDR. "10-MAR-2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936253&pc=QCB | 2023826-2020-00745 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -8.0/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY PSC CORTICAL. TRANSIENT LOSS OF BCVA IS ALSO REPORTED. REPORTEDLY, THE PATIENT DEVELOPED PSC CATARACTS ON (B)(6) 2019 THAT APPEARED TO RESOLVE BY (B)(6) 2019. THE CATARACT REAPPEARED IN (B)(6) 2019 AND BCVA WAS LIMITED TO 20/50. Manufacturer Narrative: H6 - METHOD CODE 3331. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND A CIRCULAR MARK ON THE OPTIC. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE PROBLEM WAS RESOLVED WITH S/P CATATARCT EXTRACTION AND A PCIOL IMPLANTATION." THIS STATEMENT SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4-DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. D10- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 13-MAR-2020. G4- DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476134&pc=QCB | 2023826-2020-01935 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. SEE MFR REPORT # 2023826-2020-01934 FOR INITIAL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS INSPECTED UNDER 2X MAGNIFICATION AND THERE WAS NO VISIBLE DAMAGE OR RESIDUE NOTED. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. IT IS LIKELY THAT THESE LENSES WERE INADVERTENTLY SWAPPED DURING RETURN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904423&pc=QCB | 2023826-2023-01686 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Cataract; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS OPACITY (CATARACT, ANTERIOR SUBCAPSULAR) WAS OBSERVED ON (B)(6) 2022. THERE WAS PIGMENT DISPERSION AND THE PATIENT COMPLAINED OF BLURRED VISION AND GLARE/HALOES. THE PROBLEM WAS RESOLVED. CHAFING OF IRIS WITH RESULTING SIGNIFICANT IRIS TRANSILLUMINATION BETWEEN 1 WEEK AND 1 MONTH POST OP WITH SLOW PROGRESSION; DESPITE EXCELLENT VAULT. BELIEVED TO BE MECHANICAL ISSUE RELATED TO LENS/IRIS APPROXIMATION. THE POST-OP STATUS OF THE EYE AFTER LENS REMOVAL WAS STABLE IRIS TRANSILLUMINATION. ANISOCORIA IMPROVED BUT STILL WITH MYDRIASIS AND OVOID PUPIL. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, CHAFING OF IRIS, IRIS TRANSILLUMINATION, MYDRIASIS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9874226&pc=QCB | 2023826-2020-00634 | 2020-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN-HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUES EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, IRIS ATROPHY, LENS OPACITY ASC AND PUPIL BLOCK; ELEVATED IOP AND INFLAMMATION WERE OBSERVED. SECONDARY SURGICAL INTERVENTION (YAG AND ANTERIOR CHAMBER IRRIGATION) AND MEDICAL INTERVENTION (EYE DROPS TO REDUCE INFLAMMATION) WERE PERFORMED/PRESCRIBED. THE REPORTER STATED " THE OPACITY WAS SEEN BEFORE AC IRRIGATION AND BEFORE YAG PI. PATIENT IOP IS STABLE AND VA IS IMPROVING. THE DOCTOR IS WAITING FOR PUPIL IMPROVEMENTS SINCE IT SEEMS TO BE SLIGHTLY STUCK TO THE ICL." PATIENTS LAST VISIT DATE ON (B)(6) 2020 UCVA WAS REPORTED TO BE 20/30, BCVA PIN- HOLE 20/25, ASC OPACITY REMAINS THE SAME, IN CONTINUED EVALUATION. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- INCLUDED "ELEVATED IOP AND INFLAMMATION" TO STATEMENT. H6-PATIENT CODE 1937- ELEVATED IOP. 1932- INFLAMMATION. H6-DEVICE CODE: 1401-BLURRED VISION. H6-DEVICE CODE 2993 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107370&pc=QCB | 2023826-2019-01751 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Cataract; Inflammation; Intraocular Pressure Increased; Pain; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -18.00/+1.0/177 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED. THE LENS WILL BE EXPLANTED BUT TO DATE REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1937: INFLAMMATION. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: LENS IMPLANTED UPSIDE DOWN. PATIENT CODE 1937 IN PREVIOUS MDR CORRECTED TO 1932. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16261376&pc=QCB | 2023826-2023-00326 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER (PATIENT CURRENT OBSERVING DOCTOR) STATES THAT A SURGEON AT A DIFFERENT CLINIC IMPLANTED A 13.7MM VTICMO 3.7 IMPLANTABLE COLLAMER LENS OF -7.00/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE CURRENT OBSERVING DOCTOR SAW PATIENT FOR THE FIRST TIME ON (B)(6) 2022, WHERE HE UNFORTUNATELY BECAME AWARE OF SEVERAL MEDICAL ERRORS. "TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT." COMPLAINT QUESTIONNAIRE PROVIDED REPORTS CATARACT (ASC); ANGLE CLOSURE WITH ELEVATED IOP; LENS DISLOCATION/ SUBLUXATION; PIGMENT DISPERSION; UNREACTIVE FIXED PUPIL; BLURRED VISION; AND IRIS ATROPHY WAS REPORTED. CURRENT STATUS OF THE EYE: "EYE ALREADY CALM, CORNEA CLEAR, ANTERIOR CHAMBER SHALLOW, PUPIL WIDE DILATED, ICL IN SITUATION ROTATED COUNTER CLOCKWISE TO INTENDED POSITION, ANTERIOR SUBCAPS CAT. CATARACT FORMATION." CAUSE REPORTED AS USER ERROE; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE IMPLANTING SURGEON HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10538715&pc=QCB | 2023826-2020-02049 | 2020-07-08 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Cataract; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; Vomiting; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/1.50/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE SURGEON ASSESSED AN ELEVATED IOP OF 22MMHG. THIS WAS CONTROLLED WITHIN THE FIRST 24 HOURS WITH MEDICATION. ALSO REPORTED IS PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, DILATED PUPIL, GLARE/HALOES, IRIS ATROPHY, PUPIL BLOCK, AND ANTERIOR SUBCAPSULAR OPACITIES. THE PATIENT ALSO EXPERIENCED NASEAU AND EXCESSIVE MOVEMENT DUE TO VOMITING CAUSED BY THE ANESTHESIA. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE PUPILLARY DILATION WAS "SLOWLY CORRECTED WITHIN THE FIRST MONTH, LEAVING AN ANISOCORIA OF APPROX. 2.5MM AND PRESENTING SECTORAL ATROPHY OF THE IRIS." THE PATIENT ALSO REPORTS "DECREASE IN THE INTENSITY OF THE COLORS." THE PATIENT STILL EXPERIENCES GLAUKOMFLECKEN, IRIS PIGMENT DISPERSION, AND GLARES AT NIGHT, BUT IS OVERALL "REALLY SATISFIED." THE CAUSE OF THE EVENT IS REPORTED AS "REMNANT OF OVD BEHIND THE ICL." Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844162&pc=QCB | 2023826-2023-04346 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6) 1983, A3 - M, B3 - 13-SEP-2022. CORRECTION: B5 - REPORTED AS: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND DISPHOTOPSIA ISSUES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS REPOSITIONING ON (B)(6) 2022 WAS PERFORMED DUE TO ROTATION AND GLARE/HALOES BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023 AND IN AS SECOND SURGERY ON (B)(6) 2023 A LONGER LENGTH SIMILAR POWER LENS (SPHERE/CYLINDER/AXIS) WAS IMPLANTED. PROBLEM RESOLVED. REPORTEDLY, MILD ASC AND PATIENT IS VERY HAPPY. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17844244&pc=QCB | 2023826-2023-04347 | NA | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Cataract; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - (B)(6)1983, A3 - M, B3 - 13-SEP-2022. CORRECTION B5 - REPORTED AS THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXCHANGED WITH LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) DUE TO DISPHOTOPSIA ISSUES ON (B)(6) 2023. PROBLEM RESOLVED. CORRECT TO: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXPLANTED AND A NEW LONGER LENGTH SIMILAR POWER LENS(SPHERE/CYLINDER/AXIS) WAS LATER IMPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND GLARE/HALOES ON (B)(6) 2023. PROBLEM RESOLVED. REPORTEDLY MILD ASC, PATIENT IS VERY HAPPY. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: DISPHOTOPSIA ISSUES. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941309&pc=QCB | 2023826-2019-01571 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -13.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 CATARACT SURGERY WAS PERFORMED DUE TO LENS OPACITY-CORTICAL, GLARE/HALOS, POOR VISION, AND CORTICAL CHANGES. THE SURGEON STATES, "VISION POOR WITH GLARE/HALOS, SOME DOUBLE." Manufacturer Narrative: "IT WAS LATER REPORTED AT A FOLLOW UP DATE THAT THE PATIENT STATUS IS GOOD AND STABLE WITH NO VISUAL COMPLAINTS." SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171666&pc=QCB | 2023826-2020-01273 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0/+4.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10171670&pc=QCB | 2023826-2020-01272 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+4.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE SURGEON ASSESSED LENS OPACITY ASC. ALSO REPORTED IS LOW VAULT AND LENS ROTATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON 14-JUL-2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." H6- PATIENT CODE 3191- SECONDARY SURGERY: LENS EXCHANGE CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, EXCHANGE FOR A LONGER LENGTH LENS DID NOT RESOLVE THE PROBLEM - LOW VAULT REMAINS AND THE SURGEON WILL CONTINUE TO MONITOR MILD ASC CATARACT. CORRECTIONS: B5 - "BLURRED VISION WAS ALSO REPORTED." SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 2140 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13698235&pc=QCB | 2023826-2022-00761 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Conjunctivitis; Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.00/+2.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT INTRAOCULAR PRESSURE INCREASED AND COMPLAINED OF BLURRY VISION AND PAIN ON THE DAY OF SURGERY; MEDICATION WAS PRESCRIBED AND THE SURGEON PERFORMED IRRIGATION BEFORE SENDING HER HOME. ONE DAY POST OP THE PATIENT HAD HIGH IOP AND CORNEAL EDEMA CENTRALLY WITH CONJUNCTIVAL REDNESS. PUPILLARY BLOCK 3+ CELLS/FLARES NOTED. ON THIS DAY THE SURGEON REPORTED NOTICING A THIN MEMBRANE OVER THE CENTRAL HOLE OF THE ICL; WHICH THEY PERFORMED AN ADDITIONAL IRRIGATION WITH ADDITIONAL MEDICATION PRESCRIBED. IT WAS INDICATED AFTER ANOTHER POSTOP VISIT SURGEON PERFORMED AN ADDITIONAL IRRIGATION AND THE MEMBRANE WASHED OUT DURING THIS AND THE PATIENT'S IOP BECAME STABLE. IT WAS FURTHER INDICATED 1 WEEK POSTOP THE IOP IS STABLE; NO CORNEAL EDEMA BUT PATIENT HAS PUPIL MID-DILATED DUE TO THE PUPILLARY BLOCK AND 1.2 CT VAULTING. THE PATIENT WILL CONTINUE BEING MONITORED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200502&pc=QCB | 2023826-2021-02790 | 2021-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Clouding/Hazing; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. POST-OP, THE VISION WAS HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. LENS LOOKED APPROPRIATE AND PHAKIC LENS WAS CLEAR. THE LENS REMAINED IMPLANTED . THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02534. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM LENS IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY POST-OP THE VISION WAS HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. LENS LOOKED APPROPRIATE AND PHAKIC LENS WAS CLEAR. THE LENS REMAINS IMPLANTED. THE SERIAL NUMBER FOR THIS LENS WAS NOT PROVIDED. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840739&pc=QCB | 2023826-2021-04397 | 2021-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Blurred Vision; Discomfort; Eye Pain | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO CORNEAL DECOMPENSATION. REPORTEDLY, "HOURS AFTER UNEVENTFUL SURGERY, PAIN, DECREASED VISION; ON EXAMINATION: CORNEAL DECOMPENSATION, NO DETAILS VISIBLE." THE PROBLEM IS MARKED AS RESOLVED AFTER EXPLANT HOWEVER THERE WAS STILL SOME CORNEAL EDEMA AT 7H. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE PATIENT STARTED HAVING DISCOMFORT A FEW HOURS AFTER SURGERY AND THE LENS HAD TO BE REMOVED TWO DAYS LATER. SYMPTOMS ABATED AFTER EXPLANT. THE SURGEON STATES THE ICL LOOKED FINE WITHOUT VISIBLE DEFECTS AND HE REMAINS UNCERTAIN ABOUT THE EXACT CAUSE OF THE EVENT. CLAIM# (B)(4). Manufacturer Narrative: D4: CORRECTED MODEL#, SERIAL#, AND EXPIRATION DATE. D6A: CORRECTED TO ON (B)(6) 2021. D6B: CORRECTED TO ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558781&pc=QCB | 2023826-2021-00704 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200 AND PUPIL IS DILATED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SIGNIFICANT REDUCTION OF ICA, CORNEAL DECOMPENSATION, CORNEAL EDEMA ALSO REPORTED. THE CAUSE IS REPORTED AS RETAINED VISCOELASTIC BEHIND THE ICL AND "ICL STILL A BIT OVERSIZED." REPORTEDLY THE PATIENT IS RECOVERED, VISION IS GOOD. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558806&pc=QCB | 2023826-2021-00699 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS -15.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTEDLY EXPLANTED THE LENS AND REPLACED IT WITH A SHORTER LENS BECAUSE THE LENS WAS "TOO BIG." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: EXCESSIVE VAULT-TOUCHING THE CORNEA, ELEVATED IOP, 4+ CORNEAL EDEMA, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION REPORTED. CAUSE IS REPORTED AS DEVICE: ICL TOO LARGER PER CALCULATOR. PROBLEM MARKED AS RESOLVED. HOWEVER, SEE MFR REP# 2023826-2021-00704 FOR REPLACEMENT LENS. ORIGINAL LENS IMPLANTED, REMOVED, AND REPLACED WITHIN THE SAME SURGERY. CLAIM #(B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUIED IN A VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699716&pc=QCB | 2023826-2019-00980 | NA | Injury | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/5.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK WITH ELEVATED IOP (50MMHG), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, BLURRED VISION, AND IRIS ATROPHY WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, INTRAOCULAR INJECTION AND THE LENS WAS REMOVED ON (B)(6) 2019. PER THE REPORTER, ON (B)(6) 2019 PATIENT'S EYE STATUS, "IRIS ATROPHYIC AND PUPIL DILATED. REPORTER NOTED CORNEA BECAME CLEAR AFTER 1 MONTH. EYE IS CALM. CAUSE OF THE EVENT IS REPORTED AS "URETS-ZAVALIA SYNDROME DUE TO OCCLUSION OF THE AQUAPORT FIRST WITH VISCOELASTIC THEN WITH IRIS PIGMENTS." Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE 1937: ELEVATED IOP. DEVICE CODE: 1401: MISFOCUSING; PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1401 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO "DEVICE CODE 3191: PIGMENT DISPERSION, FIXED PUPIL, IRIS ATROPHY" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13716395&pc=QCB | 2023826-2022-00768 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Corneal Edema; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.50/4.0/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS IMPLANTED UPSIDE DOWN AND WOULDN'T OPEN; WHICH THE SURGEON MANAGED TO EVENTUALLY OPEN IT AND LEFT IT IN PLACE. THERE WAS THEN CORNEAL EDEMA AND REFRACTIVE SURPRISE; WHICH THE SURGEON NOTED IT WAS UPSIDE DOWN. THE LENS WAS THEN EXPLANTED ON THE SAME DATE. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. POSTOPERATIVE REPORT INDICATED "EXCHANGED AS SUCCESSFUL. SURGEON NOW UNDERSTANDS HOW TO ADAPT TECHNIQUE TO ENSURE LENS UNFOLDS CORRECTLY" THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273272&pc=QCB | 2023826-2023-02723 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/3.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS, ON 16-JUN-2023. IT WAS NOTED THE PATIENT HAS GRADE 1 EDEMA AND BLURRY VISUAL ACUITY BUT IMPROVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17275327&pc=QCB | 2023826-2023-02810 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/4.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS, ON (B)(6) 2023. IT WAS NOTED THE PATIENT HAS GRADE 1 EDEMA AND BLURRY VISUAL ACUITY BUT IMPROVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18295971&pc=QCB | 2023826-2023-05524 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/2.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS A REFRACTIVE ERROR DUE TO NEGATIVE SIGN NOT ENTERED INTO CYLINDER PART OF REFRACTION IN OCOS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16689487&pc=QCB | 2023826-2023-01207 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5/5.0/098 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, BLURRED VISION AND CORNEAL EDEMA. CAUSE OF EVENT WAS UNKNOWN." IRIS ATROPHY IN HIGH SECTOR" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). PER UPDATED INFORMATION THE PATIENT IS IN PSEUDOPHAKIA. MEDICAL INTERVENTION ( INTENSIVE FREQUENCY OF CORTICOSTEROIDES DROPS AND ANTIEDEMA DROPS) WAS PRESCRIBED FOR CORNEAL EDEMA. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN HAPTIC. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS HAPTIC DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: ADVERSE EVENT PROBLEM - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). ADVERSE EVENT PROBLEM - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16746796&pc=QCB | 2023826-2023-01372 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.0/4.0/038 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION; REFRACTIVE CHANGE OVER TIME AFTER THE LENS HAD ROTATED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND AN INCISION ENLARGEMENT WAS ALSO REPORTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+4.0/38 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION, REFRACTIVE CHANGE OVER TIME, INCISION ENLARGEMENT AND A MILD CORNEAL OEDEMA. REPORTEDLY "AFTER CONSULTATION WITH THE SURGEON -MILD CORNEAL EDEMA WAS EXPECTED AT IT WAS AN INITIAL PATIENT SPECIFIC OF CORNEA TENDENCY OF CORNEA". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIMILAR POWER; LONGER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE WAS REPORTED AS "MILD CORNEAL EDEMA - TEMPORAL PERIPHERAL PART. OS REFRACTION +0.75/-1.75 AX 5". THE CAUSE OF THE EVENT AS UNKNOWN. H6- HEALTH EFFECT- CLINICAL CODE: 4581- " INCISION ENLARGEMENT." CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17935878&pc=QCB | 2023826-2023-04537 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Corneal Edema; Headache; Blurred Vision; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -12.5/+1.00/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION, THE SURGEON ALSO NOTED HEADACHES, CORNEAL EDEMA, BLURRY VISION AND GLARES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR - PREOP AXIS WAS NOT ACCURATE BASED ON THE TOPOGRAPHY AND AUTO-REFRACT AXIS OF ABOUT 49 DEGREES AND IT WAS A PRE-OP MEASUREMENT ERROR NOT AN ISSUE WITH THE LENS ITSELF. Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THAT THE PATIENT HAD EYE PAIN AND MEDICATION WAS USED, TIMOLOL AND DIAMOX. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473796&pc=QCB | 2023826-2024-00131 | 2023-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129041&pc=QCB | 2023826-2018-01773 | 2018-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.5/+2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY, THE PATIENT PRESENTED WITH BLURRY VISION, CORNEAL EDEMA, HIGH IOP AND EXCESSIVE VAULT. THE SURGEON PERFORMED A YAG PI AND PATIENT WAS PLACED ON IOP MEDICATIONS. CURRENTLY, THE PATIENT'S SYMPTOMS ARE RESOLVING AND VISION HAS IMPROVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9709011&pc=QCB | 2023826-2020-00272 | 2019-12-17 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.00/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. POST-OP, THERE WAS CORNEAL EDEMA AND THE INTRAOCULAR PRESSURE (IOP) INCREASED, WHICH WAS MANAGED WITH WOUND BURPING, PO DIAMOX, PREDNEFRINE FORTE AND TOPICAL G COMBIGAN. THE CORNEAL EDEMA AND THE IOP HAD SETTLED DOWN IN LESS THAN 2 WEEKS. POST-OP, THE PUPIL WAS REPORTED AS OVOID (OVAL SHAPED). THE LENS REMAINS IMPLANTED AND THE SURGEON AND PATIENT ARE WAITING FOR THE CORNEA TO SETTLE DOWN FOR A COUPLE OF MONTHS BEFORE ANY DECISIONS ARE MADE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8780421&pc=QCB | 2023826-2019-01173 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/4.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA (PRE-OP BCVA = 20/80 POST-OP BCVA = 20/500) AND ELEVATED IOP (40 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2019. LENS WAS REMOVED ON (B)(6) 2019. PER REPORTER, PATIENTS CONDITION ON LAST VISIT, (B)(6) 2019, IOP 38 MMHG, CORNEAL EDEMA, MYDRIASIS, PHOTOPHOBIA, BLURRED VISION, PAIN. REPORTER INDICATED ON (B)(6) 2019, THAT PATIENT RECEIVED TREATMENT FOR HIGH IOP, DILATING DROPS WITH STEROIDS AND NON-STEROIDS. "NO MORE CORNEAL EDEMA AFTER" LENS REMOVAL. AFTER LENS REMOVAL, SURGEON CHOSE CLE (CLEAR LENS EXTRACTION) TO RESOLVE THE REFRACTIVE ISSUE. THE PATIENT IS FINE. THE REPORTER INDICATED THAT CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. CORRECTED TO NO, DEVICE EVALUATION ANTICIPATED BUT NOT YET BEGUN, IN INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334355&pc=QCB | 2023826-2021-00359 | 2021-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ANOTHER LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. POST-OP MILD CORNEAL EDEMA. THE LENS WAS DISCARDED. Manufacturer Narrative: B5 - THE PATIENT STATED SEEING A GLARE THE LEFT EYE (OS) WHILE NIGHT DRIVING AND USES ARTIFICIAL TEARS OCCASIONALLY AS NEEDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860009&pc=QCB | 2023826-2019-01291 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Pain; Red Eye(s); Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/3.5/007 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION, "PAINFUL, AND RED EYE" WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS "THE SIZES OF THE LENS PROBABLY TOO BIG." PER PATIENTS LAST VISIT ON (B)(6) 2019, "REDUCED VISION LEFT EYE STILL RIGID BIG PUPIL, IRIS DEFECT AS 13.00, CORNEAL EDEMA 201090410, RIGID PUPIL HAS NOT RETURNED TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191 - IRIS DEFECT AND FIXED PUPIL. DEVICE CODES 3001 AND 1583 THAT WERE PREVIOUSLY REPORTED ARE NOT APPLICABLE. COMMON DEVICE NAME/ PRODUCT CODE CORRECTED. MTA PRODUCT CODE IN INITIAL MDR NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9126212&pc=QCB | 2023826-2019-01791 | 2019-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/3.0/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-ANGLES AND GLARE/HALOES. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE CORNEAL EDEMA WAS REPORTED. Manufacturer Narrative: PATIENT CODE: 1791 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350551&pc=QCB | 2023826-2019-02189 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Ulcer; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/081 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE WAS INITIATED AND A THERAPEUTIC CONTACT LENS INSERTED. PAIN WAS REPORTED AND CORNEAL ULCER. THE LENS REMAINS IMPLANTED. PATIENT CODE: 3191 MEDICATION TREATMENT WITH STEROID, ANTIBIOTIC LUBRICANT AND ACETAZOLAMIDE PRESCRIBED, THERAPEUTIC CONTACT LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11464544&pc=QCB | 2023826-2021-00625 | 2020-11-28 05:00 | Malfunction | STAAR SUGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTED COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT EXPERIENCED GLARE/HALOES. LENS REMAINS IMPLANTED,BEING MONITORED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) "MILD DRYNESS INFERIOR CORNEA" WAS REPORTED. THE REPORTER STATED "FEELS LIKE SEES DIPLOPIA IN DIMMER LIGHT AT DISTANCE AND NEAR." SURGEON IS WAITING TO SEE IF NEUROADAPTATION OCCURS OVER TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: PRODUCT PROBLEM; H1: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11464664&pc=QCB | 2023826-2021-00626 | 2020-11-28 05:00 | Malfunction | STAAR SUGICAL COMPANY | 2021-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Blurred Vision; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/1.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. PATIENT EXPERIENCED GLARE/HALOES. LENS REMAINS IMPLANTED,BEING MONITORED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) "MILD DRYNESS INFERIOR CORNEA" WAS REPORTED. THE REPORTER STATED "FEELS LIKE SEES DIPLOPIA IN DIMMER LIGHT AT DISTANCE AND NEAR." SURGEON IS WAITING TO SEE IF NEUROADAPTATION OCCURS OVER TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: PRODUCT PROBLEM. H1: MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983283&pc=QCB | 2023826-2022-01033 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.5/+2.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6: CLINICAL CODE 4581 (ROUND PUPIL). CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983288&pc=QCB | 2023826-2022-01032 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Dry Eye(s); Inflammation; Intraocular Pressure Increased; Blurred Vision; Discomfort; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.5/+1.0/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, AC INFLAMMATION CELLS +, SPK, HEAVINESS AND EYE DISCOMFORT, PHOTOPHOBIA, EYE PAIN, UNABLE TO OPEN EYE, DRY EYE. REPORTEDLY MEDICATION WAS ALSO ADMINISTERED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: BLURRED VISION AND ROUND PUPIL ALSO REPORTED. H6-CLINICAL CODE 4581 (ROUND PUPIL). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15583155&pc=QCB | 2023826-2022-03588 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -14.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AS A REPLACEMENT LENS. IT WAS REPORTER THAT IT RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) DRYNESS AND FLOATERS WERE REPORTED. FOR PATIENTS CURRENT CONDITION IT WAS REPORTED THAT THE EYE IS S/P ICL AND DOING WELL, NO INTERVENTION WAS NECESSARY AND IT HAS IMPROVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023217&pc=QCB | 2023826-2021-04898 | NA | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision | N/A | NA | 1 | Event Description: A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10.5/+2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS ENDOPHTHALMITIS AFTER ONE WEEK OF SURGERY. TREATMENT WAS PROVIDED, THE LENS WAS EXPLANTED. VITREORETINAL INJECTION AND OTHER MEDICATION WAS GIVEN. THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS INITIALLY IMPLANTED (B)(6) 2021. THE LENS WAS EXPLANTED (B)(6) 2021. ADDITIONAL SYMPTOM OF BLURRED VISION REPORTED. H6-DEVICE CODE 1494: (AGE<21YRS) . CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180458&pc=QCB | 2023826-2019-01904 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION AND THE SURGEON SUSPECTED ASEPTIC ENDOPHTHALMITIS. FIBRIN DEPOSITION AND LENS REAR SURFACE DEPOSITION WERE NOTED. THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND DECADRON INJECTION TO THE SUBCONJUNCTIVAL WAS PERFORMED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE CODE 2993 CORRECTED TO DEVICE CODE 1401 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER 12 DAYS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463780&pc=QCB | 2023826-2020-01914 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Loss of Vision; Vitrectomy; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/1.50/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ENDOPHTHALMITIS AND LOSS OF VISION WERE OBSERVED. DIAGNOSTICS (PRP OF VITREOUS SAMPLE) AND MEDICAL INTERVENTION (VITREOUS BIOPSY, CORE VITRECTOMY, INTRAVITREAL VANCOMYCIN, INTAVITREAL PIPERACILLIN TAZOBACTAM) WERE PERFORMED. VITREOUS BIOPSY SAMPLE REVEALED PSEUDOMONAS AERUGINOSA. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND ADDITIONAL INTERVENTION/TREATMENT WAS PERFORMED (ANTERIOR CHAMBER WASH, VITREOUS BIOPSY, VITRECTOMY, FLUID- AIR EXCHANGE, SILICONE OIL INJECTION, INTRAVITREAL VANCOMYCIN, INTRAVITREAL PIPERACILLIN TAZOBACTAM, INTRAVITREAL DEXAMETHSONE). VITREOUS BIOPSY WAS SENT FOR REPEAT GRAMS STAIN, KOH MOUNT, CULTURE AND SENSITIVITY FOR BACTERIA AND FUNGUS. ON (B)(6) 2020 VITREOUS CULTURE SHOWED NO GROWTH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INFORMATION RECEIVED ON 21 SEP 2020: PATIENT HAD NO GENERAL HEALTH CONDITIONS. PATIENT WAS PREPPED WITH POVIDONE IODINE. A PROPHYLACTIC INTRACAMERAL (MOXIFLOXACIN) ANTIBIOTIC WAS USED AT THE TIME OF IMPLANTAION. PATIENT HAS PL+VE VISION IN THE AFFECTED EYE. VISUAL PROGNOSIS IS NOT GOOD. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16622823&pc=QCB | 2023826-2023-01088 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO 09-FEB-2023. H6: 4629 SHOULD BE CORRECTED TO 4627. ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT IS HYPEROPIC AND SHOWING RESIDUAL CYL; HAS HEADACHES AND FATIGUE; STEROIDS WERE PRESCRIBED. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SAME MODEL AND SAME LENGTH BUT DIFFERENT DIOPTER LENS.IT IS UNCLEAR WHETHER LENS REMOVAL AND REPLACEMENT OCCURRED INTRAOPERATIVELY. RESIDUAL CYLINDER WAS PRESENT AFTER THE EXCHANGE. IT HAS NOT BEEN REPORTED IF THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 HYPEROPIC/SHOWING CYL. TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715562&pc=QCB | 2023826-2023-03867 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715631&pc=QCB | 2023826-2023-03868 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Fatigue; Nausea; Visual Disturbances; Dizziness; Halo; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOES, DIZZINESS, NAUSEA, EYE FATIGUE AND PAIN. THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS NORMAL AND GOOD NEAR AND FAR VISION. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766109&pc=QCB | 2023826-2023-01337 | NA | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Glaucoma; Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)-2023, YAG WAS PERFORMED DUE TO TRANSIENT LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY, "PATIENT HAD INTERMITTENT TRANSIENT VISION LOSS RESOLVED WITH LASER PI. HER LATEST IOP IS 12MMHG. POSSIBLE INTERMITTIENT RESULTS OF PUPILARY BLOCK PRIOR TO LASER PI IN THIS CASE. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(6). Manufacturer Narrative: CORRECTED DATA: B5- THE PATIENT ALSO EXPERIENCED GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. H6- HEALTH EFFECT- CLINICAL CODE: 4581- GRAYISH VISION, VISUAL FIELD DEFECT, ISCHAEMIC ATTACK, AND PUNCTATE EPITHELIAL EROSIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, YAG WAS PERFORMED, ELEVATED IOP: 24 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, PERIORBITAL EYE PAIN, MEDICATION WAS PRESCRIBED, GLAUCOMA, INTERMITTENT PUPILLARY BLOCK, AND HEADACHES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE PATIENT DEVELOPED OS SECONDARY GLAUCOMA POST ICL, PATIENT DECIDED TO REMOVE BOTH EYE ICL". CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153715&pc=QCB | 2023826-2022-02531 | NA | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Glaucoma; Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS GRADUAL LOSS OF VISION AND DIAGNOSIS OF BLINDNESS ((B)(6) 2022). BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION (ANTI GLAUCOMA AND ANTI INFLAMMATION) . THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- " DATE OF (B)(6) 2014 SHOULD BE CORRECTED TO (B)(6) 2014." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM 12.5 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/+6.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LOSS OF VISION; BLINDNESS DUE TO GLAUCOMA (DIAGNOSED ON (B)(6) 2022; ELEVATED IOP; SECONDARY YAG PERFORMED; LENS REPOSITIONING, LENS ROTATION; INFLAMMATION; PIGMENT; AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE LENS REMAINS IMPLANTED. THE IMPLANTED LENSES ARE NOT LENSES SOLD BY STAAR JAPAN. THE LENSES WERE PURCHASED BY PHYSICIANS THROUGH PRIVATE IMPORTATION." THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE WAS REPORTED AS "BLINDNESS DUE TO GLAUCOMA RESULTING FROM ICL IMPLANTATION. THE SPECIFIC ETIOLOGY IS UNKNOWN. THE GLAUCOMA PROGRESSED RELATIVELY QUICKLY WITH REPEATED MILD ELEVATIONS IN IOP AND SEDATION WITH MEDICATION." THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN." H6- HEALTH IMPACT-CLINICAL CODE: 4581- PIGMENT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434301&pc=QCB | 2023826-2022-03078 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND BLURRED VISION. "PAIN IN RIGHT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "OCCASIONAL CLOUDINESS RIGHT EYE ;MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: CLINICAL CODE 4581: PIGMENT DISPERSION. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15434626&pc=QCB | 2023826-2022-03079 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; Patient Device Interaction Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EYE PAIN; ANGLE CLOSURE WITH ELEVATED IOP AND GLARE/HALOES. "PAIN IN LEFT SIDE AFTER IMPLANTATION. PUSHING FORWARD ON IRIS. GLAUCOMA SUSPECT POSSIBLE; PHOTOSENSITIVITY. PIGMENTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE ELEVATED IOP; PIGMENT DISPERSION; BLURRED VISION; ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. STATUS OF EYE IS "MGD." CAUSE IS REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6-CLINICAL CODE 4581- PIGMENT DISPERSION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652117&pc=QCB | 2023826-2019-00935 | 2019-05-05 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-29 | QCB | IMPLANTABLE COLLAMER LENS | Optical Problem | Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/154 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THOUGH SURGEON REPORTS THAT THE STATUS OF THE EYE IS "ALL RIGHT," IT IS ALSO NOTED THAT THE PATIENT COMPLAINTS ABOUT HALOS, ESPECIALLY DURING NIGHT DRIVING. REFERENCE MFR REPORT# 2023826-2019-00934 FOR INITIALLY IMPLANTED LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368707&pc=QCB | 2023826-2021-00468 | 2021-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/092, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. # 2023826-2021-00097 BUT THE PROBLEM WAS NOT RESOLVED, THE PATIENT WAS NOT SATISFIED DUE TO HALOS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. EXPLANT DATE: NA. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378632&pc=QCB | 2023826-2021-03367 | 2021-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS HALOS AROUND LIGHT IN THE EVENING AND NIGHT IN THE CENTRAL VISUAL FIELD. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402700&pc=QCB | 2023826-2022-01550 | 2022-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -3.5/+1.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THE PATIENT STILL HAD HALOES. THEN ON (B)(6) 2022 THE LENS WAS EXPLANTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299128&pc=QCB | 2023826-2022-02800 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/3.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE REPORTER STATES THE PATIENT IS HAVING PROBLEMS WITH HALOES. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: UNK. IMPLANT DATE: UNK. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15495834&pc=QCB | 2023826-2022-02684 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 03-AUG-2021. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: LENS WAS EXCHANGED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL FINE! PATIENT IS VERY HAPPY! NO MORE HALOS". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15643334&pc=QCB | 2023826-2022-03625 | 2022-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.00/101 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, LENS OPACITY (CATARACT) WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057494&pc=QCB | 2023826-2022-04732 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- 1494 OFF-LABEL HISTORY OF AMBYLOPIA, BINOCULAR NYSTAGMUS. CLAIM#(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057499&pc=QCB | 2023826-2022-04731 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- OFF-LABEL:1494 SHOULD BE ADDED FOR CORRECTION; HISTORY OF BINOCULAR AMBLYOPIA, BINOCULAR NYSTAGMUS. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178026&pc=QCB | 2023826-2023-00155 | 2022-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON 18-OCT-2022. THE PATIENT EXPERIENCED HALOES AT NIGHT BY CAR DASHBOARD. THE LENS REMAINED IMPLANTED. THE VA HAD IMPROVED AND PATIENT WAS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: RACE: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. DEVICE CODE: (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178121&pc=QCB | 2023826-2023-00156 | 2022-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/073 (SPHERE,CYLINDER,AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AT NIGHT BY CAR DASHBOARD. THE LENS REMAINED IMPLANTED. THE VA HAD IMPROVED AND PATIENT WAS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250870&pc=QCB | 2023826-2023-00339 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.0/1.50/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE PROBLEM WAS REPORTED AS RESOLVED AND NOT RESOLVED; IT WAS INDICATED THE LENS WAS REMOVED DUE TO THE PATIENT'S WISH; REFRACTIVE OUTCOME UA UNCHANGED TO PRE-OP. PATIENT IS CONSIDERING LASER REFRACTIVE SURGERY IN ONE YEAR. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269395&pc=QCB | 2023826-2023-00312 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON 28-(B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269396&pc=QCB | 2023826-2023-00309 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY. "PATIENT REQUEST TO TAKE OUT TICL." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499869&pc=QCB | 2023826-2023-00784 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE, AND HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. REPORTEDLY, "DOING WELL HAPPY WITH VISION.". Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500068&pc=QCB | 2023826-2023-00783 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE, AND HALOS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. REPORTEDLY, "DOING WELL HAPPY WITH VISION." Manufacturer Narrative: PATIENT INFORMATION:UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506867&pc=QCB | 2023826-2023-00628 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/174 ((B)(6) 2023 DUE TO THE PATIENT EXPERIENCED HALOS AND THE PROBLEM WAS RESOLVED. HYALURONIC ACID WAS USED DURING THE EXPLANTATION, WHICH INDUCED A CATARACT. THE PATIENT IS BEING CLOSELY MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506879&pc=QCB | 2023826-2023-00627 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/002 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT EXPERIENCED HALOS. HYALURONIC ACID WAS USED DURING THE EXPLANTATION, WHICH INDUCED A CATARACT. THE PATIENT IS BEING CLOSELY MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922247&pc=QCB | 2023826-2023-01774 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/169 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF HALOS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922381&pc=QCB | 2023826-2023-01773 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF HALOS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253193&pc=QCB | 2023826-2023-02659 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/6.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "HALOS BECAUSE OF THR REMAINING CYLINDER." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. IT WAS ALSO REPORTED, "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/6.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED GLARES/HALOES. REPORTEDLY "HALOES BECAUSE OF THE REMAINING CYLINDER". ON (B)(6) 2022 WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. THE PATIENT IS HAPPY!". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: OFF LABEL- AGE <21. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253194&pc=QCB | 2023826-2023-02658 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/6.0/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "HALOS BECAUSE OF THR REMAINING CYLINDER." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. IT WAS ALSO REPORTED, "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. Manufacturer Narrative: A4-A6:UNK. H3: OFF LABEL- AGE <21. CLAIM# (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT: IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6: MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.50/6.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED GLARES/HALOES AND RESIDUAL ASTIGMATISM. REPORTEDLY "HALOES BECAUSE OF THE REMAINING CYLINDER". ON (B)(6) 2022 WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR IMPLANTED AN IPCL WITH A HIGHER CYLINDER +1/+8/A80. THE PATIENT IS HAPPY!". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# 729917. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088946&pc=QCB | 2023826-2023-04903 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT COMPLAINED OF LIGHT RINGS/HALOS. CYCLOGEL WAS USED. AT POST-OP VISIT THE LIGHT RINGS/HALOS HAD RESOLVED AND THE PATIENT WAS 20/15 AND HAPPY. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18088947&pc=QCB | 2023826-2023-04904 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT COMPLAINED OF LIGHT RINGS/HALOS. CYCLOGEL WAS USED. AT POST-OP VISIT THE LIGHT RINGS/HALOS HAD RESOLVED AND THE PATIENT WAS 20/15 AND HAPPY. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18322977&pc=QCB | 2023826-2023-05582 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18322978&pc=QCB | 2023826-2023-05583 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE DEVICE NOT PERFORMING AS INTENDED. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. PATIENT-RELATED-FACTOR. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337475&pc=QCB | 2023826-2023-05532 | 2023-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD DISLOCATED/SUBLUXED AND PATIENT COMPLAINED OF HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18340659&pc=QCB | 2023826-2023-05533 | 2023-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD DISLOCATED/SUBLUXED AND PATIENT COMPLAINED OF HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387277&pc=QCB | 2023826-2023-05785 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18389150&pc=QCB | 2023826-2023-05784 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING AND GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384508&pc=QCB | 2023826-2022-03045 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL.13.7; -10.0/+1.0/153 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED GLARE/HALOES; WHICH THEN WAS TREATED WITH LUMIFY EYE DROPS WHICH STARTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED AND PATIENT REMAINS ON STUDY. Manufacturer Narrative: EXPIRATION DATE: 11-30-2023. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15384832&pc=QCB | 2023826-2022-03046 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM TMICL.13.7; -11.5/+1.0/025 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED GLARE/HALOES; WHICH THEN WAS TREATED WITH LUMIFY EYE DROPS WHICH STARTED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. LENS REMAINS IMPLANTED AND PATIENT REMAINS ON STUDY. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: EXPIRATION DATE: 11-30-2023. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17680112&pc=QCB | 2023826-2023-03649 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE SURGEON NOTED A RING STRUCTURE ON THE LENS SURFACE AFTER THE LENS WAS IMPLANTED. THE LENS REMAINED IMPLANTED. THE PATIENTS REFRACTIONS AND VISION MEASUREMENTS WERE NORMAL BUT THE PATIENT MENTIONED FEELING UNCOMFORTABLE DUE TO SEEING HALOS AND A GRAY SPOT IN FRONT OF LEFT EYE. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581 - VISUAL DYSPHOTOPSIA. CLAIM#: (B)(4). Manufacturer Narrative: B3: UNK. D6A: UNK. D6B: NA . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519714&pc=QCB | 2023826-2023-00810 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -12.00/2.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2227" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"3001" SHOULD BE ADDED. B5- THE PATIENT ALSO EXPERIENCED GLARE/HALOS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/71 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. REPORTEDLY "THE NEW LENS WAS NEVER INSTALLED AGAIN AND WAS THEREFORE RETURNED. CLIENT CANCELED THE TREATMENT. HE HAD NO CONFIDENCE THAT THE NEW LENS WOULD BE ANY GOOD." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE- "WRONG WTW". CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051835&pc=QCB | 2023826-2023-02126 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF 13.5/3.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT DISCOMFORT/GLARE AND HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6MPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED DISCOMFORT AND GLARES/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FROM ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051836&pc=QCB | 2023826-2023-02127 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Discomfort | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF 15.0/3.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT DISCOMFORT/GLARE AND HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2330" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED DISCOMFORT AND GLARES/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FROM ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831603&pc=QCB | 2023826-2022-00850 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Discomfort; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/1.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: B5- PATIENT EXPERIENCED DISCOMFORT AND HALOS (RESIDUAL REFRACTIVE ERROR)- SHOULD HAVE BEEN INCLUDED ON INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED DRY IN WRAPPED CLOTH WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. DEBRIS THROUGHOUT THE LENS AND ONE HAPTIC IS FOLDED IN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577870&pc=QCB | 2023826-2022-00513 | 2022-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Insufficient Information | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT EXCESSIVE VAULT WAS OBSERVED IN THE PATIENT'S RIGHT EYE (OD).Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -10.00/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 -10.00/+1.5/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 15 JUN 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS WAS EXCHANGED ON 17 MAY 2022 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT IS HAPPY AND ALL FINE. CLAIM (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH - UNK. SEX - UNK. WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE: UNK. MODEL#/ SERIAL#/ EXP DATE/ CATALOG #/ LOT#/ UDI#- UNK. IMPLANT DATE - UNK. EXPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577873&pc=QCB | 2023826-2022-00514 | 2022-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Insufficient Information | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT EXCESSIVE VAULT WAS OBSERVED IN THE PATIENT'S LEFT EYE (OS).Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.50/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 -8.50/+1.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(3) 2021. EXCESSIVE VAULT AND GLARE/HALOES WERE OBSERVED. LENS WAS EXCHANGED ON (B)(3) 2022 FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. ALL FINE AND PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, COMMON DEVICE NAME: PRODUCT CODE: UNK. MANUFACTURER NAME, CITY AND STATE - UNKNOWN, IMPLANT DATE: UNKNOWN. EXPLANT DATE:UNKNOWN, PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830517&pc=QCB | 2023826-2022-03969 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. REPORTEDLY, "VISUAL GHOSTING," AND "TAKE OUT THE CRYSTAL." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4) 4B LOW VALT EXCHANGE FOR SAME SIZE LENS, PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 20-DEC-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC STRETCHED, BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECTS - CLINICAL SIGNS AND SYMPTOMS-2227 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015475&pc=QCB | 2023826-2023-02011 | 2023-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5: LENS WAS REMOVED DUE TO GLARE AND A STRONG APERTURE. REPORTEDLY, "THE POSTOPERATIVE GLARE AND APERTURE WERE STRONG, AND THE PATIENT COULD NOT ADAPT, SO THE CRYSTAL WAS REQUIRED TO BE REMOVED." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107825&pc=QCB | 2023826-2023-04970 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem; No Apparent Adverse Event | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/145(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT IS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/145(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT, AND HALOES WERE OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. IT WAS REPORTED THE HALOS DIMINISHED BECOMING BEARABLE FOR THE PATIENT. THE SURGEON IS WAITING FOR A POSTOP EXAM MAYBE (B)(6) 2024. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND INDICATED LENS TOO BIG. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749935&pc=QCB | 2023826-2020-02521 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749941&pc=QCB | 2023826-2020-02520 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), HEADACHES, INTERMITTENT ANGLE CLOSURE, AND BLURRED VISION. AN ENLARGEMENT OF PI-YAG WAS ALSO PERFORMED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701923&pc=QCB | 2023826-2023-03896 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701924&pc=QCB | 2023826-2023-03895 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2019. REPORTEDLY, "OVER THE PAST 4 YEARS HAS DEVELOPED AN INCREASED VAULT AND NARROWING ANGLES." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS ALSO REPORTED THERE WAS BLURRY VISION AND HEADACHES. H6 - 4581 - NARROW ANGLES, 2137, 1880 CODE SHOULD HAVE BEEN INCLUDED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916646&pc=QCB | 2023826-2023-01725 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, INFLAMMATION, ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), THE PATIENT IMPLANTED THE LENS IN JORDAN BY (B)(6) ON (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR (B)(6) INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO JORDAN FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP.". Manufacturer Narrative: CORRECTED DATA: TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/95 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(6). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916648&pc=QCB | 2023826-2023-01726 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND INFLAMATION WITH ELEVATED IOP WERE OBSERVED, AND STEROIDS WERE PROVIDED. IT WAS ALSO STATED, "AS THE PATIENT WAS TOLD BY HER DOCTOR IN USA: 'PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP."" CAUSE OF THE EVENT WAS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. REPORTEDLY" SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "AS PER THE PATIENT: INFLAMMATION BEING TREATED WITH STEROIDS BY HER CURRENT DOCTOR, (B)(6)," AND "PATIENT SEEMS TO HAVE A NARROW SULCUS, MEASUREMENTS WAS DONE AND REVIEWED SEVERAL TIMES BY (B)(6), TE PATIENT IMPLANTED THE LENS IN (B)(6) BY (B)(6) IN (B)(6) 2022, AND IN (B)(6) 2022 THE PATIENT CONTACTED HIM INFORMING HIM THAT SHE DID AN EXAMINATION IN USA AND HER DOCTOR SAMER MULKI INFORMED THAT SHE HAS AN EXCESSIVE VAULT. AS THE PATIENT IS IN THE UNITED STATES AND COULD NOT COME TO (B)(6) FOR REPLACEMENT SURGERY, SO SHE REQUESTED THAT SHE REPLACES THE ICL IN THE STATES, AS PER HER CURRENT DOCTOR, HE ADVISED THAT HE EXPLANTS AND THEN MONITOR THE PATIENT FOR 3 MONTHS AND THEN WE PROCEED ACCORDINGLY AS SHE SEEMS TO HAVE ELEVATED IOP." Manufacturer Narrative: A4-A6:UNK H6: PATIENT RELATED FACTOR. CLAIM#(B)(4). Manufacturer Narrative: H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6- HEALTH EFFECT- CLINICAL CODE: "1880" AND "2137" SHOULD BE ADDED. ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/78 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, INFLAMMATION, HEADACHE, BLURRED VISION AND MEDICATION WAS PRESCRIBED. REPORTEDLY "PERSISTENT INFLAMMATION THAT IS CAUSING HEADACHE AND BLURRING OF VISION WHICH MANDATES CONTINUOUS TREATMENT WITH TOPICAL STEROIDS THAT HAS LET TO PERSISTENT INCREASED IOP". THE LENS WAS EXPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "HE DOESN'T INTEND TO IMPLANT NEW LENSES ANY TIME IN THE IMMEDIATE FUTURE ". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS A PATIENT RELATED FACTOR AND UNKNOWN. - "SEEMS THAT THE PATIENT HAS NARROW SULCUS. KINDLY NOTE THAT IOP ELEVATION IS A REACTION TO HER CURRENT STEROID MEDICATION TO FIGHT THE INFLAMMATION SHE IS HAVING FROM THE OVERSIZING OF THE LENS". CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781744&pc=QCB | 2023826-2020-00390 | NA | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -12.0/+2.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO MORE HEADACHES. COMPLIANT WITH DROPS." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200458&pc=QCB | 2023826-2021-00183 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.00/1.5/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED PAIN, HEADACHE, LIGHT SENSITIVITY, AND BLURRY VISION. PUPIL BLOCK WITH ELEVATED IOP WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (YAG PI) AND MEDICAL INTERVENTION (TROPICAMIDE, PHENYLEPHRINE, COSOPT PF, DIAMOX, BROMINIDINE) WERE PERFORMED/PRESCRIBED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, STATUS OF EYE "NORMAL". IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: B3 - (B)(6) 2020. B5 - EVALUATION ON 19-DEC-2020: EXAMINATION DEMONSTRATED CORNEAL EDEMA, A MID-DILATED PUPIL, ELEVATED IOP, ELEVATED ICL VAULT AND IOP. YAG PERFORMED. AQUEOUS TAP DONE (B)(6) 2020. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145766&pc=QCB | 2023826-2022-01205 | NA | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Eye Pain | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -11.5/+1.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 A PI-YAG WAS PERFORMED, THE SURGEON REPORTING TRANSIENT LOSS OF BCVA, PERIORBITAL EYE PAIN, HEADACHE, ELEVATED IOP, AND POSSIBLE INTERMITTENT BOUTS OF PUPILLARY BLOCK PRIOR TO LASER PI. THE LENS REMAINS IMPLANTED. REPORTEDLY THE TRANSIENT VISION LOSS WAS RESOLVED WITH THE LASER PI; LATENT IOP IS 12MMHG. Manufacturer Narrative: B5 LENS WAS EXPLANTED ON (B)(6) 2023. STATUS OF EYE IS REPORTED AS "PATIENT DEVELOPED LEFT EYE SECONDARY GLAUCOMA POST ICL IMPLANTATION ASSOCIATED WITH HEADACHE. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. "PATIENT COMPLAINTED HEADACHE ASSOCIATED WITH EYE PAIN." CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, THIS CLAIM IS FOR THE RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070995&pc=QCB | 2023826-2022-02369 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070996&pc=QCB | 2023826-2022-02370 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Headache; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND COMPLAINED OF HEADACHES; AND EXAMINATION SHOWED NO ABNORMALITIES. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986875&pc=QCB | 2023826-2023-04601 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSTATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14649917&pc=QCB | 2023826-2022-01840 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Inflammation; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -8.00/+1.50/76 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD INFLAMMATION BEHIND THE ICL; VA POSTOP WAS 6/6 BUT BLURRY; AND IS CURRENTLY ON STEROIDS. PATIENT IS STILL BEING MONITORED. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8642522&pc=QCB | 2023826-2019-00911 | NA | Injury | STAAR SURGICAL COMPANY | 2019-05-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Inflammation; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC IMPLANTABLE COLLAMER LENS, -13.5/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), WITHOUT ANY COMPLICATIONS. ELEVEN MONTHS AFTER THE LENS WAS IMPLANTED, THE PATIENT EXPERIENCED BLURRED VISION AFTER BLUNT OCULAR TRAUMA. EXAMINATION REVEALED MILD INFLAMMATION IN THE ANTERIOR CHAMBER AND BOTH SUPERIOR AND INFERIOR TEMPORAL HAPTICS OF THE ICL WERE DISLODGED INTO THE ANTERIOR CHAMBER AND ENTRAPPED WITHIN THE PUPIL, WITHOUT TOUCHING THE CORNEAL ENDOTHELIUM. THE LENS WAS REPOSITIONED AND THE PATIENT WAS ON MEDICATION. ONE MONTH AFTER THE REPOSITIONING SURGERY, THE BLURRED VISION WAS RESOLVED AND THERE WAS NO DAMAGE TO THE EYE. TWENTY SIX MONTHS LATER, THE PATIENT AGAIN EXPERIENCED BLUNT OCULAR TRAUMA TO THE RIGHT EYE. THE LENS WAS REPOSITIONED AND ONE MONTH AFTER THERE WERE NO COMPLICATIONS ASSOCIATED TO THE REPOSITIONING SURGERY AND NO TRAUMA OBSERVED. PIGMENT DEPOSITION ON THE POSTERIOR SURFACE OF THE ICL WAS OBSERVED AFTER SURGERY BUT THE AMOUNT OF THE PIGMENT HAD NOT CHANGED ONE YEAR AFTER THE INJURY. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: EXCESSIVE VAULTING SHOULD BE ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173237&pc=QCB | 2023826-2019-01901 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION AND TASS (TOXIC ANTERIOR SEGMENT SYNDROME). THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE FOLLOWING MEDICATIONS WERE PRESCRIBED FOR ABOUT 10 DAYS AFTER SURGERY - RINDERON OPHTHALMIC SOLUTION, MYDRIN-P OPHTHALMIC SOLUTION AND PREDNISOLONE. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER 10 DAYS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549694&pc=QCB | 2023826-2023-03497 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Inflammation; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A,12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -18.0/1.00/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, AND BLURRED VISION. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "UP TO DATE THE EYE CONTINUES WITH THE INFLAMMATORY PROCESS." CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER STATED, "DUE TO A PATIENT INFLAMMATORY RESPONSE." Manufacturer Narrative: CLAIM# (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763823&pc=QCB | 2023826-2021-04230 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Iritis; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.5/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, INFLAMMATION AND IRITIS. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT INITIALLY THE DOCTOR WANTED TO EXPLANT THE LENS BUT HAS DECIDED TO WAIT AWHILE. MEDICAL INTERVENTION (PRED FORTE, PILOCARPINE, PREDNISONE) WAS PRESCRIBED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- CORRECTION: -11.50/2.5/100 SHOULD BE CORRECTED TO -11.50/2.5/108 IN THE ORIGINAL MEDWATCH REPORT. CLAIM: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465648&pc=QCB | 2023826-2022-03247 | NA | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, UNEACTIVE (FIXED) PUPIL, ANTERIOR SEGMENT INFLAMMATION, AND GLARE/HALOES WAS OBSERVED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "MILD CELL/FLARE. 6.0MM UNREACTIVE PUPIL." UNKNOWN WAS REPORTED FOR CAUSE OF EVENT AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5: PER UPDATED INFORMATION PATIENT IS SEEING WELL BUT HAS A DILATED PUPIL AND HAS BEEN USING STEROID EYE DROPS TO CONTROL INFLAMMATION. INFLAMMATION HAS BEEN IMPROVING. SURGEON WILL BE TAPERING STEROID EYE DROPS WITH NO PLANS OF SURGICAL INTERVENTION AT THE MOMENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308771&pc=QCB | 2023826-2021-00318 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (32MMHG) OBSERVED ON (B)(6) 2020, REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -15.00/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION; REFRACTIVE SURPRISE; PUPIL BLOCK; ELEVATED IOP; AND ANGLE CLOSURE WERE REPORTED. THE LENS WAS EXPLANTED AND LATER REPLACED FOR A LONGER LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. THE REPORTER ALSO NOTES "STABLE ICL VAULT, IOP, VISION, TEMPORAL SUTURE IN SITU." CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11787674&pc=QCB | 2023826-2021-01388 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ELEVATED IOP. LENS REMAINS IMPLANTED. FOR STATUS OF THE EYE " PATIENT CANNOT READ ANYMORE " WAS REPORTED. FOR ADDITIONAL INFORMATION THE REPORTER REPORTED " HIGH IOP OS: PROBLEM RESOLVED WITH DIAMOX AND PRESSURE WAS AFTER USING DIAMOX 20 MM HG. PROBABLY SOME VISCO REMAINED IN THE EYE. BUT THE PUPIL IN THIS EYE IS STILL ASYMMETRIC? THEY DONT KNOW WHAT IS THE CAUSE OF THIS REFRACTIVE SURPRISE. DOCTOR WANTS TO DO A LASER TO SOLVE THE REFRACTIVE SURPRISE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14507171&pc=QCB | 2023826-2022-01692 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/2.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ONE HAPTIC WAS NOTED TO BE TORN AFTER IMPLANTATION. PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, SUBLUXATION AND LESS VISION. MANAGEMENT WAS GIVEN FOR ELEVATED INTRAOCULAR PRESSURE. LENS REMAINS IMPLANTED. FOR ADDITIONAL INFORMATION THE REPORTER STATED, "CORNEA CLEAR, VAULT- TEMPORARILY SUBLAXATED DUE TO LESS SUPPORT (HAPTIC TEAR)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED. EVERYTHING WENT WELL, PATIENT AND THE SURGEON ARE HAPPY WITH RESULTS. VA 6/9, VAULT OK, IOP 13MMHG. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962122&pc=QCB | 2023826-2023-04570 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Intraocular Pressure Increased; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/6.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING KERATOCONUS ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE IS "STABLE." CAUSE OF THE EVENT IS A PATIENT-RELATED-FACTOR WITH THE LENS NOT PERFORMING AS INTENDED. REPORTEDLY, "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/6.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND ELEVATED IOP. ON (B)(6) 2023, THE LENS WAS EXPLANTED BUT THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "STABLE". REPORTEDLY "THE ICL WAS REMOVED BECAUSE IT WAS TO BIG FOR THE EYE OF THE PATIENT SO THE EYE PRESSURE INCREASED". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "PATIENT HAS KERATOCONUS WHICH MAKES THE RESULT UNPREDICTABLE". H6: HEALTH EFFECT- CLINICAL CODE: "1937 SHOULD BE ADDED. CLAIM#: (B)(6). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL/PATIENT RELATED FACTOR: PRE-EXISTING KERATOCONUS. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10875424&pc=QCB | 2023826-2020-02715 | 2020-07-18 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A12.6MM, VTICMO12.6, -16.0/+1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO "TILTED ICL", ELEVATED IOP, PUPIL BLOCK, LENS DISLOCATION/SUBLUXATION AND BLURRED VISION. ADDITION OF NEW PI WAS PERFORMED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2020. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11070811&pc=QCB | 2023826-2020-03177 | 2020-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND IOP (INTRAOCULAR PRESSURE), UNRESPONSIVE TO MAXIMUM MEDICAL THERAPY. DOCTOR EXPLANTED THE LENS AS HE COULD NOT REDUCE THE IOP ADEQUATELY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -15.50/1,0/118, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA AND ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H3- DEVICE EVALUATION : LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN A SPECIMEN CUP. VISULA INSPECTION FOUND ONE OF THE HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308736&pc=QCB | 2023826-2021-00317 | NA | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT "STABLE VAULT, STABLE IOP, CLEAR CORNEA. PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BENT AND DEFORMED WITH RESIDUE AND DEBRIS ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN ADDITIONAL OF NEW PI, YAG AND THE LENS WAS EXCHANGED WITH A FOR A LONGER LENGTH LENS, DUE TO PUPIL BLOCK WITH ELEVATED IOP (32MMHG), LENS ROTATION AND ANGLE CLOSURE WITH ELEVATED IOP ((32MMHG) OBSERVED ON (B)(6) 2020), REFRACTIVE SURPRISE, AND LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, "SIZING DIFFICULT, MAY BE BECAUSE OF ASIAN EYE, OR ABNORMAL SULCUS. REPORTEDLY, PATIENT HAS "STABLE VAULT, STABLE IOP, CLEAR CORNEA, AND TEMPORAL SUTURE IN SITU." PER THE REPORTER ON (B)(6) 2021, "PATIENT HAS SUCCESSFULLY TAKEN THE REPLACEMENT LENS AND IS VERY HAPPY WITH THE RESULTS WITH NO FURTHER PROBLEMS REPORTED." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# : (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366547&pc=QCB | 2023826-2021-00485 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.5/+1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PERIPHERAL ANTERIOR SYNECHIA (PAS), AND A "LITTLE" MYDRIASIS. AN ADDITION OF PI-YAG WAS PERFORMED, PILOCARPINE EYE DROPS ADMINISTERED, IRRIGATION AND ASPIRATION (IA) PERFORMED. THE PAS WAS RELEASED BY PEELING OFF THE ADHESIONS OF THE ANGLES BY A ROUND NEEDLE. THE PROBLEM IS NOTED AS RESOLVED THOUGH THE PUPIL IS REPORTED AS SLIGHTLY DILATED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939022&pc=QCB | 2023826-2021-01669 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.00/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP(32MMHG, ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: ELEVATED IOP AND BLURRINESS WAS NOTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939023&pc=QCB | 2023826-2021-01670 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL3.2, -16.50/2.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP (25MMHG, ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: BLURRY VISION AND ANGLE CLOSURE OBSERVED. H10: CODE 4581 - ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222320&pc=QCB | 2023826-2021-02802 | 2021-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SECTORAL IRIS ATROPY. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THE PROBLEM IS NOT RESOLVED THOUGH THE LENS HAS BEEN EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS . CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/3.50/95 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ELEVATED IOP WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH LENS. "THE SPECIALIST REPORTS THAT THE AGREEMENT IS FINE, WITH A LITTLE BLURRED VISION IN THE IMMEDIATE POST-OPERATIVE CONTROL, IT HAS OPEN ANGLES OF 2 PER GONIOSCOPY, VAULT OF 700 MICRONS, STABLE IOP OF 10/10 MMHG. THE PATIENT FINDS HIMSELF BETTER WITH THE EXCHANGE, SURGERY WITHOUT COMPLICATIONS." H6- HEALTH EFFCET CLINICAL CODE; 4581-OPEN ANGLES OF 2 PER GONIOSCOPY. CLAIM# (B)(4) Manufacturer Narrative: WEIGHT, ETHNICITY, RACE -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874494&pc=QCB | 2023826-2021-04515 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -14.5/2.5/62, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO ANGLE CLOSURE WITH ELEVATED IOP. IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR STATUS OF THE EYE "BLURRY VISION" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B3: 05-OCT-2021. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061327&pc=QCB | 2023826-2022-01130 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -13.0/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO BLURRED VISION, IRIS SPHINCTER LOSS AT 5 O'CLOCK, HEAVILY PIGMENTS TM, KIV EXPLANT, AND IRREGULAR PUPIL. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- ELEVATED IOP SHOULD HAVE BEEN REPORTED AS AN ADVERSE EVENT IN INITIAL MDR. H6- INTRAOCULAR PRESSURE INCREASED . CLAIM# (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15221446&pc=QCB | 2023826-2022-02760 | 2020-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15556104&pc=QCB | 2023826-2022-03511 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.00/6.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HIGH INTRAOCULAR PRESSURE WITH ANGLE CLOSURE AND CHANGE IN REFRACTION. CLOUDY CORNEA WAS ALSO REPORTED, HOWEVER THE REPORTER STATED, " REGARDING THE SYMPTOMS RELATED TO THE EXPLANTED LENS, EVERYTHING RESOLVED." IN THE REPORTER'S OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600761&pc=QCB | 2023826-2023-03684 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -3.00/+3.00/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17608279&pc=QCB | 2023826-2023-03689 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.00/+5.50/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT), KERATOCONUS. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17701985&pc=QCB | 2023826-2023-03775 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.00/0.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED WITH BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. PATIENT WAS LOST TO FOLLOW-UP. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17705120&pc=QCB | 2023826-2023-03951 | 2021-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED BLURRED VISION, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE - 4581 - PUPIL ANOMALY. CLAIM#: (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885010&pc=QCB | 2023826-2023-04315 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS HAD A LOW VAULT. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION AND BLURRED VISION. OCT PERFORMED ON (B)(6) 2023 REVEALED RNFL DAMAGE AND TEMPORAL SCOTOMA. MEDICATION WAS PRESCRIBED. THE LENS REMAINED IMPLANTED. POST-OP, GOOD IOP DURING 6 HOURS AFTER DROPS. PROSTAGLANDIN INHIBITOR WAS ADDED AT NIGHT. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, RETINAL NERVE DAMAGE, TEMPORAL SCOTOMA H6: HEALTH EFFECT IMPACT CODE: 4644 - HYPOTENSIVE DROPS - S. COSOPT (DORZOLAMIDE/TIMOLOL) X2, S. LUXFEN (BIMATOPROST) X2 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+0.5/69 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, ELEVATED IOP, PIGMENT DISPERSION, OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS, FOGGY VISION, BLURRED VISION, LASER YAG PIS PERFORMED AND MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6-HEALTH EFFECT- CLINICAL CODE: 4581- "FOGGY VISION" AND "OPTIC NERVE DAMAGE WITH VISUAL FIELD LOSS" SHOULD BE ADDED. H6-HEALTH IMPACT CODE: 4625- "'LASER YAG PI' WAS PERFORMED" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. CLAIM# (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525781&pc=QCB | 2023826-2024-00170 | 2023-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6, -11.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ANTERIOR CHAMBER IRRIGATION/VACUATION OF VISCO/FLUIDS AND ANTERIOR CHAMBER WASH ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED IOP(40MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(DATE ASSESSED (B)(6) 2023 & (B)(6) 2023), PIGMENT DISPERSION, AND BLURRED VISION. THIS DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618161&pc=QCB | 2023826-2019-00798 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 15.5/+2.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ICL WAS TOO BIG. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9282596&pc=QCB | 2023826-2019-02091 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.5/1.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND DEBRIS ON THE LENS BUT NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558794&pc=QCB | 2023826-2021-00700 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -15.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS IOP CONTROLLED WITH ALPHAGAM, HIGH VAULT, SHALLOW AC, "CRAPPY" VISION AT 20/200. PUPIL IS DILATED AND AR IS +2.75. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, HYPEROPIC SURPRISE +2.00, GLARE/HALOS, BLURRED VISION. REPORTEDLY VISION IS NOW "OK" AT +1.00 HYPEROPIC. PILOCARPINE USED TO CONSTRICT PUPIL AND SET THE ICL IN THE SULCUS. H6-CLINICAL CODE 4581 (SIGNIFICANT REDUCTION OF ICA, PIGMENT DISPERSION, UNREACTIVE/FIXED PUPIL). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664235&pc=QCB | 2023826-2022-00658 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. ADDITIONAL SYMPTOMS REPORTED: ELEVATED IOP, GLARE/HALOS, PIGMENT DISPERSION, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. TRACE PIGMENT IS PRESENT ON THE ENDOTHELIUM AT 6 O'CLOCK AFTER EXCHANGE. HOWEVER, THE EVENT IS UNLIKELY TO BE CAUSED BY THE DEVICE AND NO INTERVENTION IS PLANNED. AN ENLARGEMENT OF PI-YAG WAS PERFORMED AFTER LENS EXPLANT AND PRIOR TO REPLACEMENT LENS IMPLANT. H6-CLINICAL CODE 4581 (PIGMENT DISPERSION). CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. IMPLANT DATE, EXPLANT DATE: UNK. CLAIM# (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=QCB | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4) | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112743&pc=QCB | 2023826-2019-01626 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION, GLARE/HALOES, AND BLURRED VISION. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES. ADDITIONAL INFORMATION: PRE-EXISTING CORNEAL OR REFRACTIVE SURGERY. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140318&pc=QCB | 2023826-2021-02397 | 2021-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. GOOD POSITION AND VAULT, SYMPTOMS (GLARE) RESOLVED AND IOP NORMAL. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416136&pc=QCB | 2023826-2021-00506 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.50/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP(44MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, GLARE/HALOES, BLURRED VISION, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938959&pc=QCB | 2023825-2021-01652 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -14.50/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP(50MMHG), ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, IRIS ATOROPHY, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, "UNFORTUNATELY DURING THE ICL EXCHANGE BECAUSE OF THE LIMITED DISTANCE BETWEEN THE ICL AND THE ENDOTHELIUM SHE SUSTAINED A 1MM TEAR IN DESCEMET'S MEMBRANE. AS OF NOW WE ARE WAITING TO SEE IF HER CORNEA WILL CONTINUE TO CLEAR. WHEN THE CORNEA CLEARS ENOUGH WE INTEND TO ASSESS THE STATUS OF THE ENDOTHELIUM TO SEE HOW MUCH DAMAGE SHE HAS SUSTAINED BY THE THREE PROCEDURES ON THIS EYE. CURRENTLY HER ANGLES ARE ADEQUATE IN BOTH EYES AND THE IOP IS ACCEPTABLE." Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US, (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14809062&pc=QCB | 2023826-2022-01953 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/150/66 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PATIENT HAS EXPERIENCED EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOES AND BLURRED VISION. FOR PATIENTS CURRENT CONDITION THE REPORTER STATED " IRIS PERIPHERAL ANTERIOR SYNACHIAE WHICH REQUIRES REMOVAL OF CURRENT LENS SIZING OF 13.2MM TO 12.6MM SIZING. THE INITIAL LENS IS STILL IN PATIENT'S EYES." FOR ADDITIONAL INFORMATION THE REPORTER STATED " PATIENT HAS IRIS ATROPHY DUE TO HIGH LENS VAULTING BUT DR HAS BEEN MONITORING THIS CASE FOR 2 YEARS. NEEDS A SECONDARY INTERVENTION AS WELL DUE TO IRIS DISTORTION AND CAUSING PATIENT TO HAVE VISUAL DISTURBANCE ESPECIALLY AT NIGHTS." IN THE REPORTS OPINION THE CAUSE OF THIS EVENT IS USER ERROR. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON 28-JUN-2022 AND REPLACED WITH A SHORTER LENGTH LENS ON THE SAME DATE BUT DIFFERENT SURGERY AND THE PROBLEM RESOLVED. THE REPORTER STATED "ANTERIOR SYNECHIAE AND IRIS ATROPHY STILL PERSIST BUT VAULTING REDUCED AND EYE LOOKED BETTER, LESS GLARE." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#:(B)(6). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17060362&pc=QCB | 2023826-2023-02031 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO TRANSIENT LOSS OF BCVA (ASSESSED ON (B)(6)2023, ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2023). PIGMENT DISPERSION. GLARE, HALOES, AND BLURRED VISION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PARACENTESIS DONE AND DIAMOX GIVEN TO REDUCE IOP. PATIENT DOES NOT WANT NEW LENS IMPLANTED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6: HEALTH EFFECT CLINICAL CODE 4581: PIGMENT DISPERSION. H6:WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT - IMPACT CODE: "4625- PARACENTESIS DONE." B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-13.5/+1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED TRANSIENT LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, DILATED PUPIL, MEDICATION WAS PRESCRIBED. REPORTEDLY "PARACENTESIS DONE AND GAVE DIAMOX TO REDUCE IOP." ON (B)(6) 2023 THE LENS WAS EXPLANTED. IT WAS ALSO REPORTED, "IOP REDUCED AND PUPIL STILL DILATE AFTER REMOVAL. PATIENT NOT KEEN TO IMPLANT ANOTHER ICL ANYMORE." THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393935&pc=QCB | 2023826-2015-00168 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/+3.5/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS HAS AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF THE IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, GLARE/HALOS, BLURRED VISION AND IRIS ATROPHY. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. EXPLANTED DATE UNKNOWN. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED AND OBSERVED NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). RESUBMISSION OF INITIAL MDR REQUESTED BY FDA. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528011&pc=QCB | 2023826-2019-02495 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 6.50/+2.0/068 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH? NEEDS SMALLER LENS?) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A CONTACT CASE, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS H6 - PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECTED TO (B)(6) 2019 IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - PER UPDATED INFORMATION, THE "PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT YET FULLY RESOLVED." THE PATIENT STATES "VISION HAS IMPROVED SINCE PROCEDURE OVERALL." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528024&pc=QCB | 2023826-2019-02494 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 7.00/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH INTRAOCULAR PRESSURE (IOP) AND EXCESSIVE VAULTING. THERE WAS NO PATIENT INJURY. THE LENS WAS NOT REPLACED WITH ANOTHER LENS AT THIS TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. A YAG WAS PERFORMED TO ENLARGE THE PIS. CAUSE OF THE EVENT WAS UNKNOWN/OTHER (PI'S WIDE ENOUGH NEEDS SMALLER LENS) . THE PATIENT'S POST-OP CONDITION WAS SOME PIGMENT DISPERSION. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT CASE, WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, YAG PERFORMED TO ENLARGE PIS. PATIENT CODE: 3191 - PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: IT WAS LATER INDICATED BY THE REPORTER THAT AS OF 25/FEB/2020 "THE PATIENT'S PIGMENT DISPERSION HAS IMPROVED BUT HAS NOT FULLY HEALED. THE PATIENT'S VISION HAS IMPROVED SINCE THE PROCEDURE OVERALL THOUGH" CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140634&pc=QCB | 2023826-2020-01200 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, CONTINUOUS GRAVITATIONAL PAIN-FEELING SOMETHING PUSHING ON THE EYE, HALOS, CIRCLES, AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278375&pc=QCB | 2023826-2021-03031 | NA | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS OTHER - ICL LENS LENGTH CALCULATION ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS EXCHANGED ON (B)(6) 2021 WITH A DIFFERENT MODEL, SIMILAR POWER, SHORTER LENGTH LENS WHICH RESOLVED THE ISSUE. REPORTEDLY, DOCTOR AND PATIENT HAPPY WITH THE FINAL RESULT. THE REPORTER MARKED THE CAUSE OF THE EVENT AS PATIENT-RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 721078. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17970222&pc=QCB | 2023826-2023-04631 | NA | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) -2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE AND PIGMENT DISPERSION. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED (VAULT). Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE, PIGMENT DISPERSION. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, PIGMENT DISPERSION, PAIN, BLURRED VISION AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "IOP NORMAL WITH SINGLE DRUG GLAUCOMA THERAPY". ADDITIONAL INFORMATION PROVIDED - "WAITING FOR IOP TO LOWER FURTHER. THE ANGLES ARE OPEN NOW. PIGMENT DISPERSION STILL PRESENT". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "RAISED IOP DUE TO ANGLE CLOSER AND HIGH VAULT AND PIGMENT DISPERSION". H6: HEALTH EFFECT- CLINICAL CODE: "2137" AND "1994" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "4644" SHOULD BE ADDED. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13240884&pc=QCB | 2023826-2022-00156 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL_12.6; -10.50/+1.00/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE; ACUTE ANGLE CLOSURE SECONDARY TO PUPILLARY BLOCK FOLLOWING PHARMACOLOGIC MYDRIASIS; PAIN; DECREASED VISUAL ACUITY; NAUSEA; PUPIL DILATION. IT WAS REPORTED THAT "SUBJECT TREATED WITH ORAL ACETAZOLAMIDE AND TOPICAL TIMOLOL; PILOCARPINE; LATANOPROST AND BRINZOLAMIDE OVER THE COURSE OF APPROXIMATELY 2 HOURS. SUBJECT DISCHARGED WITH INSTRUCTIONS TO USE TIMOLOL; BRIMONIDINE; LATANOPROST AND ACETAZOLAMIDE AT HOME AND RETURN FOR FOLLOW-UP EXAM THE NEXT DAY". POST OP NOTES INDICATED THAT "THE SUBJECT RETURNED FEELING MUCH BETTER; WITHOUT PAIN OR HEADACHE. INTRAOCULAR PRESSURE MEASURED 8 MMHG; AND THE CORNEA WAS CLEAR OS. UNCORRECTED VISUAL ACUITY IMPROVED TO 20/16 OS. BRIMONIDINE WAS TAPERED; AND ALL OTHER MEDICATIONS DISCONTINUED". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS NOTED AS THE DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B5: "ON (B)(6) 2022 LENS OPACITY (CATARACT) ADC WAS OBSERVED. ANTERIOR SUBCAPSULAR OPACITY NOTED AS GRADE1/MILD. NO ACTION TAKEN, PATIENT REMAINS ON STUDY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14505054&pc=QCB | 2023826-2022-01687 | 2021-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -11.0/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; REFRACTIVE SURPRISE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; PUPIL BLOCK; ANGLE CLOSURE; ELEVATED INTRAOCULAR PRESSURE; UNREACTIVE (FIXED) PUPIL; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ON (B)(6) 2021 A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION-LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794474&pc=QCB | 2023826-2022-01950 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794499&pc=QCB | 2023826-2022-01951 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO TRANSIENT LOSS OF BCVA, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE (FIXED) PUPIL, AND IRIS ATROPHY. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE, THE REPORTER STATED "THE DOCTOR THINKS THAT THE PATIENT'S MYOPIA PROBLEM HAS NOT BEEN SOLVE." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440465&pc=QCB | 2023826-2024-00046 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OP DUE TO EXCESSIVE VAULT; REFRACTIVE SURPRISE; PUPIL BLOCK WITH ELEVATED IOP; AND ANGLE CLOSURE WITH ELEVATED IOP. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS DEVICE. SEE MFR# 2023826-2024-00047 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-CLINICAL CODE 4581: ANGLE CLOSURE WITH ELEVATED IOP. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10956516&pc=QCB | 2023826-2020-02851 | 2020-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON (B)(6) 2020 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -8.0/2.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON 06 NOV 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP, AND BLURRED VISION. ON 19 NOV 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. ON (B)(6) 2020 (LAST VISIT) PATIENTS CURRENT CONDITION WAS BCVA 20/40, UCVA 20/70 AND IOP WAS 10MMHG. "BLURRY VISION AND MILD-MOD IRRITATION" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). HOWEVER, SIGNS/SYMPTOMS WERE NOT BEYOND THAT OF TYPICAL SURGERY. IT WAS REPORTED THAT IRRITATION AND BLURRY VISION RESOLVED. PATIENT IS NOW 20/20. H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND RESIDUE ON THE LENS SURFACE. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9039904&pc=QCB | 2023826-2019-01703 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Retinal Detachment; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICMO, -18.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WITH ROTATION, ELEVATED IOP, REFRACTIVE SURPRISE, AND RETINA DETACHMENT WAS OBSERVED. SURGICAL INTERVENTION (REPOSITIONING, RETINAL REPAIR AND GAS INFUSION) AND MEDICAL INTERVENTION (ANTIBIOTICS, STEROIDS, GLAUCOMA MEDICATION) WERE PERFORMED/ PRESCRIBED. LENS CURRENT STATUS WAS REPORTED TO BE IN POSITION BUT VAULT WAS APPROXIMATELY 100 MICRONS. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED VIA EMAIL. THE LENS WAS REPOSITIONED A SECOND TIME ON (B)(6) 2019. "PATIENT HAS GAS LEAKAGE IN THE EYE. SO AGAIN RETINA SURGEON HAS GAS FILLED UP AND THEY WILL WAIT FOR FOUR WEEK FOR VISUAL RECOVERY UPDATE." CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT THE LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE ON (B)(6) 2019. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378904&pc=QCB | 2023826-2023-05715 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.50/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, BLURRED VISOIN AND PROLONGED POSTOPERATIVE ANTERIOR UVEITIS. IT WAS REPORTED THAT THE IOP WAS CONTROLLED WITH MEDICATION. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112820&pc=QCB | 2023826-2019-01627 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -8.00/1.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP, LENS DISLOCATION/SUBLUXATION AND GLARE/HALOES. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED, EYE IS QUIET. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS" INJECTOR MODEL-MSI-TF; LOT#-UNK SHOULD HAVE BEEN INCLUDED. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: NARROW ANTERIOR CHAMBER ANGLES SHOULD HAVE BEEN INCLUDED. PRE-EXISTING CONDITION: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980474&pc=QCB | 2023826-2020-02945 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -12.5/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS LEFT (OS) EYE. THE PATIENT REPORTS HALOS, DOUBLE VISION, LIGHT RINGS, POSTOP PAIN, PRESSURE, AND VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. POSTOP PAIN AND PRESSURE RELIEVED BY "PUSHING ONTO THE CUT TO FORCE FLUID OUT." Manufacturer Narrative: B5: ADDITIONAL INFORMATION: 6HRS POST-OP AC FLUID DRAINED VIA EXISTING PARACENTESIS. DIAMOX ADMINISTERED. GHOSTING OF VISION IN LOW LIGHT. ELEVATED IOP DUE TO RETAINED OCCUCOAT. POST-OP HALO/GLARE IS EXPECTED DUE TO SCOTOPIC PUPIL SIZE AND RELATIVELY HIGHT MYOPIC PRESCRIPTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980487&pc=QCB | 2023826-2020-02946 | NA | Malfunction | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: REDDIT COMPLAINT: THE PATIENT HAD A 12.6MM VTICM5_12.6 IMPLANTABLE LENS, -13.0/+2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO HIS RIGHT (OD) EYE. THE PATIENT REPORTS HALOS, LIGHT RINGS, AND POSTOP VOMITING. REPORTEDLY, THE VOMITING WAS POSSIBLY CAUSED BY THE PATIENT TAKING XANAX PRIOR TO SURGERY. Manufacturer Narrative: B5-ADDITIONAL INFO: ELEVATED IOP 6HRS POST-OP DUE TO RETAINED OCCUCOAT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: CLINICAL CODE: 4581, FOR "GHOSTING VISION". 2227, HALO IS NOT APPLICABLE. DEVICE PROBLEM CODE: 3001, GLARE/HALO IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671674&pc=QCB | 2023826-2019-01020 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: REFERENCE MFR REPORT# 2023826-2019-00886 FOR INITIAL LENS IMPLANT. THE SURGEON EXCHANGED THE INITIAL IMPLANT WITH A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/013 (SPHERE/CYLINDER/AXIS). THE SURGEON REPORTS LENS ROTATION, HALOS, STARBURST, "COMA DOUBLE IMAGE" AND PATIENT IS ON "SYMBIZA+MONOPROST". Manufacturer Narrative: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT DUE TO MEDICATION INTERVENTION. SURGEON REPORTED PER LAST VISIT DATE ON (B)(6) 2019, PATIENT STILL IS EXPERIENCING MILD HALOS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: MALFUNCTION SHOULD BE CORRECTED TO SERIOUS INJURY IN INITIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 1937 SHOULD BE CORRECTED TO PATIENT CODE: 1937, THE PATIENT'S IOP IS MAINTAINED ON TWO IOP LOWERING MEDICATIONS: SIMBRINZA (BRINZOLAMIDE/BRIMONIDINE TARTRATE OPHTHALMIC SUSPENSION) AND MONOPROST. THE ICL REMAINS IMPLANTED AS OF THIS TIME. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140593&pc=QCB | 2023826-2020-01201 | NA | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT REPORTS ELEVATED IOP WHICH HAS BEEN RESOLVED BUT PAIN PERSISTS. ALSO REPORTED: FIXED MYDRIASIS, NO RESPONSE TO LIGHT STIMULI, LASER IRIDOTOMY, PILOCARPINE ADMINISTRATION, HALOS AND CIRCLES. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6541708&pc=QCB | 2023826-2016-01557 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.0/97 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: PMA/510(K): NA, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6549621&pc=QCB | 2023826-2016-01558 | 2014-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-05-05 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.0/109 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT HAS EXPERIENCED EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED IOP AND REFRACTIVE CHANGE OVER TIME. PATIENT LIVES OUTSIDE OF COUNTRY AND WILL ARRANGE FOLLOW-UP WITH SURGEON. Manufacturer Narrative: NA- THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552490&pc=QCB | 2023826-2019-00648 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -15.5/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE VA OF THE EYE AFTER EXPLANT IS REPORTED AS "FOGGY, SMEARED." Manufacturer Narrative: PATIENT CODE 2140: VISUAL DISTURBANCE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030329&pc=QCB | 2023826-2020-01045 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE." SHOULD BE REMOVED FROM SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10030336&pc=QCB | 2023826-2020-01044 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -15.00/5.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), EXCESSIVE VAULT AND HALOS. MEDICAL INTERVENTION WAS PRESCRIBED TO TREAT IOP (INTRAOCULAR PRESSURE). THE SURGEON PLANS TO EXCHANGE THE LENS WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR." SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 2682 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - INVESTIGATION CODE 4117 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS HAS BEEN EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS. REPORTEDLY, "EVERYTHING WAS OK" AFTER THE EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652119&pc=QCB | 2023826-2019-00915 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, BLURRED VISION, PAIN, AND "MINOR TEMPORALLY ICL SHIFT." CORNEA IS CLEAR, NO INJUNCTIVAL INJECTION, IOP AT 14MMHG AT LAST VISIT. PATIENT CLAIMS NO HISTORY OF EYE INJURY. Manufacturer Narrative: PRODUCT PROBLEM: SHOULD BE CORRECTED TO ADVERSE EVENT IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE MARKED IN INITIAL MEDWATCH REPORT. DATE OF THIS REPORT: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. DATE RECEIVED BY MANUFACTURER: SHOULD BE CORRECTED TO 02/MAY/2019 IN INITIAL MEDWATCH REPORT. POST-OP PATIENT UCVA WAS RECORDED AS 20/25 AND THE VAULT WAS 458UM. REPOSITIONING WAS CONDUCTED ON A UNSPECIFIED DATE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10454797&pc=QCB | 2023826-2020-01864 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Pain; Blurred Vision; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PUPIL BLOCK WITH ELEVATED IOP (37MMHG) OBSERVED ON (B)(6) 2020. ADDITIONAL TREATMENT OF ENLARGEMENT OF PI AND A YAG WAS NECESSARY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS NON PATENT PI. REPORTEDLY, PATIENT CONDITION "IMPROVED - NO COMPLAINTS." Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED BLURRY VISION WITH EXTREME PAIN. H6 - DEVICE CODE: 2993 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311393&pc=QCB | 2023826-2022-00176 | NA | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Pupillary Block; Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT 12.6MM TMICL12.6; -7.00/+3.50/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021; AS A REPLACEMENT LENS. IT WAS NOTED THE PATIENT EXPERIENCED PUPIL BLOCK ON THE REPLACEMENT LENS WHICH WAS CAUSED FROM THE INITIAL IMPLANTED LENSES ANGLE CLOSURE. THE PATIENT THEN FELT VISION WASN'T GOOD AND STILL HAD A BLOCK PUPIL WITH RESIDUAL RX RIGHT EYE -0.25; +1.25; 145. THE SURGEON DECIDED TO PERFORM PRK SURGERY. IT WAS INDICATED THE PATIENT'S VISION IS STILL BLURRY. SEE MDR# 2023826-2022-00174 FOR INITIAL IMPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention + Secondary Operation | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010278&pc=QCB | 2023826-2022-04648 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Uveitis; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -1.50/+3.5/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS UVEITIS AFTER SURGERY WITHOUT PREVIOUS HISTORY; DEPOSITS IN ICL THAT DECREASE VISUAL ACUITY; PATIENT IS TREATED WITH CORTICOSTEROIDS; AND THERE IS PIGMENT DISPERSION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN; AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE-1494 SHOULD BE ADDED FOR CORRECTION. OFF-LABEL USAGE. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9706691&pc=QCB | 2023826-2020-00304 | 2020-01-10 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Uveitis; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PATIENT WAS DIAGNOSED WITH UVEITIS ON (B)(6) 2020 (VISUAL ACUITY 6/4.5-2) AND WAS PRESCRIBED PRED FORTE. AT PATIENT'S (B)(6) 2020 FOLLOW-UP, BLURRY VISION WAS OBSERVED (VISUAL ACUITY: 6/7.5-1, PINHOLE: 6/6-1). PATIENT WAS PRESCRIBED CIPROFLOXACIN, DEXAMETHASONE, CYCLOPENTOLATE, CRAVIT. PER REPORTER ON (B)(6) 2020, "PATIENT HAVE FULLY RECOVERED FROM THE UVEITIS." LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941512&pc=QCB | 2023826-2020-00870 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE RE4PORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/1.0/90 9 (SPHERE/CYLINDER/AXIS),IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941534&pc=QCB | 2023826-2020-00871 | 2020-01-31 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-11.0/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT IS REPORTED TO BE EXPERIENCING "DOUBLE VISION THAT CAN BE SEEN UP AND DOWN IN THE DARK PLACES". THE SYMPTOMS DO NOT OCCUR IN BRIGHT PLACES. THE SURGEON REPORTED THAT THE "PATIENT HAS NORMAL IOP, NORMAL VAULT, NORMAL OCULAR BODY AND NO COMPLICATIONS." THE REPORTER STATED " THE DOCTOR DOES NOT INTEND TO POINT OUT ANY PRODUCT DEFECTS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740194&pc=QCB | 2023826-2020-02482 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.0/058 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740224&pc=QCB | 2023826-2020-02481 | 2020-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE PATIENT WAS UNHAPPY WITH VISION (RESIDUAL MYOPIA). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098621&pc=QCB | 2023826-2020-03250 | 2020-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. REPORTEDLY, THE PATIENT IS EXPERIENCING LIGHTNING/FLASHES WHILE HE LIES IN BED AND TURNS HIS HEAD; STANDING UP THE VISION IS FINE. RETINAL DIAGNOSTIC SHOWS NO CHANGES AND THE SURGEON THINKS THIS COULD BE A "STRESS SYMPTOMATIC." Manufacturer Narrative: B5-UPDATED INFORMATION: THE HALOS HAVE BECOME LESS AND LESS. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238074&pc=QCB | 2023826-2021-00173 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. REPORTEDLY, "WTW MEASUREMENT CALLED FOR 12.6MM LENS(STAAR OPTIMIZED). UNCLEAR IF THERE WAS A AN ABNORMAL SULCUS ANATOMY CAUSING EXCESS VAULT. PER REPORTER, "VAULT IS NOW EXCELLENT, PATIENT NOTICES SIGNIFICANTLY REDUCED GLARE." Manufacturer Narrative: CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238075&pc=QCB | 2023826-2021-00213 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6, -11.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "WE ARE PLANNING TO EXCHANGE LENS IN MARCH." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: REPORTEDLY, "WE EXCHANGED HER LEFT EYE A COUPLE OF MONTHS AGO. SIGNIFICANTLY IMPROVED HER OVERALL VISION, BUT ESPECIALLY HER NIGHT VISION. SHE HAD A WTW OF 11.5 WTW OU AND I PLACED A 12.6 LENS IN EACH EYE. 300% CT VAULT OU." CLAIM (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11411793&pc=QCB | 2023826-2021-00568 | 2021-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.00/5.0/045 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WITHOUT ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DIFFICULTY WITH ACCOMMODATION FOR NEAR OBJECTS, THE REPORTER STATED " SX REPORTS DIFFICULTY WITH ACCOMMODATION POST ICL IMPLANTATION IN THE CONTEXT OF GOOD VISION AND REFRACTION DURING 1 DAY AND 1 WEEK POST OP REVIEW. BETTER WHEN OBJECTS ARE HELD FURTHER AWAY." LENS REMAINS IMPLANTED FOR NOW THE DOCTOR WILL MONITOR THE EYE. THE REPORTER STATED " PATIENTS CURRENT CONDITION IS FINE, ONLY ISSUE IS DIFFICULTY FOR READING NEAR SOMETIMES. VA/ EYE PRESSURE IS GOOD. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558804&pc=QCB | 2023826-2021-00706 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER IMPLANT, THE PATIENT FELT UNCOMFORTABLE WITH THE GLARE WHEN DRIVING AT NIGHT AND SO THE ICL WAS REMOVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563271&pc=QCB | 2023826-2021-00759 | NA | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.0/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ORIGINAL LENS TOO LARGE, SMALLER REPLACED LENS ROTATED. MFR # 2023826-2021-00652. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563284&pc=QCB | 2023826-2021-00757 | NA | Malfunction | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, LENS ROTATION IS PLANNED. WILL PROVIDE FURTHER INFORMATION ONCE A RESOLUTION IS REACHED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -16.00/4.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 201. DIPLOPIA AND ASTIGMATISM WAS OBSERVED, WITH POSSIBLE ROTATION. THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2021, "DIPLOPIA AND ASTIGMATISM HAVE RESOLVED." CORRECTED INFORMATION: B1 - REPORTED AS ADVERSE EVENT - CORRECT TO PRODUCTION PROBLEM. B2 - REPORTED AS REQUIRED INTERVENT TO PREVENT PERMANENT IMPAIRMENT/DAMAGE - NO OUTCOMES ATTRIBUTED TO ADVERSE EVENT. H1 - REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820848&pc=QCB | 2023826-2021-01410 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -11.00/1.5/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -11.00/1.5/053 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY "GLARE VISUAL QUALITY IS BAD AND NOT IMPROVING." GLARE/HALOES REPORTED. LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820914&pc=QCB | 2023826-2021-01411 | 2021-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER/PATIENT INDICATED THAT A 13.2MM, TMICL13.2, -10.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER/PATIENT INDICATED THAT A 13.2MM TMICL13.2 -10.50/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, "GLARE, VISUAL QUALITY IS BAD AND NOT IMPROVING." LENS REMAINS IMPLANTED. PROBLEM NOT RESOLVED. STATUS OF THE EYE IS QUIET - WELL HEALED. REPORTEDLY THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281705&pc=QCB | 2023826-2021-03069 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. REPORTEDLY, LENS REPOSITIONING IS PLANNED. WILL ADVISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS MISALIGNED BY 15 DEGREES. LENS REPOSITIONED ON (B)(6) 2021 DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR, MOST LIKELY CAUSE IS THAT INK AXIS MARKING WERE OFF TARGET FOR THE PROCEDURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478533&pc=QCB | 2023826-2021-03579 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND GLARE (THE PATIENTS PUPIL WAS 5MM). THE PUPIL SIZE WAS STILL 5MM FOR HALF A MONTH AFTER SURGERY, SO THE SURGEON DECIDED NOT TO IMPLANT ANOTHER ICL LENS FOR THE TIME BEING. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479095&pc=QCB | 2023826-2021-03592 | 2021-06-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479109&pc=QCB | 2023826-2021-03593 | 2021-06-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER STATED THAT THE PATIENT DEVELOPED DIPLOPIA WITH VERY LOW VAULT AND LENS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558495&pc=QCB | 2023826-2021-03810 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS GLARE/HALOS AND PILOCARPINE ADMINISTRATION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725161&pc=QCB | 2023826-2021-04177 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.5/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ABERATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION " DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " DEVICE CAUSES MILD ABERATIONS- PT VERY SENSITIVE TO ABERATIONS". Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725173&pc=QCB | 2023826-2021-04176 | 2020-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.0/4.0/089 (SPHERE / CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO GLARE/ HALOES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ ABERRATIONS IMPROVED BUT STILL PRESENT. PLAN ON DOING LASIK FOR RESIDUAL REFRACTIVE ERROR." IN THE REPORTERS OPINION "DEVICE AND PATIENT-RELATED FACTOR" WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED "DEVICE CAUSES GLARE/ABERRATIONS- PT VERY SENSITIVE TO GLARE/ABERRATIONS". Manufacturer Narrative: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057182&pc=QCB | 2023826-2021-04829 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(6). | Unknown | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057258&pc=QCB | 2023826-2021-04828 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF NIGHT GLARE, WHICH AFFECTED HIS NORMAL LIFE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082455&pc=QCB | 2023826-2021-04974 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED DUE TO "UNACCEPTABLE GLARE AFTER IMPLANTATION OF LENS." THE GLARE DISAPPEARED AFTER THE REMOVAL OF THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362150&pc=QCB | 2023826-2022-00347 | 2021-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4/+6/55 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "PATIENT SAW GLARE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14093007&pc=QCB | 2023826-2022-01195 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL_12.1; -9.00/1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT WANTED TO TRY FOR MONOVISION; IN WHICH SHE WAS UNHAPPY WITH THE DEGREE OF MONOVISION AND WOULD PREFER TO HAVE AN EXCHANGE FOR A POWER CLOSER TO HER DISTANCE REQUIREMENT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706970&pc=QCB | 2023826-2022-01749 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -13.5/3.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14706973&pc=QCB | 2023826-2022-01748 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 -14.5/5.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS OF WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY GLARES/HALOES OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE - THE PATIENT FELT THAT THE POSTOPERATIVE HALO WAS BECOMING MORE SERIOUS AND COULD NOT BEAR IT AND ASKED TO REMOVE IT. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003548&pc=QCB | 2023826-2022-02198 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6). POST-OP, THE UCVA WAS GOOD BUT THERE WAS GLARE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK .EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567406&pc=QCB | 2023826-2022-03564 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -9.50/4.0/91. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS AFTER THE INITIAL LENS WAS FOUND DAMAGE BEFORE LOADING. IT WAS REPORTED THAT POST ICL IMPLANTATION THE PATIENT EXPERIENCED GLARE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE GLARE RESOLVED. LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000576&pc=QCB | 2023826-2022-04596 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE REPORTER STATES THAT THE PATIENT HAS AN USTABLE DISTANCE VISION. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM IS RESOLVED. "THE PATIENT HAS GOTTEN USED TO THE SITUATION AND DOES NOT WANT ANY FURTHER INTERVENTIONS." CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040405&pc=QCB | 2023826-2022-04702 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.5/6.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS REMOVED. REPORTEDLY, "EXCESSIVE GLARE AND HALO, PATIENT CAN'T TOLERATE. CAUSE IS REPORTED AS THE DEVICE AND PATIENT RELATED FACTOR. REPORTEDLY, STATUS OF THE EYE IS "RETURNED TO PRE-OP STATUS." Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | Unknown | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296735&pc=QCB | 2023826-2023-00431 | NA | Injury | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -15.0/2.00/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED DECREASED VISION AND HALO GHOST IMAGE. THE REPORTER STATES THAT ICL ROTATION WAS PERFORMED IN JUNE 2020 AND THAT HE NOW PLANS TO EXCHANGE FOR LONGER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS. REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296738&pc=QCB | 2023826-2023-00429 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/5.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296741&pc=QCB | 2023826-2023-00433 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.0/64 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2022. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296747&pc=QCB | 2023826-2023-00430 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/4.5/77(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES THAT A LENS WITH WRONG LENGTH WAS ORDERED AND THAT THE CAUSE OF ROTATION IS LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2019. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296766&pc=QCB | 2023826-2023-00428 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/3.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2021. LOW VAULT WITH ROTATION WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES "CAUSE OF ROTATION IS LOW VAULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). IMPLANT DATE: (B)(6) 2021. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM SHOULD BE CORRECTED TO ADVERSE EVENT. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS SHOULD BE ADDED FOR CORRECTION. H3: SERIOUS INJURY SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16296833&pc=QCB | 2023826-2023-00432 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN SEPTEMBER 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVER GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). IMPLANT DATE: SEPTEMBER 2022 TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". H6: MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE: AGE>45 YEARS OLD AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -7.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED AND THE PATIENT EXPERIENCED SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES "HIGH VAULT WITH UNKNOWN CAUSE". CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16742900&pc=QCB | 2023826-2023-01256 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023. DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED DOUBLE VISION AND GLARE AT NIGHT. THE PATIENT WAS PRESCRIBED MEDICATION BUT THE SYMPTOMS DID NOT IMPROVE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN PIECES AND UNABLE TO PERFORM THE DIMENSIONAL INSPECTION DUE TO THE DAMAGED LENS. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16877245&pc=QCB | 2023826-2023-01587 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6)2022. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO THE PATIENT EXPERIENCED GLARE AND THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT ACCEPT THE GLARE AND REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273274&pc=QCB | 2023826-2023-02719 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED SEVERE DOUBLE VISION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS, AND LIMBAL RELAXING INCISION WAS PERFORMED TO CORRECT ITS ASTIGMATISM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302685&pc=QCB | 2023826-2023-02829 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322117&pc=QCB | 2023826-2023-02830 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279670&pc=QCB | 2023826-2023-05479 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/1 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18279671&pc=QCB | 2023826-2023-05480 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS OF -2.0/0.5/69 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6). LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# 732815. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456573&pc=QCB | 2023826-2024-00038 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456592&pc=QCB | 2023826-2024-00039 | 2023-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND GLARE/HALOS WERE OBSERVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684401&pc=QCB | 2023826-2022-00551 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -9.50/1.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE, AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, GLARE REDUCED A LOT, VISION BETTER PUPIL BACK TO NORMAL. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10852592&pc=QCB | 2023826-2020-02686 | 2020-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Discomfort | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT WAS UNCOMFORTABLE AND EXPERIENCED VISION FATIGUE. THE SURGEON WILL NOT IMPLANT ANOTHER ICL LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938663&pc=QCB | 2023826-2020-02843 | NA | Injury | STAAR SUGICAL COMPANY | 2020-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019. THE PATIENT CONTINUED TO STRUGGLE WITH GLARE AND HALO SINCE INITIAL SURGERY, POSSIBLE LENS EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THE REPORTER STATED " PUPIL SIZE GREATER OD > OS DUE TO EXCESSIVE VAULT CAUSING GLARE/HALO OD. " THIS EXCHANGE RESOLVED THE PROBLEM, "GLARE/HALO COMPLAINT HAS RESOLVED WITH 12.6MM ICL IMPROVING VAULT." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322676&pc=QCB | 2023826-2019-00229 | NA | Malfunction | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED. ON PATIENTS LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA: 20/25-2, UCVA: 20/30 AND IOP: 12 MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: " PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE & HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, VTICMO13.7, -12.0/4.0/075 (SPHERE/CYLINDE/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018 AND THE PATIENT IS COMPLAINING OF GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. ON PATIENT'S LAST VISIT DATE: (B)(6) 2019, STATUS OF THE EYE WAS REPORTED AS "NORMAL", BCVA 20/25-2, UCVA:20/30 AND IOP: 12MMHG. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION: "PATIENT EXPERIENCING RESIDUAL HYPEROPIA AFTER SURGERY, ALSO COMPLAINING OF GLARE AND HALOES". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8544980&pc=QCB | 2023826-2019-00594 | 2018-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -1.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT GLARE/HALOES WERE OBSERVED. CAUSE OF THE EVENT IS REPORTED AS CENTRAL FLOW HOLE. LENS REMAINS IMPLANTED. REPORTEDLY, LENS EXCHANGE IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9323816&pc=QCB | 2023826-2019-02164 | 2018-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. ON (B)(6) 2019 REPORTER INDICATED "THIS PATIENT HAS ALREADY EXCHANGED HIS LENS ON OS AND MIGHT NEED TO DO IT ON OD."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - DEVICE CODE 3189 CORRECTED TO DEVICE CODE 3001. (B)(4). Manufacturer Narrative: H6 - RESULT CODE 213 CORRECTED TO RESULT CODE 114. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427016&pc=QCB | 2023826-2019-02296 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO GLARE/HALOES. PER REPORTER, "PATIENT EXPERIENCED CONCENTRIC CIRCLES OF LIGHT THRU LENS HOLD - INTOLERABLE." PATIENT WILL NOT BE PROCEEDING WITH ANOTHER LENS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633749&pc=QCB | 2023826-2020-00161 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/4.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT IMPACT CODE 4629 SHOULD BE IN THE INITIAL MDR. H10- DEVICE EVALUATION: "RETURNED IN CARTRIDGE" SHOULD BE REMOVED FROM PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE AND MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS RETURNED IN THE CARTRIDGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135948&pc=QCB | 2023826-2020-01171 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - SHOULD BE CORRECTED TO MALFUNCTION. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135999&pc=QCB | 2023826-2020-01172 | 2020-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT COMPLAINED OF GLARE/HALOS AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740529&pc=QCB | 2023826-2020-02488 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY" CORRECTED INFORMATION: B5 - (B)(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10740575&pc=QCB | 2023826-2020-02487 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00/+4.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH GLARE AND HALOS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - (B))(6) 2020 IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2020. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984748&pc=QCB | 2023826-2020-02858 | 2020-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. GLARE/HALOES HAVE BEEN OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B1 - REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B2 - UPDATE TO REQUIRD INTERVENTION TO PREVENT IMPAIREMENT/DAMAGE. B5 - NOTED LENS REMAINS IMPLANTED. LENS REMOVED ON (B)(6) 2020 AND THE PROBLEM RESOLVED. D6B = (B)(6) 2020. H6- HEALTH EFFECT IMPACT CODE - REPORTED AS 2199, UPDATED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN WITH FOREIGN MATERIAL ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206966&pc=QCB | 2023826-2021-00106 | 2020-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO LOW VAULT AND GLARE/HALOES. THIS RESOLVED THE PROBLEM. REPORTEDLY, "THERE IS NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238056&pc=QCB | 2023826-2021-00172 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10/1.0/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11238066&pc=QCB | 2023826-2021-00171 | 2020-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "IRIS IS NOT FULLY JUMPING OVER OPTICAL ZONE. MESOPIC PUPIL MID-DILATED POST OP (5MM). EDGE OF THE PUPIL IS NOT IN THE WAY OF THE LENS AND THE PUPIL DOES CONSTRICT AND RELAX BUT DOES NOT FULLY CONSTRICT. PATIENT TO BE MONITORED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368705&pc=QCB | 2023826-2021-00458 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368840&pc=QCB | 2023826-2021-00457 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+4.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642581&pc=QCB | 2023826-2021-00943 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON (B)(6) 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON (B)(6) 2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11642627&pc=QCB | 2023826-2021-00944 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. THE REPORTER STATED " ONE YEAR AGO I GOT THE EVO VISIAN ICL IMPLANTED IN MY EYES (I'M MYOPIC WITH ASTIGMATISM). THE SURGERY WENT WELL AND MY EYES HEALED WELL. I DON'T HAVE ANY HEALTH PROBLEMS, BUT DESPITE HAVING 100% EYESIGHT, I HAVE STRONG SIDE EFFECTS: I SEE LARGE CIRCLES OF LIGHT AT ANY TIME OF THE DAY WHEN I AM ILLUMINATED BY LAMPS OR THE SUN. THE LIGHT CIRCLES ARE EXTREMELY BRIGHT, LARGE AND DEPEND ON HOW THE LIGHT FALLS INTO THE EYE. SOME TUNNELS OR AVENUES OF STREET LAMPS TRIGGER SEVERAL CIRCLES OF LIGHT ONE AFTER THE OTHER, WHICH ARE GETTING BIGGER AND BIGGER. FROM THE SUNLIGHT I ADDITIONALLY SOMETIMES SEE A SMALL CIRCLE OF LIGHT IN THE MIDDLE, WHICH APPEARS IN LIGHT RAINBOW COLORS. IN DAWN AND IN POORLY LIT ROOMS I SEE HIGH-CONTRAST EDGES / CONTOURS TWICE, WHEREAT THEY BLUR INTO EACH OTHER. THE DOUBLE EDGES ARE NOTICEABLY LIGHTER AND BLURRY. EVEN WHEN I WATCH TV IN THE EVENING, THE IMAGES KEEP BLURRING. DUE TO THE RISKS OF A SURGERY, I DO NOT WANT TO HAVE THE LENSES REMOVED PREMATURELY, BUT I AM VERY UNHAPPY WITH MY CURRENT CONDITION AND I REGRET THE SURGERY VERY MUCH." ON 06 APR 2021 THE REPORTER STATED " WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WILL STAY IN THE. THE REPORTER STATED " THE LENS WILL STAY IN THE EYE, BECAUSE THE UCVA 20/16 IS PERFECT , PATIENT IS TRYING TO GET USED TO HALOS." CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -6.5/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY THE PATIENT EXPERIENCED GLARES; HALOES; BLURRY VISION AND LARGE CIRCLES OF LIGHT. ON 06/04/2021 THE REPORTER STATED "WE HAVE NOW STARTED AN ATTEMPT AT THERAPY WITH EYE DROPS AND I HOPE THAT MY SYMPTOMS CAN BE ALLEVIATED WITH THEM." AFTER SOME TIME REPORTS THAT "THE LENS WILL STAY IN THE EYE BECAUSE THE UCVA 20/16 IS PERFECT; PATIENT IS TRYING TO GET USED TO HALOS." LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: SEX-RACE:UNK. DATE OF EVENT:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651557&pc=QCB | 2023826-2021-00883 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -4.50/1.5/143 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER ON (B)(6) 2021, NO PLANS TO IMPLANT ANOTHER LENS, UCVA OF THIS PATIENT WAS GREAT AFTER THE SURGERY, BUT THE PATIENT COULDN'T STAND THE HALO AND GLARE, PATIENT INSISTED TO TAKE OUT THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4), ACD <3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721921&pc=QCB | 2023826-2021-01040 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778986&pc=QCB | 2023826-2021-01325 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/1.0/139 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779037&pc=QCB | 2023826-2021-01326 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/1.00/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT HAD SERIOUS NIGHTTIME DAZZLES AFTER SURGERY AND REQUESTED LENS BE REMOVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820836&pc=QCB | 2023826-2021-01405 | 2018-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS,INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. GLARE/HALOS WERE OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.50/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, LENS NATURALLY SITS VERY SLIGHTLY SUPERIOR/TEMPORAL. EYE DROPS (ALPHAGAN) DO NOT WORK. LENS REMAINS IMPLANTED. STATUS OF EYE; "ALL WITHIN NORMAL LIMITS." H6: HEALTH IMPACT CODE- 2199 IN INITIAL MDR IS CORRECTED TO 4644. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222777&pc=QCB | 2023826-2021-02889 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/6.0/104, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR AND IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222778&pc=QCB | 2023826-2021-02888 | 2020-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.0/1.5/115, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULT, LENS/DISLOCATION/SUBLUXATION AND GLARE/HALOES. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER STATED " LENS INITIALLY HAD A SMALL VAULT AND IS OFF-CENTERED AS A RESULT. IT IS CURRENTLY OFF-CENTERED TEMPORALLY INFERIORLY DISLOCATION. AXIAL POSITION AND OVER-OVER REFRACTION THEREFORE ARE NOT RELIABLE. PATIENT IS DETECTING HALOS AND GLARE DUE TO DISLOCATION." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: B5: 12.6MM VTICM5 -13.0/1.5/104 (SPHERE/CYLINDER/AXIS). CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "THE LENS NO LONGER REMAINS IMPLANTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/ LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION, LENS DISLOCATION/SUBLUXATION, AND GLARE HALOES. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587386&pc=QCB | 2023826-2021-03898 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587379&pc=QCB | 2023826-2021-03897 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.0/075, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- CORRECTION: INITIAL LENS WAS IMPLANTED (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K : THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593044&pc=QCB | 2023826-2021-03919 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/+1.0/012, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12593615&pc=QCB | 2023826-2021-03920 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/+0.5/033, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " THE PATIENT HAS AN UCVA OF 20/20 WITH ICLS BUT MASSIVE ISSUES WITH HALOS AND GLARE AT NIGHT. SO THE SURGEON EXPLANTED THE ICL AND NOW ALL IS FINE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791947&pc=QCB | 2023826-2021-04321 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER STATED " 2 WEEKS POST-OP ICL PATIENT HAS COMPLAINTS OF HALO". PER UPDATED INFORMATION THE REPORTER STATED THAT THE PATIENT IS DOING QUITE WELL, WITH GLARE THAT CAN BE MANAGED WITH BRIMONIDINE. LENS REMIANS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DATE IMPLANTED: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -6.0/1.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. PATIENT COMPLAINTS OF HALO/GLARE POST ICL IMPLANTATION. PER (B)(6) 2021 EMAIL FROM REPORTED THE REPORTER STATES "PATIENT IS ACTUALLY DOING QUITE WELL, WITH GLARE THAT HE CAN MANAGE WITH BRIMONIDINE". PATIENT-RELATED FACTOR AND DEVICE WAS REPORTED FOR CAUSE OF EVENT. HOWEVER, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. D4: (B)(6) 2023. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872823&pc=QCB | 2023826-2021-04493 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/011 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHINICITY , RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872831&pc=QCB | 2023826-2021-04494 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED GLARE/ HALOES AND THE PROBLEM WAS RESOLVED. ANOTHER ICL LENS WAS NOT IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169183&pc=QCB | 2023826-2022-01222 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE AND HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14169248&pc=QCB | 2023826-2022-01221 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PATIENT FEELING UNCOMFORTABLE AND COULD NOT TOLERATE. AFTER REPEATED COMMUNICATION, THE PATIENT STILL INSISTED ON REMOVING THE LENS. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734140&pc=QCB | 2023826-2022-01850 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO GLARE/HALOES. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15003801&pc=QCB | 2023826-2022-02199 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT COMPLAINED OF GLARE/HALOS. THE LENS WAS REPLACED WITH AN IOL LENS AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591507&pc=QCB | 2023826-2022-03594 | 2022-08-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM13.2, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GALRE AND HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICM13.2" SHOULD BE CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2." CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM13.2 -8.50/1.0/130 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE AND HALOES. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030530&pc=QCB | 2023826-2022-04434 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -9.5/1.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTD TO BE PATIENT-RELATED FACTOR. HOWEVER, THE REPORTER STATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16130089&pc=QCB | 2023826-2023-00094 | 2022-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00/+1.5/087, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED OF GLARE/HALOES AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301533&pc=QCB | 2023826-2023-00382 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/5.5/047 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301607&pc=QCB | 2023826-2023-00383 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -75/5.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. "JUST IN CASE OF KERATOCONUS HIGH VAULT IS COMMON EVENT AND THIS CASES AS HISTORY OF KERRARING (INTRA CORNEAL RING)." Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301634&pc=QCB | 2023826-2023-00384 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2020. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301672&pc=QCB | 2023826-2023-00385 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED WITH SEVERE GLARE AND GHOST IMAGE. THE LENS REMAINS IMPLANTED. PATIENT HAS A PRE-EXISTING CONDITION OF KERATOCONUS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2021. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO : SERIOUS INJURY IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE SHOULD BE REMOVED FROM MFR REPORT #: 2023826-2023-00385-SUPP1. H1 - CORRECTED TO : MALUNCTION IN THE INITIAL MDR. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506881&pc=QCB | 2023826-2023-00625 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/168 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595253&pc=QCB | 2023826-2023-00996 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/051 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 , THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL, LONGER LENGTH LENS DUE TO LOW VAULT, LENS OPACITY, AND GLARE HALOES. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "PATIENT STILL WITH LENS OPACITY SUPRISINGLY, BUT BETTER VAULT" Manufacturer Narrative: CORRECTED DATA: H6- 1494: OFF-LABEL USAGE- PRE-EXISTING CONDITION OF PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434410&pc=QCB | 2023826-2023-03186 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 2(B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440152&pc=QCB | 2023826-2023-03185 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -13.0/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT STRONGLY REQUESTED REMOVAL AND NO REPLACEMENT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495490&pc=QCB | 2023826-2023-03323 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0/0.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "2227" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE REMOVED AND "3001" SHOULD BE ADDED. (B)(4) Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733767&pc=QCB | 2023826-2023-04036 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.00/ 3.0/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733799&pc=QCB | 2023826-2023-04037 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/145 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882968&pc=QCB | 2023826-2023-04447 | NA | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.00/1.0/120 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. HALOS/GLARE WITH NIGHTTIME DRIVING ARE REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961829&pc=QCB | 2023826-2023-04557 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067593&pc=QCB | 2023826-2023-04969 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075196&pc=QCB | 2023826-2023-04725 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075197&pc=QCB | 2023826-2023-04726 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6. -12.0/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO NIGHTTIME GLARE/HALOES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: NIGHTTIME GLARE. REPORTEDLY, PATIENT REQUEST REMOVAL OF LENS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080758&pc=QCB | 2023826-2023-04992 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13/2.5/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18117390&pc=QCB | 2023826-2023-04761 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/077 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136579&pc=QCB | 2023826-2023-04762 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+5.5/103 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT COMPLAINED OF GLARE/HALOS AND WAS NOT SATISFIED WITH RESULTS. THE PROBLEM WAS RESOLVED. POST-OP THE EYE WAS NORMAL AND WAS WITHOUT SIGNS AND SYMPTOMS. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440045&pc=QCB | 2023826-2024-00048 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440457&pc=QCB | 2023826-2024-00049 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0//055 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WITH GLARE/HALOES WAS REPORTED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129006&pc=QCB | 2023826-2021-00086 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/010(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11129010&pc=QCB | 2023826-2021-00085 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.00/1.0/170(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTFIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11368769&pc=QCB | 2023826-2021-00471 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11386930&pc=QCB | 2023826-2021-00544 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2021-02-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. FOR ADDITIONAL INFORMATION "OS- VAULT 1000 OBSERVED" WAS REPORTED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. DATE OF EVENT: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697117&pc=QCB | 2023826-2021-01089 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/+1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697124&pc=QCB | 2023826-2021-01088 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.0/+1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "UNREACTIVE PUPIL AND GLARE HALO DUE TO HIGH VAULT WERE CONFIRMED. THE LENS SIZE DID NOT FIT THE PATIENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119230&pc=QCB | 2023826-2021-02331 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS THEN EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOS, BLURRED VISION, AND SPARKLES. THE EYE "LOOKS GOOD" NOW. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: SPHERE/CYLINDER/AXIS OF THE IMPLANTED LENS IS -0.5/+1.5/049. THE LENS WAS IMPLANTED ON (B)(6) 2021. THE PATIENT CAN SEE CENTRAL RINGS. ON (B)(6) 2021 THE LENS WAS EXCHANGED. THE PROBLEM WAS NOT RESOLVED. SEE MFR REP# 2023826-2021-02969 FOR THE REPLACEMENT LENS. H6-HEALTH EFFECT CLINICAL CODE 4581 (SEES CENTRAL RINGS). H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. IMPLANT, EXPLANT DATE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MFG DATE - UNK. CLAIM# (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669520&pc=QCB | 2023826-2021-04044 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED DYSPHOTOPSIAS AT NIGHT, WITH MESOPIC PUPIL (6MM). THE PATIENT REPORTED GLARE/HALOS. THE PATIENT WAS TREATED WITH ALPHAGAN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296711&pc=QCB | 2023826-2022-02793 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: EXPIRATION DATE: 31-OCT-2024 (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307873&pc=QCB | 2023826-2022-02794 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS QUIET. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPIRATION DATE: 30-APR-2025. HEALTH IMPACT- CLINICAL CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15476840&pc=QCB | 2023826-2022-03279 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18/1.00/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON . THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE AND GLARE/HALOES. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ERROR ON LENS DIAMETER CALCULATION. ANGLE CLOSURE ON ONE CUADRANT (NASAL)". Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939032&pc=QCB | 2023826-2023-01654 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -7.00/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OD DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. " CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939036&pc=QCB | 2023826-2023-01655 | 2023-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, GLARE/HALOES, AND EXCESSIVE VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. PATIENT REMAINS UNDER OBSERVATION. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. REPORTEDLY, PATIENTS ANATOMY MAY BE ABNORMAL OR A MEASUREMENT ERROR. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED, AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -10.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) HE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND GLARES/HALOES. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "NARROW ANGLE, HIGH VAULT, NO OTHER SYMPTOMS, MONITORING FOR NOW". ADDITIONAL INFORMATION PROVIDED: "WOULD LIKE TO REQUEST A CONSULTATION WITH AN MD TO DISCUSS THE SCENARIO AND TO RECEIVE MEDICAL PERSPECTIVE ON HIS PLANT TO OBSERVE." THE CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "HIGH VAULT OS DESPITE ACCURATE PREOP MEASUREMENTS, PATIENTS ANATOMY MAY BE ABNORMAL; OR A MEASUREMENT ERROR. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138291&pc=QCB | 2023826-2023-02315 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/046 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING, UNREACTIVE (FIXED) PUPIL AND GLARE/HALOES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WERE NO COMPLAINTS, EYE WAS CALM AND THE PUPIL WAS REACTIVE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: A PI WAS PERFORMED ON (B)(6) 2023 AND WAS THE FIRST PI (NO NEED OF BEFORE PI - MYOPIC LENS). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITION OF PI. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "4581- WIDE PUPIL " SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER-18.0/+1.5/46 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, 'WIDE' PUPIL, AND GLARES/HALOES. ON (B)(6) 2023 PI WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE: "NO COMPLAINTS, CALM EYE, REACTIVE PUPIL (STATUS VERY GOOD)". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "TO BIG SIZE OF THE LENS UNUSUAL STRUCTURE OF THE EYE". CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887128&pc=QCB | 2023826-2021-01641 | 2021-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED TOGETHER WITH THE PATIENT, THE DOCTOR DECIDED NOT TO PERFORM A SURGICAL INTERVENTION. ALL IS GOOD AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: IN INITIAL MDR THE WORK ORDER SEARCH RESULTS SHOULD BE CORRECTED TO: NO SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6 MEDICAL DEVICE CODE- 2923 CODE IN INITIAL MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: AT A POST-OP EXAMINATION ON (B)(6) 2021, THE SURGEON DETECTED AN EXCESSIVE VAULTING. THERE WAS PIGMENT DISPERSION AND PATIENT COMPLAINED OF GLARE/HALOS. H6: HEALTH IMPACT CLINICAL CODE: 4581 PIGMENT DISPERSION. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION, AND GLARE/HALOS. THE PROBLEM WAS RESOLVED AND THE "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756627&pc=QCB | 2023826-2021-04227 | 2021-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; Discomfort | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7,-12.5/2.5/101, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PATIENT EXPERIENCED EXCESSIVE VAULT AND GLARE/HALOES, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN, CAUSE DETAILS " LENS IS TO BIG FOR PATIENT" WAS REPORTED. PER UPDATED INFORMATION THE REPORTER STATED " PATIENT IS JUST DISCOMFORT WITH THE VISUAL ACUITY IOP IS FINE." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "PATIENT ALSO EXPERIENCED DISCOMFORT". CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057307&pc=QCB | 2023826-2023-02178 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/4.5/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, GLARES, AND HALOES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091344&pc=QCB | 2023826-2023-02233 | NA | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/044 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR DIFFERENT MODEL, SHORTER LENGTH LENS, DUE TO EXCESSIVE VAULT AND GLARE/HALOES. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "2227" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519001&pc=QCB | 2023826-2019-00612 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WASEXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519013&pc=QCB | 2023826-2019-00608 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552494&pc=QCB | 2023826-2019-00659 | 2019-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE PROBLEM WAS RESOLVED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746386&pc=QCB | 2023826-2019-01163 | 2019-05-04 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/096 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT COULD NOT TOLERATE GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895522&pc=QCB | 2023826-2019-01453 | 2019-03-02 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM WAS RESOLVED, PATIENT CONDITION WAS STABLE. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966508&pc=QCB | 2023826-2019-01647 | 2019-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND IRIS PROLAPSE. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: EVENT DESCRIPTION CORRECTED TO STATE THAT A 13.7MM VTICM5 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.00/6.0/031 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966519&pc=QCB | 2023826-2019-01646 | 2019-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT, GLARE/HALOS, AND NEGATIVE DYSPHOTOPSIA. THE CAUSE OF THE EVENT IS REPORTED AS "NOMOGRAM ERROR." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPES WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). Manufacturer Narrative: CORRECTED TO STATE THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/1.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CLAIM# : (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED PATIENT CODE 3191: NEGATIVE DYSPHOTOPSIA. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK. MODEL #, IMPLANTED, EXPLANTED: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8969999&pc=QCB | 2023826-2019-01662 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -0.50/2.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. HALOS AND "LIGHTSPRED" WERE OBSERVED. LENS REMAINS IMPLANTED. REPORTER INDICATED THAT "THE SURGEON DELAYED POSSIBLE LENS EXCHANGE DUE TO VERY GOOD REFRACTIVE RESULT. PATIENT WOULD LIKE TO EXCHANGE LENS, STILL HAVING DAILY PROBLEMS WITH GLARE." PER THE REPORTER ON (B)(6) 2019 EXCHANGE WILL NOT HAPPEN UNTIL END OF (B)(6). Manufacturer Narrative: PATIENT GENDER: UNK SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. ADVERSE EVENT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL BUT SAME POWER/SIZE LENS AND THE PROBLEM WAS RESOLVED. PATIENT STATUS WAS REPORTED AS "GOOD". EXPLANT DATE: (B)(6) 2019. TYPE OF REPORTABLE EVENT: SERIOUS. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115540&pc=QCB | 2023826-2019-01765 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -9.00/1.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: DATE OF EVENT UPDATED TO (B)(6) 2019 IN INITIAL MDR. PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL STUCK TO A PIECE OF PAPER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, BROWN RESIDUE ON THE LENS AND THE LENS RETURNED STUCK TO A PIECE OF PAPER WITH INK. THE LENS HAS BLUISH/ PURPLE TINT POSSIBLY FROM THE INK ON THE PAPER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115541&pc=QCB | 2023826-2019-01764 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "LENS WAS REMOVED DUE TO STRONG HALOES. PATIENT WEARS GLASSES AGAIN." Manufacturer Narrative: PATIENT CODE 2227 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. DATE OF EVENT: UPDATED TO (B)(6) 2019 IN INITIAL MDR. EVENT DESCRIPTION: PER UPDATED INFORMATION, THE CAUSE OF THE EVENT IS DUE TO A PATIENT RELATED FACTOR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RED RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203032&pc=QCB | 2023826-2019-01953 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -08.00/1.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY SURGEON NOTED "HALOS/GLARE PRESENT BOTH WITH SMALL AND DILATED PUPIL. RELATED IN MY OPINION TO CENTRAL HOLE." CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE, SMALL AND DILATED PUPIL. RESULT CODE 114 CORRECTED TO RESULT CODE 213. PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LONGER LENGTH, DIFFERENT POWER LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM#(B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: RESULT CODE 114 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252703&pc=QCB | 2023826-2019-01960 | NA | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+0.5/170 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: EVENT - "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED IN THE INITIAL MDR." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252792&pc=QCB | 2023826-2019-01961 | NA | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE PUPIL SIZE WAS LARGE/REACTIVE. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED FOR 2 MONTHS. Manufacturer Narrative: CORRECTED DATA: "THE PATIENT EXPERIENCES GLARE/HALOES DURING DAY AND NIGHT TIME." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492810&pc=QCB | 2023826-2019-02418 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/091 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". Manufacturer Narrative: EVENT: "GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492867&pc=QCB | 2023826-2019-02419 | 2019-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -12.00/2.0/089 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM BUT "STILL A LITTLE HIGH VAULT". ABNORMAL SULCUS TO SULCUS WIDTH WAS REPORTED AS CAUSE OF EVENT. Manufacturer Narrative: CORRECTED DATA: B5- " GLARE AND HALOS WERE OBSERVED." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9498138&pc=QCB | 2023826-2019-02410 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARES AND HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO SIZING VARIATION IN OUTCOME. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633747&pc=QCB | 2023826-2020-00160 | 2019-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/4.5/083 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARES/HALOES WERE OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY AFTER EXPLANT (REMOVAL)." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 3191 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881022&pc=QCB | 2023826-2020-00659 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916667&pc=QCB | 2023826-2020-00735 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND LENS WAS RETURNED IN PIECES - UNABLE TO EVALUATE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916671&pc=QCB | 2023826-2020-00736 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105254&pc=QCB | 2023826-2020-01079 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/1.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO GLARE/HALOES AND LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138815&pc=QCB | 2023826-2020-01192 | 2020-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem; No Apparent Adverse Event | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED AFTER IMPLANTABLE COLLAMER LENS SURGERY. STARBURST AND FLOATERS ARE BEING OBSERVED. REPORTEDLY, HIGH REFRACTIVE POWER (AROUND -16 D) AND PATIENT NOTED STARBURST, FLOATERS AND HALOS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H6 - DEVICE CODE 3189 IS NOT APPLICABLE IN INITIAL MDR H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT/DAMAGE IN INITIAL MDR. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.50/1.5/039 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT HAS EXPERIENCED STARBURST, FLOATERS, AND GLARE/HALOES FOR THE PAST TWO YEARS. THE CLINIC SUGGESTED TO USE PILO (MEDICATION) FOR A MONTH (DURING NIGHT TIME). THE PATIENT WAS FINE FOR 3 DAYS BUT STARTED HAVING THE SAME PROBLEM BEYOND THAT. THE PATIENT HAS BEEN REGULARLY VISITING THE SURGEON SINCE IMPLANTATION. THE LENS REMAINS IMPLANTED. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL AND SUPPLEMENTAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4644 - PILO MEDS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184788&pc=QCB | 2023826-2020-01309 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184802&pc=QCB | 2023826-2020-01310 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.50/3.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO GLARE/HALOS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491985&pc=QCB | 2023826-2020-01969 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, SMALL ANTERIOR CHAMBER AND GLARE/HALOS. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON THE LENS. H6- DEVICE IMPACT CODE 4629- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: "H6- DEVICE HISTORY RECORD- BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK WORD(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." SHOULD HAVE BEEN IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10585633&pc=QCB | 2023826-2020-02222 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, "PATIENT STILL HAVING SOME GHOSTING ON THE LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -5.00/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THERE WAS LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WITH SOME "GHOSTING OF LETTERS ON THE CHART NO MATTER WHERE THE AXIS IS TURNED", AND ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE REPORT INDICATED THAT "NORMAL POST OP COURSE. PATIENT STILL HAS A SUTURE OS AT AXIS 85". CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. G3 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 28-AUG-2020. H6 - HEALTH EFFECT CLINICAL CODE: 2140 SHOULD BE CORRECTED TO 4582. H6 - HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SUTURE. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823660&pc=QCB | 2023826-2020-02592 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: IT WAS LATER CLARIFIED THAT THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE WHICH WAS ALSO RESOLVED WITH THE NOTED LENS EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: H6: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE PREVIOUSLY BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439778&pc=QCB | 2023826-2022-01596 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement; No Apparent Adverse Event | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER STATES THAT A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -8.0/+2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS FLIPPED AND WAS REMOVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO LENS BEING IMPLANTED UPSIDE DOWN. THIS RESOLVED THE PROBLEM. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED AND FLIPPED A 12.6MM VICM5_12.6 -8.00/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. ON THE SAME DAY BUT DIFFERENT SURGERY, THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. THE STATUS OF EYE IS REPORTED AS " BLURRY VISION AT NIGHT, PATIENT SEE HALOES AND GLARE AT NIGHT." CAUSE OF THE EVENT IS REPORTED AS "DEVICE". CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION- UNK. IMPLANT DATE, EXPLANT DATE-UNK. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14887046&pc=QCB | 2023826-2022-02097 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00/+4.50/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14889364&pc=QCB | 2023826-2022-02098 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/+4.50/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688274&pc=QCB | 2023826-2019-01077 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -10/0.5/75 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL AND HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8688384&pc=QCB | 2023826-2019-01078 | 2019-04-01 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -8.0/0.5/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER THE EXPLANT, PATIENTS CURRENT CONDITION IS WELL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10053827&pc=QCB | 2023826-2020-01098 | 2019-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-18.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO A PATIENT RELATED ISSUE AND THE PROBLEM RESOLVED. THE REPORTER STATED " THE PATIENT WAS UNABLE TO ADAPT TO THE GLARE AFTER THE OPERATION AND ASKED TO REMOVE THE LENS AND CANCEL THE OPERATION." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- RESULT CODE IN INITIAL MDR SHOULD BE 213. H6- RESULT CODE 114 SHOULD BE IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138800&pc=QCB | 2023826-2020-01190 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020 AND (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION, THE CAUSE OF THE EVENT IS ATTRIBUTED TO THE DEVICE. "LENS SIZING IS SMALL, CAUSING LENS TO ROTATE." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - IT WAS REPORTED PATIENT ALSO EXPERIENCED DOUBLE VISION WHICH WAS RESOLVED WITH THE LENS EXCHANGE. THIS SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10369229&pc=QCB | 2023826-2020-01678 | NA | Injury | STAAR SURGICAL COMPANY | 2020-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED HIGH GLARE IMMEDIATELY AFTER REMOVING THE DRESSING ONE WEEK AFTER SURGERY AND IS STILL SUFFERING FROM HIGH GLARE. THE REPORTER STATED "ANNOYING GLARE ESPECIALLY AT NIGHT, ESPECIALLY WHEN DRIVING. PATIENT DECLARES AN IMPROVEMENT DURING THE INSTALLATION OF THE PILOCARPINE BUT UPON STOPPING THE GLARE RESUMES."ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628804&pc=QCB | 2023826-2021-03986 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2M VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED PERMANENT LOSS OF BCVA. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (THE WRONG REFRACTION WAS TYPED INTO THE ICL FORM). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. G4: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628810&pc=QCB | 2023826-2021-03990 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/090, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED ABOUT HIS VISION AND IT WAS FOUND THE WRONG LENS WAS IMPLANTED IN ERROR. THE LENS WAS REPLACED WITH THE CORRECT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: AFTER THE LENS WAS REPLACED WITH THE CORRECT LENS, THE PROBLEM WAS RESOLVED. THERE WERE NO SYMPTOMS AND NO PATIENT COMPLAINT. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CORRECTED DATA: B3: 13-SEP-2021. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365144&pc=QCB | 2023826-2022-01479 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Impairment; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4551528&pc=MTA | 2023826-2015-00205 | 2014-12-25 | Injury | STAAR SURGICAL COMPANY | 2015-02-26 | MTA | VISIAN TICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND AN REFRACTIVE SURPRISE WAS NOTED. PATIENT EXPERIENCING BLURRED VISION. THE LENS WAS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH, DEVICE HISTORY RECORD. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: ACCORDING TO THE SURGEON, THE LENS WAS ALIGNED AT 20 DEGREES CCW, WHICH IS 5 DEGREES OFF-AXIS. THE PATIENT EXPERIENCED BLURRED VISION AND THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO ROTATE THE LENS 31 DEGREES ACCORDING TO ABERROMETRY TESTING. THIS RESOLVED THE PROBLEM. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF HE EYE'S MEASUREMENT TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ECT. THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE CAUSE OF THE EVENT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, LENS WORK ORDER SEARCH AND MEDICAL REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. THE LENS REMAINS IMPLANTED. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037536&pc=MTA | 2023826-2016-01482 | 2016-09-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS - -18/1/145 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON DECIDED TO MAKE A LASIK TOUCH-UP FOR BOTH EYES TO CORRECT REFRACTIVE SURPRISE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6231187&pc=MTA | 2023826-2017-00030 | NA | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074352&pc=MTA | 2023826-2016-00909 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+5.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789550&pc=MTA | 2023826-2015-00578 | NA | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789592&pc=MTA | 2023826-2015-00579 | NA | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS A RESIDUAL HYPEROPIC ERROR AND THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): (OTHER): LENS IMPLANTED. EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213981&pc=MTA | 2023826-2018-00135 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/+2.0/134 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. ADDITIONALLY, THE SURGEON STATES THE LENS IS "ON A DEAD VERTICAL POSITION." REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136399&pc=MTA | 2023826-2017-01940 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/X090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ONE YEAR LATER, THE REPORTER INDICATED A REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.5/3.0/090 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ONE YEAR LATER THE REPORTER INDICATED LENS ROTATION AND REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE IMPLANTATION DATE IN THE PREVIOUS SUPPLEMENTAL MDR IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2016". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6326555&pc=MTA | 2023826-2017-00153 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON PATIENT'S LAST VISIT ON (B)(6) 2015, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/20. THE STATUS OF THE EYE WAS A "QUIET EYE" AND "GOOD VAULTING". Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THE PATIENT HAS AN ANTERIOR CHAMBER DEPTH OF 2.93MM. THIS LENS MODEL IS INDICATED FOR USE WITH AN ANTERIOR CHAMBER DEPTH (ACD) OF EQUAL OR GREATER THAN 3.0 MM, AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE. CLAIM #(B)(4). LENS NOT RETURNED. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429909&pc=MTA | 2023826-2017-00487 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -17.0/+6.0/062 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE AND BLURRY VISION. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501471&pc=MTA | 2023826-2017-00653 | 2016-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6, DIOPTER -3.5/3.5/031 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: IN THE INITIAL REPORT, THE LENS MODEL IS DESCRIBED AS A VTICM12.6. THE CORRECT MODEL IS VTICMO12.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. OFF-LABEL USE: ACD < 3.0MM (2.87MM), AGE (B)(6) AT TIME OF IMPLANT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6564060&pc=MTA | 2023826-2017-00804 | 2016-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2, DIOPTER -18.0/+2.5/X178 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393933&pc=MTA | 2023826-2017-00368 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+3.50/168 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAIN IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE NOT EVALUATED BY MANUFACTURE BECAUSE LENS NOT RETURNED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7274885&pc=MTA | 2023826-2018-00224 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8/4/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.00/4.00/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -8.00/4.00/100 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275256&pc=MTA | 2023826-2018-00225 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -6.50/+3.00/087 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6545668&pc=MTA | 2023826-2017-00764 | 2017-04-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -17.00/250/125 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, HOWEVER A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PREVIOUS: LENS DIOPTER -17.00/250/125; CORRECTED: LENS DIOPTER -17.00/2.5/125. DEVICE EVALUATION, PREVIOUS: STATED THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. CORRECTED: NO CLEAR RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. (B)(4). Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6680940&pc=MTA | 2023826-2017-01087 | 2017-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ACCORDING TO THE CUSTOMER, THE SURGEON DOES NOT HAVE A PLAN TO EXCHANGE THE LENS, AND WILL FOLLOW-UP THE PATIENT'S CONDITION WITHIN THE NEXT FEW MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941344&pc=MTA | 2023826-2017-01558 | 2017-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON DURING THE POST-OP VISIT, DATE OF (B)(6) 2017. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952440&pc=MTA | 2023826-2017-01566 | 2017-08-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.5/+3.0/076 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. Manufacturer Narrative: CORRECTED DATA: LENS REMAINS IMPLANTED. DOCTOR WANTS TO OBSERVE FOR A FEW MONTHS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938985&pc=MTA | 2023826-2017-01554 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.5/+3.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970586&pc=MTA | 2023826-2017-01595 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11.5/+1/009 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE DOCTOR HAS DECIDED TO WAIT ANOTHER FEW MONTHS TO SEE WHAT HAS HAPPENED WITH THE REFRACTION. LENS IS STILL CURRENTLY IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970590&pc=MTA | 2023826-2017-01568 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11/1.5/151 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097940&pc=MTA | 2023826-2017-01835 | 2017-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.50/+2.00/080 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097943&pc=MTA | 2023826-2017-01836 | 2017-09-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10.50/+3.50/097 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987323&pc=MTA | 2023826-2017-01633 | 2017-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -3.0/4.5/X070 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE SURGEON IS PLANNING TO EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE LENS WILL BE ROTATED IN (B)(6) 2018. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146626&pc=MTA | 2023826-2017-01941 | 2017-10-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.0/+2.0/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WITH THE POST-OP MANIFEST REFRACTION, SPH. -0.1 AMD CYL. -1.75. THE SURGEON IS PLANNING TO PERFORM LASIK FOR ADDITIONAL CORRECTION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205784&pc=MTA | 2023826-2018-00103 | 2017-12-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -10.5/+6/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395325&pc=MTA | 2023826-2017-00359 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS CURRENTLY UNDER MEDICAL CONSULTATION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429982&pc=MTA | 2023826-2017-00475 | NA | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+2/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. DOCTOR IS MONITORING THE PATIENT STATUS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336540&pc=MTA | 2023826-2018-00411 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2, IMPLANTABLE COLLAMER LENS, -13.50/+2.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398188&pc=MTA | 2023826-2018-00497 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -3.50/5.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THE THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -3.50/5.5/107 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE LEFT EYE (OS) VALUES WERE USED FOR CALCULATION OF RIGHT EYE (OD) LENS REQUIREMENT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398192&pc=MTA | 2023826-2018-00498 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -4.00/3.0/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.00/ 3.0/119 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE RIGHT EYE (OD) VALUES WERE USED FOR CALCULATION OF LEFT EYE (OS) LENS REQUIREMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108230&pc=MTA | 2023826-2016-01651 | 2016-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 3.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/074 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: PREVIOUSLY, IT WAS REPORTED THAT THE CONCLUSION CODE. THIS IS INCORRECT. THE CORRECT CONCLUSION CODE WAS REPORTED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT MODEL NUMBER AND LENGTH OF THE LENS IS VTICM5 13.2 IMPLANTABLE COLLAMER LENS. REPORT SOURCE: USER FACILITY TO DISTRIBUTOR - PREVIOUS SELECTION 'USER FACILITY' IS INCORRECT, THE CORRECT SELECTION SHOULD BE 'DISTRIBUTOR.' (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223864&pc=MTA | 2023826-2017-00021 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/066 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089425&pc=MTA | 2023826-2017-01830 | 2017-11-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A 13.2 MM VTICM5_13.2, -4.5/+4.5/075 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AS WELL AS REFRACTIVE SURPRISE WAS NOTICED. THE LENS IS PLANNED TO BE REPOSITIONED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227635&pc=MTA | 2023826-2017-00016 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/78 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201639&pc=MTA | 2023826-2018-00086 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -1.5/+1.50/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND EXCESSIVE VAULT. IN ADDITION, THE ICL WAS IMPLANTED "IN A PIGGY-BACK CONDITION." REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER (B)(6) YEARS OF AGE AT TIME OF IMPLANT ((B)(6) YRS), PREVIOUS CORNEAL OR REFRACTIVE SURGERY]. (B)(4) | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336748&pc=MTA | 2023826-2018-00377 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052129&pc=MTA | 2023826-2018-01618 | 2018-10-08 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -2.50/6.0/168 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND LOSS OF BCVA; THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE PATIENT IS DOING FINE AFTER EXPLANTATION. Manufacturer Narrative: CORRECTED DATA: "KERATOCONUS" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE HAPTIC IS MISSING. (B)(4). Manufacturer Narrative: CORRECTED DATA: RESULT CODE CORRECTED TO: "114" ONLY IN THE SUPPLEMENTAL MDR#2, PREVIOUS CODE NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051045&pc=MTA | 2023826-2017-01755 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/099 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD LOW VAULTING AND LENS ROTATION. THE PATIENT COMPLAINED OF BLURRY VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER STATED THAT THE EXCHANGE IS LIKELY TO BE IN (B)(6) 2018. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877224&pc=MTA | 2023826-2017-01462 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -13.5/+1.5/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/25 AND UCVA WAS 20/25 ON (B)(6) 2016. PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135239&pc=MTA | 2023826-2017-01925 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/+5.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIROR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135248&pc=MTA | 2023826-2017-01926 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/+6.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIOR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227640&pc=MTA | 2023826-2017-00017 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/2.0/152 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. ADDITIONALLY, THE COMPLAINT FORM STATES, "CUSTOMER MIXED UP THE CALCULATION DATA BETWEEN OD AND OS." Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242804&pc=MTA | 2023826-2017-00041 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/82 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AMD THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242817&pc=MTA | 2023826-2017-00039 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/81 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: CORRECT OUTCOME ATTRIBUTED TO ADVERSE EVENT IS OTHER SERIOUS (DEVICE REMAINS IMPLANTED). (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783494&pc=MTA | 2023826-2018-01202 | 2018-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS NOT SATISFIED WITH THE VISION BECAUSE IT IS DIFFERENT FROM THE EXPECTED REFRACTION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071556&pc=MTA | 2023826-2018-01682 | 2018-09-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.00/4.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090071&pc=MTA | 2023826-2018-01702 | 2018-09-24 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+4.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.0/4.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND POST OP DATA SHOWS LOSS OF VISUAL ACUITY. THE LES REMAINS IMPLANTED WITH LASER TREATMENT PLANNED FOR THE FUTURE. DEVICE CODE 4110: LENS WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT TYPES EVENTS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090052&pc=MTA | 2023826-2018-01684 | 2018-10-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.5/+6.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DATE OF EVENT CORRECTED TO: "UNKNOWN". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372545&pc=MTA | 2023826-2019-00314 | 2018-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227420&pc=MTA | 2023826-2019-00048 | 2018-11-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. CORRECTED TO: "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241988&pc=MTA | 2023826-2019-00086 | 2018-12-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/4.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED ON (B)(6) 2018. REMOVAL OF THE LENS AND EXCHANGE IS PLANNED ONCE NEW LENS IS CALCULATED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403806&pc=MTA | 2023826-2019-00366 | 2019-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.00/6.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER STATES, "WE DON'T KNOW WHY SHE STILL HAS A REFRACTIVE CYLINDER ERROR ALTHOUGH THE LENS IS NOT ROTATED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372275&pc=MTA | 2023826-2019-00302 | 2019-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+6.0/086 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO ASTIGMATIC AXIS ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206969&pc=MTA | 2023826-2021-00112 | 2020-12-18 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE, LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS REPOSITIONED AND THE THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY". H6 - HEALTH EFFECT IMPACT CODE REPORTED AS 2199 - UPDATED/CHANGED TO 4629 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DB6 DATE: LENS NOT EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17218627&pc=MTA | 2023826-2023-02556 | 2023-04-11 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -13.50/1.5/062 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4) Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4) | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783466&pc=MTA | 2023826-2018-01203 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. DATE RECEIVED SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908155&pc=MTA | 2023826-2018-01426 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.00/+3.0/080 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908160&pc=MTA | 2023826-2018-01427 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.50/+3.0/X044 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7916208&pc=MTA | 2023826-2018-01451 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/4.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, AS OF THE DATE OF MDR REPORTING NO INTERVENTION HAS BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928342&pc=MTA | 2023826-2018-01467 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/X092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8143302&pc=MTA | 2023826-2018-01787 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/166 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND THE LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS REMAINS IMPLANTED. THE PATIENTS CURRENT CONDITION IS VISUAL ACUITY IS NOT AS GOOD AS IT SHOULD BE BUT THE PATIENT CONDITION IS OK. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459217&pc=MTA | 2023826-2022-00392 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, "PREPARE FOR FOLLOW-UP OBSERVATION AND VISUAL FUNCTION TRAINING." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY:UNK. SEX: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15440641&pc=MTA | 2023826-2022-03146 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO HAD AN IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A -6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A PATIENT WHO HAD A 13.2MM VTICMO13.2 -16.50/1.5/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A 6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010741&pc=MTA | 2023826-2022-04480 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010748&pc=MTA | 2023826-2022-04486 | NA | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: CORRECTION LENS DIOPTER IS -11.0/1.0/118 . CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920600&pc=MTA | 2023826-2021-04619 | 2021-11-03 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS DISLOCATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: RECEIVED CQ ON (B)(6) 2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -17.50/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS (13.7MM VTICMO13.7 DIOPTER -17.0/+3.0/105) ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. PROBLEM RESOLVED. REPORTEDLY, "PATIENT IS HAPPY!" THE CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. CLAIM#: (B)(4) IS REGARDING EXCESSIVE VAULTING WHICH IS UNRELATED TO THIS COMPLAINT. NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. DATE OF EVENT-UNK. PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. PRODUCT INFO, IMPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372531&pc=MTA | 2023826-2019-00316 | 2018-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+5.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824424&pc=MTA | 2023826-2018-01262 | 2018-06-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THERE WAS REDUCTION OF BEST CORRECTED VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4).. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765095&pc=MTA | 2023826-2018-01173 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.0/+2.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8351772&pc=MTA | 2023826-2019-00274 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/4.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SUPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354792&pc=MTA | 2023826-2019-00275 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/4.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SURPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284774&pc=MTA | 2023826-2017-00156 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS,-4.0/+1.5/17 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LEFT EYE WAS IMPLANTED WITH THE CALCULATED RIGHT EYE (OD) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284878&pc=MTA | 2023826-2017-00155 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+1.5/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE RIGHT EYE (OD) WAS IMPLANTED WITH THE CALCULATED LEFT EYE (OS) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4). Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508705&pc=MTA | 2023826-2017-00682 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00/+2.00/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE CUSTOMER STATED THAT THE DR. MADE A MISTAKE IN CALCULATION OF THE OPTIC. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON PERFORMED THE LIMBAL INCISION TO REDUCE ASTIGMATISM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711503&pc=MTA | 2023826-2017-01175 | 2017-04-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, DIOPTER -15.5/+2.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, REFRACTIVE SURPRISE, AND UNCOMFORTABLE VISION. THE LENS WAS REPOSITIONED THREE TIMES ON (B)(6) 2015, (B)(6) 2016, AND (B)(6) 2017 BUT UNSUCCESSFUL, THE LENS WAS ROTATED AGAIN. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135234&pc=MTA | 2023826-2017-01923 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.5/+2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525079&pc=MTA | 2023826-2016-00333 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/130 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE, BLURRED VISION AND GLARES/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559295&pc=MTA | 2023826-2017-00792 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED AT ONE DAY POST-OP, THE LENS HAD AN EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT EXPERIENCED GLARE/HALOES. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919768&pc=MTA | 2023826-2015-00838 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS OBSERVED. PATIENT EXPERIENCED BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312180&pc=MTA | 2023826-2015-01675 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE, HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH PERFORMED AND FOUND NO SIMILAR COMPLAINTS. LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO,LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685199&pc=MTA | 2023826-2016-00756 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/008 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND A REFRACTIVE SURPRISE. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831916&pc=MTA | 2023826-2016-00948 | NA | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/113 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: (DEVICE FROM SAME LOT EVALUATED); WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216613&pc=MTA | 2023826-2016-01850 | 2016-10-09 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+5.00/054 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER TWO WEEKS, THE LENS WAS REPOSITIONED. ADDITIONALLY, ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ROTATION, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE: PREVIOUS DATE WAS INCORRECT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LOSS OF BCVA, LENS EXPLANT. (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540150&pc=MTA | 2023826-2016-00375 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS DISLOCATION/SUBLUXATION AND ROTATION. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND BLURRED VISION. THE LENS WAS REPOSITIONED AND THE INCISION WAS ENLARGED. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THAT LENS ALSO TILTED AND SUBLUXED - SEE MFR REPORT # 2023826-2016-00376. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC SCRATCHED; DIMENSIONAL INSPECTION FOUND THE LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047717&pc=MTA | 2023826-2015-01115 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON 11/10/2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. LENS EXPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540385&pc=MTA | 2023826-2016-00376 | NA | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS DISLOCATED/SUBLUXED. THE LENS REMAINS IMPLANTED. THIS WAS THE REPLACEMENT LENS - SEE MFR REPORT # 2023826-2016-00375 FOR THE FIRST LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432678&pc=MTA | 2023826-2016-00150 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/052 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/25). THE LENS WAS EXCHANGED FOR A SHORTER LENS (VICMO12.6, -11.0 DIOPTER) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND SCRATCHES. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5085796&pc=MTA | 2023826-2015-01195 | 2015-07-27 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+5.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY. THE PATIENT ALSO EXPERIENCED PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), BLURRED VISION AND INFLAMMATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT DATE OF BIRTH - UNK. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671131&pc=MTA | 2023826-2016-00738 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/20). THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THE LENS DESCRIPTION IS -18.0/+2.0/X087 (SPHERE/ CYLINDER/ AXIS). PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/40. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC AND HAPTIC WERE TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867161&pc=MTA | 2023826-2016-01048 | 2016-02-12 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION/DISLOCATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE EXCHANGED LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CORRECTED DATA: THE CORRECT REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834798&pc=MTA | 2023826-2016-00996 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+5.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE - FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. THE PATIENT IS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834802&pc=MTA | 2023826-2016-00997 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+5.0/125 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001950&pc=MTA | 2023826-2016-01394 | 2016-08-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem; Insufficient Information | Blurred Vision; Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/001 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/SIZE LENS BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, BUT FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT. DATE OF BIRTH AND WEIGHT: - UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15341923&pc=MTA | 2023826-2022-03007 | 2022-07-27 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT HAS BEEN REPORTED THE LENS HAS ROTATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.5/171 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHORT DIAMETER". H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402698&pc=MTA | 2023826-2022-01529 | 2022-01-19 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | NA | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN OVERSIZED ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B3 - DATE OF EVENT (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -17.00/2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE LOW VAULT WITH ROTATION, LENS "DISLOCATION"/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER MARKED THE CAUSE OF THE EVENT AS THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. D6A - IMPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - DEVICE CODE: 1401 -2923 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - TYPE IF INVESTIGATION CODE: 4110-41111 :" LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE ADDED TO INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113371&pc=MTA | 2023826-2018-01753 | 2016-05-02 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -5.00/+5.0/150 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA) AND BLURRED VISION. THE SURGEON HAS NOT DECIDED AT THIS TIME IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THERE WAS NO CATARACT, NO ICL ROTATION AND NO PROGRESSION OF KERATOCONUS. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426150&pc=MTA | 2023826-2017-00460 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/X009 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. AT ONE DAY POST-OP THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED - THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20 AND THE ISSUE WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635312&pc=MTA | 2023826-2015-00382 | 2014-11-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THERE WAS AN INCIDENT OF REFRACTIVE SURPRISE AS WELL AS LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT):UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600864&pc=MTA | 2023826-2015-00312 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (REFRACTIVE SURPRISE); (CHOROIDAL EFFUSION-SUSPECT); BLURRED VISION. (LENS IMPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089928&pc=MTA | 2023826-2016-01601 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+6.0/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720709&pc=MTA | 2023826-2016-00373 | NA | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WAS REPOSITIONED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED AND IS PLANNED TO EXCHANGE WITH ANOTHER LENS. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5509447&pc=MTA | 2023826-2016-00286 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF BLURRED VISION AND THE LENS EXCHANGE. (B)(4). THERE WAS NO BREAK IN THE MATERIAL DUE TO MATERIAL INTEGRITY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION NARRATIVE WAS INCORRECT. VISUAL INSPECTION DID NOT FIND THE OPTIC TORN. VISUAL INSPECTION FOUND THE HAPTIC TORN. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918200&pc=MTA | 2023826-2017-01534 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 DIOPTER -15.0/+4.0/091 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON NOTED LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND LENS REPOSITIONING. THE CLAIM REPORT STATES THERE IS A PLAN TO EXCHANGE THE LENS ("RE-IMPLANTATION"), AND THE SAME REPORT STATES THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: LENS REMAINS IMPLANTED; PLAN TO EXCHANGE. LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: PREVIOUS: LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. - UPDATED: THE LENS WAS EXCHANGED FOR A LARGER, DIFFERENT DIOPTER LENS ON (B)(6) 2016. THE LENS EXCHANGE RESOLVED THE PROBLEM. EXPLANTED DATE: IT IS CORRECTED FROM (B)(6) 2016 TO (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS REPOSITIONING. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609744&pc=MTA | 2023826-2017-00917 | 2016-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WERE NOTED. THE LENS WAS REPOSITIONED ON (B)(6) 2016, BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER LENS OF A DIFFERENT MODEL/DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/100 ON (B)(6) 2016. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC TORN/BROKEN. (B)(4). CONCLUSION CODE AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE AT THE TIME OF IMPLANTATION. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6618987&pc=MTA | 2023826-2017-00910 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION. THE LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2016, BUT UNSUCCESSFUL. ON (B)(6) 2016 THE LENS WAS EXCHANGED. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144234&pc=MTA | 2023826-2017-01947 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS REPORTED AS HAVING ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED POOR VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS SUCCESSFULLY REPOSITIONED AND REMAINS IMPLANTED. AT FIRST DAY POST-OP, THE PATIENT HAD ENLARGED PUPILS BUT IS MUCH HAPPIER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970600&pc=MTA | 2023826-2017-01597 | 2017-07-16 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -13/+4.0/154 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, IT WAS NOTICED THAT THE LENS ROTATED. ON (B)(6) 2017, THE SURGEON REDIALED THE LEFT EYE. THE PATIENT WAS HAPPY FOR A WEEK, THEN THE LENS ROTATED AGAIN ON (B)(6) 2017. AS OF (B)(6) 2017, THE VAULT IS 410 MICRONS, BCVA IS 20/80 AND UCVA IS LESS THAN 20/200. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE SURGEON STATED THAT PATIENT DIDN`T WANT EXCHANGE FOR A NON-TORIC ICL. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089423&pc=MTA | 2023826-2017-01829 | 2017-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.0/+5/086 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE WERE NOTICED. THE LENS WAS THEN REPOSITIONED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180692&pc=MTA | 2023826-2018-00039 | 2017-11-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/+3.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE, LENS ROTATION, BLURRED VISION, AND "RESIDUAL POWER." Manufacturer Narrative: THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. "PATIENT IS COMFORTABLE WITH 6/6 VISION." (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180665&pc=MTA | 2023826-2018-00035 | 2017-12-03 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+6.0/099 (SPHERE/ CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REPOSITIONED. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT IS REPORTED. LENS REPOSITIONING DID NOT RESOLVE THE PROBLEM. SURGEON STATES THAT THE LENS ROTATED POST-OPERATIVELY. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED REFRACTIVE SURPRISE IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136378&pc=MTA | 2023826-2017-01938 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/006 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, THE LENS HAD ROTATED AND WAS REPOSITIONED AGAIN. THE LENS WAS NOT STABLE AND TO DATE, REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/6.0/006 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, IT WAS REPORTED THAT THE LENS WAS NOT STABLE AND "ROTATED AGAIN BUT LESS". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701661&pc=MTA | 2023826-2017-01156 | 2012-01-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -12.0/+1.0/91 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME. THE CUSTOMER STATED CURRENTLY THE LENS REMAINS IMPLANTED, AND THE DR. PLANS TO EXCHANGE THE LENS WITH HIGHER REFRACTIVE POWER TO RESOLVED THE PROBLEM FOR THE PATIENT. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.0/91 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR SAME MODEL DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: SHOULD HAVE BEEN RECORDED. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: FUNCTIONAL INSPECTION WAS PERFORMED BUT DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6307761&pc=MTA | 2023826-2017-00166 | 2012-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT HAS EXPERIENCED REFRACTIVE CHANGE OVERTIME. AS OF THE DATE OF THIS MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428720&pc=MTA | 2023826-2017-00464 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/105 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE SURGEON ATTEMPTED TO REPOSITION THE LENS BUT THE RESULT WAS NOT STABLE. THE LENS WAS THEN EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE SURGEON IS PLANNING TO PERFORM CORNEAL REFRACTIVE SURGERY TO REDUCE THE ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED WAS RETURNED IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, LENS HAPTIC BROKEN, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DRIED, DISCOLORED MATERIAL). (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442677&pc=MTA | 2023826-2017-00516 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR ANOTHER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2017 WAS 20/25. DEVICE EVALUATION: THE DEVICE WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393825&pc=MTA | 2023826-2017-00364 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+0.5/86 DIOPTER IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR THE SAME LENS MODEL WITH A DIOPTER -7/ +0.5/84 DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE NOT APPLICABLE, PATIENT'S ACD IS ABOVE 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429971&pc=MTA | 2023826-2017-00470 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+1.5/31 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND ENLARGEMENT OF INCISION SURGERY. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS THEN EXCHANGED WITH THE SAME SIZE, AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED.Event Description: THERE WAS NO ENLARGEMENT OF THE INCISION ON (B)(6) 2016. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THERE WAS NO ENLARGEMENT OF INCISION ON 08/08/2016 AS PREVIOUSLY REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (MORE THAN 45 YEARS OF AGE AT THE TIME OF LENS IMPLANT). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6381058&pc=MTA | 2023826-2017-00320 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, DIOPTER -11.50/+3.00/0.87 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ALSO ON (B)(6) 2016, THE SURGEON RE-POSITIONED THE LENS. ON (B)(6) 2016 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (PREVIOUS): REPOSITIONING WAS PERFORMED ON (B)(6) 2016. (CORRECTED) REPOSITIONING WAS PERFORMED ON (B)(6) 2016. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6414543&pc=MTA | 2023826-2017-00433 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE DOCTOR ALSO STATED THE PATIENT EXPERIENCED RESIDUAL MIXED ASTIGMATISM, AND BELIVES THAT THE LENS MAY BE OVERROTATED. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4) NO CODE AVAILABLE (RESIDUAL MIXED ASTIGMATISM). RESULT CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395339&pc=MTA | 2023826-2017-00362 | 2017-02-06 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.0/76 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND SIMILAR DIOPTER DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND THICK, STICKY, CLEAR RESIDUE ON PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487511&pc=MTA | 2023826-2017-00629 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+0.5/X013 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS.(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473540&pc=MTA | 2023826-2017-00593 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+3.0/X085 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6667615&pc=MTA | 2023826-2017-01034 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, DIOPTER -2.0/+3.5/005 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. UPDATED BASED ON PRODUCT EVALUATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701659&pc=MTA | 2023826-2017-01166 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/83 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULT. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657557&pc=MTA | 2023826-2017-01018 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.5/20 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE AND DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674235&pc=MTA | 2023826-2017-01075 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/080 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT ALSO EXPERIENCES HYPEROPIA. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725634&pc=MTA | 2023826-2017-01205 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.00/+3.0/X096 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON PLANS TO PERFORM AN ADDITIONAL LENS ROTATION. Manufacturer Narrative: STATES THAT AN ADDITIONAL LENS ROTATION IS PLANNED. UPDATED/CORRECTED: REPORTEDLY, THE LENS WAS REPOSITIONED ON (B)(6) 2017. AS OF (B)(6) 2017 THE "PATIENT IS HAPPY AFTER REPOSITIONING." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725726&pc=MTA | 2023826-2017-01206 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.50/+3.50/X192 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON WANTS TO KNOW THE AMOUNT OF REQUIRED ROTATION FOR PLANO REFRACTION. Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE LENS ROTATED BACK AGAIN AFTER REPOSITIONING, SO THE SURGEON PLANS TO MAKE "BIOPTIC LVC" TO "CORRECT RX." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711486&pc=MTA | 2023826-2017-01182 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -0.5/+6/109 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, MISALIGNMENT. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ATTEMPTS TO CONTACT TO THE CUSTOMER FOR ADDITIONAL INFORMATION HAVE BEEN ON GOING. Manufacturer Narrative: THE IMPLANTED VTICM5_12.6 LENS WAS REPOSITIONED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE SURGEON DECIDED NOT TO DO A SECOND REPOSITIONING, THE PATIENT IS HAPPY WITH THE SLIDE CYLINDER AT THE MOMENT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/22. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970595&pc=MTA | 2023826-2017-01603 | 2017-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -1.5/1.5/175 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS THEN EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS PRESENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BROKEN. METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -1.5/1.5/175 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A SIMILAR LENS WAS THEN IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE DATE OF EXPLANT IS CORRECTED FROM (B)(6) 2017 TO (B)(6) 2017. THE DEVICE MANUFACTURED DATE IS CORRECTED FROM 25-JUN-2015 TO 26-JUN-2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050900&pc=MTA | 2023826-2017-01707 | 2017-07-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.50/+5.0/070 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335039&pc=MTA | 2023826-2018-00340 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.50/+4.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6) 2018, THE LENS WAS REMOVED DUE PATIENT'S DISSATISFACTION WITH HIS VISUAL ACUITY WHICH WAS 20/30 AFTER THE SURGERY. HIS PRE-OP VISUAL ACUITY IS 20/20. THE LENS WAS EXCHANGED FOR LONGER DIFFERENT LENS. THE PATIENT CURRENT VISUAL ACUITY IS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND HAD PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952444&pc=MTA | 2023826-2017-01564 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+1.0/087 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS. THE PATIENT'S UCVA IS 20/30 AS OF (B)(6) 2017 AND MANIFEST REFRACTION IS +0.50/-1.25/30 AND 20/20 AS OF (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855452&pc=MTA | 2023826-2017-01420 | 2017-08-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.50/4.0/91 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ADDITIONAL REFRACTIVE TREATMENT (ENCHANCEMENT) WAS DONE. THE SURGEON DECIDED TO WAIT A FEW WEEKS AFTER THE SURGERY TO ALLOW THE EYE AND CORNEA TO HEAL AND GET STABLE. THE SURGEON DECIDED THAT NO EXCHANGE IS NEEDED, BUT WILL PERFORM A LASER TOUCH UP IF NECESSARY. Manufacturer Narrative: ADDITIONAL REFRACTIVE TREATMENT (ENHANCEMENT) WAS NOT PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6912388&pc=MTA | 2023826-2017-01496 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2017-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.0/+1.5/111 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: LENS EXCHANGE WAS PERFORMED ON (B)(6) 2017. PATIENT VISUAL ACUITY WAS NOT CORRECTED COMPLETELY AFTER EXCHANGE. THE DOCTOR HAS DECIDED TO PERFORM LASIK FOR ADDITIONAL CORRECTION. PATIENT POST-OP, PRE-LASIK, MR IS SPH. -0.1 AND CYL. -1.75. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941354&pc=MTA | 2023826-2017-01559 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/079 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON, INCIDENT DATE REPORTED AS (B)(6) 2017. THE SURGEON IS CONSIDERING LENS EXCHANGE OR ROTATION. Manufacturer Narrative: UPDATED INFORMATION: LENS REPOSITIONING PERFORMED. SURGEON STATES, "VISUS IS NOW 1.25." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). [ANTERIOR ENDOTHELIUM CHAMBER DEPTH LESS THAN 3.0MM. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135226&pc=MTA | 2023826-2017-01920 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/180 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT CURRENTLY IS NOT PLANNED TO HAVE ANY ADDITIONAL SURGERY. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050082&pc=MTA | 2023826-2017-01757 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -0.50/+5.5/089 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". THE DIOPTER VALUE OF THE INTIAL LENS IMPLANTED IS CORRECTED FROM "-0.50/+5.5/089 SPHERE/ CYLINDER/ AXIS" TO " -0.5/+5.5/084 SPHERE/ CYLINDER/ AXIS". DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050099&pc=MTA | 2023826-2017-01756 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -1.50/+4.5/089 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM " 26-OCT-2017" TO "25-OCT-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) IS CORRECTED FROM "-1.50/ +4.5/ 089 SPHERE/ CYLINDER/ AXIS" TO "-2.0/ +4.5/ 072 SPHERE/ CYLINDER/ AXIS". (B)(4) Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156010&pc=MTA | 2023826-2017-01970 | 2017-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484262&pc=MTA | 2023826-2018-00640 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+6.0/152 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), DECREASED VISUAL ACUITY, BLURRED VISION, ENLARGEMENT OF PI OR ADDITION OF NEW PI PERFORMED VIA YAG, AND ADMINISTRATION OF TOPICAL 2% PILOCARPINE. THE STATUS OF THE EYE IS LISTED AS SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER. ADDITIONAL INFORMATION-FIRST DAY POST-OP THE VISUAL ACUITY WITHOUT CORRECTION WAS "ALMOST 20/20," BUT "A FEW DAYS AFTER" IT WAS 20/50. POST-YAG LASER AND TOPICAL PILOCARPINE THE VA WAS 20/30 BUT "TODAY" IS 20/80 WITHOUT CORRECTION. THE LAST VISIT DATE IS LISTED AS (B)(6) 2018. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED LOSS OF BCVA. "PELLUCID DEGENERATION, PREVIOUS CORNEAL SURGERY (NON- REFRACTIVE)" IS CORRECTED TO "PELLUCID DEGENERATION". (B)(4) Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT: ON (B)(4) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS. THE PROBLEM IS RESOLVED. (SECONDARY SURGERY, LENS EXCHANGE). CLAIM # (B)(4). Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: ON (B)(6) 2019 THE SURGEON REPORTS THE PATIENT'S BCVA IS 20/30 IN HIS RIGHT (OD) EYE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA, ENLARGEMENT OR ADDITION OF PI-YAG, ADMINISTRATION OF TOPICAL 2% PILOCARPINE, SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER). CONCLUSION CODE: OFF-LABEL USE [PREVIOUS HISTORY: PELLUCID DEGENERATION, CORNEAL SURGERY (NON-REFRACTIVE)] (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403701&pc=MTA | 2023826-2018-00514 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/050 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2018 WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY THE LENS REMAINS IMPLANTED AND "VISION IS STILL BLURRY." Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED AGAIN ON (B)(6) 2018. AFTER REPOSITIONING, THE PATIENT IS "VERY HAPPY WITH THE OUTCOME" AND "HE WAS READING WITH NO EFFORT AT ALL." THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE : NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594283&pc=MTA | 2023826-2018-00894 | 2018-05-21 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/+5.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. REPORTEDLY, THE SURGEON PERFORMED THE "WRONG CALCULATION." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE : USER ERROR (SURGEON PERFORMED WRONG CALCULATIONS) (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314688&pc=MTA | 2023826-2017-00158 | NA | Injury | STAAR SURGICAL COMPANY | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AFTER THE SURGERY WITH AN UNCORRECTED VISUAL ACUITY (UCVA) OF 6/36 AND A VAULT VALUE OF 250 MICROMETERS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. SECONDARY SURGICAL INTERVENTION, EXPLANT, LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673303&pc=MTA | 2023826-2017-01052 | NA | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.5/+2.0/X174 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ASTIGMATISM WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4). Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT: "MALFUNCTION" INSTEAD OF "SERIOUS INJURY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674075&pc=MTA | 2023826-2017-01054 | NA | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.0/+6.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BLURRY VISION WAS REPORTED. THE LENS WAS THEN ROTATED 10 DEGREES ANTI CLOACKWISE ON (B)(6) 2017 WHICH IMPROVED UNAIDED VISION FROM 0.7 TO 0.9. Manufacturer Narrative: CORRECTED DATA: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED. | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795668&pc=MTA | 2023826-2017-01318 | NA | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6)2016 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS, AND THE PROBLEM WAS RESOLVED. THE DR. STATED THAT WHEN HE ORDERING THE PRIMARY LENS, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS PRODUCT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA 510(K):_THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE SURGEON STATED THAT WHEN ORDERING PRIMARY ICL, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201657&pc=MTA | 2023826-2018-00093 | NA | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): -THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201662&pc=MTA | 2023826-2018-00095 | NA | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5/+4.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419742&pc=MTA | 2023826-2018-00526 | NA | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. (B)(4). AGE IS (B)(6) YEARS, NOT (B)(6) YEARS AS PREVIOUSLY REPORTED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE IN SPECIFICATION. CONCLUSION: - THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF KERATOCONUS AND PREVIOUS CORNEAL SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557296&pc=MTA | 2023826-2018-00803 | NA | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED PATIENT HAD A 12.5MM, TICM125V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT RECENTLY CAME BACK WITH REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADD'L INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE IMPLANTED DATE IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2007". ADDITIONAL DATA: THE REPORTER INDICATED THAT THE PATIENT HAD A 12.5MM TICM145V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2007. THE PATIENT RECENTLY HAD REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, EXPLANT AND NEW LENS IMPLANTED LATER. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4903447&pc=MTA | 2023826-2015-00807 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 AND REFRACTIVE CHANGE WAS NOTED ON (B)(6) 2015, DUE TO MYOPIA PROGRESSION. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE FOR A DIFFERENT POWER. THE PATIENT'S LAST BCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) REFRACTIVE SURPRISE. CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT YET EXCHANGED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6117561&pc=MTA | 2023826-2016-01637 | 2016-10-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/046 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGE FOR THE SAME MODEL, SIZE BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR THE SAME MODEL, SIZE, DIFFERENT DIOPTER LENS. (B)(4): BLURRED VISION/REFRACTIVE SURPRISE. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6682110&pc=MTA | 2023826-2017-01090 | 2017-05-30 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -1.5/+3.0/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. APPROXIMATELY TWO MONTHS AFTER RECOVERY, PATIENT WAS HIT IN THE EYE BY A CHILD. THE PATIENT'S VISION WORSENED AND PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. PHYSICIAN REPOSITIONED THE LENS TWICE WHICH DID NOT RESOLVE THE ISSUE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. SURGEON PLANS TO DO ANOTHER REPOSITION AFTER CONSULTING WITH THE MEDICAL AFFAIRS TEM Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS REPOSITIONED ACCORDING TO RE-CALCULATION RESULTS. THE PATIENT IS HAPPY, THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4) CONCLUSION CODE:OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.82MM)]. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5250502&pc=MTA | 2023826-2015-01535 | 2015-01-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/X075 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. REFRACTIVE SURPRISE WAS NOTED IN (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL.THIS RESOLVED THE PROBLEM. POST-OP, UCVA WAS 20/20 SEE MFR. #2023826-2015-01534 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). CORRECTED DATA: BEST ESTIMATE IS (B)(6) 2015 ((B)(6) 2015). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00/1.0/X075 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325949&pc=MTA | 2023826-2015-01712 | 2015-05-19 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/063 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) . DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK EVENT PROBLEM AND EVALUATION CODES: PATIENT: NEW INCISION, SECONDARY SURGERY, REFRACTIVE SURPRISE. DEVICE: EXPLANTED. METHOD: WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654713&pc=MTA | 2023826-2016-00656 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/122 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE DIOPTER MEASUREMENT IS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596424&pc=MTA | 2023826-2016-00489 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.5/+5.0/076 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT HAD LENS EXCHANGE, AND PROBLEM WAS RESOLVED. PATIENT IS SLIGHTLY MONOVISION. PREVIOUS REPORT SOURCE, USER FACILITY, WAS INCORRECT. CORRECT INFORMATION IS DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596425&pc=MTA | 2023826-2016-00490 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+5.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE DISCARDED BY USER. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917266&pc=MTA | 2023826-2016-01182 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/118 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/14. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: DATE ON FOLLOW-UP #1 IS INCORRECT. CORRECT DATE WAS 08/29/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5324056&pc=MTA | 2023826-2015-01717 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+3.5/026 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT'S VISION WAS BLURRED, WITH SHADOW. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH ANOTHER LENS (ICL), SAME MODEL, LENGTH AND -7.5 DIOPTER, THIS RESOLVED THE PROBLEM. THE PATIENT WAS DOING WELL WITH BEST CORRECTED VISUAL ACUITY (BCVA) 6/6. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (HAIR,FIBERS) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089924&pc=MTA | 2023826-2016-01600 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+6.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/16. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5730070&pc=MTA | 2023826-2016-00818 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BLURRED VISION WAS REPORTED. LENS REPOSITIONING WAS PERFORMED AND THE PROBLEM WAS RESOLVED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074344&pc=MTA | 2023826-2016-00908 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+5.0/113 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ON (B)(6) 2016, THE REPORTER STATED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE ASSOCIATED WITH LENS MISALIGNMENT. FIVE MONTHS AFTER SURGERY, THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO REPOSITION THE LENS. TICL WAS STILL FOUND TO BE OFF-AXIS AFTER REPOSITIONING AND THE SURGEON DECIDED TO REMOVE THE LENS ON (B)(6) 2016 AND EXCHANGE THE LENS FOR LONGER ONE WITH THE SAME POWER. THIS SOLVED THE PROBLEM. DEVICE EVALUATION - RETURNED PRODUCT EVALUATION FOUND THE LENS WAS RETURNED DRY AND BENT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). REPORT SOURCE: "USER FACILITY" TO "DISTRIBUTOR". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799981&pc=MTA | 2023826-2016-00931 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/042 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972072&pc=MTA | 2023826-2016-01271 | 2016-08-21 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972074&pc=MTA | 2023826-2016-01270 | 2016-08-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059254&pc=MTA | 2023826-2016-01537 | 2016-09-19 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223834&pc=MTA | 2023826-2017-00013 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2010. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EXPLANT DATE: N/A. (B)(4). LENS NOT RETURNED. | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867090&pc=MTA | 2023826-2016-01047 | NA | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACDTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978094&pc=MTA | 2023826-2017-01614 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/X063 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS NOTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. (B)(4). ADDITIONAL CONCLUSION CODE FOR OFF-LABEL USE ADDED (PREVIOUS CORNEAL/REFRACTIVE SURGERY). (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A". (B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6927297&pc=MTA | 2023826-2017-01530 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/175 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED WITH LOW VAULTING), THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY AND THE SURGEON NOTED A REFRACTIVE SURPRISE. THE LENS WAS ALSO NOTED TO HAVE TORN DURING INSERTION. THE LENS WAS REPOSITIONED ON (B)(6) 2017 BUT HAD ROTATED AGAIN, WITH LOSS OF BEST-CORRECTED VISUAL ACUITY AFTER REPOSITIONING INTERVENTION. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/60. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DOCTOR HAS REMOVED THE LENS AFTER MAKING SURE THAT THE TORN HAPTIC AFFECTS THE STABILITY OF LENS IN SULCUS. THE DOCTOR WILL RE-EVALUATE IN A MONTH AFTER REMOVAL. WORK ORDER SEARCH: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS REMOVED ON (B)(6) 2017. THE REPORTER STATED THAT THE DOCTOR IS EXPECTED TO RE-ASSESS THE PRE-OP REFRACTION AND MEASUREMENTS. THIS WAS SCHEDULED TO HAPPEN IN (B)(6) BUT MAYBE PATIENT DID NOT COME BACK YET. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6940698&pc=MTA | 2023826-2017-01493 | 2017-08-28 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.0/+2.5/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Repositioning | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955554&pc=MTA | 2023826-2017-01578 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/X135 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED A REFRACTIVE SURPRISE. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER IS CORRECTED FROM "(B)(6)" TO "(B)(6)" IN THE INITIAL. THE SURGEON NOTED A REFRACTIVE SURPRISE CAUSED DUE TO "WRONG ROTATION OF LENS". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: "CURRENTLY, THE LENS REMAINS IMPLANTED." ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR POWER LENS. REPORTEDLY, "THE OUTCOME IS VERY GOOD. UCVA 20/20." (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487562&pc=MTA | 2023826-2018-00673 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED 3 TIMES BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT HAD LASER TOUCH UP ON (B)(6) 2017. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A LONGER LENS. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED INFORMATION: INJECTION MODEL SYSTEM: FTP; INJECTOR MODEL: MSI-PF; CARTRIDGE MODEL: SFC-45 NOT IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594290&pc=MTA | 2023826-2018-00890 | 2018-05-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+2.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ADDITIONALLY, THE SURGEON STATES THE LENS WAS REPOSITIONED ON (B)(6) 2018; EACH TIME THE LENS WAS REPOSITIONED IT ROTATED AT A 90 DEGREE ANGLE AGAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978090&pc=MTA | 2023826-2017-01613 | NA | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.5/X096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS REPORTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS, VTICMO13.2 IMPLANTED IS CORRECTED FROM -5.0/+2.5/096 TO -6.0/+2.5/096 DIOPTER. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A".(B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S.. SECONDARY SURGERY, LENS REPOSITIONING. KERATOCONUS EYE. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020606&pc=MTA | 2023826-2017-01695 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/002 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020855&pc=MTA | 2023826-2017-01696 | 2017-02-11 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144255&pc=MTA | 2023826-2017-01912 | 2017-07-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+5.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.6MM) BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SAME MODEL/DIFFERENT DIOPTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5521849&pc=MTA | 2023826-2016-00325 | 2016-02-13 | Injury | STAAR SURGICAL COMPANY | 2016-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS ROTATED AND HAD A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS ALSO HAD REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A LONGER LENS, SAME MODEL AND SIMILAR DIOPTER. THE PATIENT IS DOING FINE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. (B)(4). THE DATES ON THE INITIAL WERE INCORRECTLY REPORTED AS (B)(6) 2016. THE CORRECT DATE IS (B)(6) 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832781&pc=MTA | 2023826-2017-01389 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -9.5/+1.5/102 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY AND ALL IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832815&pc=MTA | 2023826-2017-01390 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -13.5/+2.5/078 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY AND EVERYTHING IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365387&pc=MTA | 2023826-2017-00289 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.0/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING WITH REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND OTHER UNKNOWN ADVERSE EVENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A CENTRIFUGE TUBE, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645753&pc=MTA | 2023826-2017-00988 | 2016-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -11.5/+2/107 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, GLARES/HALOES, DECENTRALIZATION, AND "FUNNY VISION." THE LENS WAS REPOSITIONED.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. THE CURRENT LENS HAS BEEN EXCHANGED FOR A LONGER, SIMILAR DIOPTER LENS ON (B)(6) 2017. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR IN THE LENS OPTIC AND PIECE OF HAPTIC MISSING FROM THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703130&pc=MTA | 2023826-2017-01160 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.0/+3.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703148&pc=MTA | 2023826-2017-01161 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.5/+2.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT (NO ROTATION) AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438818&pc=MTA | 2023826-2018-00562 | 2016-06-15 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+1.5/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT ON 12-SEP-2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443489&pc=MTA | 2023826-2017-00512 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/X089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.8MM)] (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443524&pc=MTA | 2023826-2017-00513 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/X087 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS AND SIMILAR DIOPTER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6842513&pc=MTA | 2023826-2017-01337 | 2016-10-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12/1/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, AND HYPEROPIA SIDE-EFFECT AFTER A WEEK. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP STATED THAT THE BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/13 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/16. Manufacturer Narrative: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US BUT NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404238&pc=MTA | 2023826-2017-00392 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/5/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, REFRACTIVE SURPRISE, AND BLURRED VISION. AS THE DATE MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION-THE PATIENT COMPLAINED OF BLURRED VISION, "AS IF SEEING EVERYTHING THROUGH WATER." THE SURGEON SUSPECTED ISSUES MAY BE RELATED TO A VERY DIFFICULT PERSONAL PERIOD THE PATIENT IS GOING THROUGH. THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415314&pc=MTA | 2023826-2017-00434 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/144 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, FLUCTUATIONS IN VISUAL ACUITY, AND ACCOMMODATIVE SPASM. ON (B)(6) 2017 A SECONDARY SURGERY FOR IRIDECTOMY WAS PERFORMED, AND THE PATIENT STARTED PILOCARPINE TREATMENT. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING THE PATIENT'S STATUS. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. CONCLUSION: OFF LABEL USE (KERATOCONUS EYE). (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832836&pc=MTA | 2023826-2017-01391 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -8.0/+1.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION WAS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -8.0/+1.0/087 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES); WAS NOT SELECTED IN PREVIOUS MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -8.0/+1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 03/06/2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832866&pc=MTA | 2023826-2017-01392 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -7.5/+1.0/079 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION IS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -7.5/+1.0/079 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -7.5/+1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(4) 2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6)2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6847031&pc=MTA | 2023826-2017-01408 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.0/2.0/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS EXCHANGE FOR A SHORTER LENS RESOLVED THE PROBLEM. CORRECTED DATA: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139572&pc=MTA | 2023826-2017-01930 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/87 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH - NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139744&pc=MTA | 2023826-2017-01929 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/91 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 11 APR 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7359037&pc=MTA | 2023826-2018-00338 | 2018-02-07 | Injury | STAAR SURGICAL COMPANY | 2018-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -12.00/+1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: CORRECTED DATA: "ADVERSE EVENT" IS CORRECTED TO "PRODUCT PROBLEM". "OTHER SERIOUS" IS REMOVED FROM THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639247&pc=MTA | 2023826-2018-00978 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, -11.0/+2.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696918&pc=MTA | 2023826-2018-01031 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/066 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE CAUSE IS LISTED AS A PATIENT-RELATED FACTOR AND STATES THE LENS WAS DIFFICULT TO MANIPULATE INTO THE CORRECT POSITION BECAUSE THE "PATIENT WAS VERY NERVOUS AND HECTIC." THE SURGEON ALSO STATES THE SULCUS SIZE DOES NOT CORRELATE WITH THE WHITE TO WHITE MEASUREMENT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED/CORRECTED INFORMATION: ON (B)(4) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM IS RESOLVED AND THE "PATIENT IS HAPPY WITH 20/20 VISION." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6060685&pc=MTA | 2023826-2016-01533 | NA | Injury | STAAR SURGICAL COMPANY | 2016-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND UNREACTIVE PUPIL. THE LENS WAS REPOSITIONED, AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, UNREACTIVE PUPIL, AND IRIS PUSHED TOWARDS THE CORNEA. THE LENS WAS REPOSITIONED AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. (B)(4). CORRECTED DATA: EXPLANTED DATE: N/A. (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199927&pc=MTA | 2023826-2015-01454 | 2014-07-05 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -5.5/+2.0/X081 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT, LENS ROTATION, BLURRED VISION AND REFRACTIVE CHANGE WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND WAS EXCHANGED FOR A LONGER LENS LENGTH. THE PROBLEM WAS RESOLVED. PATIENT'S LAST VISIT ON (B)(6) 2014 - UCVA 20/32. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -5.5/+2.0/X081. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806370&pc=MTA | 2023826-2015-00588 | 2014-08-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND NOTED EXCESSIVE VAULT, WITH A REFRACTIVE SURPRISE/REFRACTIVE CHANGE OVERTIME, IN THE POST-OPERATIVE PERIOD. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740720&pc=MTA | 2023826-2015-00489 | 2014-10-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE NOTED, PATIENT EXPERIENCED A REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193772&pc=MTA | 2023826-2015-01444 | 2015-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/112 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING ON EXPLANTING THE LENS. THE LENS REMAINS IMPLANTED.Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. CLAIM #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4) METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464177&pc=MTA | 2023826-2016-00200 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+0.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089943&pc=MTA | 2023826-2016-01603 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23.Event Description: THE PROBLEM WAS RESOLVED AFTER LENS EXCHANGE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CONCLUSION CODE: (OFF-LABEL USE): PATIENT ACD WAS BELOW 3.00 MM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC BROKEN; DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND WITHIN SPECIFICATIONS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247436&pc=MTA | 2023826-2017-00065 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0/+2.5/098, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT REPORTING LOW VAULT, ROTATION OF THE LENS, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE REFRACTIVE CHANGE OVER TIME IS DUE TO THE LOW VAULT AND ICL ROTATION. THE LENS HAS BEEN RETURNED, BUT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED AND HAS BEEN EVALUATED. (B)(4). THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (HAIRS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245220&pc=MTA | 2023826-2018-00205 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245224&pc=MTA | 2023826-2018-00207 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH WHITE LIQUID ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268474&pc=MTA | 2023826-2018-00194 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -12.5/+1.5/65 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT, LENS ROTATION, AND REFRACTIVE CHANGE OVER TIME WAS NOTICED. A SURGERY FOR A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND A NEW LONGER LENGTH LENS WAS IMPLANTED LATER ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. EXPLANT DATE: (B)(6) 2018. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM" IN THE INITIAL MDR. "EXCESSIVE VAULT" IS CORRECTED TO "LOW VAULT" IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673242&pc=MTA | 2023826-2017-01044 | 2017-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE DR. MENTIONED THAT THE LENS IS DECENTERED IN PUPIL TEMPORALLY, HE WONDERED IF A TEMPORAL HAPTIC COULD BE DAMAGED/FOLDED. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. ACCORDING TO THE DR., THE PATIENT IS MANAGING THE PROBLEM WITH CONTACT LENSES, AND THE DR. ALSO PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595410&pc=MTA | 2023826-2017-00892 | 2017-05-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -7.00/+6.0/118 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND LENS ROTATION. THE LENS WAS REPOSITIONED (B)(6) 2017.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. CURRENT: ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER LENS. VAULT IS NORMAL, PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6865845&pc=MTA | 2023826-2017-01453 | 2017-05-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -9.0/+3.5/095 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED BECAUSE LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE WERE NOTED BY THE SURGEON. THE REPLACEMENT LENS WAS OF LONGER LENGTH AND SIMILAR DIOPTER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC BENT AND DEFORMED. (B)(4). DEVICE MANUFACTURER DATE: IT IS CORRECTED TO "27-MAY-2016". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169880&pc=MTA | 2023826-2018-00012 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT; THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE AND DEBRIS ON SURFACE OF THE LENS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169893&pc=MTA | 2023826-2018-00013 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481819&pc=MTA | 2023826-2018-00628 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/2.0/082 (SPHERE/AXIS/SPHERE), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(46 2018 DUE TO LOW VAULTING, LENS ROTATION, AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT ALONG WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS ALSO MISSING THE TORN PIECE OF ITS HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639140&pc=MTA | 2023826-2018-00977 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/+2.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589138&pc=MTA | 2023826-2017-00867 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | INPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: ON (B)(6) 2017 THE REPORTER STATED THAT HE NEEDED TO RETURN A LENS THAT WAS "REMOVED FROM THE EYE DUE TO SIZING ISSUES." ATTEMPTS TO RETRIEVE ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED/CORRECTED INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -16.0/+4.5/115 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN EXPERIENCING LOW VAULT, LENS ROTATION, BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM IS RESOLVED. ADDITIONAL INFORMATION ADDED. LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS: 26-APR-2017. _CORRECTED: 25-APR-2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114673&pc=MTA | 2023826-2017-01860 | NA | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.00/+1.50/89 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING THAT CAUSE LENS ROTATION TWICE AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769268&pc=MTA | 2023826-2017-01284 | NA | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INPLANTED AN ICL IMPLANTABLE COLLAMER LENS, AND THERE WAS A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/180 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT STILL HAS SOME REFRACTION LEFT, BUT THE SURGEON DECIDED NOT TO ROTATE THE LENS. THE SURGEON IS PLANNING TO DO A LASER TOUCH-UP. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675761&pc=MTA | 2023826-2017-01083 | 2016-12-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+4.5/062 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE. THE SURGEON STATED THAT THE LENS SEEMS NOT TO BE ROTATED AT SLIT LAMP, AND REFRACTION WAS CHANGING THE FIRST WEEKS BUT IT IS STABLE SINCE 2 MONTHS. THE DOCTOR PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: THE DOCTOR PLANS TO EXCHANGE THE LENS. UPDATED ON 21-AUG-2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS, SIMILAR DIOPTER. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- NO CODE AVAILABLE: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4621347&pc=MTA | 2023826-2015-00356 | 2014-11-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE WITH AN INCIDENT OF REFRACTIVE SURPRISE. AN EXCHANGE OF THE LENS IS PLANNED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE LENS WAS EXPLANTED 5 MONTHS AFTER THE INITIAL SURGERY. THE LENS WAS EXCHANGED FOR A SMALLER LENS WITH DIFFERENT POWER TO ADJUST THE ERROR AND THE PROBLEM WAS RESOLVED. EXPLANTED: (B)(6) 2015. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5865428&pc=MTA | 2023826-2016-01033 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). LENS REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5838759&pc=MTA | 2023826-2016-00993 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL WITH PARTICULATES ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PREVIOUSLY REPORTED EVENT DESCRIPTION IS INCORRECT. CORRECT EVENT DESCRIPTION IS THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135217&pc=MTA | 2023826-2017-01919 | 2012-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/097 (SPHERE/CYLINDAR/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. REPORTEDLY, THE SURGEON NOTES EXCESSIVE VAULT AND REFRACTIVE CHANGE OVER TIME. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: REPORTEDLY, WHEN THE SURGEON EXAMINED THE EYE PRIOR TO PLANNED LENS EXCHANGE SURGERY, IT WAS DISCOVERED THAT THE LENS HAD ROTATED. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594258&pc=MTA | 2023826-2018-00888 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.5/+1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. THE EXCHANGE WAS SUCCESSFUL. PROBLEM RESOLVED. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND VISUAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO (B)(6) 2018 IN THE SUPPLEMENTAL MDR#1. DEVICE EVALUATION CORRECTED TO: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635334&pc=MTA | 2023826-2015-00381 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THE LENS DESCRIPTION IS -10.0/+3.0/063 (SPHERE/CYLINDER/AXIS). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CONCLUSION: AT THE TIME OF IMPLANTATION, PATIENT WAS (B)(6) (I.E. UNDER THE AGE OF 21 YEARS) AND HAD PRE-EXISTING PROGRESSIVE MYOPIA OR HYPEROPIA, BOTH OF WHICH ARE CONTRAINDICATED AS PER THE DFU OF THIS LENS MODEL. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5678984&pc=MTA | 2023826-2016-00743 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT, ROTATION, AND A REFRACTIVE SURPRISE, AND WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.84MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4634764&pc=MTA | 2023826-2015-00380 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). NOT RETURNED. Manufacturer Narrative: PER MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5968182&pc=MTA | 2023826-2016-01309 | 2016-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ROTATION AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (SMALL FIBERS). (B)(4). METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5018700&pc=MTA | 2023826-2015-01065 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+1.50X45 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT COMPLAINT OF BLURRED VISION. ON EXAMINATION, THE PATIENT'S UCVA: 6/15 WITH -0.50/+1.75X115. A REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S) BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4)- NO CODE AVAILABLE (REFRACTIVE SURPRISE), UNLABELED DEVICE CODE: (B)(4)- NO INFORMATION. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.(B)(4). OTHER TEXT : LENS REMAINS IMPLANTED | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817519&pc=MTA | 2023826-2017-01367 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE REPORTER INDICATED THE REASON FOR THE PROBLEM WAS NOT A ROTATION OF THE LENS CAUSED THROUGH A SIZING PROBLEM, IT WAS THE SURGEON WHO IMPLANTED THE LENS 5º CW (CLOCKWISE) INSTEAD OF 5º CCW(COUNTER CLOCKWISE). THE PATIENT'S CURRENT CONDITION IS UCVA - 20/20, AND PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589758&pc=MTA | 2023826-2017-00773 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A PHAKIC IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE HAPTIC BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL TESTING WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 +1, THE EYE IS QUIET. DUANE'S RETRACTION SYNDROME. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164819&pc=MTA | 2023826-2018-01818 | 2017-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/1.0/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENTS IS RPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7700430&pc=MTA | 2023826-2018-01066 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14/1.5/24 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC TORN/BROKEN/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824451&pc=MTA | 2023826-2018-01256 | 2018-06-06 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCING A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PER FDA REQUEST. THIS REPORT WAS PREVIOUSLY FILED WITH INCORRECT MFR NUMBER. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696916&pc=MTA | 2023826-2018-01035 | 2018-06-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY THE LENS REMAINS IMPLANTED. THE SURGEON STATES THERE IS A "SIGN ERROR ON ASTIGMATISM" THAT OCCURRED DURING LENS CALCULATION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED STATEMENT: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS. THE PROBLEM IS RESOLVED. THE STATUS OF THE EYE IS REPORTED AS "PERFECT". (B)(4). Manufacturer Narrative: PREVIOUS MDR SHOULD INCLUDE THE FOLLOWING STATEMENT: "USE/HANDLING AND ADDITIONAL PROCESSING IMPACTS LENS DIMENSIONS/FUNCTIONS; RETURNED LENS MAY NOT MEET ORIGINAL SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE CYLINDER AND AXIS TO BE WITHIN SPECIFICATIONS BUT FOUND THE SPHERE TO BE 0.07D OFF SPECIFICATIONS. FUNCTIONAL INSPECTION AT THE TIME OF FINAL INSPECTION PRIOR TO PRODUCT RELEASE FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765576&pc=MTA | 2023826-2018-01175 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/4.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7961068&pc=MTA | 2023826-2018-01509 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO12.6,-18.00/3.0/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 WITH AN OBSERVATION OF REFRACTIVE SURPRISE. REPOSITIONING OF THE LENS WAS CONDUCTED AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT 'THE PATIENT IS ABSOLUTELY FINE' AND THAT 'VN IS 6/9'. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS DUE TO A USER ERROR, SURGEON ADMITTED TO IMPLEMENTING INCORRECT DATA IN OCOS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7914678&pc=MTA | 2023826-2018-01450 | 2018-06-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/+3.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906882&pc=MTA | 2023826-2018-01446 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906886&pc=MTA | 2023826-2018-01443 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, IMPLANTABLE COLLAMER LENS, -6.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN VIAL WITH CLEAR SURGICAL DEBRIS/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860735&pc=MTA | 2023826-2018-01248 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: DIOPTER CORRECTED TO "-9.00/+1.0/167" IN THE INITIAL MDR. THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/+1.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860748&pc=MTA | 2023826-2018-01249 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO: "-8.00/0.5/002" IN THE INITIAL MDR. EXPIRATION DATE CORRECTED TO: 30-APR-2021 IN THE INITIAL MDR. MANUFACTURE DATE CORRECTED TO: 11-MAY-2018 IN THE INITIAL MDR. ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -8.00/0.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113360&pc=MTA | 2023826-2018-01756 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+1.5/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC BENT AND DEFORMED. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. DIMENSIONAL AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227438&pc=MTA | 2023826-2019-00059 | 2018-09-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM IS RESOLVED WITH THE EXCHANGE; "ALL IS FINE!". (B)(4). Manufacturer Narrative: ADDED APPLICABLE OUTCOME/ PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL, SPHERE, CYLINDER, AXIS LENS. REPORTEDLY, POST-OP RESULTS ARE NOT CURRENTLY AVAILABLE. DISTRIBUTOR CHECKED IN ERROR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- (B)(4). CLAIM# COMPLAINT-(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975584&pc=MTA | 2023826-2018-01531 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -8.50/1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975592&pc=MTA | 2023826-2018-01532 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -9.00/0.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229422&pc=MTA | 2023826-2019-00023 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8268673&pc=MTA | 2023826-2019-00018 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270834&pc=MTA | 2023826-2019-00017 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE , "SIMILARE" (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. WITHIN THE SAME SURGERY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8089154&pc=MTA | 2023826-2018-01716 | 2018-10-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON PLANS TO DO A PHOTOREFRACTIVE KERATECTOMY (PRK) TO RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT "AFTER PRK REFRACTION NOW IS -0,25/-0,25 WITH A BCVA OF 20/20". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8276381&pc=MTA | 2023826-2019-00156 | 2018-10-24 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS BUT DIFFERENT SPHERE/CYLINDER/AXIS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8166385&pc=MTA | 2023826-2018-01798 | 2018-11-03 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/X166 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (I.T. ERROR AT THE SURGEON'S OFFICE. THE PATIENTS REFRACTION DATA WAS INPUT INCORRECTLY AS PLUS CYLINDER INSTEAD OF A MINUS CYLINDER). THE PATIENT HAD RESIDUAL ASTIGMATISM POST-OPERATIVE AND A REFRACTIVE TREATMENT (LASIK) IS PLANNED TO CORRECT PROCEDURE. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAD LASIK SURGERY ON (B)(6) 2018 FOR CORRECTION AND THE PATIENT IS HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149333&pc=MTA | 2023826-2018-01801 | 2018-11-08 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS NOT RESOLVED. REPORTEDLY, THE "PATIENT IS DISSATISFIED WITH THE MYOPIA REMAINING." THE CAUSE IS LISTED AS UNKNOWN. SEE MFR REPORT# 2023826-2018-01804 FOR CASE CONTINUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7891566&pc=MTA | 2023826-2018-01420 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/+1.0/69 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER DUE TO OBSERVATION OF REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. IT WAS REPORTED THAT THE DESIRED DEGREE OF MYOPIA WAS NEVER ACHEIVED AND OVER TIME IT DRIFTED EVEN LESS MYOPIC. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION WAS RECEIVED. REPORTER INDICATED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM. LAST REPORT STATES THAT VISION IS IMPROVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7918080&pc=MTA | 2023826-2018-01453 | NA | Injury | STAAR SURGICAL COMPANY | 2018-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO13.2, -8.50/5.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED AND ROTATION/REPOSITIONING WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE DOCTOR DID THE REPOSITIONING LAST WEEK AND HE IS WAITING FOR PATIENT'S FOLLOW UP TO SEE THE RESULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE ICL WAS ROTATED ANTI-CLOCKWISE TO THE DESIRED AXIS. REFRACTION IN HER LAST FOLLOW UP VISIT (3 WEEKS FOLLOWING THE PROCEDURE) IS AS FOLLOWS: -0.25 -0.75 AXIS 70 WITH UCVA OF 20/25. PATIENT IS HAPPY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052105&pc=MTA | 2023826-2018-01626 | 2018-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE SURGEON DECIDED TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). POST SURGERY RESULTS ARE PENDING AT THE TIME OF THIS REPORT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATENT IS FINE AT THE MOMENT, STRUGGLING WITH RX DUE TO REFRACTIVE SUPRISE. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR. SURGEON USED THE WRONG AXIS OF THE RX. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -7.50/1.5/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) -2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON USED THE WRONG AXIS OF THE RX. ON (B)(6) 2018, REPORTER STATED: "PATIENT CONDITIONS ARE NOT BAD, BUT THE UCVA IS ONLY 20/25 BECAUSE OF THE REFRACTIVE SURPRISE". ON (B)(6) 2018 IT WAS STATED: "PATIENT IS FINE-AT THE MOMENT, HE'S ONLY A BIT STRUGGLING WITH HIS RX ON OS (BECAUSE OF REFRACTIVE SURPRISE)". THE SURGEON PLANS TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). "THE ROTATION HAS NOT BEEN DONE DUE TIME PROBLEMS OF THE PATIENT, YET IT IS PLANNED FOR MID OF DECEMBER". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS REPOSITIONED ON (B)(6) 2019, "STABLE REFRACTION AFTER REPOSITIONING." AND THE PROBLEM WAS RESOLVED. PATIENT CODE ADDITIONAL INFORMATION: 3191 - SECONDARY SURGICAL INTERVENTION, REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974926&pc=MTA | 2023826-2018-01528 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/+2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974933&pc=MTA | 2023826-2018-01527 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710867&pc=MTA | 2023826-2018-01071 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+6.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT THE "DOCTOR DECIDED TO ROTATE THE LENS BUT AFTER SIX DAYS THE LENS RETURNED TO ITS INITIAL POSITION OBTAINING THE SAME REFRACTIVE SURPRISE. THE DOCTOR WILL ORDER THE SAME LENS WITH THE SAME POWER BUT A BIGGER SIZE BECAUSE THE VAULT AFTER THESE DAYS HAS DECREASED." (B)(4). Manufacturer Narrative: ADVERSE EVENT" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE)" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. (B)(4). Manufacturer Narrative: LENS WAS EXCHANGED FOR A SIMILAR POWER AND A LARGER SIZE ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7851176&pc=MTA | 2023826-2018-01292 | 2018-07-24 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -13.00+/4.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018, LENS REMAINS IMPLANTED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182343&pc=MTA | 2023826-2018-01887 | 2018-11-23 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE "LENS WAS NOT STABLE" ON (B)(6) 2019. ON (B)(6) 2019 THE REPORTER STATED, "THE SURGEON TRIED TO CORRECT THE REFRACTIVE SURPRISE WITH A REPOSITIONING OF BOTH LENSES. FOR THIS HE ROTATED THE MISALIGNED TICLS TO HE AXIS OF THE IODS. AFTER THAT (1DAY POSTOP) FIRST THE REFRACTIONS WERE GOOD AND EVERYTHING WAS FINE. AT THE NEXT EXAMINATION MID OF (B)(6) THE TICLS HAVE ROTATED AGAIN TO THE SAME MISALIGNED AXIS AS BEFORE. IN THIS POSITION THE LENSES WERE NOW STABLE AND SO THE SURGEON DECIDED TO MAKE A LASER TOUCH UP TO CORRECT THE RESIDUAL REFRACTION ON BOTH EYES". PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: (B)(4). NOT INCLUDED IN THE SUPPLEMENTAL MDR#2 CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT HAD DECIDED NOT TO HAVE A LVC PROCEDURE TO CORRECT CURRENT REFRACTIVE ERROR AFTER TICL IMPLANTATION. THE PLAN IS TO EXPLANT THE LENS. CORRECTED DATA: PATIENT CODE: 3191 - (SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP) - PREVIOUS CODE NOT APPLICABLE, PATIENT DECIDED NOT TO HAVE LASER TOUCH UP (LVC). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272716&pc=MTA | 2023826-2019-00152 | NA | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+2.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD ROTATED AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, THE LENS ROTATED AGAIN. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725096&pc=MTA | 2023826-2017-01217 | 2017-04-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: CORRECTED DATA: CONCLUSION: PREVIOUS CODE IS NOT APPLICABLE, THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF ANY OF THE FOLLOWING CIRCUMSTANCES AND/OR CONDITIONS SUCH AS FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: IT WAS REPORTED ON (B)(6) 2019 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589617&pc=MTA | 2023826-2018-00878 | 2018-04-25 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) INSPECTION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SAME MODEL/DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227439&pc=MTA | 2023826-2019-00056 | 2009-03-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/029 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY AND EXCHANGE WENT WELL!" (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773875&pc=MTA | 2023826-2018-01187 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICMO13.2, -13.0/+5.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT. ONE DAY AFTER REPOSITIONING UCVA/BCVA WAS 20/40 AND 20/20, IOP WAS 15 MMHG, VAULT- 902 MICRONS. RX WAS THE SAME AS BEFORE MANIPULATION. IT WAS REPORTED THAT THE SURGEON PLANS TO REPLACE THE ICL DUE TO THE RESIDUAL REFRACTIVE ERROR AND HIGH VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021583&pc=MTA | 2023826-2018-01572 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2018. THE "VAULT WAS GREATLY IMPROVED" AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND BLUE RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7790898&pc=MTA | 2023826-2018-01237 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6,-5.5/+1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED AND THE SURGEON IS NOT SURE IF THE LENS WILL BE EXPLANTED OR EXCHANGED, THE SURGEON WILL BE MONITORING OVER THE NEXT MONTHS. PATIENT'S CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765084&pc=MTA | 2023826-2018-01171 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL DRY WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS REMOVED ON AN UNKNOWN DATE. ON (B)(6) 2019 A LONGER LENGRTH REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. H6- PATIENT CODE 3191: SECONDARY SURGERY; LENS EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711588&pc=MTA | 2023826-2017-01176 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -15.5/+2.0/65 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS WAS REPOSITIONED ON(B)(6)2015. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, AND WILL NOT RETURN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2 -15.5/+2.0/065 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPOSITIONED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2018 IT WAS REPORTED THAT REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION WERE OBSERVED IN THE PATIENT; THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 WHICH RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194494&pc=MTA | 2023826-2018-01903 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS REPOSITIONED. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: EVENT- LENS WAS EXPLANTED. EXPLANT DAT- UNK. PATIENT CODE 3191: NO CODE AVAILABLE - SECONDARY SURGERY, EXPLANT. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #: (B)(4). Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM: ADDITIONAL INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED, THE PATIENT IS "VERY HAPPY!" DISTRIBUTOR PREVIOUSLY CHECKED IN ERROR. PATIENT CODE 3191 PREVIOUSLY STATED LENS EXPLANT UPDATED, 3191 SHOULD STATE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227441&pc=MTA | 2023826-2019-00057 | 2018-09-26 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." Manufacturer Narrative: ADDITIONAL INFORMATION: REFER TO MFR REPORT # 2023826-2019-01063 FOR CLAIM# (B)(4). FOR FURTHER INVESTIGATION DETAILS. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL INFORMATION: PROBLEM IS NOT RESOLVED DUE TO THERE STILL BEING REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED LENS WAS EXPLANTED AT A LATER DATE. UPDATED INFORMATION: LENS WAS EXCHANGED ON 15-JAN-2019 WITH A LONGER, DIFFERENT MODEL LENS. PROBLEM IS RESOLVED. DISTRIBUTOR WAS CHECKED IN ERROR. PREVIOUSLY STATED LENS EXPLANT. UPDATED-LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172154&pc=MTA | 2023826-2018-01843 | 2018-11-09 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+2.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS DISLOCATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. SURGEON ALSO STATES THE VAULT IS 200UM AND PLANS TO IMPLANT A LONGER LENS. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019, THE LENS WAS REPOSITIONED A 2ND TIME. REPORTEDLY, THE "PATIENT AND SURGEON ARE HAPPY WITH THE RESULTS." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172170&pc=MTA | 2023826-2018-01851 | 2018-11-13 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+3.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS ROTATION, PIGMENT DISPERSION, AND BLURRED VISION. REPORTEDLY, THE ICL HAD ROTATED TO THE VERTICAL POSITION. IT IS NOTED THAT THE PATIENT DOES NOT WANT TO UNDERGO ANOTHER OPERATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE STATUS OF THE EYE IS NOTED AS "CALM" AFTER THE EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488984&pc=MTA | 2023826-2018-00677 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REFRACTTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR ANOTHER LENS. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT MALFUNCTION TO SERIOUS INJURY (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL VERIFICATION WAS PERFORMED ON THE LENS. INVESTIGATION FOUND LENS WAS WITHIN SPECIFICATION AT THE TIME OF RELEASE. RESULTS OF RETURNED LENS FUNCTIONAL EVALUATION ARE INCONCLUSIVE; LENS MAY HAVE BEEN MISSHAPEN DURING PERIOD LENS REMAINED IMPLANTED OR DUE TO HANDLING DURING EXPLANT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201647&pc=MTA | 2023826-2018-00091 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Shelf Life Exceeded; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -4.50/+1.5/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT, LENS ROTATION, AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654727&pc=MTA | 2023826-2016-00659 | 2010-12-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Information; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND ROTATION AND REMAINS IMPLANTED. IT IS PLANNED TO EXPLANT THE LENS. Manufacturer Narrative: THE SURGEON PERFORMED A SECONDARY INTERVENTION TO REPOSITION THE LENS TO THE ORIGINAL AXIS. AN EXCHANGE IS PLANNED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. OUTCOME ATTRIBUTED TO ADVERSE EVENT: UNK. EXPLANT DATE: N/A. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8219629&pc=MTA | 2023826-2019-00045 | 2018-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available; Missing Value Reason | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2, -9.5/+1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS A PATIENT RELATED FACTOR. SURGEON PLANS TO EXCHANGE THE LENS IN (B)(6); IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: ADVERSE EVENT SHOULD BE MARKED IN PREVIOUS MDR. REQUIRED INTERVENTION SHOULD BE MARKED IN PREVIOUS MDR. REFRACTIVE SURPRISE WAS ALSO REPORTED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH, STRONGER POWER LENS. SHOULD BE ADDED TO THE INITIAL MDR. DATE RECEIVED BY MFR: SHOULD BE CORRECTED TO "OCTOBER 14,2019" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: "LENS REMAINS IMPLANTED" SHOULD BE CORRECTED TO "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. "THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR"SHOULD BE CORRECTED TO " THE CAUSE OF THE EVENT WAS USER ERROR." IN THE INITIAL MDR. ADDITIONAL INFORMATION: SECONDARY SURGICAL INTERVENTION - EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARK. ETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361178&pc=MTA | 2023826-2017-00262 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; Overwear Syndrome | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15/0.5/149 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED EXCESSIVE VAULT, GLARE/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED AS OF (B)(6) 2017. Manufacturer Narrative: CORRECTED DATA: EXCESSIVE VAULT, GLARES/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCHES: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). LENS NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021086&pc=MTA | 2023826-2018-01581 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021095&pc=MTA | 2023826-2018-01582 | NA | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | NA | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. EVENT: "VTICMO13.2, -16.00/2.0/90" SHOULD BE CORRECTED TO "VTICMO13.2, -16.00/2.0/94" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210270&pc=MTA | 2023826-2023-00247 | 2021-11-05 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.50/+3.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 06-SEP-2021. THE PATIENT EXPERIENCED LENS ROTATION; UNREACTIVE (FIXED) PUPIL; AND IMMEDIATELY AFTER SURGERY; GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNREACTIVE (FIXED) PUPIL. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8379600&pc=MTA | 2023826-2019-00351 | 2018-09-22 | Injury | STAAR SURGICAL COMPANY | 2019-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/039 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, WITH GLARE/HALOS AND BLURRED VISION. THE PATIENT HAD DISCOMFORT WITH THEIR VISION IN COMPARISON WITH THE ICL LENS IN THE RIGHT EYE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. METHOD CODE 4117 NO LONGER APPLICABLE. INITIAL MDR SHOULD HAVE INCLUDED DEVICE CODE 1494: OFF-LABEL USE, PRE-EXISTING HISTORY OF KERATOCONUS AND CORNEAL CROSS-LINKING. INITIAL MDR SHOULD HAVE INCLUDED METHOD CODE 4111. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE PATIENT IS STABLE AND HAS RETURNED TO PREOPERATIVE REFRACTION. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. D7 - UPDATED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876813&pc=MTA | 2023826-2017-01461 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -14.5/+2.0/071 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2017" FROM "(B)(6) 2017" IN THE INITIAL. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752257&pc=MTA | 2023826-2017-01144 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.5/+1.0/92 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752269&pc=MTA | 2023826-2017-01143 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -9/+2/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135224&pc=MTA | 2023826-2017-01924 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/+4.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY, LENS REPOSITIONING H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320740&pc=MTA | 2023826-2018-00322 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+1.5/180 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED. REPORTEDLY, THE PATIENT EXPERIENCED "PROBLEMS READING AND DOING COMPUTER WORK." THE SURGEON STATES THE EYE IS "CALM AND QUIET" AFTER THE EXCHANGE AND THAT THE EYE IS 0.5 MORE MYOPIC THAN EXPECTED. THE SURGEON ALSO REPORTS THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "PROBLEMS READING AND DOING COMPUTER WORK." (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6520730&pc=MTA | 2023826-2017-00707 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -7.0/+1.5/089 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE ICL IS ON THE AXIS. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: STATED LENS REMAINS IMPLANTED. LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LARGER, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. FINAL POST-OP RESULTS: VAULT 500UM, UCVA 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE), AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6461375&pc=MTA | 2023826-2017-00573 | 2015-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -16.5/+5.5/088 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. IN ADDITION, THE PATIENT SUFFERED SIGNIFICANT VISION LOSS AND GLARE WHILE DRIVING AT NIGHT THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6211496&pc=MTA | 2023826-2016-01841 | 2016-10-06 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Impairment; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12/03/78 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). THE LENS WAS EXCHANGED FOR ANOTHER LENS VTICM 13.2 MM -12/03/79 DIOPTER. THE PROBLEM WAS RESOLVED AFTER THE LENS WAS EXCHANGED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THE LENS WAS EXCHANGED FOR A 13.2 MM VTICM5_13.2, NOT A VTICM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629094&pc=MTA | 2023826-2015-00368 | NA | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THERE WAS A REFRACTIVE SURPRISE AND THE VISUAL ACUITY WAS NOT GOOD ENOUGH. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS IMPLANTED ON (B)(6) 2014. SEVEN MONTHS AFTER THE INITIAL SURGERY, THE SURGEON DECIDED TO REPOSITION THE LENS 30 CW TO THE TARGET POSITION OF 3 CCW. HOWEVER, THE LENS ROTATED AGAIN TO THE SAME POSITION. THE SURGEON IS PLANNING TO REMOVE AND REPLACE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: LOW VAULT WAS REPORTED. LENS WAS EXPLANTED SOME TIME AFTER REPOSITIONING THE LENS. DEVICE EVALUATION: AN EVALUATION WAS PERFORMED; ADDITIONAL INVESTIGATION FOUND THE DEVICE WAS IN SPECIFICATION AT TIME OF RELEASE. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AS WELL AS PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, RUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: EVALUATION METHOD: DEVICE HISTORY RECORD REVIEW RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Repositioning | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495703&pc=MTA | 2023826-2016-00263 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5496053&pc=MTA | 2023826-2016-00264 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873717&pc=MTA | 2023826-2016-01069 | 2015-07-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.50/+1.50/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR/POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/50 +1. Manufacturer Narrative: CONCLUSION: THE PRE-EXISTING "GRAINY PERIPHERAL CORTICAL CLOUDING" WAS LISTED ON THE FORM AND THAT CONDITION MIGHT HAVE CAUSED THE CATARACT DEVELOPMENT SINCE IT IS WELL KNOWN THAT OPACITIES START AS VACUOLES AND CLEFTS BETWEEN THE LENS FIBERS AND PROGRESS TOWARDS ANTERIOR, POSTERIOR OR EQUATORIAL CORTEX. IT SHOULD BE NOTED THAT AT THE TIME OF SURGERY THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. GIVEN ALL THIS INFORMATION, OBSERVED LENS OPACITY AND SUBSEQUENT CATARACT DEVELOPMENT WAS AGE RELATED (PATIENT FACTOR) AND NOT CAUSED BY THE DEVICE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC BROKEN OFF. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873711&pc=MTA | 2023826-2016-01070 | 2004-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.75/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/30 -2. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955552&pc=MTA | 2023826-2017-01577 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Activation Failure | Cataract, Induced; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5/+2.0/097 INTO THE PATIENT'S LEFT EYE (OS), THE LENS WOULD NOT PROPERLY UNFOLD. THE SURGEON STATES THAT, "AFTER ALL POSSIBLE MANEUVERS" THE LENS DIDN'T UNFOLD PROPERLY AND THE LENS WAS EXPLANTED. DURING EXPLANT THE LENS WAS SCRATCHED. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THE SURGEON THEN PEFORMED PHACO-EMULSIFICATION ON THE EYE AND IMPLANTED AN EXCHANGE IOL ON (B)(6) 2017 "WITHOUT COMPLICATIONS." Manufacturer Narrative: BEST CORRECTED VISUAL ACUITY (BCVA), ON (B)(6) 2017, WAS 20/30 WHICH WAS A DECREASE FROM BASELINE BCVA (20/25). (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PHACO-EMULSIFICATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176008&pc=MTA | 2023826-2016-01763 | 2014-09-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract, Induced; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/042 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2014. THE PATIENT EXPERIENCED BLURRED VISION AND CLOUDY VISUAL ACUITY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25 +2. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. THE LENS WAS REMEASURED AND WAS FOUND TO BE IN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16474967&pc=MTA | 2023826-2023-00725 | 2022-07-20 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.0/1.0/99 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT FOR A PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THIS WAS DUE TO A EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 12.6, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/99 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE PATIENT EXPERIENCED EARLY LENS OPACITY. REPORTEDLY "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 13.2, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS BECAME CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721163&pc=MTA | 2023826-2017-01196 | 2017-06-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/4.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT, BLURRED VISION, AND LENS REPOSITIONING. LENS WAS "REMOVED AND CATARACT PROCEDURE HAVE BEEN DONE DUE TO CHANGES ON NATURAL LENS". THE PATIENT'S POST-OP CONDITION IS "WITH APHAKIC LENS AND HIS VISION IS 6/12". Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS MISSING PIECE OF HAPTIC. THE PATIENT'S POST-OP VISION IS 6/12. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874039&pc=MTA | 2023826-2017-01447 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 13.2 MM VTICMO13.2, -7.0/+4.5/107 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2015, THE LENS WAS EXPLANTED DUE TO POST TRAUMATIC DISLOCATION IN A/C AND TRAUMATIC CATARACT DUE TO FIREWORKS MUCH LATER THAN THE IMPLANTATION. IT HAD NO CONNECTION TO THE SURGERY. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND PRESENCE OF FIBERS ON LENS SURFACE. THE PATIENT ALSO HAD BLURRED VISION. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE:TRAUMA CAUSED BY FIRWORK. NO CONNECTION TO SURGERY. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6681259&pc=MTA | 2023826-2017-01089 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/40 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND THE DEVELOPMENT OF A CATARACT. A REPOSITIONING OF THE LENS, ENLARGEMENT OF THE INCISION (IN ORDER TO PERFORM CATARACT SURGERY), PHACO-EMULSIFICATION, AND IOL IMPLANTATION WAS PERFORMED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND THE PATIENT'S POST OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25 WITH STATUS OF THE EYE BEING PSEUDOPHAKIC. Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. Manufacturer Narrative: "HOSPITALIZATION-INITIAL OR PROLONGED" WAS CHECKED BY MISTAKE. THERE WAS NO LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5762131&pc=MTA | 2023826-2016-00817 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -11.00/ +6.0/100 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL) BLURRED VISION, THE CATARACT WAS REMOVED, AND A TORIC IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513549&pc=MTA | 2023826-2015-00149 | 2014-11-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LOW VAULT AND LENS OPACITY WAS OBSERVED AND THE PATIENT EXPERIENCED BLURRED VISON. THE LENS WAS EXCHANGED FOR A LONGER LENGTH WHICH RESOLVED THE VAULTING ISSUE, HOWEVER, THE CATARACT PROGRESED. THE SECOND LENS WAS REMOVED, THE CATARACT WAS EXTRACTED AND AN IOL WAS IMPLANTED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: .(B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. THERE WAS EVIDENCE OF A CLEAR SURGICAL RESIDUE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). IN CASES OF POSTOPERATIVE LOW VAULTING STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS OR ROTATION ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPUSLAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE MOST PROBABLE CAUSE OF THIS EVENT IS LOW VAULT SECONDARY TO A SHORT LENS. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247963&pc=MTA | 2023826-2017-00071 | 2011-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+1.50/094 DIOPTER, IN THE PATIENT'S LEFT EYE. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, REFRACTIVE SURPRISE, LENS OPACITY, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS, GLARE/HALOES, AND BLURRED VISION. THE PATIENT UNDERWENT PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6129538&pc=MTA | 2023826-2016-01682 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract; Blurred Vision; Visual Impairment; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND SURGEON REPOSITIONED THE NEW LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE SURGEON EXCHANGED THE LENS FOR A LONGER LENS AND REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS OPACITY, LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. REPOSITIONING OF THE LENS WAS PERFORMED AND THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. (B)(4). Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LENS ROTATION, DISLOCATION, REFRACTIVE CHANGE, AND BLURRED VISION. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE SURGEON REPORTED LENS OPACITY (ASCO) AND LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694861&pc=MTA | 2023826-2016-00799 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Loss of Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED ASC (ANTERIOR SUBCAPSULAR CATARACT) ON (B)(6) 2016 AND SUBSEQUENTLY LOST BCVA (BEST-CORRECTED VISUAL ACUITY). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572298&pc=MTA | 2023826-2016-00432 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/X010 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015, AND THE PATIENT EXPERIENCED LOSS OF BCVA. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/200. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). LENS NOT RETURNED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987375&pc=MTA | 2023826-2018-01533 | 2018-08-08 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Corneal Edema; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/060 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), THE LAST WEEK OF (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A FIXED DILATED PUPIL POST-OP, SEVERE CORNEAL EDEMA, WHICH HAS BEEN PERSISTENT UNTIL NOW AND AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, BLURRED VISION AND PUPIL/IRIS ATROPHY. THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED ONE MONTH AFTER THE SURGERY AND AN IOL WAS IMPLANTED. THE PUPIL REMAINED DILATED AND THE CORNEA WAS EDEMATOUS AND WHITE. THE PATIENT HAS HAD A CORNEAL TRANSPLANT. THE REPORTER INDICATED THE INTRAOCULAR PRESSURE HAS BEEN STABLE AND NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENTS VISION IS STILL HAND MOTION POST OPERATIVELY AND INTRAOCULAR PRESSURE IS 12MMHG. THE FIRST CHECKUP AFTER KERATOPLASTY INDICATED ALL IS GOOD. CORRECTED DATA: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937371&pc=MTA | 2023826-2015-00893 | 2015-03-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY, WHICH WAS FIRST NOTED ON (B)(6) 2015. THE REPORTER INDICATED IT WAS AN ANTERIOR SUBCAPSULAR OPACITY. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARE/HALOS AND BLURRED VISION. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926670&pc=MTA | 2023826-2016-00374 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED AT FOLLOW-UP VISITS, THE LENS HAPTIC WAS FOLDED AND THE PUPIL WAS OVALISED. THE LENS WAS TILTED ON ONE SIDE AND WAS RE-POSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP #4 FOR MDR # 2023826-2017-00374 WAS SUBMITTED IN ERROR. THE REPORTED INFORMATION WAS MEANT TO BE FOLLOW-UP #4 FOR MDR # 2023826-2016-00374. AS A RESULT, ONLY THE INITIAL MDR # 2023826-2017-00374 PERTAINS TO CLAIM# 710106. FOLLOW-UP #4 FOR MDR # 2023826-2016-00374 WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704665&pc=MTA | 2023826-2015-00425 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, INFLAMMATION, ANTERIOR UVEITIS, DISCOMFORT, LIGHT SENSITIVITY, PAIN AND TEARING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300916&pc=MTA | 2023826-2018-00281 | 2017-12-13 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Low Blood Pressure/ Hypotension; Inflammation; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, INFLAMMATION, AND "CHOROIDAL DETACHMENT, SEVERE HYPOTENSION." REPORTER STATES THAT HYPOTENSION IS "BEING RESOLVED" AND "IN FACT THE PATIENT IS VERY GOOD." Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. MEASUREMENT OF THE HYDRATION FOUND MOISTURE CONTENT ON A HIGH LEVEL; BUT NOT EXCEEDING LIMITS. THE OVERALL LENS DIAMETER WAS THEN FOUND OUT OF TOLERANCE. THE LENS SN: (B)(4) WAS A 12.6MM IMPLANTABLE LENS THAT IS 0.22MM LARGER THAN THE THEORETICAL; THE TOLERANCE IS +/-0.20MM. SINCE THE WORK ORDER SEARCH DID NOT FIND ADDITIONAL SIMILAR ERRORS TO THIS COMPLAINT EVENT, IT CAN BE CONCLUDED THAT THE OOS CONDITION IS NOT LIKELY DUE TO MANUFACTURING OR PACKAGING. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: MEASUREMENT OF THE HYDRATION FOUND THE LENS TO BE OVERHYDRATED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378746&pc=MTA | 2023826-2018-00466 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.5/+3.5/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOPLEGIC REFRACTION FOR THE LENS CALCULATION. LENS REMAINS IMPLANTED. LENS EXCHANGE IS PLANNED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, IMPLANTABLE COLLAMER LENS OF -5.5/+3.5/066 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. REFRACTIVE SURPRISE ALONG WITH BLURRED VISION WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOGENIC REFRACTION FOR LENS CALCULATION. LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS "LITTLE PERIPHERAL CORNEAL OEDEMA". EXPLANTED DATE: (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267530&pc=MTA | 2023826-2018-00214 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; Blurred Vision; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+2.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE AND CORNEAL EDEMA; INCIDENT DATE STATED AS (B)(6) 2018. REPORTEDLY, THE CORNEAL EDEMA DISAPPEARED AT 1-WEEK POST-OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. CLAIMANT STATES THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MMVTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. IMPLANTED DATE: IT IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017". EXPLANTED DATE: (B)(6) 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: IT IS CORRECTED TO "THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6760836&pc=MTA | 2023826-2017-01145 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Corneal Edema; Edema; Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -4.5/1.5/86 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE NEXT DAY, SIGNS OF TOXIC ANTERIOR SEGMENT SYNDROME, INCLUDE TEARING, PAIN, BLURRY VISION, EYELID EDEMA, HYPEREMIA, CORNEAL EDEMA, FIXED PUPIL AND FIBRIN CLOT IN THE ANTERIOR CHAMBER WAS NOTICED. THERAPY WAS STARTED WITH ORAL PREDNISOLONE (60 MG/DAY) AND ACETAZOLAMIDE 250 MM (1+1+1), TROPICAL PREDNISOLONE 1 DROP PER HOUR, TRAPICAMIDE, OFLOXACINE, AND BROMOFENAC. THE FOLLOWING DAY, THE PATIENT EXPERIENCE LESS PAIN, INTENSE HYPEREMIA, LESS EYELID EDEMA BUT A VERY CLOUDY WHITISH CORNEA. CORNEAL PERIPHERY WAS PARTIALLY PARED AND A BLOOD CLOT WAS SEEN IN THE INFERIOR A.C. AN ANTERIOR CHAMBER WASHING WAS PERFORMED AS WELL AS THE TEMPORAL INCISION WITCH BEING SUTURED WITH 2 10/00 NYLON STICHES AND INTRAVITREOUS INJECTION OF VANCOMICINE 1MG/0.1 ML AND CEFTAZIDIME 1.0 MG/0.1 ML. ACCORDING TO THE SURGEON, THE CORNEA IS CLEARING UP AND WOULD NEED TO MAKE ANOTHER SURGERY TO REMOVE THE MEMBRANE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531883&pc=MTA | 2023826-2017-00727 | NA | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Corneal Edema; Inflammation; Intraocular Pressure Increased; Pupillary Block; Uveitis; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED GLARE/HALOS, BLURRED VISION AND INFLAMMATION. ALSO, MILD UVEITIS, PIGMENT DISPERSION AND HAD DEVELOPED URETTS ZAVALIA SYNDROME (PERSISTENT PUPIL DILATION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP SHOWED CORNEAL EDEMA, SHALLOW ANTERIOR CHAMBER AND NORMAL INTRAOCULAR PRESSURE. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389853&pc=MTA | 2023826-2017-00347 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Edema; Inflammation; Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS, VAULTED FORWARD. IN ADDITION, THE PATIENT EXPERIENCED IRIS ATROPHY AND A DESCEMET MEMBRANE TEAR. ON (B)(6) 2017, THE LENS WAS EXCHANGED AND REPOSITIONED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE LENS WILL BE RETURNED. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING PATIENT'S PROGRESS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS WHICH CAUSED THE LENS TO VAULT FORWARD SO LENS REPOSITIONING WAS PERFORMED. IN ADDITION, PATIENT EXPERIENCED IRIS ATROPHY AND DESCEMET MEMBRANE TEAR. ON (B)(6)2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. AS OF (B)(6)2017 THE PATIENT WAS STILL EXPERIENCING POSTOPERATIVE SEQUELAE BUT CORNEA WAS CLEAR. THE LENS WILL NOT BE RETURNED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594294&pc=MTA | 2023826-2018-00898 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/156 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "SHALLOW AC," DECREASE IN UA, AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER "TOW TIME AC WASH" AND "AC REFORMATION." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN UPDATED INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION: LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.76MM)]. (B)(4) | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152455&pc=MTA | 2023826-2017-01962 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Cyst(s); Blurred Vision; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.5/+3.5/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, LENS ROTATION, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) (ANGLE CLOSURE), HALOES, VISUAL DISCOMFORT, AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. ALSO REPORTED IS SURGICAL ENLARGEMENT/ADDITION OF NEW PI, LOW BCVA AND UCVA, AND LENS REPOSITIONING. THE SURGEON STATES THAT "PROBABLY THE LENS IS TOO BIG FOR THE EYE" BUT BELIEVES THAT THE VISUAL PROBLEMS OF THE PATIENT ARE NOT RELATED TO THE LENS BUT TO THE CORNEA. HE BELIEVES THAT CHANGING THE ICL WILL NOT SOLVE THE VISUAL DISCOMFORT AND THAT THE LENS IS NOT DOING ANY HARM TO THE PATIENT. THE LENS REMAINS IMPLANTED. THE PATIENT IS "NOT HAPPY WITH HIS VISUAL ACUITY." Manufacturer Narrative: THE SURGEON NOTED THAT "THE IRIS CYST WAS NOT SUSPECTED BEFORE THE SURGERY. THE SUSPICION WAS RISEN AFTER THE FIRST BY THE FACT THAT THE VAULT WAS UNEVEN AND THAT WITH DILATION ALWAYS GETS MORE EVEN AND WITH A NORMAL HEIGHT. AND THE CYST SUSPICION WAS EVEN STRONGER AFTER THE SECOND SURGERY SINCE RE-ACCOMMODATION OF THE ICL DID NOT CHANGE THE BEHAVIOR OF THE LENS." WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PREVIOUS CODE IS NOT APPLICABLE, KERATOCONUS IS NO LONGER CONSIDERED A CONTRAINDICATION. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Retention | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375453&pc=MTA | 2023826-2016-00050 | NA | Injury | STAAR SURGICAL COMPANY | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Cyst(s); Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND LENS ROTATION. THE PATIENT EXPERIENCED GLARES/HALOS. THE REPORTER INDICATED THE LENS ROTATION MAY HAVE BEEN CAUSED BY A CILIARY CYST. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676112&pc=MTA | 2023826-2018-01030 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Endophthalmitis; Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SUSPECTED ENDOPHTHALMITIS, AND PIGMENT DISPERSION. THE PROBLEM WAS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS WAS EXCHANGED DUE TO ENDOPHTHALMITIS, EXCESSIVE VAULT, AND PIGMENT DISPERSION. CORRECTED-THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT AND PIGMENT DISPERSION ONLY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [PRE-EXISTING CONDITION-NARROW ANTERIOR CHAMBER ANGLES (I.E. GRADE II OR LESS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4) | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419890&pc=MTA | 2023826-2018-00531 | 2018-03-09 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; Visual Impairment; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/+2.5/166 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION, UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. A YAG PI WAS PERFORMED ON (B)(6) 2018. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC USED IN SURGERY, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS EXPLANT, THE PATIENT WAS ADVISED TO TAKE PRED/GATI 3 TIMES DAILY. AFTER DISCONTINUING, THE PATIENT DEVELOPED AN INCREASE IN IOP AND INFLAMMATION IN THE ANTERIOR CHAMBER. CURRENTLY, THE PATIENT IS BEING TREATED WITH PREDNISOLONE AND COMBIGAN, AZOPT, AND TRAVATAN Z FOR PRESSURE SPIKES. IT IS BELIEVED THAT THE SPIKES ARE DUE TO "STEROID USE VS INFLAMMATORY CAUSES." REPORTEDLY, NO INFECTION HAS OCCURRED WITHIN THE EYE. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT REPORTS "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA, AND LOSS OF VISION. ALSO STATED IS THAT THE IRIS IS PERMANENTLY FIXED AND DILATED. GLAUCOMA SURGERY PERFORMED. G3-REPORT SOURCE ADDED. H6-ADDITIONAL 3191: (BLOWN IRIS, FIXED IRIS, GLAUCOMA SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, LENS EXPLANT, FIXED PUPIL) OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT (46YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5440040&pc=MTA | 2023826-2016-00152 | 2014-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -16.0/5.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO GLARES/HALOS AND THE PATIENT WAS NOT SATISFIED WITH THE BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN AND BENT. FOREIGN MATERIAL IN THE FORM OF FIBERS WERE PRESENT ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043822&pc=MTA | 2023826-2016-01111 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043825&pc=MTA | 2023826-2016-01112 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/010 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909288&pc=MTA | 2023826-2016-01176 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/005 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767214&pc=MTA | 2023826-2015-00519 | NA | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HE HAS BAD NIGHT VISION (GLARES/HALOS). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404331&pc=MTA | 2023826-2017-00389 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/35 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND CLEAR RESIDUE AND DEBRIS ON LENS SURFACE. LENS IS FOLDED AT THE HAPTICS AND OPTIC IS TWISTED DUE TO POSITION IN VIAL. NO VISIBLE TEARS OR SCRATCHES. UNABLE TO FURTHER EVALUATE DUE TO LENS SHAPE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577529&pc=MTA | 2023826-2016-00440 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.5/+2.50/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00441 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN MATERIAL ON LENS SURFACE (FIBERS) AND A SCRATCH. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.5/+2.50/X095. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577591&pc=MTA | 2023826-2016-00441 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.00/+3.00/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00440 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CHANGE "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.00/+3.00/X091. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938281&pc=MTA | 2023826-2016-01247 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/005 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, NO DAMAGE TO THE LENS. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085625&pc=MTA | 2023826-2016-00882 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/073 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING OF THE LENS, GLARE/HALOES, AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN/ BENT. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701679&pc=MTA | 2023826-2017-01154 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, DIOPTER -5.5/+1.0/089 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, HALOES, AND WIDE PUPIL. THE SURGEON COMMENTED THAT THE PATIENT'S PUPIL > 8 MM ON (B)(6) 2017, AND PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS). (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP INDICATED NORMAL PUPIL SIZE, BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836431&pc=MTA | 2023826-2017-01376 | 2017-07-24 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/140 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE PATIENT IS OK. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724366&pc=MTA | 2023826-2015-00456 | 2015-02-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. LOW VAULT WAS NOTED, AS WELL AS REFRACTIVE CHANGE OVER TIME (INCREASED MYOPIA), AND GLARE/HALOS. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO - 81(OTHER): LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788339&pc=MTA | 2023826-2016-00883 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/13. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5462012&pc=MTA | 2023826-2016-00199 | NA | Injury | STAAR SURGICAL COMPANY | 2016-02-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600863&pc=MTA | 2023826-2015-00311 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, HEADACHE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444673&pc=MTA | 2023826-2017-00527 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+6.0/085 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED (B)(6) 2016 AND THE SURGEON NOTICES A GRADUAL INCREASE IN IOP. THE PROBLEM IS NOT RESOLVED AND A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694745&pc=MTA | 2023826-2017-01119 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN MICRO CENTRIFUGE VIAL WITH LITTLE MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694750&pc=MTA | 2023826-2017-01120 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WITH A CURVED SHAPE. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987323&pc=MTA | 2023826-2018-01529 | 2018-09-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA), PIGMENT DISPERSION, CELL DEPOSIT AND INFLAMMATION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983668&pc=MTA | 2023826-2022-01035 | 2022-03-07 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), DISTANCE AND NEAR VISION SLIGHTLY WORSE THAN THE OD EYE, BLEPHARITIS, IRIS BOMBE WITH PREVIOUS APPARENT TOUCH POINT SUPERIOR TO PUPIL, "VH 1/0," NIGHTTIME HALOS WHEN DRIVING AND ANY CEILING LIGHTS, AND IRIS THINNED WITH PIGMENT ON ICL. THE SURGEON STATES, "VERY NARROW ANGLE...ANGLE SHUT OFF." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO BIG." Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422330&pc=MTA | 2023826-2017-00444 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDE: ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. ON (B)(6) 2017, THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR THICK AND REDDISH BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDED ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. AFTER LENS EXPLANT, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100 AND THE PATIENT EXPERIENCED A LITTLE INFLAMMATION OF THE ANTERIOR SEGMENT. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. CONCLUSION: OFF-LABEL USE-KERATOCONUS EYE. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Extraction + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721249&pc=MTA | 2023826-2018-01084 | 2018-06-13 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP, ANGLE CLOSURE, INFLAMMATION, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE PATIENT IS "DOING WELL." AZETAZOLAMIDE 250MG GIVEN TO MANAGE SYMPTOMS. THE SURGEON LISTS THE CAUSE OF THE EVENT AS DEVICE. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DATE RECEIVED BY MFR: -IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422283&pc=MTA | 2023826-2017-00442 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/091 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN FORTE (ALLERGEN). UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: UPDATED INFORMATION-ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422303&pc=MTA | 2023826-2017-00443 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/109 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN COLLYRIUM. UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE (B)(4) NO CODE AVAILABLE (CYCLOPEGIA, TREATMENT WITH PREDNEFRIN COLLIRIUM) WORK ORDER SEARCH PERFORMED-NO SIMILAR COMPLAINTS FOUND. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795682&pc=MTA | 2023826-2017-01320 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/159 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT IOP IS 20.6MMHG, AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN, AND DIAMOX. MULTIPLE ATTEMPS TO GET MORE INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT'S IOP IS 20.6MM HG AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465307&pc=MTA | 2023826-2018-00607 | 2018-02-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+3.0/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2018. TO DATE, THE PATIENT HAS NOT DECIDED WHETHER TO REPLACE WITH ANOTHER LENS OR NOT. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, THE PATIENT DEVELOPED RAISED IOP AFTER TICL IMPLANT, WHICH WAS "REFRACTORY TO TOPICAL ANTI-GLAUCOMA MEDICATIONS." THE TICL WAS EXPLANTED AND A TRABECULECTOMY PERFORMED. THE CORNEA IS CURRENTLY CLEAR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876795&pc=MTA | 2023826-2017-01460 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/142 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876796&pc=MTA | 2023826-2017-01459 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641431&pc=MTA | 2023826-2017-00976 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER. | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4725289&pc=MTA | 2023826-2015-00428 | NA | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE LENS ROTATED (NOT ASSOCIATED TO A LOW VAULT). THE INCISION WAS ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED AT THE FIVE MONTH VISIT, THE SURGEON CONFIRMED THE OFF-AXIS ALIGNMENT AND DECIDED TO REPOSITION THE LENS. THE PATIENT EXPERIENCED BLURRED VISION. AFTER THIS INTERVENTION, THE PATIENT EXPERIENCED AN IOP RISE, WHICH WAS SOLVED BY ANTERIOR CHAMBER ASPIRATION AND IRRIGATION. FOR THIS SECONDARY INTERVENTION A NEW INCISION WAS CREATED, LEADING TO INCREASED CORNEAL ASTIGMATISM. THE LENS REMAINS IMPLANTED AND IS STABLE AFTER REPOSITIONING. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS IMPLANTED METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS UNDER-CORRECTING THE CYLINDER AND SURGICAL INDUCED ASTIGMATISM SECONDARY TO A NEW INCISION. THE LENS WAS OFF-AXIS AT ONE DAY POSTOPERATIVELY THEREFORE THIS COULD BE DUE TO SURGICAL MISALIGNMENT OR EARLY MOVEMENTS OF THE LENS. (B)(4). | Lens Repositioning | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589772&pc=MTA | 2023826-2017-00878 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS. (B)(6). (B)(4). Manufacturer Narrative: WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5766297&pc=MTA | 2023826-2016-00871 | 2016-05-12 | Injury | STAAR SURGICAL COMPANY | 2016-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Improper or Incorrect Procedure or Method; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/123 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, UNREACTIVE PUPIL (FIXED), GLARE/HALOS AND BLURRED VISION. THE SURGEON ENLARGED/ADDITION THE PERIPHERAL IRIDOTOMIES (YAG) AND ADDED A SUTURE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -1 AND THE PATIENT HAS PHOTOPHOBIA. Manufacturer Narrative: (B)(4): IMPROPER OR INCORRECT PROCEDURE OR METHOD - THE SURGEON REALIZED THAT THE IMPLANTED ICL MODEL REQUIRED PREOPERATIVE PERIPHERAL IRIDOTOMIES BUT HE FAILED TO DO IT DUE TO CONFUSION WITH THE ICL MODEL ORDERED. CORRECTION: FROM "13.7MM VTICMO13.7" TO "13.7MM VTICM13.7". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. (B)(4). | Lens Retention + Secondary Operation | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734555&pc=MTA | 2023826-2017-01130 | 2016-08-22 | Injury | Unknown Manufacturer | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.50/+3.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP WAS NOTICED. THE LENS WAS THEN REPOSITIONED. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/+3.5/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP WAS NOTICED. THE LENS WAS THEN REPOSITIONED. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PRIOR TO LENS EXCHANGE, BCVA WAS DECREASED TO 20/30 COMPARED TO BASELINE (20/20). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686682&pc=MTA | 2023826-2016-00778 | 2016-04-21 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pupillary Block; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/064 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT EXPERIENCED PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037544&pc=MTA | 2023826-2016-01480 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS 16.50/25/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.5/2.5/74 DIOPER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. PATIENT ALSO EXPERIENCED DIPLOPIA IN LOW LIGHT. THE LENS WAS EXCHANGED ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: DATE OF EVENT LISTED ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2016. DATE RETURNED TO MANUFACTURER ON THE INITIAL REPORT IS INCORRECT; CORRECT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037545&pc=MTA | 2023826-2016-01481 | 2016-06-11 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS - -16.00/25/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOS. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. DIPLOPIA EXPERIENCED IN LOW LIGHT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). METHOD: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.00/2.5/133 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARES/HALOES. DIPLOPLIA WAS ALSO EXPERIENCED IN LOW LIGHT. THE ELEVATED IOP WAS CONTROLLED WITH TWO GLAUCOMA MEDICATIONS. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. AS OF (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED AND UNCORRECTED VISUAL ACUITIES WERE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254868&pc=MTA | 2023826-2019-00062 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, GLARE/HALOS, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE OF LENS BECAUSE OF RIGHT EYE ANATOMY". Manufacturer Narrative: THE LENS WAS RETURNED MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7409133&pc=MTA | 2023826-2018-00523 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119131&pc=MTA | 2023826-2015-01310 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/X057 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE STRONG CENTRAL DEPOSITS ON THE CRYSTALLINE LENS (VISCO). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - 1 THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. 2 UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124740&pc=MTA | 2023826-2015-01307 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/X094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE DEPOSITS ON THE SURFACE OF THE CRYSTALLINE LENS (VISCO). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE (B)(4) REPORTED CASES, ALONG WITH RINGER'S SOLUTION. IN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. CLAIM # (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS IMPLANTED. WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151863&pc=MTA | 2023826-2015-01256 | 2015-09-15 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/X071 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. POST-OP. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS. SEE MFR #2023826-2015-01257 LEFT EYE. Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFÌCANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -6.5/+0.5/X071. (B)(4). METHOD: EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4672087&pc=MTA | 2023826-2015-00420 | NA | Injury | STAAR SURGICAL COMPANY | 2015-04-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND A REFRACTIVE SURPRISE WAS NOTED. PATIENT ALSO EXPERIENCED A LOSS OF BCVA. THE REPORTER INDICATED THE EVENT WAS THE RESULT OF A SURGEON'S ERROR WHEN ENTERING THE PRE-OP DATA IN OCOS. THE LENS REMAINS IMPLANTED, HOWEVER, THE SURGEON PLANS TO EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (REFRACTIVE SURPRISE03); VISION, LOSS OF; NO KNOWN PROBLEM WITH THE DEVICE. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW- WHEN A WRONG LENS POWER IS IMPLANTED AND EXCHANGED WITH THE CORRECT LENS POWER RESOLVES THE PROBLEM. THE SECONDARY INTERVENTION IS USUALLY DONE BY MEANS OF SAME INCISION WHICH WOULD NOT REQUIRE INCISION ENLARGEMENT AND/OR SUTURING DUE TO THE FOLDABLE PROPERTIES OF THE COLLAMER MATERIAL (ICL). (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4786463&pc=MTA | 2023826-2015-00594 | 2014-04-11 | Injury | STAAR SURGICAL COMPANY | 2015-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE PATIENT EXPERIENCED A LOSS OF VISION (PRE-OP 20/20,POST-OP 20/50. THE EVENT WAS THE RESULT OF A SURGEONS ERROR WHEN ENTERING CYLINDER SIGN IN OCOS. THE LENS WAS EXCHANGED FOR A DIFFERENT POWER, SAME SIZE ON (B)(6) 2014 AND THIS RESOLVED THE PROBLEM. UCVA ON (B)(6) 2014 WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE COMPLAINT FILE INDICATES THAT THE SURGEON IMPLANTED A WRONG LENS POWER BY MISTAKE. THE SURGEON CONFIRMS THE ERROR IN THE COMPLAINT FORM OR ADDITIONAL COMMUNICATIONS TO STAAR. THE POWER OF THE IMPLANTED LENS IS DIFFERENT FROM THE POWER LENS USED FOR EXCHANGE. THIS RESOLVED THE PROBLEM. SYMPTOMS REPORTED COULD INCLUDE: BLURRED VISION, DECREASED UCVA. WHEN A WRONG LENS POWER IS IMPLANTED AN EXCHANGE WITH THE CORRECT LENS POWER RESOLVES THE PROBLEM. THE SECONDARY INTERVENTION IS USUALLY DONE BY MEANS OF SAME INCISION WHICH WOULD NOT REQUIRE INCISION ENLARGEMENT AND/OR SUTURING DUE TO THE FOLDABLE PROPERTIES OF THE COLLAMER MATERIAL (ICL). (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5352440&pc=MTA | 2023826-2016-00013 | 2015-12-04 | Injury | STAAR SURGICAL COMPANY | 2016-01-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -13.00/+1.5/X069 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE IENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS. Manufacturer Narrative: CORRECT SN IS (B)(4). METHOD: LENS WORK ORDER SEARCH/DEVICE HISTORY REPORT REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. DEVICE HISTORY REPORT REVIEW: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.50/X069. DEVICE EVALUATED BY MANUFACTURER: DEVICE NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119125&pc=MTA | 2023826-2015-01308 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/X084 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON THE CRYSTALLINE LENS AFTER ICL EXPLANTATION. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. IN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, INDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119128&pc=MTA | 2023826-2015-01305 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/179 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE POST-OP DEPOSITS ON CRYSTALLINE LENS (CENTRAL) WHICH DECREASED VISUAL ACUITY AFTER ICL IMPLANTATION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE (B)(4) REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, (B)(4) HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124750&pc=MTA | 2023826-2015-01306 | NA | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.0/074 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WERE DEPOSITS ON CRYSTALLINE LENS (VISCO) WHICH DECREASED VISUAL ACUITY AFTER ICL IMPLANTATION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - 1 THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408629&pc=MTA | 2023826-2016-00114 | NA | Injury | STAAR SURGICAL COMPANY AG | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/036 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704671&pc=MTA | 2023826-2015-00426 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Loss of Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LOSS OF BEST-CORRECTED VISUAL ACUITY. ALSO REFRACTIVE CHANGE OVERTIME AND AN AXIS GAP OF ABOUT 90 DEGREES. THE LENS WAS RE-POSITIONED. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) Manufacturer Narrative: METHOD: - MEDICAL REVIEW. RESULTS.: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7381601&pc=MTA | 2023826-2018-00491 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Macular Edema; Inflammation; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.5/+0.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) OF 2017. THE SURGEON REPORTS MACULAR EDEMA WITHOUT OTHER SIGNS. IT WAS ALSO REPORTED THAT THIS WAS A MYOPIC CASE. THE SURGEON IS CONSIDERING REMOVAL OF ICL BUT IT HAS NOT BEEN CONFIRMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, LOSS OF BCVA, INFLAMMATION, AND MACULAR EDEMA. ON (B)(6) 2018 THE REPORTER STATED "DISAPPEARANCE OF MACULAR EDEMA OR POSTERIOR SYNECHIA WITH PUPILLARY DEFORMITY RIGHT EYE." SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction + Secondary Operation | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7890073&pc=MTA | 2023826-2018-01388 | 2018-07-18 | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Macular Edema; Inflammation; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BCVA (POST-OP BCVA 20/50) AND CYSTOID MACULAR EDEMA. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS INFLAMMATION. NO INFECTION OCCURED. IT WAS REPORTED THAT THE EXPLANT RESOLVED THE PROBLEM. PATIENT OK AFTER EDEMA ADSORPTION. NO PLANS TO IMPLANT AN ICL LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT - IS ADDED TO INITIAL MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15/+0.5/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BCVA (POST-OP BCVA 20/50) AND CYSTOID MACULAR EDEMA. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS INFLAMMATION. NO INFECTION OCCURRED. IT WAS REPORTED THAT THE EXPLANT RESOLVED THE PROBLEM. PATIENT OK AFTER EDEMA ADSORPTION. NO PLANS TO IMPLANT AN ICL LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867141&pc=MTA | 2023826-2017-01464 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Nausea; Blurred Vision; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. INVESTIGATION CONFIRMED THE LABELLING ISSUE. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.0/+5.5/158 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THAT THE PATIENT COULD SEE "NOTHING WITH OD AND CANNOT MOVE DUE TO NAUSEA WITH HIS ASTIGMATISM." THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS, VTICMO13.7, -6.0/+5.5/166 DIOPTER, AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON FELT THE LENS WAS MISLABELED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721248&pc=MTA | 2023826-2018-01078 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Pain; Blurred Vision; Halo; Discomfort; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, BLURRED VISION, PAIN, AND "NO COMFORT FOR NEAR AND MIDDLE VISION." THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403798&pc=MTA | 2023826-2019-00367 | 2019-02-04 | Injury | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; No Apparent Adverse Event | Pain; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN (B)(6) 2017. THE REPORTER INDICATED THAT AN ACCIDENT TO THE PATIENTS EYE HAS RESULTED IN BLUR VISION WHEN READING OR CONCENTRATING, AND PAIN WITH LIGHT CHANGES AND WHILE LOOKING AT A MOVING OBJECT. REPORTEDLY, THE PATIENT HAD A VISUAL CHECK WITH THE SURGEON AND THE VISUAL CHECK WAS "PERFECT." THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2017. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527426&pc=MTA | 2023826-2017-00698 | 2015-03-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Retinal Detachment; Blurred Vision; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -21.50/+4.5/100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED RETINAL DETACHMENT, AND LOST OF BCVA. THE PATIENT WAS TREATED WITH "EXTERNAL CERCLAGE" AND RETINAL REPAIR WAS PERFORMED ON (B)(6) 2015. THE CUSTOMER STATED HE THOUGH THE RETINAL DETACHMENT IS INDEPENDENT OF THE LENS. ON (B)(6) 2015 THE LENS WAS EXPLANTED. AT PRESENT, THE PATIENT REFUSES TO IMPLANT ANOTHER LENS. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4923027&pc=MTA | 2023826-2015-00828 | 2015-06-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WILL BE EXCHANGED FOR ANOTHER LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO -(OTHER): LENS REMAINS IMPLANTED. EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS -(OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS -(OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/1/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. EXPLANT DATE: - UNK. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CONCLUSION CODE: THE PATIENT IS UNDER 21 YEARS OLD AND HAS AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD AND ACD BELOW 3.0MM. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6158460&pc=MTA | 2023826-2016-01736 | 2016-11-11 | Injury | STAAR SURGICAL COMPANY | 2016-12-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.00/40.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. THE LENS REMAINS IMPLANTED AT THE TIME OF THIS REPORT. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -7.00/0.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2015. THE PATIENT IS EXPERIENCING ISSUE WITH VISION WHEN IT IS DARK. ON (B)(4) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, DIFFERENT LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7287284&pc=MTA | 2023826-2018-00243 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-02-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS TORN IN THE HAPTIC AREA, DURING DELIVERY INTO THE EYE AND REMAINS IMPLANTED. THE PATIENT COMPLAINS OF GLARE AND HALOS, ESPECIALLY AT NIGHT. THE SURGEON PLANS TO EXPLANT THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954143&pc=MTA | 2023826-2018-01500 | 2018-04-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574193&pc=MTA | 2023826-2017-00831 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.00/2.5/085 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED TO A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574194&pc=MTA | 2023826-2017-00830 | NA | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS; -16.50/2.5/097 DIOPTER; IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THIS WAS A 2ND LENS IMPLANTED IN THIS PATIENT (EXCHANGED FOR A SHORTER LENS) AND THE PATIENT REPORTED THE EXCHANGE DID NOT RESOLVE THE PROBLEM. THE PATIENT REPORTED VISION WAS POOR IN BAD LIGHTING AND HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6218810&pc=MTA | 2023826-2017-00006 | 2015-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; Discomfort | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+5.0/090 DIOPTER IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED VISION DISCOMFORT, GLARES/HALOES, AND BLURRED VISION. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT DATE OF IMPLANTATION IS (B)(6) 2015. (B)(4). Manufacturer Narrative: HIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | Unknown | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442665&pc=MTA | 2023826-2017-00511 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; Halo; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/80 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED GLARE/HALOES. THE ISSUE WAS RESOLVED, HOWEVER, THE PATIENT KEEPS EXPRESSING SUBJECTIVE DISCOMFORT. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (ACD<3.00MM). (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442671&pc=MTA | 2023826-2017-00510 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; Halo; Discomfort; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/96 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED GLARE/HALOES. THE ISSUE WAS RESOLVED, HOWEVER THE PATIENT KEEPS EXPRESSING SUBJECTIVE DISCOMFORT. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740001&pc=MTA | 2023826-2018-01147 | 2016-07-13 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -1.5/+5.5/060 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR DIOPTER LENS DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE SURGEON NOTES "HAPPY PATIENT". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824090&pc=MTA | 2023826-2018-01288 | 2018-07-30 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050569&pc=MTA | 2023826-2017-01751 | NA | Injury | STAAR SURGICAL COMPANY AG | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN 2014. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE AND HALOS, THE PATIENT HAD LARGE PUPIL SIZE. THE LENS WAS EXCHANGED FOR ANOTHER SAME SIZE (13.2 MM) BUT DIFFERENT DIOPTER (11.0/+2.0/109) LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6454953&pc=MTA | 2023826-2017-00563 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17/3.0/79 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. SURGEON STATED THAT THIS WAS DUE TO "BIG PUPIL," WHICH HE TOLD THE PATIENT WILL GET BETTER ON ITS OWN. AFTER EXPLANT, THE EYE IS FINE AND THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SCRATCHES ON A HAPTIC. THERE WAS EVIDENCE OF CLEAR AND DARK BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6454975&pc=MTA | 2023826-2017-00564 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/096 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES. SURGEON STATED THAT THIS WAS DUE TO "BIG PUPIL," WHICH HE TOLD PATIENT WILL GET BETTER ON ITS OWN. AFTER EXPLANT, THE EYE IS FINE AND THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201664&pc=MTA | 2023826-2018-00094 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/001 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "READING PROBLEMS" AND LENS WAS EXPLANTED "DESPITE VERY GOOD VA." Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF FOREIGN MATERIAL (RESIDUE) ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201666&pc=MTA | 2023826-2018-00096 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/160 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "READING PROBLEMS" AND LENS WAS EXPLANTED "DESPITE VERY GOOD VA." Manufacturer Narrative: SHOULD HAVE BEEN RECORDED IN INITIAL MDR. THE PATIENT EXPERIENCED GLARE/HALOES ALONG WITH "READING PROBLEMS." (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7456987&pc=MTA | 2023826-2018-00595 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -9.50/+1.00/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO GLARES AND HALOES. IT WAS REPORTED THAT THE PATIENT HAS RECOVERED AND THERE IS NO MORE PROBLEMS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108232&pc=MTA | 2023826-2016-01649 | 2016-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.00/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT HAS SIGNIFICANT PROBLEMS WITH GLARE/HALOS. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT/EXCHANGE FOR A DIFFERENT MODEL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FOREIGN MATERIAL (FIBERS) WAS PRESENT ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED FROM USER FACILITY TO DISTRIBUTOR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108233&pc=MTA | 2023826-2016-01650 | 2016-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT HAD SIGNIFICANT PROBLEMS WITH GLARE/HALOS. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT/EXCHANGE FOR A DIFFERENT MODEL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: THE LENS WAS EXCHANGED FOR ANOTHER LENS BUT THE PROBLEM WAS NOT RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701687&pc=MTA | 2023826-2017-01152 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/63 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED EXCESSIVE VAULT, GLARE/HALOES, SLOW PUPILLARY REACTION, OVERCORRECTION, AND SEES "CENTERFLOW" REFLECTIVES EVEN ON DAYTIME. THE CUSTOMER STATED THAT THE SURGEON PLANNED TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801631&pc=MTA | 2023826-2017-01301 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED GLARE AND HALOS. THE PROBLEM WAS RESOLVED AFTER THE EXCHANGE. THE PATIENT'S POST-OP SHOWED UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS SLIGHTLY CURVED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440720&pc=MTA | 2023826-2017-00503 | 2017-02-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/005 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. ALTHOUGH THE PATIENT STILL COMPLAINED OF GLARE AND "READING PROBLEMS," THE SURGEON BELIEVED AFTER "ADAPTION," TIME, AND HEALING THE PROBLEM WOULD BE RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440729&pc=MTA | 2023826-2017-00504 | 2017-02-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/154 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ICA, AND GLARE/HALOES. ALTHOUGH THE PATIENT STILL COMPLAINED OF "READING PROBLEMS" AND "STRANGE" VISUAL EFFECTS AT DISTANCE, THE SURGEON STATED THAT OBJECTIVELY EVERYTHING WAS FINE AND HE BELIEVED THAT AFTER TIME AND HEALING THE PROBLEM WOULD BE RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA:WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER. | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657591&pc=MTA | 2023826-2017-01020 | 2017-03-01 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/1.5/74 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 01 MAR 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, AND GLARE/HALOES. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). RESULT CODE: PREVIOUS CODE NOT APPLICABLE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS), THE LENS MISSING A PIECE OF HAPTIC, AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320737&pc=MTA | 2023826-2018-00328 | 2018-02-02 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+3.00/119 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLARE/HALOES. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.85MM)]. (B)(4). | Lens Retention | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516771&pc=MTA | 2023826-2018-00745 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLARE/HALOS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740014&pc=MTA | 2023826-2018-01141 | 2018-06-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596417&pc=MTA | 2023826-2016-00500 | 2015-12-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/+2.00/X019 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF INDO-CORNEAL ANGLES, GLARE/HALOES WAS OBSERVED ON (B)(6) 2015 THE ICL WAS EXPLANTED ON (B)(6) 2016, THE INCISION WAS ENLARGED TO REMOVE THE ICL. ANTERIOR CHAMBER IRRIGATION/ELEVATION OF THE REMAINING VISCO/FLUIDS AND INTRAOCULAR INJECTION (MEDICATION) WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. A SUTURE WAS REQUIRED. LAST VISIT ON (B)(6) 2016, BCVA: 20/20. Manufacturer Narrative: (B)(4). PATIENT IS ABOVE THE AGE OF 45 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS ABOVE THE AGE OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. NEW INCISION. SECONDARY SURGERY. ADDITIONAL SUTURE. LENS EXPLANTED. VAULTING. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908149&pc=MTA | 2023826-2017-01499 | 2017-09-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -13/+5/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO LOW VAULTING, LENS ROTATION, AND GLARES/HALOES. A NEW LENS IS PLANNED TO BE IMPLANTED NEXT MONTH. THE PATIENT IS CURRENTLY FINE. Manufacturer Narrative: THE NEW LENS WAS NOT IMPLANTED YET DUE TO PATIENT WORKING ISSUES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A VTICM5 13.2 IMPLANTABLE COLLAMER LENS -12.5/+5 /93 DIOPTER. THE SURGEON NOTES THAT THE VISUAL ACUITY AND VAULT ARE OKAY, AND NO ROTATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXPLANT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740026&pc=MTA | 2023826-2018-01150 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/174 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), GLARE/HALOES, AND A "SLIGHT TILT." THE CAUSE IS LISTED AS "SIZING." THE LENS REMAINS IMPLANTED AND PATIENT IS UNDER EVALUATION. Manufacturer Narrative: STATES THE LENS REMAINS IMPLANTED. CORRECTED INFORMATION ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8191674&pc=MTA | 2023826-2018-01877 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -5.0/+1.50/5 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND GLARE/HALOES. A REPLACEMENT LENS, SHORTER LENGTH WAS LATER IMPLANTED IN A SECOND SURGERY. IT WAS REPORTED THAT THE REPLACEMENT RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTION PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. PATIENT CODE 2140 - 2227 IN THE PREVIOUS MDR SHOULD BE INCLUDED IN THE ORIGINAL MDR. DEVICE CODE 3001 IN THE PREVIOUS MDR SHOULD BE INCLUDED IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853013&pc=MTA | 2023826-2018-01336 | 2018-08-07 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS EXPERIENCING GLARE AND "HALOS DUE TO THE AQUAPORT". THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -3.00/+4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED GLARES AND "HALOS DUE TO AQUAPORT". THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853023&pc=MTA | 2023826-2018-01333 | 2018-08-07 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5/+2.5/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS EXPERIENCING GLARES AND "HALOS DUE TO THE AQUAPORT." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -2.50/+2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REPORTEDLY THE PATIENT EXPERIENCED GLARES AND "HALOS DUE TO AQUAPORT". THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION OF THE LENS FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8365953&pc=MTA | 2023826-2019-00268 | 2019-01-04 | Injury | STAAR SURGICAL COMPANY | 2019-02-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Visual Disturbances; Halo; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -13.50/3.0/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT, PIGMENT DISPERSION, UNREACTIVE(FIXED) PUPIL, AND GLARE/HALOES WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6110075&pc=MTA | 2023826-2016-01634 | 2016-04-06 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Optical Decentration; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES, BLURRED VISION, AND THE LENS BEING DECENTERED. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL AND SIZE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE PREVIOUS DEVICE CODE (B)(4) WAS INCORRECT; REPLACED BY (B)(4) METHOD CODE WORK ORDER SEARCH PERFORMED AND NO SIMILAR COMPLAINTS FOUND CONCLUSION CODE PATIENT ACD WAS BELOW 3.0 MM. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176968&pc=MTA | 2023826-2018-01855 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS (VTICM). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: EVENT: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM. DEVICE EVALUATED BY MFR : DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE/RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: DATE OF EVENT CORRECTED FROM (B)(6) 2016 TO UNK. DEVICE EVALUATION- LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE- 2227 ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255318&pc=MTA | 2023826-2019-00107 | 2018-06-25 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXCHANGED FOR A PHAKIC INTRAOCULAR LENS (PIOL). THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON 12-FEB-2019 STATING THE REPLACEMENT LENS WAS AN "ARTIFLEX". DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.; (B)(4). | Lens Exchange | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8255338&pc=MTA | 2023826-2019-00109 | 2018-07-30 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Visual Disturbances; Halo; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+6.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS DUE TO BIG PUPIL SIZE. THE LENS WAS NOT REPLACED WITH ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320749&pc=MTA | 2023826-2018-00324 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/041 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED. REPORTEDLY, THE PATIENT EXPERIENCED "PROBLEMS READING AND DOING COMPUTER WORK." THE SURGEON STATES THE EYE IS "CALM AND QUIET" AFTER THE EXCHANGE AND THAT THE PROBLEM IS RESOLVED. Manufacturer Narrative: - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR AXIS, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. WORK ORDER SEARCH FOUND NO OTHER SIMILAR DEFECT WITHIN THE SAME LOT. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "PROBLEMS READING AND DOING COMPUTER WORK." (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6598926&pc=MTA | 2023826-2017-00901 | 2017-03-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/84 IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED GLARE AND UNREACTIVE PUPIL (FIXED). AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320720&pc=MTA | 2023826-2018-00321 | 2017-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO OCCASIONAL "FLASHING LIGHTS." ATTEMPTS TO OBTAIN PATIENT'S CURRENT STATUS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE PATIENT COMPLAINS OF CHROMATIC ABERRATION AND COLOR DISTORTION. REPORTEDLY THIS DID NOT OCCUR PRIOR TO ICL IMPLANT. PATIENT STATES THESE SYMPTOMS, INCLUDING "FLASHES OF LIGHT," PERSIST AFTER ICL EXPLANT. THE PATIENT STATES THESE DISTURB HER DAILY LIFE. OCT AND FUNDUS CAMERA TESTS WERE TAKEN TO ENSURE THE BACK OF THE EYE IS IN GOOD SHAPE. APPARENTLY, THE PATIENT'S FEEDBACK KEEPS CHANGING AND INVOLVED SURGEONS BELIEVE THE PATIENT MAY BE "MENTALLY UNSTABLE." THE SURGEON REPORTS THAT HE HAS DISCHARGED THE PATIENT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: (B)(4). CONCLUSION: IT IS ADDED TO THE PREVIOUS SUPPLEMENTAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320758&pc=MTA | 2023826-2018-00323 | 2017-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/+1.5/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO OCCASIONAL "FLASHING LIGHTS." ATTEMPTS TO OBTAIN PATIENT'S CURRENT STATUS HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE PATIENT COMPLAINS OF CHROMATIC ABERRATION AND COLOR DISTORTION. REPORTEDLY THIS DID NOT OCCUR PRIOR TO ICL IMPLANT. PATIENT STATES THESE SYMPTOMS, INCLUDING "FLASHES OF LIGHT," PERSIST AFTER ICL EXPLANT. THE PATIENT STATES THESE DISTURB HER DAILY LIFE. OCT AND FUNDUS CAMERA TESTS WERE TAKEN TO ENSURE THE BACK OF THE EYE IS IN GOOD SHAPE. APPARENTLY, THE PATIENT'S FEEDBACK KEEPS CHANGING AND INVOLVED SURGEONS BELIEVE THE PATIENT MAY BE "MENTALLY UNSTABLE." THE SURGEON REPORTS THAT HE HAS DISCHARGED THE PATIENT. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.94MM)]. (B)(4). | Lens Extraction | Not_Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6917932&pc=MTA | 2023826-2017-01501 | 2017-08-31 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED A YEAR AFTER THE IMPLANTATION AS THE PATIENT WAS BOTHERED BY THE CENTRAL HOLE OF THE LENS AS THE LIGHT DISRUPTION FROM THE HOLE DIDN'T GET BETTER. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/105 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED AS THE PATIENT WAS BOTHERED BY THE CENTRAL HOLE OF THE LENS AS THE LIGHT DISRUPTION FROM THE HOLE DIDN'T GET BETTER. THERE IS NO PLAN TO IMPLANT ANOTHER LENS. PATIENT POST-OP BCVA WAS 20/20 ON (B)(6) 2017. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4). | Lens Extraction | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152459&pc=MTA | 2023826-2017-01966 | 2017-11-28 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/072 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "VERY STRONG GLARES," AND GHOST IMAGES. AFTER EXPLANT, IT IS NOTED THAT "AS BEFORE SURGERY CRYSTALLINE LENS IS CLEAR." Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE LENS EXPLANTATION RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152539&pc=MTA | 2023826-2017-01965 | 2017-11-28 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/093 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. REPORTEDLY, THE PATIENT EXPERIENCED "VERY STRONG GLARES," AND GHOST IMAGES. AFTER EXPLANT, IT IS NOTED THAT "AS BEFORE SURGERY...LENS IS CLEAR." Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE. (B)(4) Manufacturer Narrative: THE LENS EXPLANT RESOLVED THE PROBLEM. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245216&pc=MTA | 2023826-2018-00198 | 2017-12-29 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.50/+2.50/105 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH THE SAME SIZE, DIFFERENT DIOPTER LENS. REPORTEDLY, THE PATIENT WAS UNHAPPY POSTOPERATIVELY AND THE SURGEON DECIDED TO CHANGE THE LENS "BASED ON EITHER WRONG ASIGMATS AND/OR WRONG AXIS" OF THE LENS. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE OR DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND PRESENCE OF CLEAR AND WHITE RESIDUE AND DEBRIS ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245212&pc=MTA | 2023826-2018-00197 | 2018-01-02 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/091 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH THE SAME SIZE, DIFFERENT DIOPTER LENS. REPORTEDLY, THE PATIENT WAS UNHAPPY POSTOPERATIVELY AND THE SURGEON DECIDED TO CHANGE THE LENS "BASED ON EITHER WRONG ASIGMATS AND/OR WRONG AXIS" OF THE LENS. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND WHITE RESIDUE AND DEBRIS). (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.95MM)] (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7395672&pc=MTA | 2023826-2018-00502 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/89 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND EXCHANGED DUE TO GLARE/HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7897512&pc=MTA | 2023826-2018-01435 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND GLARE/HALOES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6562054&pc=MTA | 2023826-2017-00799 | 2017-03-07 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Migration or Expulsion of Device; Inadequacy of Device Shape and/or Size | Visual Disturbances; No Code Available | N/A | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -805/+1.0/070 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND GHOSTING (DOUBLE VISION). THE SURGEON DECIDED TO IMPLANT THE LENS HORIZONTALLY, STATING THE LENS OPTICAL ZONE HAD MIGRATED DOWNWARD. ALTHOUGH THERE IS STILL LOW VAULT, THE PATIENT AND SURGEON ARE HAPPY WITH THE RESULTS. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: PREVIOUS DIOPTER: -805/+1.0/070; CORRECT DIOPTER: -8.5/+1.0/070. THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PREVIOUS DIOPTER: -805/+1.0/070; CORRECT DIOPTER: -8.5/+1.0/070. THE LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC MISSING A PIECE. CLAIM# (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006956&pc=MTA | 2023826-2017-01685 | 2017-08-31 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Operates Differently Than Expected | Visual Disturbances; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/067 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO DYSPHOTOPSIA; THE PATIENT WAS "BOTHERED BY THE CENTRAL HOLE CONSTANTLY, WHICH HE COULD NOT TOLERATE." THE EXPLANT RESOLVED THE PROBLEM. NO LENS EXCHANGE IS PLANNED. Manufacturer Narrative: (B)(4). RESULT: NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). | Lens Extraction | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119133&pc=MTA | 2023826-2015-01309 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/X145 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY POST-OP. THE LENS REMAINS IMPLANTED. THERE WERE POST-OP DEPOSITS ON THE SURFACE OF THE CRYSTALLINE LENS (VISCO) WHICH DECREASED THE VISUAL ACUITY . Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF (B)(6) 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS". THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, INDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION (BSS, ALCON) AND OCUCOAT (BAUSCH & LOMB) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. CLAIM # (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). | Lens Retention | Unknown
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024707&pc=MTA | 2023826-2016-01453 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material | Visual Impairment | NA | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/092 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016, AND THE LENS TORE/BROKE DURING INJECTION INTO THE EYE. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WILL BE TAKEN OUT AT A LATER TIME.Event Description: THE LENS WAS EXCHANGED FOR A SAME SIZE LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN, AND DISCOLORED MATERIAL ON SURFACE. PREVIOUS PATIENT CODE (B)(4) (VISION, IMPAIRED) WAS INCORRECT; ONLY LOSS OF BCVA WAS REPORTED. METHOD CODE (DEVICE FROM SAME LOT EVALUATED); A WORK ORDER SEARCH WAS PERFORMED. (B)(4). | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119420&pc=MTA | 2023826-2015-01311 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; No Code Available | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/X132 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS WITH A DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP ((B)(6) 2015) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. SECONDARY INTERVENTION, EXCHANGE AND PROBLEM RESOLVED. AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OLD. (B)(4). Manufacturer Narrative: EVALUATION RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. Manufacturer Narrative: (B)(4): METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS IMPLANTED. | Lens Exchange | Resolved
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5365600&pc=MTA | 2023826-2016-00009 | 2015-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2016-01-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Fracture; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Visual Impairment; Visual Disturbances; Halo | NA | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/+2.5/X099 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS TORE DURING INSERTION INTO THE EYE. A SIGNIFICANT TEAR WAS NOTICED AFTER IMPLANT. LENS WAS LEFT IN THE EYE WITH SUBSEQUENT SYMPTOMS (HALOS AND OCCASIONAL DIPLOPIA) AND DISPLACEMENT (LENS ROTATION). PATIENT LAST OFFICE VISIT ON (B)(6) 2015, UCVA, 20/50. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.50/+2.50/X099. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE REMAINS IMPLANTED. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). DEVICE REMAINS IMPLANTED. | Lens Retention | Unknown