## Original_MTA
Web Address | Report Number | Event Date | Event Type | Manufacturer | Date Received | Product Code | Brand Name | Device Problem | Patient Problem | PMA/PMN Number | Number of Events | Event Text
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13765543&pc=MTA | 2023826-2022-00787 | 2022-01-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event |  |  | 1 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 12.6MM VICM5_12.6; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022; THE LENS TORE. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11713110&pc=MTA | 2023826-2021-01141 | 2018-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT HAD SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS OVER A PERIOD OF 3 YEARS. THE SURGEON HAS DECIDED TO LEAVE THE LENS IMPLANTED AND THE PATIENT WILL BE MONITORED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738872&pc=MTA | 2023826-2021-01192 | 2018-10-31 | Injury | STAAR SURGICAL COMPANY | 2021-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PIGMENT DISPERSION. THIS RESOLVED THE PROBLEM. REPORTEDLY, IN ORDER TO PREVENT THE ADVERSE EFFECTS CAUSED BY LONG-TERM EXISTENCE IN THE FUTURE, THE LENS WAS REMOVED FROM POSTERIOR CHAMBER INTRAOCULAR LENS WITH LENS. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900025&pc=MTA | 2023826-2020-02749 | 2020-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). APPROXIMATELY 2 WEEKS LATER IT WAS FOUND THE PATIENTS ENDOTHELIAL VALUE HAD DECREASED SIGNIFICANTLY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS, AND THE PATIENT IS BEING MONITORED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507492&pc=MTA | 2023826-2021-00661 | 2021-01-29 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS AN UNREACTIVE (FIXED) PUPIL. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS OTHER: THE SURGEON THINKS IT'S RELATED TO OVD REMANANT. THE PUPIL IS UNREACTIVE, EVEN WITH PIL TREATMENT THE PUPIL DOES NOT CLOSE. THE PATIENT IS UNCOMFORTABLE WITH THE LIGHT. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443637&pc=MTA | 2023826-2021-03481 | 2021-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPECTED A BETTER NEAR VISION OUTCOME. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS BUT THE PROBLEM WAS NOT RESOLVED. POST-OP THE PATIENT HAD A SMALL ANTERIOR SUBCAPSULAR CATARACT AND VISION IS WORSE. SEE MFR. REPORT # 2023826-2021-03482 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. THERE WAS DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. (B)(4). CLAIM # : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478420&pc=MTA | 2023826-2021-03565 | 2021-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) 2 TIMES, DURING THE SAME SURGERY DUE TO THE LENS WOULD NOT UNFOLD. THERE WAS SOME IRIS PROLAPSE DURING THE PROCEDURE. AFTER THE SECOND TRY, THE SURGEON REPLACED THE LENS WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE VAULT IS NORMAL AND SOME CORNEAL EDEMA. CONTINUE VIGAMOX/ACULAR/PF QID OU. THE LENS WAS DISCARDED. Manufacturer Narrative: B1- CORRECTED TO "PRODUCT PROBLEM (MALFUNCTION)." H1- CORRECTED TO "MALFUNCTION." CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107551&pc=MTA | 2023826-2022-01197 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A SPHERICAL NON-TORIC ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT UNDERWENT PRK SURGERY FOR ASTIGMATISM ON (B)(6) 2022 AND IS NOW DOING WELL. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE: UNK. EXPLANT DATE: NA. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). Manufacturer Narrative: B5: THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14310615&pc=MTA | 2023826-2022-01395 | 2022-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, IRIS ATROPHY, FIXED PUPIL WITH DIFFUSE TRANSILLUMINATION IRIS DEFECTS (TID) WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IT WAS ALSO REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED, "GOOD SURGICAL OUTCOME, HAPPY VISION, VAULT GOOD 560MICRONS." IT WAS REPORTED THAT NO MEDICAL OR SURGICAL INTERVENTION WAS NECESSARY. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PERSISTENT, MILD DILATED, FIXED PUPIL WITH DIFFUSE TID OD ONLY. ACUITY GOOD." FOR ADDITIONAL INFORMATION THE REPORTER STATED "THIS IS THE SECOND CASE OF FIXED PUPIL WITH DIFFUSE TIDS. KEEPING LENS IN PLACE, AS I DON'T THINK IT IS THE ICL. WILL NEED TO MONITOR IOP DUE TO PIGMENT DISPERSION. I AM UNCLEAR AS TO ETIOLOGY, BUT THINK IT MAY HAVE TO DO WITH INTRACAMERAL MOXI-DEX (5MG/CC MOXI). HOWEVER, NO PROBLEMS WITH CATARACT CASES. OR OTHER EYES FOR ICLS." Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H1 - TYPE OF REPORTABLE EVENT SHOULD HAVE BEEN SELECTED AS MALFUNCTION IN PREVIOUS SUPPLEMENTAL MFR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937259&pc=MTA | 2023826-2022-04286 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL;SIZE LENS BUT DIFFERENT DIOPTER DUE TO PATIENT DISSATISFACTION DUE TO UNDER CORRECTION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- UNDER CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DEBRIS THROUGHOUT THE LENS AND CLEAR RESIDUE IN OPTIC. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032620&pc=MTA | 2023826-2022-04640 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 -7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL LOSS/COUNT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032644&pc=MTA | 2023826-2022-04641 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 -7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL LOSS/COUNT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018654&pc=MTA | 2023826-2022-04627 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018234&pc=MTA | 2023826-2022-04628 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16034514&pc=MTA | 2023826-2022-04643 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE REPORTER STATED "WE EXPLANTED THE LENS BECAUSE THE PATIENT WAS NOT PHYSICALLY SATISFIED. SHE SAW ONE PUPIL BIGGER THAN THE OTHER. SHE NOTICED THE PUPIL OF THE RIGHT EYE WAS LARGER THAN THAT OF THE LEFT EYE." IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS PATIENT-RELATED FACTOR. HOWEVER, IT WA REPORTED THAT IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18014407&pc=MTA | 2023826-2023-04782 | 2023-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO PATIENT DISSATISFACTION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-CLINICAL CODE: 4681- PATIENT DISSATISFACTION. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739732&pc=MTA | 2023826-2023-01369 | 2023-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTES THE PATIENT HAD "BIGGER PUPIL DIAMETER POSTOP THAN PREOP". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON NOTES THE PATIENT HAD "BIGGER PUPIL DIAMETER POSTOP THAN PREOP". LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. D4: MODEL: SHOULD BE CORRECTED TO VICM5_13.2. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - BIGGER PUPIL. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17225357&pc=MTA | 2023826-2023-02469 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116582&pc=MTA | 2023826-2023-04949 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO PATIENT DISSATISFACTION-VISUAL ACUITY. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-CLINICAL CODE 4581: PATIENT DISSATISFACTION WITH VISUAL ACUITY. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18249912&pc=MTA | 2023826-2023-05298 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, 09.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON PERFORMED AN LRI TO CORRECT ASTIGMATISM BUT THE ASTIGMATISM WAS NOT FULLY CORRECTED. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT WAS DISSATISFIED WITH VISUAL ACUITY. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DISSATISFACTION WITH VISUAL ACUITY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655765&pc=MTA | 2023826-2023-03799 | 2023-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON NOTES THERE WAS INTRAOCULAR HEMORRHAGE AND INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "DURING THE OPERATION THE PATIENT HAD INTRAOCULAR HEMORRHAGE AT THE TEMPORAL ANGLE, WHICH WAS MANAGED AND THE LENS WAS REMOVED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - INTRAOCULAR HEMORRHAGE. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18376653&pc=MTA | 2023826-2023-05568 | 2023-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -02.75 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT DISSATISFACTION WITH NEAR VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (02.25) LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PATIENT DISSATISFACTION NEAR VISION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10659707&pc=MTA | 2023826-2020-02350 |  | Injury | STAAR SUGICAL COMPANY | 2020-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6,-5.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2020. ON (B)(6) 2020 A REMOVAL AND EXCHANGE WAS PERFORMED DUE TO SIZING AND ANGLE CLOSURE. THE REPORTER STATED " NO PERMANENT PATIENT INJURY ALTHOUGH THERE HAD BEEN TEMPORARY ANGLE CLOSURE NECESSITATING ADDITIONAL PI'S AND INTERVENTION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE ZIPLOC BAG. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BENT AND RESIDUE ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ADDITIONAL PI SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091994&pc=MTA | 2023826-2020-03227 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE PATIENT COMPLAINED OF CENTRAL CIRCULAR DYSPHOTOPSIA. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - CENTRAL CIRCULAR DYSPHOTOPSIA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11092002&pc=MTA | 2023826-2020-03228 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT COMPLAINED OF CENTRAL CIRCULAR DYSPHOTOPSIA. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT- CLINICAL CODE: 4581 - CENTRAL CIRCULAR DYSPHOTOPSIA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651548&pc=MTA | 2023826-2021-00887 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, LENS IS TOO BIG AND CAUSING NARROW ANGLES. SURGEON WOULD LIKE TO EXCHANGE FOR A SHORTER LENGTH LENS. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLAUCOMA, PIGMENT DISPERSION, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "WRONG SIZE". PER REPORTER ON (B)(6) 2021, "THE ANGLE HAS OPENED MORE AND THE PRESSURE IS STABILIZED TO NORMAL. THE DOCTOR IS CLOSELY WATCHING THE PATIENT AND CHECKING HER ANGLES TO SEE IF SHE IMPROVES OVER TIME AS THE EYE QUIETS DOWN." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6: HEALTH EFFECT - CLINICAL CODE - 4581 - PIGMENT DISPERSION. H6: HEALTH EFFECT - IMPLANT CODE - REPORTED 2199 - UPDATE TO 4629. H6: MEDICAL DEVICE PROBLEM CODE - REPORTED 2993 - UPDATE TO 1583. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11738951&pc=MTA | 2023826-2021-01202 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2021. NARROW ANGLES ARE BEING OBSERVED. LENS REMAINS IMPLANTED AND PATIENT IS UNDER OBSERVATION. PER REPORTER ON (B)(6) 2021, "ANGLE HAS OPENED MORE AND THE PRESSURE IS STABILIZED TO NORMAL. THE SURGEON IS CLOSLY WATCHING THE PATIENT AND CHECKING ANGLES TO SEE IF THERE ARE IMPROVEMENTS OVERTIME AS THE EYE QUIETS DOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT IMPACT CODE (B)(4): NARROW ANGLES. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13002075&pc=MTA | 2023826-2021-04802 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT#: 2023826-2021-03077. DUE TO IRIS PROLAPSE DURING THE EXCHANGE, THE PUPIL WAS DISTORTED TEMPORALLY OUTWARD. THE IRIS PROLAPSE HAS BEEN RESOLVED AND NO SURGICAL INTERVENTION WILL BE PERFORMED. THE LENS REMAINED IMPLANTED AND THE PATIENT WILL BE STRICTLY MONITORED. Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, DATE OF EVENT: UNK, EXPLANT DATE: NA, PMA/510K : THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13006655&pc=MTA | 2023826-2021-04886 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -2.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED "MINOR VOLCANO EFFECT" AND NOTED "BUT THE CHAMBER IS DEEP". LENS REMAINS IMPLANTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -2.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED "MINOR VOLCANO EFFECT" AND NOTED "BUT THE CHAMBER IS DEEP". LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14150295&pc=MTA | 2023826-2022-01240 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE A PATIENT EXPERIENCED RETINAL DETACHMENT AFTER ICL IMPLANTATION 3 OR 4 MONTHS POST OP. REPORTEDLY IT HAS BEEN FIXED AND DOCTOR BELIEVES IT IS OKAY. Manufacturer Narrative: PATIENT IDENTIFIER-RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE, IMPLANT DATE, EXPLANT DATE: UNK. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220985&pc=MTA | 2023826-2022-01152 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: REPORTS FOR REDUCTION IN ENDOTHELIAL CELL COUNT IN PATIENTS RIGHT EYE. SEE MFR. REP#2023826-2022-01153 FOR FELLOW EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14221078&pc=MTA | 2023826-2022-01153 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: REPORTS FOR REDUCTION IN ENDOTHELIAL CELL COUNT IN PATIENTS LEFT EYE. SEE MFR. REP#2023826-2022-01152 FOR FELLOW EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153632&pc=MTA | 2023826-2022-02509 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM6_12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED PERMANENT LOSS OF BCVA AND THE PROBLEM WAS RESOLVED. THE PATIENT HAD A REFRACTIVE LENS EXCHANGE WITH A TRIFOCAL LIO. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI) # (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15414373&pc=MTA | 2023826-2022-03108 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THE PATIENT IS HAVING DIFFICULTY READING/SEEING; AND WAS PRESCRIBED MULTIPLE MEDICATIONS WITH A PLANNED FOLLOW UP VISIT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS. REPORTEDLY, "PATIENT TOLERATED THE PROCEDURE WELL" WITHOUT COMPLICATIONS. H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498836&pc=MTA | 2023826-2022-03340 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. THE LOW VAULT ISSUE WAS RESOLVED BUT " MILD MYDRIASIS DUE TO SECTORAL ATROPHY OF THE PUPILLARY SPHINCTER " WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555886&pc=MTA | 2023826-2022-03513 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -5.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. IT WAS REPORTED THAT THE VAULT ISSUE RESOLVED BUT FOR SATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " PERSISTENT RESIDUAL ASTIGMATISM." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16603911&pc=MTA | 2023826-2023-00946 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED AN ICL IMPLANTABLE COLLAMER LENS, DUE TO THE SURGEON NOTED PIGMENT DISPERSION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED PIGMENT DISPERSION. REPORTEDLY "DR. (B)(6), NOTED TODAY THAT HE BELIEVES HE IS SEEING PIGMENT DISPERSION ON A CASE THAT THE LENS WAS REMOVED AND THEN IMPLANTED. HE WOULD LIKE A CONSULTATION." THE LENS WAS EXPLANTED AND REPLACED. IT WAS LATER REPORTED "THE PATIENT IS DOING BETTER AND THE PIGMENT DISPERSION DOES NOT APPEAR ACTIVE ANYMORE. THE LIBERATED PIGMENT MOST LIKELY CAME FROM SURGERY. OCT SHOWED THE ICL VAULT TO BE AT NORMAL AMOUNT. I CAN'T RECALL EXACTLY. I THINK IT WAS RIGHT AT 500 MICRONS. IOP HAS REMAINED LOW. VISION EXCELLENT." CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DARE: UNK. DEVICE MANUFACTURE DATE: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16769941&pc=MTA | 2023826-2023-01373 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: RECEIVED A COPY OF AN ARTICLE IN REGARD TO "A HYPOTHESES TO EXPLAIN ACUTE ADJUSTMENT DISORDER AFTER PHAKIC INTRAOCULAR LENS IMPLANTATION IN A HIGHLY MYOPIC PATIENT: POTENTIAL CAUSES AND SUGGESTED REMEDIES". THE ARTICLE INVOLVED A 35 YEAR OLD FEMALE WITH HIGH MYOPIA FOLLOWING POST BILATERAL ICL PLACEMENT. AT ONE WEEK POST-OP, THE PATIENT WAS HAVING DIFFICULTY ADJUSTING TO THE VISION PROVIDED BY HER IMPLANTS PARTICULARLY RELATED TO IMAGE SIZE. THE PATIENT WAS HAVING WAXING AND WANING SYMPTOMS OF PANIC ATTACKS AND WAS PRESCRIBED ANTI-ANXIETY MEDICATION. THIS LESSENED SOME OF HER SYMPTOMS BUT SHE CONTINUED TO HAVE DIFFICULTY ADJUSTING. THE PATIENT WAS STARTED ON SEVERAL DIFFERENT SELECTIVE SEROTONIN REUPTAKE INHIBITORS WITH MILD IMPROVEMENT. SHE THEN SOUGHT COGNITIVE THERAPY TO HELP WITH HER DIFFICULTY ADJUSTING WITH SUBJECTIVE IMPROVEMENT AT HER 2 MONTH FOLLOW-UP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - DIFFICULTY ADJUSTING TO VISION, IMAGE SIZE (PANIC ATTACKS). HEALTH EFFECT IMPACT CODE: 4644 - SEROTONIN REUPTAKE INHIBITORS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16925028&pc=MTA | 2023826-2023-01809 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY MEDICAL DEVICE REPORT FROM FDA- MDR REPORT # MW5116464 ON 13-APR-2023. THE PATIENT REPORTED ON (B)(6) 2023, A "SERIOUS INJURY" EVENT REPORT WITH OUTCOMES OF "DISABILITY OR PERMANENT DAMAGE". THE PATIENT STATED "HAD STAAR MEDICAL'S IMPLANTABLE COLLAMER LENS, EVO ICL IMPLANTED INTO BOTH RIGHT AND LEFT EYE. NOTICED NEXT DAY A LOSS OF PERIPHERAL VISION. RETURNED TO DOCTOR, WAS TOLD TO WAIT A WEEK. RETURNED TO HAVE RETINA CHECKED, RETINA WAS OKAY". PATIENT STATES THAT ISSUE IS STILL PRESENT AND THAT THIS SIDE EFFECT WAS NOT LISTED AS A POSSIBLE SIDE EFFECT. PATIENT IS UNSATISFIED WITH POST-OP OUTCOME. PER THE SAME REPORT HEALTH EFFECT-CLINICAL AND MEDICAL DEVICE PROBLEM CODES: FAILURE OF IMPLANT, LOSS OF VISION AND PRODUCT QUALITY PROBLEM WERE REPORTED. ON (B)(6) 2023 THE PATIENT WAS ABLE TO PROVIDE US WITH ONE ICL SERIAL NUMBER, NAME OF THE IMPLANTING SURGEON, AND DATE OF IMPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM SURGEON, BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16925033&pc=MTA | 2023826-2023-01810 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY MEDICAL DEVICE REPORT FROM FDA- MDR REPORT # MW5116465 ON 13-APR-2023. THE PATIENT REPORTED ON (B)(6) 2023, A "SERIOUS INJURY" EVENT REPORT WITH OUTCOMES OF "DISABILITY OR PERMANENT DAMAGE". THE PATIENT STATED "HAD STAAR MEDICAL'S IMPLANTABLE COLLAMER LENS, EVO ICL IMPLANTED INTO BOTH RIGHT AND LEFT EYE. NOTICED NEXT DAY A LOSS OF PERIPHERAL VISION. RETURNED TO DOCTOR, WAS TOLD TO WAIT A WEEK. RETURNED TO HAVE RETINA CHECKED, RETINA WAS OKAY". PATIENT STATES THAT ISSUE IS STILL PRESENT AND THAT THIS SIDE EFFECT WAS NOT LISTED AS A POSSIBLE SIDE EFFECT. PATIENT IS UNSATISFIED WITH POST-OP OUTCOME. PER THE SAME REPORT HEALTH EFFECT-CLINICAL AND MEDICAL DEVICE PROBLEM CODES: FAILURE OF IMPLANT, LOSS OF VISION AND PRODUCT QUALITY PROBLEM WERE REPORTED. ON (B)(6) 2023 THE PATIENT WAS ABLE TO PROVIDE US THE NAME OF THE IMPLANTING SURGEON AND DATE OF IMPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM SURGEON, BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: LOSS OF PERIPHERAL VISION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222300&pc=MTA | 2023826-2021-02817 | 2021-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THE PATIENT IS UNHAPPY WITH THE VISUAL OUTCOME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B1: UPDATED; B5: UPDATED INFORMATION: ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS. AND THE PROBLEM WAS RESOLVED. PREOP BCVA VALUE 20/20. POSTOP BCVA PRIOR TO EXCHANGE WAS 20/32. AFTER EXCHANGE, REPORTED AS 20/20. INITIAL LENS IMPLANTED ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) (DISSATISFACTION WITH VISUAL OUTCOME). INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457497&pc=MTA | 2023826-2021-03549 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS IRIS TRANSILLUMINATION DEFECTS. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE PATIENT IS SATISFIED WITH THEIR VISION; NO OTHER SYMPTOMS. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795209&pc=MTA | 2023826-2022-00794 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE PATIENT WAS NOT HAPPY WITH HER QUALITY OF VISION AS HER DISTANCE WAS DOWN 2 LINES. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- PER ADDITIONAL INFORMATION RECEIVED ON 09-MAY-2022, THE INITIAL LENS HAS BEEN EXPLANTED AND EXCHANGED. NO ADDITIONAL INFORMATION HAS BEEN GIVEN. D6B: UNK. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT WAS NOT HAPPY WITH HER QUALITY OF VISION AS HER DISTANCE WAS DOWN 2 LINES. THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14221081&pc=MTA | 2023826-2022-01154 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: REPORTS OF LENS ROTATION AND BLOCKING PI IN RIGHT EYE. PLANNED LENS EXCHANGE. SEE MFR. REP# 2023826-2022-01155 FOR FELLOW EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT IDENTIFIER: UNKNOWN, AGE OR DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. DATE OF EVENT: UNKNOWN. MODEL, IMPLANT DATE, EXPLANT DATE: UNKNOWN. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14221127&pc=MTA | 2023826-2022-01155 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: REPORTS OF LENS ROTATION AND BLOCKING PI IN LEFT EYE. PLANNED LENS EXCHANGE. SEE MFR. REP#2023826-2022-01154 FOR FELLOW EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878443&pc=MTA | 2023826-2021-01611 | 2017-06-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.2, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS AND EXCESSIVE VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878372&pc=MTA | 2023826-2021-01612 | 2017-06-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7, -1700 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THERE ARE PLANS TO REPLACE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. MODEL #:UNK THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405440&pc=MTA | 2023826-2021-00554 | 2020-04-23 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10619476&pc=MTA | 2023826-2020-02276 | 2020-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10639546&pc=MTA | 2023826-2020-02304 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOUISTURE, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A DEFORMED HAPTIC. RESIDUE ON LENS SURFACE WAS OBSERVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755792&pc=MTA | 2023826-2020-02465 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICH13.2, +02.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755801&pc=MTA | 2023826-2020-02466 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICH13.2, +04.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980484&pc=MTA | 2023826-2020-02940 | 2020-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -2.75 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980525&pc=MTA | 2023826-2020-02939 | 2020-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -3.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405412&pc=MTA | 2023826-2021-00553 | 2020-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10989738&pc=MTA | 2023826-2020-02967 | 2020-11-11 | Injury | STAAR SURGICAL COMPANY | 2020-12-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. MEDICAL INTERVENTION (DIAMOX) WAS PRESCRIBED DUE TO HIGH VAULT ONE DAY POST OP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT PERIPHERAL ANTERIOR CHAMBER IS SHALLOW AND THAT THE DOCTOR WILL MONITOR CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11161162&pc=MTA | 2023826-2021-00129 | 2020-11-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " NORMAL - SHALLOW AC, AC DEPTH PREOPERATIVE 3.87 AC DEPTH POSTOPERATIVE 2.6". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SHALLOW AC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102590&pc=MTA | 2023826-2021-02135 | 2020-11-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND NARROW ANGLE ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, "BIOLOGIC VARIATION." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11556080&pc=MTA | 2023826-2021-00746 | 2020-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271707&pc=MTA | 2023826-2021-03037 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +05.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11722565&pc=MTA | 2023826-2021-01210 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6- WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. INVESTIGATION FOR CLAIM# (B)(4) FOUND NO INDICATION THAT MANUFACTURING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNCITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938943&pc=MTA | 2023826-2021-01643 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. REPORTEDLY, "SOME CORNEAL EDEMA(ONE DAY POST OP. OK NOW)." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938948&pc=MTA | 2023826-2021-01642 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, NOT ACCURATE MEASUREMENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855169&pc=MTA | 2023826-2021-01505 | 2021-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO LOW VAULT IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, 'AFTER SURGERY PATIENT'S ARCH WAS TOO HIGH AND TOO LOW." AFTER EXPLANTING THE LENS, IT WAS FOUND THAT PATIENT'S SUSPENSORY LIGAMENT WAS NOT EXAMINED AND IT WAS RELAXED, SO LENS WAS EXPLANTED. PATIENT IS UNWILLING TO HAVE ANOTHER OPERATION TO REPLACE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855176&pc=MTA | 2023826-2021-01506 | 2021-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO LOW VAULT IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, 'AFTER SURGERY PATIENT'S ARCH WAS TOO HIGH AND TOO LOW." AFTER EXPLANTING THE LENS, IT WAS FOUND THAT PATIENT'S SUSPENSORY LIGAMENT WAS NOT EXAMINED AND IT WAS RELAXED, SO LENS WAS EXPLANTED. PATIENT IS UNWILLING TO HAVE ANOTHER OPERATION TO REPLACE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE, RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167107&pc=MTA | 2023826-2021-02524 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, TOO LARGE/LONG. Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664247&pc=MTA | 2023826-2022-00662 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- PER ADDITIONAL INFORMATION RECEIVED ON 06-MAY-2022, SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES WAS REPORTED IN ERROR. THERE WAS NO REDUCTION IN IRIDOCORNEAL ANGLES. H6- HEALTH EFFECT CLINICAL CODE: 4581 (SIGNIFICANT REDUCTION OF ICA) NO LONGER APPLIES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443620&pc=MTA | 2023826-2021-03493 | 2021-07-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON SAW A FEMALE PATIENT WITH PLUS POWERED ICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN RIGHT EYE (OD) IN 2004 - 2005. THE LENS WAS REPORTED AS HAVING A LARGE VAULT, WITH SHALLOW ANTERIOR CHAMBER. THE PIS WERE PATENT AND IOP WAS NORMAL. THE ENDOTHELIAL CELL COUNT WAS LOW (1200). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12449010&pc=MTA | 2023826-2021-03494 | 2021-07-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON SAW A FEMALE PATIENT WITH PLUS POWERED ICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN LEFT EYE (OS) IN 2004 - 2005. THE LENS WAS REPORTED AS HAVING A LOW VAULT. THE ENDOTHELIAL CELL COUNT WAS LOW (1500). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE, WEIGHT, ETHNICITY, RACE: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15504589&pc=MTA | 2023826-2022-03343 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS OBSERVED POST ICL IMPLANTATION. THE LENS WAS EXPLANTED ON THE SAME DATE AS IMPLANTATION BUT IN A DIFFERENT SURGERY. ON (B)(6) 2021 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION: BURPING OF THE WOUND WAS PERFORMED TO RELEASE THE PRESSURE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599725&pc=MTA | 2023826-2021-03916 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS LOW VAULT AND PIGMENT DISPERSION. REPORTEDLY, "THE REFRACTIVE RESULT HAS BEEN ACHIEVED. DURING THE NEXT CONTROL OBSERVATION IT WAS NOTED THAT THE LENS HAS DROPPED AND THE VAULT IS A CRITICAL MARK OF 120UM. IMMEDIATELY, A DECISION WAS MADE TO EXPLANT THE LENS AND REPLACE IT WITH A LARGER MODEL. THE PATIENT DID NOT COMPLAIN." THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LOW VAULT AND PIGMENT DISPERSION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599739&pc=MTA | 2023826-2021-03915 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LOW VAULT AND PIGMENT DISPERSION. REPORTEDLY, "THE REFRACTIVE RESULT HAS BEEN ACHIEVED. DURING THE NEXT CONTROL OBSERVATION IT WAS NOTED THAT THE LENS HAS DROPPED AND THE VAULT IS A CRITICAL MARK OF 105UM. IMMEDIATELY, A DECISION WAS MADE TO EXPLANT THE LENS AND REPLACE IT WITH A LARGER MODEL. THE PATIENT DID NOT COMPLAIN." THE CAUSE IS REPORTED AS UNKNOWN-"PRE-OPERATE DATA, CALCULATION AND ALL MEASUREMENTS HAVE DONE AS USUAL IN CORRESPONDENCE WITH STANDARDS. IMPLANTATION WAS ALSO DONE WITH STANDARDS. WE CAN'T NOTICE ANY INCORRECT REASON ABOUT THAT." Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS LOW VAULT AND PIGMENT DISPERSION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716124&pc=MTA | 2023826-2021-04134 | 2021-09-14 | Injury | STAAR SURGICAL COMPANY | 2021-10-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619778&pc=MTA | 2023826-2021-03984 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13442385&pc=MTA | 2023826-2022-00315 | 2021-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF CORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. "DURING THE OPERATION, THE DOCTOR USED THE OPERATING MICROSCOPE TO SEE THAT THERE WAS A CRACK BESIDE THE CENTER HOLE OF THE LENS. DR.(B)(6) USED THE SPARE LENS FOR THE LEFT EYE." CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13006620&pc=MTA | 2023826-2021-04885 | 2021-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -2.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY IT WAS REMOVED DUE TO EXCESSIVE VAULTING. ON THE SAME DATE A REPLACEMENT LENS OF SHORTER LENGTHS WAS IMPLANTED. POSTOP RESULTS NOTED "MINOR VOLCANO EFFECT BUT CHAMBER IS DEEP". CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "TOO LONG FOR EYE".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -2.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY IT WAS REMOVED DUE TO EXCESSIVE VAULTING. ON THE SAME DATE A REPLACEMENT LENS OF SHORTER LENGTHS WAS IMPLANTED. POSTOP RESULTS NOTED "MINOR VOLCANO EFFECT BUT CHAMBER IS DEEP". CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "TOO LONG FOR EYE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030248&pc=MTA | 2023826-2022-01090 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (THE PATIENT DID NOT HAVE THE EXPECTED RESULT). THE LENS WAS REPLACED WITH A SMALLER SIZE LENS (12.1) WITH THE SAME DIOPTER. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND MYDRIASIS. A REPLACEMENT LENS OF A SHORTER LENGTH WAS IMPLANTED ON THE SAME DAY SEPARATE SURGERY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6- HEALTH IMPACT- CLINICAL CODE: 4581- MYDRIASIS. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ANTERIOR SYNECHIAE. THE LENS WAS REPLACED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANTERIOR SYNECHIAE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082670&pc=MTA | 2023826-2021-04922 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS WAS REPORTED AS EXCESSIVE VAULTING, WITH NARROW ANGLES AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (NARROW ANGLES). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO A SIZING PROBLEM. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639226&pc=MTA | 2023826-2022-01791 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND PERMANENT LOSS OF BCVA. THE LENS WAS REPLACED WIH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14647871&pc=MTA | 2023826-2022-01792 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND PERMANENT LOSS OF BCVA. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13553238&pc=MTA | 2023826-2022-00539 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -2.75 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13553240&pc=MTA | 2023826-2022-00540 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -2.25 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE FOOT PLATE. - THIS INFORMATION SHOULD HAVE BEEN INCLUDED IN PREVIOUS SUPPLEMENTAL MFR FOLLOW UP #1. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726657&pc=MTA | 2023826-2022-00757 | 2022-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND VARIATION IN ELEVATED INTRAOCULAR PRESSURE (IOP) (NO ELEVATED IOP PEAKS MEASURABLE). THE LENS WAS REPLACED WITH A SHORTER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WERE LITTLE FIBERS THOUGHOUT THE LENS. CLAIM #(B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE IMPLANT OF THE SHORTER LENS RESOLVED THE PROBLEM. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13796099&pc=MTA | 2023826-2022-00785 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. ON (B)(6)2022 THE LENS WAS REPOSITIONED TO VERTICAL ORIENTATION AND THE PROBLEM WAS RESOLVED FOR VAULT. "(ECL) NO CHANGE FROM PRE-ROTATION AND IS BEING MONITORED CLOSELY. CASE REPORTED AS UNKNOWN. "COULD HAVE BEEN PERIPHERAL CORNEAL ENDOTHELIAL TOUCH CAUSING DAMAGE TO CELLS." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14287983&pc=MTA | 2023826-2022-01305 | 2022-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER STATED "THE PATIENT CANCELLED ALL APPOINTMENTS. HE DO NOT WANT AN EXCHANGE ANYMORE. THE PATIENT IS DIFFICULT TO READ." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK, THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14292696&pc=MTA | 2023826-2022-01306 | 2022-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: HE REPORTER STATED "THE PATIENT CANCELLED ALL APPOINTMENTS. HE DO NOT WANT AN EXCHANGE ANYMORE. THE PATIENT IS DIFFICULT TO READ." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598572&pc=MTA | 2023826-2022-01740 | 2022-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS OF +7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA WERE REPORTED. THE LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS OF +7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE IS UNKNOWN. (B)(4) Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT CLINICAL CODE: (B)(4) (SIGNIFICANT REDUCTION OF ICA). (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985565&pc=MTA | 2023826-2022-02143 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010000&pc=MTA | 2023826-2022-02231 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. "OVERSIZE LENS" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010097&pc=MTA | 2023826-2022-02232 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. " OVERSIZE LENS" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15135244&pc=MTA | 2023826-2022-02454 | 2022-06-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND NARROW ANGLES WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- " THE PATIENT ALSO EXPERIENCED SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND SHALLOW PERIPHERAL ACD. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SHALLOW PERIPHERAL ACD. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15583834&pc=MTA | 2023826-2022-03523 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND A SHORTER LENS WAS IMPLANTED DUE TO SHALLOW AC AND EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE AND PATIENT RELATED FACTOR. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584122&pc=MTA | 2023826-2022-03522 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND A SHORTER LENS WAS IMPLANTED DUE TO SHALLOW AC AND EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE AND PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: G2: OTHER: SAUDI ARABIA SHOULD BE REMOVED AND REPLACED WITH EGYPT FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620686&pc=MTA | 2023826-2023-01029 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULT. A LONGER LENGTH LENS WAS IMPLANTED ON 05/OCT/2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16877955&pc=MTA | 2023826-2023-01617 | 2022-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16877979&pc=MTA | 2023826-2023-01616 | 2022-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830802&pc=MTA | 2023826-2022-04016 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. Manufacturer Narrative: CLINICAL CODE (B)(4) - SIGNIFICAT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759256&pc=MTA | 2023826-2022-03827 | 2022-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. Manufacturer Narrative: RACE- UNK. ETHNICITY - UNK. WEIGHT - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035511&pc=MTA | 2023826-2022-04577 | 2022-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A LONGER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330310&pc=MTA | 2023826-2023-00473 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO A PREVIOUS ISSUE. THE SURGEON NOTED "PRO2H; VOMIT; THERE IS EXUDATIVE MEMBRANE IN THE POSTERIOR CHAMBER". ON THIS SAME DATE THE LENS WAS EXPLANTED IN A SECOND SURGERY AND IT WAS NOTED "PATIENT FEELSL GOOD AFTER SURGERY". THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - POSTOPERATIVE 2H, VOMIT, EXUDATIVE MEMBRANE. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIUDE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15991899&pc=MTA | 2023826-2022-04413 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT. CAUSE REPORTED AS UNKNOWN. REPORTER STATED " PATIENT GOT COVID, SO THE SECONDARY IMPLANT IS POSTPONED." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED; HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: D6A: CORRECTED TO: IMPLANT DATE: ON (B)(6) 2022 IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: NEW INFORMATION- B5 - THE REPORTER LATER REPORTED THAT ON (B)(6) 2023 A REPLACEMENT LENS OF A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210216&pc=MTA | 2023826-2023-00206 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THIS EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS. THE REPORTER ALSO STATES THAT THE DEVICE FAIL TO PERFORM AS INTENDED. FOR STATUS OF THE EYE AND ADDITIONAL INFORMATION THE REPORTER STATES " NORMAL CORNEA AND PUPIL, MILD ANTERIOR SUBCAPSULAR LENS OPACITIES. PATIENT WILL BE FOLLOWED TO SEE IF MILD LENS CHANGES RESOLVE AND BCVA RETURNS TO BASELINE. PATIENT DOES NOT WANT REPEAT EVO ICL, BACK IN SOFT CONTACT LENSES. Manufacturer Narrative: HEALTH EFFECT- (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301127&pc=MTA | 2023826-2023-00434 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OR IRIDOCORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17061378&pc=MTA | 2023826-2023-02030 | 2023-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEL ANGLES AND EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4, A5, A6: UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16562713&pc=MTA | 2023826-2023-00883 | 2023-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023.THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "WITHOUT SIGNS AND SYMPTOMS". ADDITIONAL INFORMATION PROVIDED - "THE EXPLANTED LENS HAS A SMALL NOTCH PRODUCED DURING REMOVAL". CAUSE DETAILS- ""THE ESTIMATION OF THE LENGTH LENS BY THE CALCULATOR WAS 13.2MM. JUST BY WTW OF 0.1MM WOULD HAVE GIVEN A SMALLER LENGTH (12.6MM) AND MORE CORRECT FOR THIS CASE". H6- HEALTH EFFECT- CLINICAL CODE: "4581- PIGMENT DISPERSION" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713974&pc=MTA | 2023826-2023-01216 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620714&pc=MTA | 2023826-2023-01025 | 2023-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770485&pc=MTA | 2023826-2023-04102 | 2023-03-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -9.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS LOW VAULTING AND THERE WAS DECREASED NEAR VISION IN BRIGHT LIGHT. LENS REMAINS IMPLANTED WITH A PLANNED LENS EXCHANGED. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER AND NOTED "BRIGHT LIGHT CONSTRICTS THE PUPIL, PRESSING THE ICL TOWARDS THE CRYSTALLINE LENS INHIBITING THE INCREASE IN ANTERIOR CURVATURE WHICH IS NEEDED FOR NEAR VISION. THE ICL FLATTENS THE ANTERIOR LENS CURVATURE CAUSING PRESBYOPIA ONLY OCCURRING IN BRIGHT LIGHT". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DECREASED VISION IN BRIGHT LIGHTING. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341176&pc=MTA | 2023826-2023-05644 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULTING WAS NOTED AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES DESCRIBED AS "ANGLE 19.7°" WAS REPORTED. IT WAS ALSO MARKED THAT LENS REPOSITIONING DID NOT RESOLVE THE ISSUE; HOWEVER, NO REPOSITIONING DATA WAS PROVIDED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18080759&pc=MTA | 2023826-2023-04993 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17081937&pc=MTA | 2023826-2023-02273 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER +9.5 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTED DID NOT GIVE CAUSE FOR THIS EVENT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17167038&pc=MTA | 2023826-2023-02367 | 2023-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE CORNEAL ENDOTHELIUM WAS EDEMATOUS AND THE CENTRAL ENDOTHELIUM WAS TURBID. THE AQUEOUS HUMOR WAS SPARKLING (++). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL ENDOTHELIUM EDEMATOUS, CENTRAL ENDOTHELIUM TURBID, AQUEOUS HUMOR SPARKLING. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17321985&pc=MTA | 2023826-2023-02898 | 2023-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 03-JUN-2020. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SHALLOW ANGLE TEMPORAL. THE SURGEON NOTES THERE ARE NO SYMPTOMS AND PRESSURE IS NORMAL, SHALLOW TEMPORAL COMPARED TO NASAL. THE PATIENT IS HAPPY BUT VAULT IS HIGH AND TEMPORAL COMPARED TO NASAL IN THE LEFT EYE. LENS REMAINS IMPLANTED. THE SURGEON INDICATED THERE MAY BE FUTURE SURGICAL INTERVENTION. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW ANGLE TEMPORAL. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17654640&pc=MTA | 2023826-2023-03440 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -14.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES.. THE LENS REMAINS IMPLANTED. REPORTER STATES. "PATIENT SCHEDULED FOR REPLACEMENT OF ICL." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - LENS WAS IMPLANTED ON (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS EXCHANGED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- CLINICLAL CODE 4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17932866&pc=MTA | 2023826-2023-04401 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414575&pc=MTA | 2023826-2023-05821 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT REPORTED AS OTHER. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS . CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414605&pc=MTA | 2023826-2023-05820 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT REPORTED AS OTHER. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166351&pc=MTA | 2023826-2023-05180 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.0 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6- CLICICAL CODE4581- SIGNIFICAT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414574&pc=MTA | 2023826-2023-05824 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IN A SEPARATE SURGERY THAT DAY A LENS TWO SIZES SHORTER IN LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472104&pc=MTA | 2023826-2024-00113 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM, THE REPORTER STATED " VAULT REDUCED TO NORMAL AND PATIENT IS DOING GOOD". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11010454&pc=MTA | 2023826-2020-03078 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2,-5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119581&pc=MTA | 2023826-2021-00005 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND NARROW ANTERIOR CHAMBER ANGLES BUT WITH NORMAL INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK EXPIRATION DATE: UNK. IMPLANT DATE: UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119599&pc=MTA | 2023826-2021-00006 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING HIGH VAULTING AND NARROW ANTERIOR CHAMBER ANGLES BUT WITH NORMAL INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK EXPIRATION DATE: UNK. IMPLANT DATE: UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119242&pc=MTA | 2023826-2021-02325 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS EXCESSIVE VAULT AND SHALLOW AC. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH EFFECT CLINICAL CODE: (B)(4) (SHALLOW AC). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082714&pc=MTA | 2023826-2021-04988 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED FOR "URRETS-ZAVALIA SYNDROME AFTER IMPLANTATION OF A PHAKIC INTRAOCULAR LENS". THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS OVERSIZED AND WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS LENGTH. A FEW WEEKS LATER, THE PATIENT HAD AN ATONIC, MID-DILATED PUPIL, WITH NO DIRECT OR CONSENSUAL RESPONSE TO LIGHT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: AN ARTICLE WAS RECEIVED FOR "URRETS-ZAVALIA SYNDROME AFTER IMPLANTATION OF A PHAKIC INTRAOCULAR LENS". THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS OVERSIZED AND WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS LENGTH. A FEW WEEKS LATER, THE PATIENT HAD AN ATONIC, MID-DILATED PUPIL, WITH NO DIRECT OR CONSENSUAL RESPONSE TO LIGHT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. D4: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911479&pc=MTA | 2023826-2022-00971 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS SHALLOW VAULT AND RESIDUAL AMETROPIA. REPORTEDLY, THE LENS REMAINS IMPLANTED AND WILL BE MONITORED, "SINCE THERE ARE SOME MILD ANTERIOR CAPSULE CHANGES AND I DO NOT WANT TO GO BACK IN." Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911490&pc=MTA | 2023826-2022-00970 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LOW VAULT AND RESIDUAL AMETROPIA. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14079181&pc=MTA | 2023826-2022-01109 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. 1300 MICRON VAULT WITH SOME DEGREE OF IRIS CHAFING WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14524935&pc=MTA | 2023826-2022-01697 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525054&pc=MTA | 2023826-2022-01698 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15048522&pc=MTA | 2023826-2022-02304 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED HIGH VAULT AND NARROW ANGLES, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENT IS DOING WELL, IOP 17/18 EXCELLENT VISION 20/20. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352616&pc=MTA | 2023826-2022-03014 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOW REPORTS THE LENS IS OVERSIZED AND CAUSING NARROW ANGLES AND LARGE VAULTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352617&pc=MTA | 2023826-2022-03013 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOW REPORTS THE LENS IS OVERSIZED AND CAUSING NARROW ANGLES AND LARGE VAULTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588726&pc=MTA | 2023826-2022-03598 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -5.0DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591717&pc=MTA | 2023826-2022-03485 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM ICL IMPLANTABLE COLLAMER LENS AND THE LENS WAS REPORTED AS HIGH VAULTING, WITH NARROW ANGLES. THE LENS REMAINED IMLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17061119&pc=MTA | 2023826-2023-02040 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPOSITIONED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES AND EXCESSIVE VAULT AND DID NOT RESOLVE THE PROBLEM. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297497&pc=MTA | 2023826-2023-02782 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2023. HIGHER VAULT OUTCOME AND PIGMENTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1- UNK, A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, D6B - UNK, H4 - UNK, H6 - HEALTH EFFECT CLINICAL CODE 4581: IRIS PIGMENTATION. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297511&pc=MTA | 2023826-2023-02781 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2023. HIGHER VAULT OUTCOME AND PIGMENTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: IRIS PIGMENTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619768&pc=MTA | 2023826-2023-03600 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619877&pc=MTA | 2023826-2023-03601 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741311&pc=MTA | 2023826-2023-04039 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THE PATIENT HAS LOW VAULTING WITH CLOSED ANGLES. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750760&pc=MTA | 2023826-2023-04068 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THAT THE PATIENT HAS LOW VAULTING WITH CLOSED ANGLES. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17865574&pc=MTA | 2023826-2023-04411 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE LENS BEING TOO LARGE AND THE ANTERIOR CAMBER IS TOO SHALLOW. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANTERIOR CAMBER TOO SHALLOW. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17865575&pc=MTA | 2023826-2023-04412 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE LENS BEING TOO LARGE AND THE ANTERIOR CHAMBER IS TOO SHALLOW. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANTERIOR CAMBER TOO SHALLOW. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17978644&pc=MTA | 2023826-2023-04565 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067949&pc=MTA | 2023826-2023-04758 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPORTED AS LOW VAULTING AND PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234786&pc=MTA | 2023826-2023-05389 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS HIGH VAULT AND NARROW ANGLES. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234792&pc=MTA | 2023826-2023-05390 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS HIGH VAULT AND NARROW ANGLES. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360961&pc=MTA | 2023826-2023-05555 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING AND PATIENT EXPERIENCED A MILD DISCORIC PUPIL (MILD DILATED 5MM AND SLOWLY REACTIVE TO LIGHT). THE INTRAOCULAR PRESSURE HAS BEEN MEASURED WITHIN NORMAL LIMITS. THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DISCORIC (DILATED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472279&pc=MTA | 2023826-2024-00116 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT WITH VERY SHALLOW ANGLES. A REPLACEMENT LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SHALLOW ANGLES. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10905481&pc=MTA | 2023826-2020-02730 | 2020-11-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: "EXCESSIVE VAULT AND NARROWING OF THE ANTERIOR CHAMBER ANGLES WERE ALSO REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. MEDICAL DEVICE PROBLEM CODE 2993 CORRECTED TO 1583 IN INITIAL MDR. HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, NARROWING OF ANTERIOR CHAMBER ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA H6 - HEALTH EFFECT - IMPACT CODE: 4629 CORRECTED TO 4627. B5 - CORRECTED TO "THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. LENS REMAINED IMPLANTED. HOWEVER, ANTERIOR CHAMBER NARROWING WAS REPORTED AGAIN ON (B)(6) 2021. THE LENS WAS EXPLANTED AND ANOTHER LENS WITH A SHORTER LENGTH WAS IMPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. ADDITIONAL MANUFACTURE NARRATIVE: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "HOWEVER, ANTERIOR CHAMBER NARROWING WAS REPORTED AGAIN ON 01/APR/2021. THE LENS WAS EXCHANGED FOR ONE OF A SHORTER LENGTH ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "CAUSE OF THE EVENT WAS DUE TO THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15142507&pc=MTA | 2023826-2022-02572 | 2022-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED CORNEAL EROSION (WHICH HEALED EXPECTEDLY AS THE SURGEON REPORTS); PATIENT COMPLAINS OF MULTIPLE IMAGES/REFLECTIONS/SHADOWS (MONOCULAR DIPLOPIA); LENS ROTATION AND LOW VAULTING. LENS REMAINS IMPLANTED IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573131&pc=MTA | 2023826-2021-00766 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED OVER VAULT, NARROW ANGLE AND LENS ROTATION. IT WAS REPORTED THAT THE LENS WAS EXCHANGED AND NEW VAULT IS GOOD, PATIENT IS SEEING WELL AFTER EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808009&pc=MTA | 2023826-2021-04322 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAUL, LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND SIGNIFICANT REDUCTION OF ICA WAS OBSERVED. ON (B)(6) 2021 IT WAS REPORTED THAT THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. "OTHER" WAS REPORTED TO THE CAUSE OF THE EVENT, FOR CAUSE DETAILS THE REPORTER STATED "INADEQUATE SIZE OF LENS." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND OVERALL LENGTH TO BE WITHIN SPECIFICATION. WIDTH VERIFICATION COULD NOT BE PERFORMED DUE TO HAPTIC DAMAGE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13896064&pc=MTA | 2023826-2022-00890 |  | Injury | STAAR SURGICAL COMPANY AG | 2022-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2011. THE PLAN IS TO EXPLANT THE LENS DUE TO PROGRESSIVE MYOPIA. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022, DUE TO HAVING A SLIGHTLY HIGH VAULT AND WAS REPLACED FOR A SHORTER LENS. AT 1 DAY POST-OP, THE UNCORRECTED VA WAS 20/50. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577515&pc=MTA | 2023826-2020-02193 | 2018-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE/ DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697138&pc=MTA | 2023826-2021-01094 | 2020-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND IN A SECOND SURGERY A SHORTER LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT/SIGNIFICANT ENDOTHELIAL CELL LOSS, AND THE CORNER OF THE EYE BEING PARTIALLY CLOSED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT PATIENT RELATED FACTOR WAS REPORTED IN ERROR AND CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459218&pc=MTA | 2023826-2022-00379 | 2022-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE SURGEON REPORTS VAULTING AND CORNEAL TOUCH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: SEX, WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON ALSO REPORTS EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. THE SURGEON REPORTS MILD PIGMENT DISPERSION WITHOUT SYMPTOMS AFTER EXCHANGE THAT HAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR; THE LENS DID NOT FAIL TO PERFORM AS INTENDED: "POOR ESTIMATE OF CILIARY SULCUS DIAMETER RELATIVE TO HORIZONTAL WTW." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459239&pc=MTA | 2023826-2022-00378 | 2022-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE SURGEON REPORTS VAULTING AND CORNEAL TOUCH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: SEX, WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: EXCESSIVE VAULTING IS REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS AND THE PROBLEM IS RESOLVED. MILD PIGMENT DISPERSION WITHOUT SYMPTOMS AFTER EXCHANGE IS REPORTED, WHICH HAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR; THE LENS DID NOT FAIL TO PERFORM AS INTENDED: "POOR ESTIMATE OF CILIARY SULCUS DIAMETER RELATIVE TO HORIZONTAL WTW." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16413989&pc=MTA | 2023826-2023-00605 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. EYE 'ACHE' DUE TO CILIARY BODY COMPRESSION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLINICAL CODE: 4581-SIGNIFICANT REDUCTION OF ICA. EYE 'ACHE' (ASSUME DUE TO CILIARY BODY COMPRESSION). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14178237&pc=MTA | 2023826-2022-01223 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS OF -16.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(4) 2022 THE LENS WAS REMOVED DUE TO LOW VAULT. UPON REMOVAL OF THE INITIAL LENS IT WAS DETERMINED TO FLIPPED INSIDE THE EYE. REPORTER STATES THERE WAS DIFFICULTY TUCKING HAPTICS UNDER THE IRIS AS THE MANIPULATOR CONSTANTLY STUCK TO THE ICL. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF SAME MODEL/SIZE AND DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE REPORTED AS MODERATE ANTERIOR CELL. MILD CORNEAL EDEMA. PATENT PI. PER REPORTER "IT IS TYPICAL FOR ANY POST OPERATIVE EYE TO HAVE SOME CELL AND CORNEAL EDEMA, ESPECIALLY AFTER BEING OPERATED ON A SECOND TIME. THE PATENT PI IS ALSO NORMAL SINCE WE DID NOT HAVE AN EVO LENS SO WE HAD TO DO THE PI." Manufacturer Narrative: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478820&pc=MTA | 2023826-2021-00634 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND MONOCULAR DIPLOPIA (DOUBLE VISION). THE LENS WAS EXCHANGED FOR A TORIC ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN, THERE WAS SOME RESIDUAL REFRACTIVE ERROR, BUT EVEN WITH GLASSES THE PATIENT HAD MONOCULAR DIPLOPIA, VERY SMALL PERIPHERAL PI UNLIKELY CAUSE. THE EYE IS NORMAL AND ACUITY AND SYMPTOMS IMPROVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416102&pc=MTA | 2023826-2021-00508 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -03.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PROBLEMS WITH DISTANCE UCVA POST 1 MONTH WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, UCVA: 0.5, BCVA: 1.0 RX: 1.25/0.00/000". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -03.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. PROBLEMS WITH DISTANCE UCVA POST 1 MONTH WAS OBSERVED. THE LENS WAS EXCHANGED WITH A SAME LENGTH ,LESSER POWER LENS. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, UCVA: 0.5, BCVA: 1.0 RX: 1.25/0.00/000". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416142&pc=MTA | 2023826-2021-00509 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM6 13.7, -02.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. PROBLEMS WITH DISTANCE UCVA POST 1 MONTH WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, UCVA: 0.5, BCVA: 1.0 RX: 1.25/0.00/000". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.7MM, VICM6 13.7, -02.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. PROBLEMS WITH DISTANCE UCVA POST 1 MONTH WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, UCVA: 0.5, BCVA: 1.0 RX: 1.25/0.00/000" CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850227&pc=MTA | 2023826-2022-04268 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. "AFTER IMPLANTATION OF VICMO LENS, THE PATIENT COMPLAINED OF POOR VISION AND DOUBLE IMAGE, WHICH AFFECTED HIS LIFE. THE DOCTOR DECIDED TO CHANGE TO VTICMO LENS." ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC MODEL ICL AND SAME LENGTH. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820442&pc=MTA | 2023826-2023-01504 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS REPLACED WITH A DIFFERENT MODEL (TORIC) BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301926&pc=MTA | 2023826-2023-00388 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. REPORTEDLY THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955610&pc=MTA | 2023826-2022-04522 | 2022-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT THE PATIENT HAS "WRONG REFRACTION" (OVER CORRECTION). LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED THE DEVICE FAILED TO PERFORM AS INTENDED; "WRONG INITIAL REFRACTION". Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION PROVIDED: " THE PATIENT HAS NOT BEEN AVAILABLE SINCE DECEMBER 2022 AND HAS NOT SPOKEN. I WANT TO CLOSED THE CASE". CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17785989&pc=MTA | 2023826-2023-04118 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A CASE REPORT WAS PUBLISHED IN FRONTIERS IN MEDICINE ON (B)(6) 2023. TITLED " CASE REPORT: SPONTANEOUS LENS ABSORPTION AFTER IMPLANTATION OF AN IMPLANTABLE COLLAMER LENS." AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2016. A 23-YEAR OLD PATIENT COMPLAINED OF GRADUAL LOSS OF VISION IN THE LEFT EYE (OS). OBSERVATION ON MARCH,2018 SHOWED ABSORPTION OF THE TRANSPARENT LENS, IRIS ATROPHY AND LOSS OF ENDOTHELIAL CELL COUNT. THE ICL WAS EXPLANTED ON AN UNKNOWN DATE AND A IMPLANTATION OF AN INTRAOCULAR (JOHNSON & JOHNSON) LENS WAS IMPLANTED. Manufacturer Narrative: DATE OF EVENT- (B)(6) 2018 CLINICAL CODE: 4581- LENS ABSORPTION; IRIS ATROPHY; ENDOTHELIAL CELL LOSS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13015001&pc=MTA | 2023826-2021-04903 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. ON (B)(6) 2021. THE PATIENT EXPERIENCED EXTREME MYOPIA AND MIOSTAT WAS USED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1401. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13015005&pc=MTA | 2023826-2021-04902 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. ON (B)(6) 2021. THE PATIENT EXPERIENCED EXTREME MYOPIA AND MIOSTAT WAS USED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. ON (B)(6) 2021. THE PATIENT EXPERIENCED EXTREME MYOPIA AND MIOSTAT WAS USED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1401. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11091320&pc=MTA | 2023826-2020-03223 | 2020-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING (THE LENS VAULTED FORWARD), AND ANGLE CLOSURE. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S EYE IS QUIET AND THE PATIENT IS HAPPY WITH THE OUTCOME. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED ON SAME DAY DIFFERENT SURGERY. THIS WAS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE AND HYPEROPIC OUTCOME. THE REPLACEMENT LENS RESOLVED THE PROBLEMS. H6: HEALTH EFFECT- CLINICAL CODE: 4581- HYPEROPIC REFRACTIVE OUTCOME. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14458457&pc=MTA | 2023826-2022-01625 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN SIZE AND MODEL WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). PATIENT EXPERIENCED A LARGE VAULT AND WAS QUITE HYPEROPIC. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14458485&pc=MTA | 2023826-2022-01626 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN SIZE AND MODEL WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). PATIENT EXPERIENCED A LARGE VAULT AND WAS QUITE HYPEROPIC. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17878138&pc=MTA | 2023826-2023-04427 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "EFFECT OF LONG-TERM ROTATION ON ASTIGMATISM FOLLOWING EVO-TORIC INTRAOCULAR COLLAMER LENS IMPLANTATION" INDICATED THAT OUT OF 40 EYES OF 22 PATIENTS IT NOTES THREE YEARS POSTOPERATIVELY THREE EYES HAD LENS ROTATION AND SLIGHTLY SMALLER ICL SIZE MIGHT HAVE CAUSED THE SIGNIFICANT ROTATION. IT ADDITIONALLY NOTES THE ASTIGMATISM OF THREE EYES HAD EXPERIENCED UDVA LOSS BUT THEY HAD NO IMPACT ON THE PATIENT'S DAILY LIVES. LENS REMAINS IMPLANTED. THE PATIENT CHOSE TO UNDERGO CLOSE MONITORING INSTEAD OF A SECOND OPERATION. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - UDVA LOSS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558373&pc=MTA | 2023826-2023-03544 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE AMERICAN JOURNAL OF CASE REPORTS TITLED, 'PRESUMED INDICATED IMPLANTATION OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS'. THE ARTICLE REPORTS " A 23-YEAR-OLD WOMAN WAS REFERRED FOR POOR VISION FOLLOWING AN IMPLANTATION OF BILATERAL TORIC IMPLANTABLE COLLAMER LENS (ICL) IMPLANTATION WHEN SHE WAS 18 YEARS OLD FOR HIGH MYOPIC ASTIGMATISM AND ANISOMETROPIA. ON PRESENTATION, THE BEST CORRECTED VISUAL ACUITY WAS 4-6/200 AND 2-3/200 IN THE RIGHT EYE AND LEFT EYE, RESPECTIVELY. SLIT LAMP EXAMINATION INDICATED A CLEAR CORNEA WITH PIGMENT DEPOSITION ON THE ENDOTHELIUM, HIGH ICL VAULT, SHALLOW ANTERIOR CHAMBER, AND BOWING OF THE IRIS BILATERALLY. THE PATIENT UNDERWENT BILATERAL REMOVAL OF THE ICLS ON SEPARATE VISITS, YET THE VISION REMAINED UNCHANGED. DIAGNOSTICS REVEALED THAT THE PATIENT HAD BULL'S EYE MACULOPATHY WITH ATROPHY DUE TO HER CONE-ROD DYSTROPHY, WHICH CAUSED THE POOR VISION". THE REPORT EMPHASIZES THE NEED FOR CAREFUL AND APPROPRIATE PATIENT AND INTRAOCULAR SIZE SELECTION FOR REFRACTIVE SURGERY. IT INDICATES THE IMPORTANCE OF THOROUGH MEDICAL EXAMINATION, INCLUDING GENETIC TESTING, FUNDUS EXAMINATION, AND OPTICAL COHERENCE TOMOGRAPHY WHEN RETINAL DYSTROPHY IS SUSPECTED. GENERALLY, CLOSE FOLLOW-UP IS CRUCIAL IN THE SETTING OF HIGH VAULTING FOLLOWING ICL IMPLANTATION TO PREVENT SECONDARY COMPLICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: PIGMENT DEPOSITION ON THE ENDOTHELIUM, SHALLOW ANTERIOR CHAMBER, AND BOWING OF THE IRIS BILATERALLY. H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442548&pc=MTA | 2023826-2022-03180 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -15.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT WAS UNHAPPY ABOUT HER PERIPHERAL VISION AND WAS NAUSEOUS AND VOMITING FOR A MONTH. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "POOR PERIPHERAL VISION - OBJECTIVELY". Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442549&pc=MTA | 2023826-2022-03181 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -15.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT LEFT (OS) ON (B)(6) 2022. THE PATIENT WAS UNHAPPY ABOUT HER PERIPHERAL VISION AND WAS NAUSEOUS AND VOMITING FOR A MONTH. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "POOR PERIPHERAL VISION". Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784668&pc=MTA | 2023826-2022-04079 | 2018-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPORTED AS REFRACTIVE CHANGE OVERTIME AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022, WITH A VICM5_12.6 LENS, 07.50 DIOPTER. THE PROBLEM WAS RESOLVED. THE INCISION IS HEALING WELL, ICL IS POSITIONED WELL WITH GOOD VAULT. VA IS 20/15. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443631&pc=MTA | 2023826-2021-03484 | 2021-01-01 | Malfunction | STAAR SURGICAL COMPANY AG | 2021-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. A REFRACTIVE CHANGE OVERTIME WAS NOTED AND THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVERNT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308430&pc=MTA | 2023826-2021-03076 | 2021-04-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, ANGLES NARROW, WILL CONTINUE TO OBSERVE. PER THE REPORTER ON (B)(6) 2021, NO TREATMENT OR INTERVENTION HAS BEEN DECIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE (B)(4): SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182561&pc=MTA | 2023826-2022-00034 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS SPONTANEOUS MYDRIASIS SOME DAYS AFTER SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS SPONTANEOUS MYDRIASIS SOME DAYS AFTER SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 -13.0D INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. SPONTANEOUS MYDRIASIS WAS OBSERVED SEVERAL DAYS AFTER SURGERY. LENS REMAINS IMPLANTED. REPORTEDLY, "OVERALL PATIENT IS NORMAL, GOOD VISUAL ACUITY, IOP AND LENS WELL POSITIONED." CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY MODEL #: IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY MODEL #: IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14059267&pc=MTA | 2023826-2022-01027 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -9.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO A PREVIOUS REFRACTIVE SURPRISE. IT WAS INDICATED THIS DID NOT RESOLVE THE PROBLEM, AND THE SURGEON REPORTS "VAULT CORRECT. HYPEROPIA REMAINS". LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-01026 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON 10-MAY-2022 THE LENS WAS EXPLANTED AND NOTED "AFTER RE-GRADING THE PATIENT AND SEEING THAT THERE WERE NO ERRORS, THE DOCTORS DECIDED TO IMPLANT THE NEW LENS WITH THE SAME PRESCRIPTION, BUT THERE WAS THE SAME REFRACTIVE SURPRISE". CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THICK CLEAR RESIDUE ALL OVER THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15294323&pc=MTA | 2023826-2022-02807 | 2022-02-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 OF -9.0 DIOPTER IMPLANTABLE COLLAMER LENS HAD LINEAR STAINS ON THE LENS. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE THROUGHOUT THE LENS. INDENTATION IN FOOTPLATE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15294350&pc=MTA | 2023826-2022-02806 | 2022-02-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 OF -10.50 DIOPTER IMPLANTABLE COLLAMER LENS HAD LINEAR STAINS ON THE LENS. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND DIOPTER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14750793&pc=MTA | 2023826-2022-01925 | 2022-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. FOREIGN MATERIAL CONTAMINATION (PLASTIC PIECES) WERE OBSERVED. THE PLASTIC PIECES WERE REMOVED FROM INSIDE THE EYE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE EVALUATION: A MICROCENTRIFUGE VIAL WAS RETURNED WITH NO LENS. VISUAL INSPECTION FOUND CLEAR PARTICULATE IN VIAL. DUE TO PARTICULATE BEING NOTICEABLE FLOATING IN LIQUID, VIAL WAS NOT OPENED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT ,ETHNICITY, RACE: UNK. CLINICAL CODE: (B)(4): PLASTIC PIECES WERE REMOVED FROM INSIDE THE EYE. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CODE (B)(4): DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584542&pc=MTA | 2023826-2022-03619 | 2022-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THERE WAS FOREIGN MATTER ON THE ICL PERI-OPTICAL HOLE OF THE LENS AND FILAMENTOUS FOREIGN BODY IN THE VIAL. IT WAS INDICATED THE LENS REMAINS IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15781982&pc=MTA | 2023826-2022-04000 | 2022-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022 . DOCTOR FOUND THERE WERE FLOCCULENT FLOATING OBJECTS ON THE SURFACE OF THE LENS DURING THE SURGERY. AT FOLLOW UP CHECK UPS THE FOREIGN BODY WAS STILL THERE. SINCE THE PATIENT IS IN NORMAL CONDITION, THE DOCTOR DECIDED NOT TO REPLACE THE LENS TEMPORALLY. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLINICAL CODE-4581-FLOCCULENT FLOATING OBJECTS ON THE SURFACE OF THE LENS. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669074&pc=MTA | 2023826-2023-03742 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS EXCHANGED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS STUCK IN CARTRIDGE, FOOTPLATE RUPTURE, LENS IMPLANTED AND THEN EXCHANGED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885023&pc=MTA | 2023826-2023-04311 | 2023-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED PIGMENT DISPERSION AND DILATED PUPIL (NON REACTIVE). THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK D6B: UNK H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION, DILATED PUPIL (NON-REACTIVE). H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057149&pc=MTA | 2023826-2021-04908 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE PATIENT REPORTED HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN 2008. THE PATIENT REPORTED HAS NOW BEEN EXPERIENCING VISION LOSS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082674&pc=MTA | 2023826-2021-04923 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS WAS REPORTED AS HAVING NARROW ANGLES AND REMAINED IMPLANTED. THE REPORTER INDICATED THE EYE IS NOW FINE AND NO EXCHANGE IS NEEDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS WAS REPORTED AS HAVING NARROW ANGLES AND REMAINED IMPLANTED. THE REPORTER INDICATED THE EYE IS NOW FINE AND NO EXCHANGE IS NEEDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. SERIAL #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. SERIAL #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13538837&pc=MTA | 2023826-2022-00526 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021; AS A REPLACEMENT LENS TO A PREVIOUS LENS TEAR/BREAK DURING LOADING ISSUE. IT WAS INDICATED THAT THE PATIENT IS UNHAPPY WITH THE POSTOPERATIVE RESULTS AND UCVA IS 20/50. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552666&pc=MTA | 2023826-2022-00428 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN A PATIENT. THE REPORTER INDICATED THE PATIENT IS SEEING A "LINE OF LIGHT" POST PIS FOR ICL SURGERY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN A PATIENT. THE REPORTER INDICATED THE PATIENT IS SEEING A "LINE OF LIGHT" POST PIS FOR ICL SURGERY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. OTHER #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. OTHER #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726557&pc=MTA | 2023826-2022-00759 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED EVO VIVA LENSES INTO 14 EYES. THE SURGEON REPORTED 28.6% OF THE EYES (TOTAL OF 4 EYES) LOST 2 OR MORE LINES OF BEST CORRECTED VISION (BCVA) AT 3 MONTHS POST-OP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED EVO VIVA LENSES INTO 14 EYES. THE SURGEON REPORTED 50% OF THE EYES (TOTAL OF 7 EYES) LOST 2 OR MORE LINES OF BEST CORRECTED VISION (BCVA) AT 3 MONTHS POST-OP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: AGE AT TIME OF EVENT, DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE:UNKNOWN, DATE OF EVENT: UNKNOWN, MODEL: EXPIRATION DATE UNKNOWN, IMPLANT DATE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNKNOWN. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598569&pc=MTA | 2023826-2022-01741 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A VICMO/5 LENS WAS RETURNED DAMAGE. Manufacturer Narrative: PATIENT ID, AGE OR DOB, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANT/EXPLANT DATE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BENT WITH DEBRIS THROUGHOUT THE LENS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14734242&pc=MTA | 2023826-2022-01853 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2012. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED/EXCHANGED DUE TO PROGRESSIVE MYOPIA AND ASTIGMATISM. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15129629&pc=MTA | 2023826-2022-02449 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER REPORTED A 13.2MM VICMO13.2 -15.00 DIOPTER IMPLANTABLE COLLAMER LENS. THE REPORTER STATED "DEFECTIVE OPTIC, USED BACKUP LENS". PATIENT CONTACT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.00 INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED AND REPLACED WITH BACK UP LENS DUE TO DEFECTIVE OPTIC. REPORTEDLY THERE WAS A "DEFECTIVE OPTIC; USED BACK UP LENS". CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15397093&pc=MTA | 2023826-2022-03020 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THEY ORDERED A 12.6MM VICM5 12.6 OF -2.50 DIOPTER AND RECEIVED A 13.2MM VICM5 13.2 OF -2.50 DIOPTER. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS BOX RETURNED UNOPENED. LENS BOX WRAP WAS INTACT, BOX NOT OPENED. LABELS ON LENS BOX AND LABEL MATCH. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO THE USER'S FAILURE TO FOLLOW THE APPROPRIATE STANDARD OPERATING PROCEDURE. THERE IS NO INDICATION THAT SUGGESTS A DEFICIENCY IN THE MANUFACTURING PROCESS OR A TREND OF PROCESS ERRORS. FURTHER INVESTIGATION OR ACTION IS NOT REQUIRED AT THIS TIME. CLAIM # (B)(4). Manufacturer Narrative: CLINICAL CODE (B)(4). INVESTIGATION TYPE (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498809&pc=MTA | 2023826-2022-03319 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT COMPLAINED OF LOSS OF PERIPHERAL VISION AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498900&pc=MTA | 2023826-2022-03318 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE PATIENT COMPLAINED OF LOSS OF PERIPHERAL VISION AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567539&pc=MTA | 2023826-2022-03583 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS INDICATED THAT THE LENS WAS SITTING SLIGHTLY LOWER THAN ANTICIPATED. THE LENS REMAINS IMPLANTED AND IT WAS NOTED THEY ARE NOT EXCHANGING THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309074&pc=MTA | 2023826-2023-00444 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE WAS A HYPEROPIC OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16417451&pc=MTA | 2023826-2023-00610 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022 AS AN EXCHANGE LENS TO RESOLVE EXCESSIVE VAULT. ALTHOUGH HE PROBLEM OF EXCESSIVE VAULT WAS RESOLVED PATIENT HAS POOR NAKED EYE VISION DUE TO INDUCED ASTIGMATISM. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2023-00607 FOR INITIAL LENS. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16603892&pc=MTA | 2023826-2023-00947 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS. THE SURGEON NOTED PIGMENT DISPERSION ON THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT DATE OF BIRTH: UNK. PATIENT SEX: UNK. PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. HEALTH IMPACT: CLINICAL CODE: (B)(4) - PIGMENT DISPERSION. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED PIGMENT DISPERSION. REPORTEDLY "SEEING PIGMENT DISPERSION ON A CASE THAT THE LENS WAS REMOVED AND THEN IMPLANTED. HE WOULD LIKE A CONSULTATION". THE LENS REMAINS IMPLANTED. IT WAS LATER REPORTED "THE PATIENT IS DOING BETTER AND THE PIGMENT DISPERSION DOES NOT APPEAR ACTIVE ANYMORE. THE LIBERATED PIGMENT MOST LIKELY CAME FROM SURGERY. OCT SHOWED THE ICL VAULT TO BE AT NORMAL AMOUNT. I CAN'T RECALL EXACTLY. I THINK IT WAS RIGHT AT 500 MICRONS. IOP HAS REMAINED LOW. VISION EXCELLENT." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515448&pc=MTA | 2023826-2023-03416 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICLIMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS) IN (B)(6) 2023. THE PATIENT HAD MINIMAL CYLINDER, HAZE TO HER LEFT CORNEA (UNCLEAR ETIOLOGY). SHE ENDED UP IN LEFT EYE 20/25 BUT IS BOTHERED BY A LACK OF CLARITY. REFRACTION IS PL-1.25X 175. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: PER THE DOCTOR, AT THE LAST FOLLOW-UP, THE PATIENT HAD NO COMPLAINTS. THE LENS REMAINED IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: B3: UNK. D4: EXPIRATION DATE UNK. D6B: UNK. H4: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - MINIMAL CYLINDER, HAZE TO CORNEA, LACK OF CLARITY. CLAIM(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17574481&pc=MTA | 2023826-2023-03340 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2023. REPORTEDLY EXCESSIVE VAULT AND RESIDUAL CYLINDER WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17574503&pc=MTA | 2023826-2023-03341 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2023. REPORTEDLY EXCESSIVE VAULT AND RESIDUAL CYLINDER WAS OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17792198&pc=MTA | 2023826-2023-03988 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. POST-OP THERE WAS 1+ PIGMENT IN ANTERIOR CHAMBER, IRIS WAS RESOLVING. THE PROBLEM WAS RESOLVED. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT IN AC. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17813875&pc=MTA | 2023826-2023-03989 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. POST-OP THERE WAS 1+ PIGMENT IN ANTERIOR CHAMBER, IRIS WAS RESOLVING. THE PROBLEM WAS RESOLVED. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954491&pc=MTA | 2023826-2023-04628 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. MILD SHALLOWING OF THE ACD WAS REPORTED, HOWEVER IT WAS ALSO REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. "LENS SIZE A BIT OFF" WAS REPORTED FOR CAUSE OF EVENT. THE REPORTER STATED THAT THE PATIENTS CURRENT CONDITION IS GOOD. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: MILD SHALLOWING OF ACD. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18249789&pc=MTA | 2023826-2023-05292 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS. THE PATIENT COMPLAINED OF "WAVY VISION" AT 3 WEEKS POST-OP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - WAVY VISION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13896093&pc=MTA | 2023826-2022-00891 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2022-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2011. THE PLAN IS TO EXPLANT THE LENS DUE TO PROGRESSIVE MYOPIA. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. D6B: NA. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725921&pc=MTA | 2023826-2023-04005 | 2003-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860704&pc=MTA | 2023826-2021-01598 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.00 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT WAS NOT SATISFIED WITH HIS POSTOPERATIVE VISUAL ACUITY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT CLINICAL CODE: (B)(4) (DISSATISFACTION WITH VISUAL ACUITY). MEDICAL DEVICE PROBLEM CODE (B)(4): OFF-LABEL USE: AGE: (B)(6). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860730&pc=MTA | 2023826-2021-01597 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -6.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT WAS NOT SATISFIED WITH HIS POSTOPERATIVE VISUAL ACUITY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234535&pc=MTA | 2023826-2023-05297 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON PERFORMED AN LRI TO CORRECT ASTIGMATISM BUT THE ASTIGMATISM WAS NOT FULLY CORRECTED. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE PATIENT WAS DISSATISFIED WITH VISUAL ACUITY. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - DISSATISFACTION WITH VISUAL ACUITY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE: 2993. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725918&pc=MTA | 2023826-2023-04006 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES . H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582968&pc=MTA | 2023826-2022-00596 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2021 AS A REPLACEMENT LENS AND " FEW IRIS PIGMENTS IN ENDOTHELIUM" WAS OBSERVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT (OS) EYE ON (B)(6) 2021 AS A REPLACEMENT LENS AND " FEW IRIS PIGMENTS IN ENDOTHELIUM" WAS OBSERVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHINICITY:UNK. DATE OF EVENT:UNK. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHINICITY:UNK. DATE OF EVENT:UNK. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13674056&pc=MTA | 2023826-2022-00747 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS AN EXCHANGE LENS TO RESOLVE EXCESSIVE VAULT. AFTER LENS IMPLANTATION THE VAULT IS GOOD, BUT THE PATIENT HAS SLIGHT IRIS TIDS. THIS IS DUE TO INTRAOPERATIVE FLOPPY IRIS SYNDROME. THE REPORTER STATES THAT "THE PATIENT DID NOT HAVE THIS ISSUE DURING THE FIRST SURGERY, IT'S POSSIBLE SOME COMPONENT OF IT IS HIS ANATOMY." THE LENS REMAINS IMPLANTED AND NO FURTHER INTERVENTION WILL BE DONE. SEE MFR # 2023826-2022-00527 FOR INITIAL LENS. Manufacturer Narrative: DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639426&pc=MTA | 2023826-2022-01776 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS HAVING A COMPLICATION DUE TO VISCO. THE PATIENT HAD AN ISCHEMIC OPTIC NEUROPATHY DUE TO VISCO. AT SOME POINT THE PATIENT ONLY HAD LIGHT PERCEPTION BUT IS NOW RECOVERING. PATIENT HAS MYDRIASIS DUE TO IRIS ISCHEMIA AND HAS PIGMENT DISPERSION. THE DOCTOR WANTS TO FOCUS ON RECOVERY AND THEN LOOK FOR CORRECTION OF THE ASTIGMATISM WITH PRK. THE LENS REMAINED IMPLANTED. THIS IS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-00756. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. (B)(4). CLAIM # 724553 Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391267&pc=MTA | 2023826-2023-03042 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO PUPIL DISCLOSURE AND CATARACT. PHACO-EMULSIFICATION(CATARACT SURGERY) & IOL IMPLANTATION PERFORMED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF-LABEL: AGE>60 AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: PUPIL DISCLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16864757&pc=MTA | 2023826-2023-01596 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS REPORTED AS REFRACTIVE CHANGE OVERTIME AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/LENGTH AND DIOPTER POWER LENS (VICM5_13.2, -12.00) AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16871546&pc=MTA | 2023826-2023-01597 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2013. THE LENS WAS REPORTED AS REFRACTIVE CHANGE OVERTIME AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER POWER (-14.50) LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135886&pc=MTA | 2023826-2021-00057 | 2020-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-CORRECTED INFO: THE LENS WAS IMPLANTED IN THE RIGHT (OD) EYE AND NOT THE LEFT (OS) EYE AS ORIGINALLY REPORTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12242102&pc=MTA | 2023826-2021-02923 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12246275&pc=MTA | 2023826-2021-02924 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK , ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891206&pc=MTA | 2023826-2021-01704 | 2020-12-31 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11522549&pc=MTA | 2023826-2021-00684 | 2021-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12108819&pc=MTA | 2023826-2021-02190 | 2021-05-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112713&pc=MTA | 2023826-2021-02191 | 2021-05-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443623&pc=MTA | 2023826-2021-03480 | 2021-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND UNREACTIVE (FIXED) PUPIL. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT IS HAPPY. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND UNREACTIVE (FIXED) PUPIL. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138554&pc=MTA | 2023826-2022-01180 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +06.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE VAULT IS OK AND NORMAL PIO. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985509&pc=MTA | 2023826-2022-02124 | 2022-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B)(4). CLAIM # B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14750533&pc=MTA | 2023826-2022-01923 | 2022-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 16-MAY-2022. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED. POST-OP ON 07-JUL-2022 THE PATIENT IS ALL FINE. THE VAULT HAS IMPROVED AND NORMALIZED. THE PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) SHOULD BE REMOVED FROM INITIAL MDR. CLAIM (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629417&pc=MTA | 2023826-2022-03505 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15652542&pc=MTA | 2023826-2022-03506 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED, HAPTIC TORN/BROKEN/BENT/DEFORMED AND STRETCHED OUT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035464&pc=MTA | 2023826-2022-04560 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673955&pc=MTA | 2023826-2023-01060 | 2022-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170878&pc=MTA | 2023826-2023-05170 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(4) 2023 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16550264&pc=MTA | 2023826-2023-00958 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587801&pc=MTA | 2023826-2023-01007 | 2023-02-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -3.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND SHALLOW AC. THE SURGEON NOTES THEY LIKELY WILL REPLACE THE ICL WITH A SMALLER LENS; EITHER 13.2 OR 12.7. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND INDICATED "VAULT IN THE RIGHT EYE IS CAUSING A DEFINITE SHALLOWING OF AC AND WILL LIKELY NEED TO BE REMOVED NEXT WEEK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW AC; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16967963&pc=MTA | 2023826-2023-01915 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(6) 2023 AND IN ANOTHER SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18307785&pc=MTA | 2023826-2023-05437 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICH 13.2 IMPLANTABLE COLLAMER LENS OF +3.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-CLINICAL SIGNS 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18307809&pc=MTA | 2023826-2023-05438 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICH 13.2 IMPLANTABLE COLLAMER LENS OF +2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WITH SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-CLINICAL SIGNS 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086570&pc=MTA | 2023826-2023-04884 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387744&pc=MTA | 2023826-2023-05773 | 2023-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 12-OCT-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180097&pc=MTA | 2023826-2023-05210 | 2023-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS ALSO NOTED. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18248706&pc=MTA | 2023826-2023-05409 | 2023-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -13 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862020&pc=MTA | 2023826-2023-04065 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5 13.7, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, LENS DISLOCATION, AND PIGMENT DISPERSION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5 13.7, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, LENS DISLOCATION, AND PIGMENT DISPERSION. PROBLEM RESOLVED. REPORTEDLY, LENS PERFECT, LENS IN GOOD CONDITION, AND VAULT IS AS DESIRED. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS, PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, LENS DISLOCATION AND PIGMENT DISPERSION. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "THE LENS PERFECT, LENS IN GOOD CONDITION AND ITS VAULT IS DESIRED". D6B.- "10/11/2023" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- HEALTH EFFECT - IMPACT CODE: "2199 SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1494- OFF-LABEL USE (ACD<3.0MM)" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272109&pc=MTA | 2023826-2022-02875 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICH12.1; +4 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "LENS TOO LONG. ICL PROLAPSED FORWARD WITH CORNEAL TOUCH". IT WAS INDICATED THERE IS NO FURTHER INTERVENTION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). EXPIRATION DATE : 04-30-2024. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT THE CAUSE OF THE EVENT WAS THE DEVICE AND NOTED "LENS TOO BIG". H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18221440&pc=MTA | 2023826-2023-05373 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INTERNATIONAL JOURNAL OPHTHALMOL TITLED, 'ACUTE MICRO-MACULAR HOLE ASSOCIATING WITH EXTENSIVE INTRAOPERATIVE ROTATION OF IMPLANTABLE COLLAMER LENS WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE ASSISTANCE: A CASE REPORT'. THE ARTICLE REPORTS " A MICRO-MACULAR HOLE (MMH) FOLLOWING THE INTRAOPERATIVE ROTATION OF IMPLANTABLE COLLAMER LENS WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE (OVD) ASSISTANCE AND EVENTUALLY THIS MMH CLOSED SPONTANEOUSLY". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ACUTE MICRO-MACULAR HOLE. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17054417&pc=MTA | 2023826-2023-02176 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT "APPEARS GREAT AND STABLE POSTOP COURSE WITH VISION 20/40 IN THE RIGHT AND 20/25 IN THE LEFT WITH RESIDUAL MYOPIA (EXPECTED). HIS VAULT IS ESTIMATED 1 TO 1.5 CCT DEPTHS. HE HAS UPDATED GLASSES WHICH HE USES AND WEARS THAT HE STATES DOES NOTHING FOR THE GLARE. IT HELPS HIM SEE BETTER BUT NOTHING FOR THE GLARE WHICH IS THE MAIN BOTHER. THE GLASSES ARE ABOUT -2. HIS ICL'S WERE EVO AND -16.5 / 1.0 CYL AND SIZE 12.6. HE COMPLAINS OF SATISFACTORY DAY TIME VISION BUT POOR QUALITY / POOR CONTRAST WHEN INSIDE AND ESPECIALLY POOR AT NIGHT TIME. HE STATES IT IS BETTER WITH BRIMONIDINE. WHEN PRESSED HE STATES THAT ITS ACCEPTABLE. HE IS SO BOTHERED BY THE GLARE AND POOR VISION ESPECIALLY INSIDE / NIGHT THAT HE DESIRES THE ICL'S REMOVED OVER USING BRIMONIDINE. I HAVE NOT TRIED OR RECOMMENDED PILO BECAUSE BRIMONIDINE SEEMS TO HELP." LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - BAD VISION, POOR QUALITY. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17054419&pc=MTA | 2023826-2023-02175 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2023. THE SURGEON REPORTS THE PATIENT "APPEARS GREAT AND STABLE POSTOP COURSE WITH VISION 20/40 IN THE RIGHT AND 20/25 IN THE LEFT WITH RESIDUAL MYOPIA (EXPECTED). HIS VAULT IS ESTIMATED 1 TO 1.5 CCT DEPTHS. HE HAS UPDATED GLASSES WHICH HE USES AND WEARS THAT HE STATES DOES NOTHING FOR THE GLARE. IT HELPS HIM SEE BETTER BUT NOTHING FOR THE GLARE WHICH IS THE MAIN BOTHER. THE GLASSES ARE ABOUT -2. HIS ICL'S WERE EVO AND -16.5 / 1.0 CYL AND SIZE 12.6. HE COMPLAINS OF SATISFACTORY DAY TIME VISION BUT POOR QUALITY / POOR CONTRAST WHEN INSIDE AND ESPECIALLY POOR AT NIGHT TIME. HE STATES IT IS BETTER WITH BRIMONIDINE. WHEN PRESSED HE STATES THAT ITS ACCEPTABLE. HE IS SO BOTHERED BY THE GLARE AND POOR VISION ESPECIALLY INSIDE / NIGHT THAT HE DESIRES THE ICL'S REMOVED OVER USING BRIMONIDINE. I HAVE NOT TRIED OR RECOMMENDED PILO BECAUSE BRIMONIDINE SEEMS TO HELP." LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - BAD VISION, POOR QUALITY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13052015&pc=MTA | 8040449-2021-00004 |  | Injury | OPHTEC BV | 2021-12-21 05:00 | MTA | VERISYSE MYOPIA | Output Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030028 | 1 | Event Description: PATIENT'S IOL WAS EXPLANTED, SINCE IT WAS CAUSING THE RIGHT EYE'S CELL CONE TO DIMINISH. IOL INFORMATION: VERISYSE MYOPIA IOL 6/8.5, DIOPTER -6.0, SERIAL NUMBER (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065550&pc=MTA | 2023826-2020-03086 | 2020-11-24 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. SPECULAR MICROSCOPY DIAGNOSTICES PERFORMED. NO TREATMENT PERFORMED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. "THE LONGER CASE TIME AND ADDITIONAL REQUISITE MANIPULATION ARE LIKELY ASSOCIATED WITH THE ENDOTHELIAL CELL LOSS, EVENT RELATED TO THE PROCEDURE BUT NOT RELATED TO THE DEVICE." REPORTEDLY, "THE SURGEON WILL CONDUCT A REPEAT SPECULAR MICROSCOPY IN (B)(6) 2021 TO DETERMINE WHETHER FURTHER ACTION IS NEEDED." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17493392&pc=MTA | 2023826-2023-03434 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO PATIENT SUFFERING FROM PSYCHOLOGICAL ANXIETY, COULD NOT SLEEP, AND SUFFERED FROM SEVERE NEURASTHENIA. REPORTEDLY AFTER THE ICL LENS IMPLANTATION, THE PATIENT HAD A FOREIGN OBJECT IN THE EYE, AND FELT ANXIOUS AND SLEEPLESS! AFTER THE LENS WAS REMOVED THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15280270&pc=MTA | 2023826-2022-02877 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: REPORTEDLY, "A PATIENT... WHO RECENTLY HAD SURGERY; EVO ICL WITH GREAT RESULTS. PATIENT NOW HAS A NON-REACTIVE PUPIL OD AFTER NORMAL CONSTRICTION UP TO 2 WEEKS POST-OP. VAULTS ARE GOOD, OS UNREMARKABLE. PILOCARPINE NOT CONSTRICTING PUPIL." Manufacturer Narrative: COMMON DEVICE NAME: CANNOT BE LEFT BLANK DUE TO UNKNOWN SERIAL NUMBER. CLAIM#: (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT: IMPACT CODE 4644: PILOCARPINE USAGE. MEDICAL DEVICE PROBLEM CODE 2622- REPORTER STATED CAUSE OF EVENT IS PATIENT RELATED FACTOR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023012&pc=MTA | 2023826-2021-04905 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED AND THERE WAS PATIENT INJURY. DURING THE LENS REMOVAL THE LENS HAD RUPTURED AND DAMAGED THE CORNEAL ENDOTHELIUM AND IT WAS ALSO REPORTED THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF ENDOTHELIA CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS. IT WAS INDICATED THE PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15828414&pc=MTA | 2023826-2022-04012 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE ALFONSO BARTOLOZZI ET AL. EYE AND VISION ON (B)(6) 2022 ENTITLED, 'THREE YEAR FOLLOW UP OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS WITH A CENTRAL PORT DESIGN AFTER DEEP ANTERIOR LAMELLAR KERATOPLASTY'. THE ARTICLE REPORTS ADVERSE EVENTS FOLLOWING ICL IMPLANTATION. IN ONE EYE THE ICL WAS ROTATED BETWEEN THE 6- AND 12-MONTH FOLLOW-UP VISITS. THIS ICL ROTATION INDUCED MIXED ASTIGMATISM (+1.50 - 2.00 × 20°), LEADING TO A DECREASE OF 2 LINES-× OF UDVA, BUT CDVA REMAINED UNCHANGED. THE VAULT IN THE 12-MONTH VISIT WAS 700 M. AS THE CDVA WAS MAINTAINED, AND THE INDUCED MIXED WAS WELL-TOLERATED WITH SPECTACLES, NO RE-CENTERING MANEUVERER OR ICL EXCHANGE FOR A LARGER SIZE WAS PERFORMED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316755&pc=MTA | 2023826-2022-02956 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2; -07.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE. IT WAS REPORTED THAT THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. EXPIRATION DATE: 04-30-2025. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. B5 - THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2; -07.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "WW WAS 12; COMPROVED AGAIN IN POSTOPERATIVE WITH RETINAL COMPASS". CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973019&pc=MTA | 2023826-2021-04657 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO ONE OF HIS PATIENTS AND HAS NOW REPORTED A PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENT HAD SOME PRESSURE ISSUES AFTER ICL IMPLANTATION AND WAS REFERRED TO A GLAUCOMA SPECIALIST. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PRESSURE ISSUES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16147830&pc=MTA | 2023826-2023-00059 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IN THE ARTICLE, "RETAINED PRESUMED INTRAOCULAR COTTON FIBER AFTER IMPLANTABLE COLLEMER LENS IMPLANTATION: A CASE REPORT," THE STUDY OBSERVES A PATIENT WITH A THREAD BEHIND AN IMPLANTED ICL. REPORTEDLY, "A TINY WHITE THREAD WAS ATTACHED TO THE POSTERIOR SURFACE OF THE ICL WITH NO CATARACTOUS CHANGES. UPON DISCHARGE, THE PATIENT WAS PRESCRIBED. MOXIFLOXACIN 0.5% 4 TIMES DAILY FOR A WEEK AND PREDNISOLONE ACETATE 1% 4 TIMES DAILY TAPERED OVER 4 WEEKS DURATION. THIS PATIENT WAS FOLLOWED UP CLOSELY IN THE 2 POSTOPERATIVE WEEKS AND DID NOT HAVE ANY NEW COMPLAINTS OR CHANGES ON EXAMINATION. 1 YEAR AFT ER THE PROCEDURE, THE PATIENT HAD A QUIET EYE WITH A VA OF 20/20; THE COTTON FIBER WAS STILL IN THE SAME PLACE AND POSITION WITH NOT APPARENT SURROUNDING CHANGES." Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- INTRAOCULAR COTTON FIBER. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17393776&pc=MTA | 2023826-2023-03126 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "DOUBLE FENESTRATION IN AN EVO ICL." THIS WAS NOTED ONE-DAY POST-OP. Manufacturer Narrative: A4-A6:UNK. B3: UNK D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: REPORTEDLY, "THE EYE WITH THE DEFECTIVE LENS IMPLANT IS CURRENTLY WITH VISION 20/40-50." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258206&pc=MTA | 2023826-2023-05307 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. AC SHALLOWING, IRIS BOWING, AND PIGMENT ON THE ANTERIROR SURGFACE OF THE LENS WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY IOP IS NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. D6B - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258208&pc=MTA | 2023826-2023-05308 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. AC SHALLOWING, IRIS BOWING, AND PIGMENT ON THE ANTERIROR SURGFACE OF THE LENS WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY IOP IS NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK A2 - UNK A3 - UNK A4 - UNK A5 - UNK A6 - UNK B3 - UNK D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY) D4 - UNK D6A - UNK D6B - UNK H4 - UNK (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16873135&pc=MTA | 2023826-2023-01664 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Patient Device Interaction Problem | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTED COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2023. IT WAS INDICATED THAT THE LENS WAS DISPLACED THROUGH THE PUPIL. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: HEALTH EFFECT- CLINICAL CODE: 4581- "IRIS ATROPHY." B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/3.0/141 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS DISLOCATION AND IRIS ATROPHY. REPORTEDLY "HALF OF THE LENS DISPLACED THROUGH THE PUPIL" AND "PLANNING TO REPOSITION". THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. STATUS OF THE EYE: "BACK TO NORMAL". THE CAUSE OF THE EVENT: PATIENT RELATED FACTOR - "PATIENT INADVERTENTLY RUBBED EYE". CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4628" SHOULD BE ADDED. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10881155&pc=MTA | 2023826-2020-02757 | 2020-01-15 05:00 | Injury | STAAR SUGICAL COMPANY | 2020-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED "INITIAL LENS 13.2 WAS OVERSIZED AND THE EDGES MAY HAVE INCISED THE CAPSULE." LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SITTING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - CAPSULAR TEAR, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12862895&pc=MTA | 2023826-2021-04280 | 2020-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM_13.2; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411351&pc=MTA | 2023826-2021-03461 | 2021-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -5.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED PATIENT HAD "OU CONJUNCTIVAL WITHOUT CONGESTION,CORNEA CLEAR,ANTERIOR CHAMBER DEPTH WAS OK,TYNDALL, ICL CENTER LOCATION" H6 - HEALTH EFFECT CLINICAL CODE: 4581 - TYNDALL (ANTERIOR CHAMBER FLARE). ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT TYNDALL WAS NOTED ON REEXAMINATION ON THE FIRST DAY AFTER THE OPERATION AND THE ISSUE HAS RESOLVED THERE IS NO ANTERIOR CHAMBER FLARE. CLAIM#: (B)(4) Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118195&pc=MTA | 2023826-2022-02411 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6: CLINICAL CODE:4581- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741579&pc=MTA | 2023826-2021-04220 | 2021-10-02 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122991&pc=MTA | 2023826-2021-04777 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. CLINICAL CODE- 4581-SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256293&pc=MTA | 2023826-2022-00137 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.7MM MICL13.7, -13.0 DIOPTER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "THE RECOMMENDATION FROM DFU FOR 13.7 WAS TOO LARGE FOR THE PATIENT." REPORTEDLY THE LENS DID NOT FAIL TO PERFORM AS INTENDED. H6-HEALTH CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. MODEL, IMPLANT DATE: UNK. MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952754&pc=MTA | 2023826-2022-04351 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, FROM THE PATIENTS RIGHT EYE (OD) DUE TO THE BIG SIZE OF THE LENS. THE LENS WAS REPLACED WITH A SMALLER SIZE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS REPLACED WITH A SMALLER SIZE LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17195743&pc=MTA | 2023826-2023-02534 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE NOTED. THE LENS WAS REMOVED ON (B)(6) 2023 AND A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036041&pc=MTA | 2023826-2023-04894 | 2023-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056304&pc=MTA | 2023826-2023-02129 | 2023-05-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/053 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LOW VAULT WAS OBSERVED. PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. B7: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. OTHER: PREVIOUS POSTERIOR CHAMBER LENS IMPLANTATION. H6: HEALTH IMPACT- CLINICAL CODE:4581 - REFRACTIVE CHANGE OVERTIME. H6: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. Manufacturer Narrative: A4-A6:UNK; B3: UNK; D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK; H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056429&pc=MTA | 2023826-2023-02130 | 2023-05-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS OBSERVED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LOW VAULT WAS OBSERVED. PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN. B7: PREVIOUS CORNEAL OR REFRACTIVE SURGERY. OTHER: PREVIOUS POSTERIOR CHAMBER LENS IMPLANTATION. H6: HEALTH IMPACT- CLINICAL CODE:4581 - REFRACTIVE CHANGE OVERTIME. H6: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. Manufacturer Narrative: A4-A6:UNK B3: UNK D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131350&pc=MTA | 2023826-2023-02394 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -13.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2023 AND THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "THE DOCTOR ROTATED THE LENS 90 DEGREES VAULT IS OK". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH EFFECT - CLINICAL CODE: "4581- UNREACTIVE (FIXED) PUPIL" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2023, THE LENS WAS REPOSITIONED. IT WAS LATER REPORTED THE PATIENT ALSO EXPERIENCED AN UNREACTIVE (FIXED) PUPIL . THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "THE PUPIL DOES NOT CLOSE COMPLETELY". ADDITIONAL INFORMATION PROVIDED: "THE SUPPOSED HIGH VAULT WAS REDUCED BY A ROTATION. THE CASE HAS BEEN UPDATED. I DONT HAVE A NUMBER YET". REPORTEDLY "THE DOCTOR HAS NOT YET DECIDED HOW TO PROCEED WITH THE PATIENT." THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS: " THE DOCTOR USED HYALURONIC ACID". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "THE PUPIL CLOSES BETTER". ADDITIONAL INFORMATION PROVIDED: "THERE ARE CURRENTLY NO MEASURES PLANNED. THE PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036034&pc=MTA | 2023826-2023-04841 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036037&pc=MTA | 2023826-2023-04842 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670775&pc=MTA | 2023826-2022-00614 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686336&pc=MTA | 2023826-2022-03835 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT. EXCESSIVE VAULT, PIGMENT DISPERSION, SHALLOW CHAMBER, AND VERY SHALLOW ANGLE WAS OBSERVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: HEALTH EFFECT-CLINICAL CODE: 4581 ADDED FOR ANGLE CLOSURE. CORRECTED DATA: DEVICE CODE 2199 SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM SHOULD BE ADDED FOR CORRECTION. H1: MALFUNCTION SHOULD BE SELECTED FOR CORRECTION. H6: 2199: NO HEATLH OUTCOMES SHOULD BE ADDED FOR CORRECTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678123&pc=MTA | 2023826-2023-01162 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLANDER LENS, INTO THEIR EYE. THE REPORTER INDICATED THE LENS WAS TOO LARGE AND THEIR IRIS NO LONGER CONSTRICTED ADEQUATELY FOR LIGHT CONDITIONS, LEAVING THE REPORTER WITH LESS THAN IDEAL VISION. THE REPORTER USED EYE DROPS WHICH HELPED BUT PROVED TO NOT BE A LONGER TERM SOLUTION. THE REPORTER WOULD LIKE THE LENS REMOVED AND REPLACED FOR CORRECT SIZE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B1 SHOULD BE CORRECTED TO ADVERSE EVENT. B2 SHOULD BE CORRECTED TO OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS. H1 SHOULD BE CORRECTED TO SERIOUS . H6: CODE 4581 SHOULD BE ADDED FOR UNREACTIVE PUPIL. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFOMATION: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18183647&pc=MTA | 2023826-2023-05219 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: AN CASE REPORT WAS PUBLISHED IN INT OPHTHALMOL IN THE YEAR 2017; TITLED: "VERTICAL IMPLANTABLE COLLAMER LENS (ICL) ROTATION FOR THE MANAGEMENT OF HIGH VAULT DUE TO LENS OVERSIZING." IT REPORTS AN EXCESSIVE VAULT AND LENS REPOSITIONED 90 DEGREES TO A VERTICAL ORIENTATION. THE OUTCOME WAS FAVORABLE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18287974&pc=MTA | 2023826-2023-05443 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. AT 2 DAYS POST-OP, THE LENS HAD A VAULT OF 2.0. THE ANGLE IS SLIGHTLY NARROW. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT ON ANTERIOR SURFACE. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO HIGH VAULT, WITH IRIS RIDING ON THE ICL. IOP WAS NORMAL. THE LEFT EYE WAS PERSISTENTLY ACHEY AND THERE WAS SOME PIGMENT ON THE ANTERIOR SURFACE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PATIENT IS HAPPY. PATIENT IS 20/20, WITH A NORMAL VAULT AND NO IRIS TOUCH. CLAIM # (B)(4). Manufacturer Narrative: A2: UNK A3: UNK A4: UNK A5: UNK A6: UNK B3: UNK D4: EXPIRATION DATE UNK D6A: UNK D6B: UNK H4: UNK (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724511&pc=MTA | 2023826-2021-04164 | 2021-10-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14962931&pc=MTA | 2023826-2022-02177 | 2022-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. AFTER IMPLANTATION THE SURGEON NOTICED THERE WAS A FOREIGN SUBSTANCE THAT LOOKED LIKE TINY PLASTIC DEBRIS WAS FOUND ON THE LENS. THE SURGEON REMOVED THE DEBRIS WITH THE FORCEPS. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: VIAL WAS RETURNED WITH NO LENS. VISUAL INSPECTION FOUND A FOREIGN PARTICULATE NOTED IN THE VIAL. PARTICULATE WAS NOT REMOVED FROM THE VIAL. ADDITIONAL INFORMATION: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - FOREIGN BODY ADHESION. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION THE FOREIGN DEBRIS ON THE LENS WAS DETERMINED TO BE A PIECE OF ICL FOOTPLATE. THE PIECE OF THE ICL FOOTPLATE LIKELY BROKE OFF ANOTHER ICL DURING THE MACHINING PROCESS AND CAME INTO CONTACT WITH THE SUBJECT LENS DURING THE MACHINING OR TUMBLING PROCESS. HOWEVER, A DEFINITIVE ROOT CAUSE FOR THIS EVENT WAS UNABLE TO BE DETERMINED. OPERATORS INVOLVED IN THE MANUFACTURING OF THE NONCONFORMING LENS HAVE BEEN RETAINED AND MADE AWARE OF THE IMPORTANCE OF VERIFYING THAT NO PARTICULATES ARE PRESENT ON THE LENS. THERE IS NO INDICATION THAT SUGGESTS A DEFICIENCY IN THE MANUFACTURING PROCESS OR TREND OF PROCESS ERRORS. FURTHER INVESTIGATION IS NOT REQUIRED AT THIS TIME. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18121662&pc=MTA | 2023826-2023-04971 | 2023-10-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. "VERY SHALLOW AC POST ICL" WAS REPORTED. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16809011&pc=MTA | 2023826-2023-01476 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "A POST ICL PATIENT, WHO VISION IS GOOD, PRESSURE IS GOOD, BUT IS CONCERNED ABOUT THE ANGLES BEING NARROW COMPARED TO PRE-SURGERY." Manufacturer Narrative: D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: H6-HEALTH EFFECT- CLINICAL CODE: "4581- NARROW ANGLES" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED NARROW ANGLES. REPORTEDLY "POST ICL PATIENT, VISION IS GOOD, PRESSURE IS GOOD, BUT HE IS CONCERNED ABOUT THE ANGLES BEING NARROW COMPARED TO PRE SURGERY". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11782065&pc=MTA | 2023826-2021-01381 | 2020-10-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT (REPORTED AS 1000 UM OR 2 CT). THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5:THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS LATER EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT (REPORTED AS 100 UM). THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM#(B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT (REPORTED AS 100 UM). THIS RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552903&pc=MTA | 2023826-2022-00448 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18480555&pc=MTA | 2023826-2024-00090 | 2023-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP: 40MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4551528&pc=MTA | 2023826-2015-00205 | 2014-12-25 | Injury | STAAR SURGICAL COMPANY | 2015-02-26 | MTA | VISIAN TICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND AN REFRACTIVE SURPRISE WAS NOTED. PATIENT EXPERIENCING BLURRED VISION. THE LENS WAS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH, DEVICE HISTORY RECORD. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: ACCORDING TO THE SURGEON, THE LENS WAS ALIGNED AT 20 DEGREES CCW, WHICH IS 5 DEGREES OFF-AXIS. THE PATIENT EXPERIENCED BLURRED VISION AND THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO ROTATE THE LENS 31 DEGREES ACCORDING TO ABERROMETRY TESTING. THIS RESOLVED THE PROBLEM. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF HE EYE'S MEASUREMENT TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ECT. THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE CAUSE OF THE EVENT. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, LENS WORK ORDER SEARCH AND MEDICAL REVIEW, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. THE LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6037536&pc=MTA | 2023826-2016-01482 | 2016-09-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS - -18/1/145 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON DECIDED TO MAKE A LASIK TOUCH-UP FOR BOTH EYES TO CORRECT REFRACTIVE SURPRISE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153422&pc=MTA | 2023826-2016-01745 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER VISION WAS SLIGHTLY BLURRY AT NIGHT AND IN DIM LIGHTING. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY CONFIRMED THE INFORMATION AND INDICATED THE EVENT WAS DUE TO THE PATIENT HAD LARGE PUPILS AND WAS NOT DEVICE RELATED. NO OTHER COMPLICATIONS WERE REPORTED. PATIENT'S POST-OP VISUAL ACUITY WAS 20/15 AND THE OUTCOME WAS EXCELLENT. Manufacturer Narrative: ON (B)(6) 2012. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153430&pc=MTA | 2023826-2016-01744 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HER VISION WAS SLIGHTLY BLURRY AT NIGHT AND IN DIM LIGHTING. THE LENS WAS IMPLANTED ON (B)(6) 2010 AND REMAINS IMPLANTED. THE FACILITY CONFIRMED THE INFORMATION AND INDICATED THE EVENT WAS DUE TO THE PATIENT HAD LARGE PUPILS AND WAS NOT DEVICE RELATED. NO OTHER COMPLICATIONS WERE REPORTED. PATIENT'S POST-OP VISUAL ACUITY WAS 20/20 AND THE OUTCOME WAS EXCELLENT. Manufacturer Narrative: MFR DATE- 05-25-2012. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887955&pc=MTA | 2023826-2019-01354 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY, "PATIENT WANTED BETTER NEAR VISION IN EYE AND SOME ASTIGMATISM WAS FOUND TO BE RELATED TO IT AS WELL." ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A TORIC ICL. THIS RESOLVED THE PROBLEM. PER REPORTER ON (B)(6) 2019, "PATIENT IS HAPPY WITH VISION AND SAYS CONTINUALLY IMPROVING." Manufacturer Narrative: RESULT CODE: 213 REPORTED IN INITIAL MDR IS CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS EXPLANT DATE OF (B)(6) 2019 IS NOT APPLICABLE. THE EXPLANT DATE WAS ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN CONTACT LENS CASE, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6231187&pc=MTA | 2023826-2017-00030 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074352&pc=MTA | 2023826-2016-00909 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00/+5.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064967&pc=MTA | 2023826-2016-01486 | 2016-06-14 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +7,00/+5,00/97 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. IT WAS INDICATED THE PATIENT DEVELOPED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON HAS DECIDED TO CORRECT THE RESIDUAL REFRACTION WITH LASER VISION CORRECTION. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: N/A. PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513718&pc=MTA | 2023826-2015-00143 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS). THE PATIENT REPORTED THE LENS WAS EITHER REMOVED, REPLACED OR REPOSITIONED? ALSO ASTIGMATLISM IS RETURNING AND THERE IS SORENESS AND BLURRINESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON INDICATED HE DID NOT AGREE WITH THE PATIENT REPORT AND AS OF THE DATE OF THE LAST POST-OP VISIT ON (B)(6) 2015, THE PATIENT HAS NOT HAD ANY ADDITIONAL PROCEDURES. THE SURGEON INDICATED THE ASTIGMATISM WAS PRE-EXISTING AND MINOR. THE PATIENT IS 20/20 OU WITHOUT COMPLICATION. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING THE DIRECTION FOR USE (DFU), THE MICL LENS IS USED TO ADDRESS MILD TO MODERATE MYOPIA. ASTIGMATISM COULD NOT BE FULLY CORRECTED WITH A MICL LENS. MOST LIKELY THE CAUSE OF THE EVENT IS THE PATIENT'S PREEXISTED CONDITION: ASTIGMATISM. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE EVENT IS THE PATIENT'S PREEXISTED CONDITION: ASTIGMATISM. (B)(4). Manufacturer Narrative: (B)(4): EVALUATION(B)(6): METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789550&pc=MTA | 2023826-2015-00578 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4789592&pc=MTA | 2023826-2015-00579 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS A RESIDUAL HYPEROPIC ERROR AND THE PATIENT WAS HAVING DIFFICULTIES WITH NEAR VISION (BLURRED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): (OTHER): LENS IMPLANTED. EVALUATION CODES: METHOD -(OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: BASED ON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213979&pc=MTA | 2023826-2018-00137 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6, -5.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE SURGEON NOTES LENS ROTATION, LOSS OF BCVA, AND BLURRED VISION. REPORTEDLY, LENS REMAINS IMPLANTED. SURGEON REPORTS THE "ICL ON DEAD VERTICAL POSITION." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213981&pc=MTA | 2023826-2018-00135 | 2017-11-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/+2.0/134 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE SURGEON REPORTS LENS ROTATION AND BLURRED VISION. ADDITIONALLY, THE SURGEON STATES THE LENS IS "ON A DEAD VERTICAL POSITION." REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438655&pc=MTA | 2023826-2023-03285 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REQUESTED A VECTOR ANALYSIS AND LATER IT WAS REPORTED THAT "THE MEAN CORNEAL POWER INCREASED BY 0.12D. THE MAGNITUDE OF THE CORNEAL ASTIGMATISM DECREASED BY -0.12D, THE STEEP MERIDIAN ROTATED CLOCKWISE BY 57 DEGREES. THE CORNEA FLATTENED BY 0.24D AND STEEPENED BY 0.48D ALONG THE 174 AND THE 84 DEGREE MERIDIANS. THE SURGICALLY INDUCED ASTIGMATISM IS 0.72X174. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438656&pc=MTA | 2023826-2023-03286 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REQUESTED A VECTOR ANALYSIS AND LATER IT WAS REPORTED THAT THE MEAN CORNEAL POWER DECREASED BY -0.22 D. THE MAGNITUDE OF THE CORNEAL ASTIGMATISM INCREASED BY 0.81 D, AND THE STEEP MERIDIAN ROTATED COUNTERCLOCKWISE BY 50 DEGREES. THE CORNEA FLATTENED BY 0.78 D AND STEEPENED BY 0.35 D ALONG THE 176 AND THE 86 DEGREE MERIDIANS, RESPECTIVELY. THE SURGICALLY INDUCED ASTIGMATISM IS 1.13 X 176. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136399&pc=MTA | 2023826-2017-01940 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/X090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ONE YEAR LATER, THE REPORTER INDICATED A REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.5/3.0/090 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ONE YEAR LATER THE REPORTER INDICATED LENS ROTATION AND REFRACTIVE SURPRISE. THE PATIENT IS HAPPY AND THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE IMPLANTATION DATE IN THE PREVIOUS SUPPLEMENTAL MDR IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2016". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6326555&pc=MTA | 2023826-2017-00153 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON PATIENT'S LAST VISIT ON (B)(6) 2015, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/20. THE STATUS OF THE EYE WAS A "QUIET EYE" AND "GOOD VAULTING". Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THE PATIENT HAS AN ANTERIOR CHAMBER DEPTH OF 2.93MM. THIS LENS MODEL IS INDICATED FOR USE WITH AN ANTERIOR CHAMBER DEPTH (ACD) OF EQUAL OR GREATER THAN 3.0 MM, AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205806&pc=MTA | 2023826-2018-00105 | 2015-11-30 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTICED BY THE SURGEON. THE PATIENT HAS UNSTABLE REFRACTION AND THE SURGEON IS CONTROLLING HIS REFRACTION PERIODICALLY. THE CONTACT REPORTED THAT WITHOUT A STABLE REFRACTION LENS EXCHANGE CANNOT TAKE PLACE. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429909&pc=MTA | 2023826-2017-00487 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -17.0/+6.0/062 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE AND BLURRY VISION. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501471&pc=MTA | 2023826-2017-00653 | 2016-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6, DIOPTER -3.5/3.5/031 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: IN THE INITIAL REPORT, THE LENS MODEL IS DESCRIBED AS A VTICM12.6. THE CORRECT MODEL IS VTICMO12.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. OFF-LABEL USE: ACD < 3.0MM (2.87MM), AGE (B)(6) AT TIME OF IMPLANT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6564060&pc=MTA | 2023826-2017-00804 | 2016-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2, DIOPTER -18.0/+2.5/X178 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393933&pc=MTA | 2023826-2017-00368 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -0.50/+3.50/168 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAIN IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE NOT EVALUATED BY MANUFACTURE BECAUSE LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6366306&pc=MTA | 2023826-2017-00276 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE LENS, DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. ADDITIONAL INFO: 03/15/2017. ADDITIONAL INFO: DEVICE RETURNED 04/05/2017; NOT YET EVALUATED. ADDITIONAL INFO: DEVICE RETURNED, EVALUATION ANTICIPATED; NOT YET PERFORMED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) , AND LENS STUCK IN VIAL. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7274885&pc=MTA | 2023826-2018-00224 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8/4/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.00/4.00/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -8.00/4.00/100 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275256&pc=MTA | 2023826-2018-00225 | 2017-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) . THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.50/+3.00/87 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. CORRECTED FROM (B)(6) 2017 TO UNK CLAIM# (B)(4). Manufacturer Narrative: CORRECTION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS; -6.50/+3.00/087 (SPHERE/CYLINDER/AXIS); INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6545668&pc=MTA | 2023826-2017-00764 | 2017-04-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -17.00/250/125 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED, HOWEVER A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PREVIOUS: LENS DIOPTER -17.00/250/125; CORRECTED: LENS DIOPTER -17.00/2.5/125. DEVICE EVALUATION, PREVIOUS: STATED THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. CORRECTED: NO CLEAR RESIDUE/DEBRIS WAS FOUND ON THE PRODUCT. (B)(4). Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR DIOPTER LENS. UCVA IS 20/20, VAULT IS ~500UM. THE PATIENT IS HAPPY, PROBLEM IS RESOLVED. METHOD CODE, PREVIOUS:(NO TESTING METHODS PERFORMED). UPDATED: (DEVICE FROM SAME LOT EVALUATED). RESULT CODE, PREVIOUS: (OPERATIONAL PROBLEM). UPDATED: (NO FAILURE DETECTED). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6680940&pc=MTA | 2023826-2017-01087 | 2017-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ACCORDING TO THE CUSTOMER, THE SURGEON DOES NOT HAVE A PLAN TO EXCHANGE THE LENS, AND WILL FOLLOW-UP THE PATIENT'S CONDITION WITHIN THE NEXT FEW MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941344&pc=MTA | 2023826-2017-01558 | 2017-07-15 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON DURING THE POST-OP VISIT, DATE OF (B)(6) 2017. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952440&pc=MTA | 2023826-2017-01566 | 2017-08-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.5/+3.0/076 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. Manufacturer Narrative: CORRECTED DATA: LENS REMAINS IMPLANTED. DOCTOR WANTS TO OBSERVE FOR A FEW MONTHS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879884&pc=MTA | 2023826-2017-01477 | 2017-08-24 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON. ON 28-08-2017, DURING THE PATIENT'S LAST REPORTED VISIT, THE UCVA WAS MEASURED AT 20/20 AND THE RX +0.25 SPH. PER THE SURGEON NO LENS EXCHANGE IS NEEDED. THE CLAIM REPORTER STATED THAT THERE POSSIBLY WAS SOME VISCO LEFT IN THE EYE AT 1-DAY POST-OP, WHICH MAY HAVE CAUSED THE REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6938985&pc=MTA | 2023826-2017-01554 | 2017-08-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.5/+3.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970586&pc=MTA | 2023826-2017-01595 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11.5/+1/009 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE DOCTOR HAS DECIDED TO WAIT ANOTHER FEW MONTHS TO SEE WHAT HAS HAPPENED WITH THE REFRACTION. LENS IS STILL CURRENTLY IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970590&pc=MTA | 2023826-2017-01568 | 2017-09-25 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2 , -11/1.5/151 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097940&pc=MTA | 2023826-2017-01835 | 2017-09-27 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.50/+2.00/080 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097943&pc=MTA | 2023826-2017-01836 | 2017-09-28 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -10.50/+3.50/097 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6987323&pc=MTA | 2023826-2017-01633 | 2017-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2017-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -3.0/4.5/X070 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE SURGEON IS PLANNING TO EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE LENS WILL BE ROTATED IN (B)(6) 2018. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146626&pc=MTA | 2023826-2017-01941 | 2017-10-19 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.0/+2.0/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE WITH THE POST-OP MANIFEST REFRACTION, SPH. -0.1 AMD CYL. -1.75. THE SURGEON IS PLANNING TO PERFORM LASIK FOR ADDITIONAL CORRECTION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205784&pc=MTA | 2023826-2018-00103 | 2017-12-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -10.5/+6/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268465&pc=MTA | 2023826-2018-00179 | 2018-01-15 | Malfunction | STAAR SURGICAL COMPANY AG | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0 MM ICM130V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S (OS) ON (B)(6) 2001. ON (B)(6) 2018, REFRACTIVE CHANGE OVER TIME WAS NOTICED. THE SURGEON IS MAKING PLANS TO EXCHANGE THE LENS. THE REPORTER STATED THAT THE REASON FOR THE REFRACTIVE CHANGE OVER TIME WAS BECAUSE OF PENETRATING KERATOPLASTY (PKP) PERFORMED 18 YEARS AGO. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.0MM ICM130V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2001. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484263&pc=MTA | 2023826-2018-00639 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660090&pc=MTA | 2023826-2018-01003 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS >2-LINE DROP OF BCVA AT THE 1-MONTH POSTOPERATIVE VISIT. UPON THE REPEAT EXAM OF THE BCDVA DURING A FOLLOW UP VISIT, IT WAS DETERMINED THAT, ALTHOUGH THERE WAS A 1-LINE IMPROVEMENT IN THE BEST CORRECTED VISION OF THE EYE, IT IS STILL A 2-LINE DROP FROM PREOPERATIVE VALUES. THE LENS REMAINS IMPLANTED AND INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IN THE SURGEON'S OPINION, THE CAUSE OF THIS EVENT IS THE DEVICE. THE STATUS OF THE PATIENT'S EYE IS "RESOLVING." (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT IS DETERMINED TO BE RESOLVED. SUBJECT COMPLETED THE STUDY ON (B)(6) 2018 WITH BCDVA OF 20/20. IT IS THE MEDICAL MONITOR'S OPINION THAT THE TRANSIENT DECREASE AND SUBSEQUENT IMPROVEMENT IN BCDVA REFLECT'S THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395325&pc=MTA | 2023826-2017-00359 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+6.0/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS CURRENTLY UNDER MEDICAL CONSULTATION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429982&pc=MTA | 2023826-2017-00475 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+2/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. DOCTOR IS MONITORING THE PATIENT STATUS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE REPORTER CONFIRMED THAT THE DOCTOR MADE A MISTAKE WHEN WRITING LENS POWER AND PRE-OP K VALUES IN THE OS IN THE OCOS FORM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531866&pc=MTA | 2023826-2017-00659 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND FELT EMBARRASSED. THE CUSTOMER STATED THE LENS WAS CORRECTLY IMPLANTED. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULTING, ATTEMPTS TO GET ADDITIONAL INFORMATION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6839063&pc=MTA | 2023826-2017-01374 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED AS THE PATIENT IS NOW APPROXIMATELY +1.25D SEQ POST-OP. THE SURGEON CONFIRMED THAT THE ICL IS SITTING IN THE CORRECT POSITION AND NOT INVERTED. HE ALSO VERIFIED THAT THE CYCLO WAS NOT SIGNIFICANTLY DIFFERENT THAT THE MANIFEST REFRACTION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: METHOD: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6839072&pc=MTA | 2023826-2017-01375 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL12.1, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978049&pc=MTA | 2023826-2017-01625 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -12.00 INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY, IMMEDIATELY AFTERWARD, THE PATIENT COMPLAINED OF A CIRCULAR "CLOUD" (BLURRED SPOT) IN HER FIELD OF VIEW JUST OFF-CENTER, ADJACENT TO HER CENTRAL VISION. THE SURGEON STATES THE "HOLE" OF THE LENS IS "VERY CLOSE TO CENTER." UPON EXAM, PATIENT'S CORNEA IS CLEAR WITH NO CATARACT FORMATION PRESENT. SURGEON STATES THAT "ON PAPER" THE PATIENT HAS A GREAT RESULT AND SUGGESTS POSSIBLE "NON-ORGANIC" EXAGGERATION OF SYMPTOMS. STAAR WAS NOTIFIED OF THIS CASE BY A SECOND SURGEON WHO PERFORMED LASIK SURGERY ON SAME PATIENT IN (B)(6) 2017. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336508&pc=MTA | 2023826-2018-00414 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336510&pc=MTA | 2023826-2018-00413 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336522&pc=MTA | 2023826-2018-00410 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336540&pc=MTA | 2023826-2018-00411 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2, IMPLANTABLE COLLAMER LENS, -13.50/+2.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7085101&pc=MTA | 2023826-2017-01850 | 2017-10-31 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +4.0/+3.5/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: EXPLANTED DATE: IT IS CORRECTED FROM UNK TO N/A. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398188&pc=MTA | 2023826-2018-00497 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -3.50/5.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THE THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -3.50/5.5/107 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE LEFT EYE (OS) VALUES WERE USED FOR CALCULATION OF RIGHT EYE (OD) LENS REQUIREMENT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398192&pc=MTA | 2023826-2018-00498 | 2018-02-20 | Malfunction | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -4.00/3.0/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.00/ 3.0/119 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. IT WAS NOTICED THAT THE "REFRACTIVE IS STILL THE SAME AS BEFORE DURING SURGERY" AND THERE IS ALMOST NO CORRECTION. LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. FROM THE (B)(4) CALCULATION FORM WE CAN NOTICE THAT THE RIGHT EYE (OD) VALUES WERE USED FOR CALCULATION OF LEFT EYE (OS) LENS REQUIREMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6108230&pc=MTA | 2023826-2016-01651 | 2016-10-14 | Injury | STAAR SURGICAL COMPANY | 2016-11-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 3.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/074 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION CODE: PREVIOUSLY, IT WAS REPORTED THAT THE CONCLUSION CODE. THIS IS INCORRECT. THE CORRECT CONCLUSION CODE WAS REPORTED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT MODEL NUMBER AND LENGTH OF THE LENS IS VTICM5 13.2 IMPLANTABLE COLLAMER LENS. REPORT SOURCE: USER FACILITY TO DISTRIBUTOR - PREVIOUS SELECTION 'USER FACILITY' IS INCORRECT, THE CORRECT SELECTION SHOULD BE 'DISTRIBUTOR.' (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223864&pc=MTA | 2023826-2017-00021 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/066 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089425&pc=MTA | 2023826-2017-01830 | 2017-11-07 | Malfunction | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A 13.2 MM VTICM5_13.2, -4.5/+4.5/075 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AS WELL AS REFRACTIVE SURPRISE WAS NOTICED. THE LENS IS PLANNED TO BE REPOSITIONED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227635&pc=MTA | 2023826-2017-00016 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/78 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542565&pc=MTA | 2023826-2015-00174 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME (OVERCORRECTION) AND BLURRED VISION. THE ICL WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT A DIFFERENT DIOPTER POWER (-15.5), WHICH RESOLVED THE PROBLEM. THE PATIENT IS HEALING WELL OU. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED BUT WITH THE RE-MEASUREMENT OF THE LENS, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201639&pc=MTA | 2023826-2018-00086 | 2017-11-16 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -1.5/+1.50/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND EXCESSIVE VAULT. IN ADDITION, THE ICL WAS IMPLANTED "IN A PIGGY-BACK CONDITION." REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER (B)(6) YEARS OF AGE AT TIME OF IMPLANT ((B)(6) YRS), PREVIOUS CORNEAL OR REFRACTIVE SURGERY]. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180685&pc=MTA | 2023826-2018-00036 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE; INCIDENT DATE OF (B)(6) 2017. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180690&pc=MTA | 2023826-2018-00037 | 2017-12-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE; INCIDENT DATE OF (B)(6) 2017. Manufacturer Narrative: CORRECTED INFORMATION: THE LENS REMAINS IMPLANTED. THE LENS IS PLANNED TO BE EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. CONCLUSION CODE: OFF-LABEL USE [UNDER 21 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336748&pc=MTA | 2023826-2018-00377 | 2018-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052129&pc=MTA | 2023826-2018-01618 | 2018-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -2.50/6.0/168 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND LOSS OF BCVA; THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE PATIENT IS DOING FINE AFTER EXPLANTATION. Manufacturer Narrative: CORRECTED DATA: "KERATOCONUS" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE HAPTIC IS MISSING. (B)(4). Manufacturer Narrative: CORRECTED DATA: RESULT CODE CORRECTED TO: "114" ONLY IN THE SUPPLEMENTAL MDR#2, PREVIOUS CODE NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524328&pc=MTA | 2023826-2020-02041 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.00, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WERE REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MANUFACTURED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10524391&pc=MTA | 2023826-2020-02040 | 2020-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT AND REFRACTIVE CHANGE OVER TIME WERE REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030056&pc=MTA | 2023826-2022-04436 | 2021-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029739&pc=MTA | 2023826-2021-01917 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VICM5_ 12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THIS EXCHANGED RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7051045&pc=MTA | 2023826-2017-01755 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/099 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD LOW VAULTING AND LENS ROTATION. THE PATIENT COMPLAINED OF BLURRY VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER STATED THAT THE EXCHANGE IS LIKELY TO BE IN (B)(6) 2018. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877224&pc=MTA | 2023826-2017-01462 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -13.5/+1.5/080 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT POST-OP BCVA WAS 20/25 AND UCVA WAS 20/25 ON (B)(6) 2016. PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135239&pc=MTA | 2023826-2017-01925 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/+5.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIROR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135248&pc=MTA | 2023826-2017-01926 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/+6.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE AND STATES "IN SPITE OF NORMAL VAULT POST-OP DATA" HE BELIEVES THE CURRENT LENS IS LARGER THAN NEEDED. THE SURGEON NOTES THE OPTIC PART OF THE LENS IN THE INFERIOR SIGMENT IS "A BIT OBTRUSIVE." NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: CORRECTED DATA: THE SURGEON NOTES THAT AFTER THE DILATION, "THE OPTIC PART OF THE LENS IN THE INFERIOR SEGMENT IS A BIT PROTRUSIVE". CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227617&pc=MTA | 2023826-2017-00022 | 2016-11-29 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 ,THE PATIENT REPORTED LOW VAULT, ROTATION OF THE LENS, AND REFRACTIVE SURPRISE.. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULT: NO SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT WEIGHT-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED; REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592022&pc=MTA | 2023826-2023-03621 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS EXPLANTED. THE LENS WAS LATER REPLACED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS: "THE ARCH HEIGHT IS TOO LOW". CLAIM# (B)(4). Manufacturer Narrative: H6: PATIENT RELATED FACTOR: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227640&pc=MTA | 2023826-2017-00017 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/2.0/152 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. ADDITIONALLY, THE COMPLAINT FORM STATES, "CUSTOMER MIXED UP THE CALCULATION DATA BETWEEN OD AND OS." Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242804&pc=MTA | 2023826-2017-00041 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/82 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AMD THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: PATIENT WEIGHT-UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6242817&pc=MTA | 2023826-2017-00039 | 2016-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/81 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: CORRECT OUTCOME ATTRIBUTED TO ADVERSE EVENT IS OTHER SERIOUS (DEVICE REMAINS IMPLANTED). (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784581&pc=MTA | 2023826-2022-04080 | 2018-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022, WITH A VICM5_12.6 LENS, 08.00 DIOPTER. THE PROBLEM WAS RESOLVED. THE INCISION IS HEALING WELL, ICL IS POSITIONED WELL WITH GOOD VAULT. VA IS 20/25. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823161&pc=MTA | 2023826-2021-01483 | 2018-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751185&pc=MTA | 2023826-2021-01281 | 2018-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3. DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696925&pc=MTA | 2023826-2018-01042 | 2018-06-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS A 2-LINE DROP IN BCDVA COMPARED TO THE PREOPERATIVE EXAM. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IN THE SURGEON'S OPINION, THE CAUSE OF THIS IS EVENT IS THE DEVICE. THE PATIENT'S CONDITION IS "RESOLVING." (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT DETERMINED TO BE RESOLVED BY INVESTIGATOR. ON (B)(4) 2018. THE SUBJECT COMPLETED THE STUDY WITH BCDVA OF 20/20. THE MEDICAL MONITOR BELIEVES THE TRANSIENT DECREASE AND SUBSEQUENT IMPROVEMENT IN BCDVA REFLECTS THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783494&pc=MTA | 2023826-2018-01202 | 2018-07-18 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE PATIENT IS NOT SATISFIED WITH THE VISION BECAUSE IT IS DIFFERENT FROM THE EXPECTED REFRACTION. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7799794&pc=MTA | 2023826-2018-01251 | 2018-07-26 | Injury | STAAR SURGICAL COMPANY | 2018-08-20 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THERE WAS A >/= 2-LINE DROP IN BCDVA NOTED AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. IN THE SURGEON'S OPINION, THE CAUSE OF THIS EVENT IS THE DEVICE. AT THE PATIENT'S LAST VISIT THE SURGEON STATES THERE IS A SIGNIFICANT IMPROVEMENT IN BINOCULAR UNCORRECTED VISUAL ACUITY. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: AT 3 MONTHS POSTOPERATIVE VISIT THE PATIENT'S BCDVA WAS RECORDED AS 20/20. AT 6 MONTHS POSTOPERATIVE VISIT UCVA WAS RECORDED AS 20/40. THE EVENT STATUS IS NOTED AS RECOVERED/RESOLVED. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: AT THE 6-MONTH VISIT THE CDVA DROPPED AGAIN TO 20/25. AT AN UNSCHEDULED VISIT APPROXIMATELY 2 MONTHS AFTER STUDY EXIT, THE VA RETURNED TO 20/20 AND THE NEW EVENT WAS CLASSIFIED AS RESOLVED BY THE SURGEON. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE RECURRENT TRANSIENT DECREASE IN BCDVA AND SUBSEQUENT IMPROVEMENT REFLECT THE SUBJECT'S PROCESS OF NEUROADAPTATION TO THE EDOF OPTIC. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351020&pc=MTA | 2023826-2021-00491 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351101&pc=MTA | 2023826-2021-00492 | 2018-08-01 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721899&pc=MTA | 2023826-2021-01018 | 2018-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A FOR A SAME LENGTH TORIC LENS DUE TO REFRACTIVE SURPRISE ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED/BROKEN AND THE HAPTIC BROKEN/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8071556&pc=MTA | 2023826-2018-01682 | 2018-09-04 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.00/4.0/155 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090071&pc=MTA | 2023826-2018-01702 | 2018-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/+4.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.0/4.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND POST OP DATA SHOWS LOSS OF VISUAL ACUITY. THE LES REMAINS IMPLANTED WITH LASER TREATMENT PLANNED FOR THE FUTURE. DEVICE CODE 4110: LENS WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT TYPES EVENTS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452978&pc=MTA | 2023826-2019-00489 | 2018-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-26 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721905&pc=MTA | 2023826-2021-01025 | 2018-10-14 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE CHANGE OVER TIME ON (B)(6) 2019. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090052&pc=MTA | 2023826-2018-01684 | 2018-10-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.5/+6.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DATE OF EVENT CORRECTED TO: "UNKNOWN". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372545&pc=MTA | 2023826-2019-00314 | 2018-10-23 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227420&pc=MTA | 2023826-2019-00048 | 2018-11-10 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -4.5/+4.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. CORRECTED TO: "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241988&pc=MTA | 2023826-2019-00086 | 2018-12-17 | Malfunction | STAAR SURGICAL COMPANY | 2019-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/4.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED ON (B)(6) 2018. REMOVAL OF THE LENS AND EXCHANGE IS PLANNED ONCE NEW LENS IS CALCULATED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403806&pc=MTA | 2023826-2019-00366 | 2019-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.00/6.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER STATES, "WE DON'T KNOW WHY SHE STILL HAS A REFRACTIVE CYLINDER ERROR ALTHOUGH THE LENS IS NOT ROTATED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372275&pc=MTA | 2023826-2019-00302 | 2019-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+6.0/086 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO ASTIGMATIC AXIS ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11351898&pc=MTA | 2023826-2021-00494 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6 -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8740571&pc=MTA | 2023826-2019-01152 | 2019-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, - 08.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764776&pc=MTA | 2023826-2019-01147 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VICM5_12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: FILLED IN KNOWN LOT NUMBERS. CONCLUSION CODE 4310: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11322922&pc=MTA | 2023826-2021-00422 | 2019-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VIICMO12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION FOR WIDTH DUE TO BROKEN HAPTIC. DIMENSIONAL INSPECTION FOR OVERALL AND FUNCTIONAL VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846158&pc=MTA | 2023826-2019-01210 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY, "PATIENT WANTS TO WAIT AND WEAR GLASSES AGAIN." Manufacturer Narrative: LENS RETURNED TO MANUFACTURER IN PREVIOUS MDR IS INCORECT. THE CORRECT DATE IS 09-AUG-2019. DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR IS INCORRECT. THE CORRECT DATE IS 02-OCT-2019. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS NOT REMOVED ON (B)(6) 2019, IT WAS REMOVED ON THE (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY- UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492339&pc=MTA | 2023826-2019-02405 | 2019-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "LENS WAS EXCHANGED ON (B)(6) 2020 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED, " H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: IS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814218&pc=MTA | 2023826-2020-00408 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AS A REPLACEMENT LENS, SEE MFR REPORT # 2023826-2020-00407 FOR INITIAL LENS. THE REPORTER INDICATED POST-OP, THE PATIENT HAD LOSS OF VISUAL ACUITY. TO DATE, THE BCVA HAS NOT INCREASED BUT VAULT IS BETTER. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: CORRECTED DATA: D4: SERIAL # (B)(6). CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988384&pc=MTA | 2023826-2020-00955 | 2020-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS IMPLANTED AS AN EXCHANGE LENS AND IT WAS UNKNOWN IF THE PROBLEM WAS RESOLVED, IF THE PATIENT STILL EXPERIENCED A REFRACTIVE SURPRISE. SEE MFR REPORT # 2023826-2020-00954 FOR THE INITIAL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002861&pc=MTA | 2023826-2020-00977 | 2020-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -2.25 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: ON (B)(6) 2021; THE LENS WAS REMOVED AND A SAME MODEL; SAME LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. REPORTEDLY; "THE PATIENT DID NOT SHOW UP FOR THE FOLLOW-UP EXAMINATIONS. SHE IS CURRENTLY UNAVAILABLE." CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459229&pc=MTA | 2023826-2022-00385 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: "ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION." H6 - HEALTH EFFECT - CLINICAL CODE: 2137 BLURRED VISION SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453113&pc=MTA | 2023826-2021-00603 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO 12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH TORIC LENS ( -6.00/1.0/082 (SPHERE/CYLINDER/AXIS)) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. RESUT CODE 213 SHOULD HAVE BEEN REPORTED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10577512&pc=MTA | 2023826-2020-02194 | 2020-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXPLANTED ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPLACED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE REPORTER STATES "POST-OPERATIVE ASTIGMATISM DUE TO MULTIPLE SURGERIES REMAINS." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE / DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH STAINS / FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098617&pc=MTA | 2023826-2020-03256 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098631&pc=MTA | 2023826-2020-03255 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171163&pc=MTA | 2023826-2021-02584 | 2020-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE CHANGE OVERTIME. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403865&pc=MTA | 2023826-2020-01724 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE SURGEON ALSO REPORTS BLURRED VISION. REPORTEDLY, THE PATIENT'S ACD WAS CALCULATED AND ENTERED INCORRECTLY, HOWEVER THE REMAINING REFRACTION IS STILL NOT EXPLAINABLE. THE LENS WAS A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2020 THE LENS WAS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNKNOWN ENTRY. MODEL# -UNK. DEVICE DATE-UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148282&pc=MTA | 2023826-2021-02429 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: B5: PREVIOUSLY STATED THE LENS WAS EXPLANTED AND THEN REPLACED ON (B)(6) 2020. CORRECTION: THE LENS WAS EXCHANGED ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620598&pc=MTA | 2023826-2021-00877 | 2020-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SHORTER LENGTH DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620605&pc=MTA | 2023826-2021-00876 | 2020-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE REC¿D BY MFR: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171147&pc=MTA | 2023826-2021-02585 | 2020-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322376&pc=MTA | 2023826-2021-03118 | 2020-08-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -11.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2015. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM IS RESOLVED. THE SURGEON REPORTS, "TRACE CELLS, GOOD VAULT AT 1-WEEK POSTOP" AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322379&pc=MTA | 2023826-2021-03117 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -11.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2015. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM IS RESOLVED. THE SURGEON REPORTS, "TRACE CELLS, GOOD VAULT AT 1-WEEK POSTOP" AFTER EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755799&pc=MTA | 2023826-2020-02463 | 2020-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A LONGER LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED TO: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.. CLAIM# 717567.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10938834&pc=MTA | 2023826-2020-02844 | 2020-09-28 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-12-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10812799&pc=MTA | 2023826-2020-02609 | 2020-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788143&pc=MTA | 2023826-2021-01393 | 2020-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTICS BROKEN,BENT,DEFORMED AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788188&pc=MTA | 2023826-2021-01394 | 2020-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536792&pc=MTA | 2023826-2021-00734 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL VERFICATION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10937930&pc=MTA | 2023826-2020-02779 | 2020-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON APR/2015. REFRACTIVE SURPRISE WITH GLARE/HALOES WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WITH GLARE/HALOES WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2015; CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11563239&pc=MTA | 2023826-2021-00762 | 2020-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH LENS WITH A -5.5 DIOPTER TO RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 -04.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE SURPRISE REPORTED. THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS/ HIGHER POWER DUE TO REFRACTIVE SURPRISE ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317799&pc=MTA | 2023826-2022-00205 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11710516&pc=MTA | 2023826-2021-01107 | 2020-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTER AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416066&pc=MTA | 2023826-2021-00564 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. AN "UNEXPECTED REFRACTIVE OUTCOME" AND A 25 DEGREES ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED, THAT A 13.2MM, MICL 13.2, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED. A NEW LENS IMPLANT IS PLANNED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -4.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT EXPERIENCED REFRACTIVE SURPRISE AND SHADOWING OF LETTERS. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE AND REPORTEDLY, A NEW LENS IMPLANT IS PLANNED. D6B- EPLANT DATE IS CORRECTED TO UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416418&pc=MTA | 2023826-2021-00513 | 2020-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, HIGHER POWER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "THERE IS NO PROBLEM WITH THE LENS MODEL AND LENS SIZE. THE PROBLEM WAS SOLVED BY SLIGHTLY CHANGING THE POWER OF THE LENS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544735&pc=MTA | 2023826-2021-03791 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -2.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135923&pc=MTA | 2023826-2021-00059 | 2020-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE SURPRISE WAS OBSERVED." SHOULD BE IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855179&pc=MTA | 2023826-2021-01509 | 2020-12-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO 12.1, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DATE RECD BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366536&pc=MTA | 2023826-2021-00479 | 2020-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) STATEMENT IN INITIAL MDR IS NOT APPLICABLE. SHOULD HAVE BEEN REPORTED IN PREVIOUS MDR FOR DIMENSIONAL/ FUNCTIONAL INVESTIGATION CONCLUSION. ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256327&pc=MTA | 2023826-2022-00151 | 2020-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11206969&pc=MTA | 2023826-2021-00112 | 2020-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE, LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS REPOSITIONED AND THE THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY". H6 - HEALTH EFFECT IMPACT CODE REPORTED AS 2199 - UPDATED/CHANGED TO 4629 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DB6 DATE: LENS NOT EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178442&pc=MTA | 2023826-2021-02711 | 2020-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207945&pc=MTA | 2023826-2021-02747 | 2021-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED AND STRETCHED OUT WITH FIBER ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942684&pc=MTA | 2023826-2021-01770 | 2021-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308762&pc=MTA | 2023826-2021-00308 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SAME LENGTH, LESSER POWER LENS ON (B)(6) 2021 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY PRODUCT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343195&pc=MTA | 2023826-2021-00398 | 2021-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REFRACTIVE CHANGE OVERTIME WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER LENS (-10.50 DIOPTER) DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942674&pc=MTA | 2023826-2021-01779 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11400353&pc=MTA | 2023826-2021-00514 | 2021-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON THE SAME DATE IN A SEPARATE SURGERY, A SAME SIZE DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11751223&pc=MTA | 2023826-2021-01284 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207967&pc=MTA | 2023826-2021-02748 | 2021-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620582&pc=MTA | 2023826-2021-00867 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378621&pc=MTA | 2023826-2021-03375 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207976&pc=MTA | 2023826-2021-02749 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683173&pc=MTA | 2023826-2021-01074 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR. (B)(6)) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID ALREADY BEFORE TRANS-PRK ELSWHERE)". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR. BECHMANN,SMILEYES) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID ALREADY BEFORE TRANS-PRK ELSEWHERE)". CLAIM # (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION CORRECTED TO "THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -00.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND PROBLEM NOT RESOLVED. THE REPORTER STATED "PATIENT HAS TRIED TO GET AN ENHANCEMENT THERAPY ELSEWHERE AND HAS GOT A TRANS-PRK WIT -1.00 (DR.BECHMANN,SMILEYES) THIS THERAPY HAD NO REFRACTIVE EFFECT SO SHE RETURNED TO US, ASKING FOR AN ENHANCEMENT. WE RECOMMENDED AN EXCHANGE (AS WE DID BEFORE TRANS-PRK ELSEWHERE.)". REFERENCE MFR FILE#(B)(4) OF THE REPLACEMENT LENS CLAIM." IN SUPPLEMENTAL MDR#1. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11781991&pc=MTA | 2023826-2021-01385 | 2021-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT , IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11782012&pc=MTA | 2023826-2021-01386 | 2021-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A MODEL AND DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PROODUCT , IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942685&pc=MTA | 2023826-2021-01769 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK: RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457492&pc=MTA | 2023826-2021-03552 | 2021-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN WITH FIBERS ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SDU A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12208015&pc=MTA | 2023826-2021-02759 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524257&pc=MTA | 2023826-2021-03720 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A2: DOB: (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524286&pc=MTA | 2023826-2021-03719 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: A2: DOB (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072245&pc=MTA | 2023826-2021-02086 | 2021-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072278&pc=MTA | 2023826-2021-02085 | 2021-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544805&pc=MTA | 2023826-2021-03793 | 2021-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, LENGTH AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072479&pc=MTA | 2023826-2021-02033 | 2021-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS LATER EXPLANTED DUE TO VISUAL ACUITY NOT BEING AS EXPECTED AFTER SURGERY. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE OUT OF SPECIFICATION. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4): NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986341&pc=MTA | 2023826-2021-01874 | 2021-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED NIGHT VISION DUE TO LARGE PUPIL SIZE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK, DATE OF EVENT-UNK, PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK, CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791894&pc=MTA | 2023826-2021-04319 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A REPLACEMENT LENS OF DIFFERENT POWER WAS LATER IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS AS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION WAS PERFORMED, AND THE DIAMETER VERIFICATION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION PERFORMED FOUND THE LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257193&pc=MTA | 2023826-2021-02977 | 2021-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED BLURRED VISION. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524316&pc=MTA | 2023826-2021-03722 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7 -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION:LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170553&pc=MTA | 2023826-2021-02698 | 2021-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT DIOPTER AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. TYPE OF INVESTIGATION 3331 DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222316&pc=MTA | 2023826-2021-02809 | 2021-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/ BROKEN/BENT AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872779&pc=MTA | 2023826-2021-04511 | 2021-06-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS "OTHER", FOR CAUSE DETAILS THE REPORTER STATED " WRONG REFRACTION BY OPTOMETRIST." IT WAS REPORTED THAT THERE WAS AN ERROR IN THE PRE-OPERATIVE REFRACTION BEING THE PATIENT POST-PKP. OPTOMETRIST REPEATED REFRACTION TWICE BEFORE ORDERING THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178441&pc=MTA | 2023826-2021-02712 | 2021-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12644985&pc=MTA | 2023826-2021-04019 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM6_12.6,-8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. FOR PATIENTS CURRENT CONDITION "FAR SIGHT NOT GOOD" WAS REPORTED. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT, FOR CAUSE DETAILS THE REPORTED STATED " DROP OF VISION AFTER VICM6 SURGERY." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATE INFORMATION THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. ON (B)(6) 2021 A REPLACEMENT LENS OF DIFFERENT MODEL AND DIOPTER WAS IMPLANTED AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: SHOULD BE CORRECTED TO STATE: FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378568&pc=MTA | 2023826-2021-03363 | 2021-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT DIOPTER LENS-BOTH REFRACTIVE CHANGE OVER TIME AND REFRACTIVE SURPRISE ARE REPORTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY REPORTED REFRACTIVE CHANGE OVER TIME IN ERROR. THIS IS A CASE OF REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222324&pc=MTA | 2023826-2021-02814 | 2021-06-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725063&pc=MTA | 2023826-2021-04166 | 2021-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082453&pc=MTA | 2023826-2021-04969 | 2021-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A LONGER LENGTH, DIFFERENT SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426117&pc=MTA | 2023826-2021-03444 | 2021-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER STATES A VISIAN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'SEYE. THE SURGEON REPORTS A POSTOP REFRACTION OF +2.25. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: A 13.2MM MICL13.2 WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED "PATIENT COULD IMMEDIATELY TELL THAT VISION WAS MUCH BETTER." THE SURGEON REPORTS HYPEROPIC ENDPOINT AFTER CORRECT CALCULATION. REPORTEDLY, "HYPEROPIC REFRACTIVE SURPRISE. ALL CALCULATIONS DONE AS INDICATED. IMPLANTED CORRECT ICL. CYCLO REFRACTION SHOWED +2.25 SPHERICAL WITH A DRY REFRACTION OF +1.75 SPHERICAL OS." H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725238&pc=MTA | 2023826-2021-04178 | 2021-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL WITH . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566510&pc=MTA | 2023826-2021-03842 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, INSIDE MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523480&pc=MTA | 2023826-2021-03690 | 2021-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER STATES A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE ERROR. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND OTHER: "CLINICAL REFRACTIONS" AND "OVER-MINUSED IN CLINIC WITH REFRACTION." REPORTEDLY THE DEVICE DID NOT FAIL TO PERFORM AS EXPECTED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725195&pc=MTA | 2023826-2021-04172 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -14.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT LENGTH AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12909815&pc=MTA | 2023826-2021-04616 | 2021-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -3.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840513&pc=MTA | 2023826-2021-04393 | 2021-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION : THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATONS. CLAIM # (B)(4). Manufacturer Narrative: B5: LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246071&pc=MTA | 2023826-2022-01311 | 2021-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246554&pc=MTA | 2023826-2022-01312 | 2021-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. UNABLE TO IDENTIFY, WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN. UNABLE TO IDENTIFY, WHICH LENS PERTAINS TO, WHICH CLAIM AS THEY WERE RETURNED IN THE SAME BAG. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14659120&pc=MTA | 2023826-2022-01745 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH AN IDENTICAL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146229&pc=MTA | 2023826-2022-00053 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER, REFRACTIVE ERROR".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER, REFRACTIVE ERROR". Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911555&pc=MTA | 2023826-2022-00961 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -15.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182540&pc=MTA | 2023826-2022-00029 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182541&pc=MTA | 2023826-2022-00030 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE LENS WAS REPLACED WITH A TICL VERTICALLY PLACED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983927&pc=MTA | 2023826-2022-01048 | 2021-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE PATIENT IS WEARING GLASSES. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS IMPLANTED ON (B)(6) 2021 AND NOT ON (B)(6) 2021 AS INITIALLY REPORTED. IT WAS REPORTED THAT THE PATIENT CONTINUES WITH UNEXPECTED RESIDUAL OF -1.25 IN HER LAST CONTROL. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973349&pc=MTA | 2023826-2021-04699 | 2021-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "ASTIGMATISM AND KERATOMETRY FROM THE START." THE REPORTER STATES THAT POSSIBLY THE SURGEON UNDERESTIMATED THE ASTIGMATISM-KERATOMETRY GIVES A CYLINDER OF 1.4. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459224&pc=MTA | 2023826-2022-00389 | 2021-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR REP# 2023826-2022-00392 FOR REPLACEMENT LENS. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317681&pc=MTA | 2023826-2022-00204 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -14.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862585&pc=MTA | 2023826-2022-01899 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173031&pc=MTA | 2023826-2022-00099 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND ALSO REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15738720&pc=MTA | 2023826-2022-03992 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THE PATIENT HAD PERMANENT LOSS OF BCVA AND BLURRED VISION. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "RESIDUAL ASTIGMATISM REMAINS". THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT MODEL (TORIC) LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE COMMENTS INDICATED "GOOD APPEARANCE OF THE ANTERIOR CHAMBER MENISCUS NORMAL BUT NORMAL". Manufacturer Narrative: ADDITIONAL INFORMATION: D6B: EXPLANT DATE: (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862904&pc=MTA | 2023826-2022-01905 | 2021-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -2.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13981610&pc=MTA | 2023826-2022-01046 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998315&pc=MTA | 2023826-2022-00978 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE AND GHOSTING. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862403&pc=MTA | 2023826-2022-02091 | 2022-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504752&pc=MTA | 2023826-2022-01688 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN MOISTURE WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795182&pc=MTA | 2023826-2022-00797 | 2022-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRY/FOGGY VISION WHEN THE LIGHT CONDITIONS GO FROM DARK TO LIGHT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944339&pc=MTA | 2023826-2022-01025 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE WITH A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-02285 FOR REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944425&pc=MTA | 2023826-2022-01024 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2021. THE SURGEON REPORTS REFRACTIVE SURPRISE WITH A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH AND MODEL BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL AND FUNCTIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: "THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND "BLOOD STAINS" ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS." CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365578&pc=MTA | 2023826-2022-01518 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT -9.0 DIOPTER. THE EXCHANGE LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16720053&pc=MTA | 2023826-2023-01357 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON 11-FEB-2023 THE LENS WAS EXCHANGED WITH THE SAME LENGTH LENS BUT TORIC MODEL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220682&pc=MTA | 2023826-2022-01098 | 2022-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM6_12.1 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE REPORTED WITH -1.50 SPHERE RESIDUAL ERROR. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220740&pc=MTA | 2023826-2022-01099 | 2022-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM6_12.1 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE REPORTED WITH -1.75 SPHERE RESIDUAL ERROR. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862698&pc=MTA | 2023826-2022-02025 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801811&pc=MTA | 2023826-2022-03954 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022, THE LENS WAS REMOVED AND REPLACED DUE TO REFRACTIVE SURPRISE. Manufacturer Narrative: DEVICE EVALUATED BY MFR: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND EXCHANGED ON (B)(6) 2022. IT WAS REPORTED AS USER ERROR AND, "WRONG LENS ORDERED" WAS STATED. ALSO, STATUS OF THE EYE WAS REPORTED AS, "ISSUE RESOLVED." CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR THE SAME MODEL/ LENGTH LENS; MORE MYOPIC POWER AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS " ISSUE RESOLVED". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY "WRONG LENS ORDERED". DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14417748&pc=MTA | 2023826-2022-01591 | 2022-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS PERMANENT LOSS OF BCVA AND IT WAS NOTED "PATIENT HAS COMPLAINTS OF POOR FAR AND NEAR VISION". LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAS PERMANENT LOSS OF BCVA AND IT WAS NOTED "PATIENT HAS COMPLAINTS OF POOR FAR AND NEAR VISION". ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14455997&pc=MTA | 2023826-2022-01621 | 2022-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT HAS COMPLAINTS IN LEFT EYE (OS) OF POOR FAR AND NEAR VISION POST VIVA IMPLANTABLE COLLAMER LENS IMPLANTATION. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VICM6 12.6 -6.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. PERMANENT LOSS OF BCVA WAS OBSERVED. LENS REMAINS IMPLANTED, PATIENT HAS NOT DECIDED FOR AN EXCHANGE. IN THE REPORTER OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM #(B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6: INVESTIGATION TYPE 4110 ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153652&pc=MTA | 2023826-2022-02582 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN/BROKEN HAPTIC. BLOOD STAINS WERE OBSERVED ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE: SHOULD BE CORRECT TO 31-JUL-2024. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15061062&pc=MTA | 2023826-2022-02329 | 2022-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGE WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND IT WAS NOTED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423702&pc=MTA | 2023826-2023-00619 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC MODEL AND SAME SIZE LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862210&pc=MTA | 2023826-2022-01907 | 2022-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE REPORTED. REPORTEDLY RESIDUAL ASTIGMATISM WAS CORRECTED BY INCISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697668&pc=MTA | 2023826-2022-03892 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO PRODUCT NON-CONFORMITY. REPORTEDLY, "FELLOW EYE LENS VAULTING WAS EXPLANTED DUE TO EXCESSIVE INFLAMMATION AND CELLS POST OP, PATIENT IS TRAUMATISED WITH THE WHOLE EXPERIENCE, PATIENT CURRENTLY ON GLAUCOMA MEDICATIONS G GANFORT ON AND G SIMBRINZA TDS ON RE." CAUSE OF THE EVENT IS UNKNOWN. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: MODEL NUMBER- SHOULD BE CORRECTED TO STATE VICM5_12.1. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298835&pc=MTA | 2023826-2022-02920 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 03-31-2025. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237726&pc=MTA | 2023826-2022-02810 | 2022-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS OF 4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION : LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17228874&pc=MTA | 2023826-2023-02567 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE TREATMENT WAS PERFORMED ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15276405&pc=MTA | 2023826-2022-02828 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: B5: CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE- 11/30/2024. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822417&pc=MTA | 2023826-2023-04280 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (TORIC LENS) AND POWER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. PER REPORTER THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED "PATIENT DIDN'T REACH EXPECTED VA" WAS REPORTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921202&pc=MTA | 2023826-2022-04343 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18101043&pc=MTA | 2023826-2023-05054 | 2022-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, SLIGHTLY MORE HYPEROPIC THAN PLANNED, UDVA IS 1.0. PATIENT IS DISSATISFIED WITH NEAR VISION AND COMPUTER DISTANCE, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15160364&pc=MTA | 2023826-2022-02614 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT POWER. IT WAS ADDITIONALLY NOTED AFTER INITIAL IMPLANTATION OF POSTOPERATIVE VISION; IT COULD NOT MEET THE DAILY WORK AND LIFE NEEDS; AND WAS DECIDED TO BE EXCHANGED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882900&pc=MTA | 2023826-2022-04253 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -14.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712752&pc=MTA | 2023826-2022-03897 | 2022-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712753&pc=MTA | 2023826-2022-03898 | 2022-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -8.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728673&pc=MTA | 2023826-2023-01334 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR DUE TO REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712749&pc=MTA | 2023826-2022-03896 | 2022-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO TRANSIENT LOSS OF BCVA. IN A SUBSEQUENT SURGERY LATER THAT DAY, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: CORRECTED TO - 07640145865114 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849481&pc=MTA | 2023826-2022-04262 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0. DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH A DIFFERENT DIOPTER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. D6B: ON (B)(6) 2022. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937239&pc=MTA | 2023826-2022-04285 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE WAS REPORTED AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT A DIFFERENT DIOPTER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3: DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR IS CORRECTED TO 07-NOV-2022. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577434&pc=MTA | 2023826-2022-03585 | 2022-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE DURING THE POSTOPERATIVE EXAMINATION WITH THE PATIENT. ON (B)(6) 2022 THE SURGEON EXCHANGED THE LENS WITH A SAME LENGTH/MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE SURGEON SUSPECTS THE LENS DIOPTER WAS INCORRECT WITH THE INITIAL IMPLANTED LENS AND WANTS INVESTIGATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, RETURNED WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301418&pc=MTA | 2023826-2023-00379 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: LENS RETURNED 20-FEB-2023. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154822&pc=MTA | 2023826-2023-00129 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODE AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9:(B)(6)2022. DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154821&pc=MTA | 2023826-2023-00131 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE BUT DIFFERENT DIOPTER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 29-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16414710&pc=MTA | 2023826-2023-00520 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REPLACED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "POSTOPERATIVELY, THE PATIENT HAS DIFFICULTY OPENING HER EYES EVEN AFTER MYDRIASIS, AND ALTHOUGH THE OD HAS NOT REPLACED LENSES, PATIENT'S VISUAL ACUITY HAS INCREASED FROM 0.8 TO 1.2 WHICH I WONDER MAY HAVE A PSYCHOLOGICAL COMPONENT." ALSO, "I CONSIDER THAT POOR CORNEAL CONDITION SEEMS TO HAVE HAD A SIGNIFICANT IMPACT ON LENS POWER." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16828266&pc=MTA | 2023826-2023-01532 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17338886&pc=MTA | 2023826-2023-02982 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -09.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THE PATIENT HAD VISUAL ACUITY DISSATISFACTION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344588&pc=MTA | 2023826-2023-00474 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT - CLAIM# 727957 FOR FELLOW EYE (OD). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16356403&pc=MTA | 2023826-2023-00475 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT CLAIM#: (B)(4) FOR FELLOW EYE (OS). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464988&pc=MTA | 2023826-2023-00719 | 2022-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16150999&pc=MTA | 2023826-2023-00139 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119689&pc=MTA | 2023826-2023-02165 | 2022-12-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6 12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 27/NOV/2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. REPORTEDLY, LENS MAYBE EXCHANGED. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119700&pc=MTA | 2023826-2023-02166 | 2022-12-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. REPORTEDLY, LENS MAYBE EXCHANGED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250313&pc=MTA | 2023826-2023-00320 | 2023-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. IT WAS ADDITIONALLY NOTED "PATIENT HAPPY WITH DISTANCE VISION BUT DISSATISFIED WITH INTERMEDIATE AND NEAR VISION". LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16250339&pc=MTA | 2023826-2023-00319 | 2023-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -10.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE; IT WAS ADDITIONALLY NOTED "PATIENT HAPPY WITH DISTANCE VISION BUT DISSATISFIED WITH INTERMEDIATE AND NEAR VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17121541&pc=MTA | 2023826-2023-02331 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2023. THE PATIENT EXPERIENCED BLURRED VISION WITH VISION QUALITY ISSUES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281081&pc=MTA | 2023826-2023-00403 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -5.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED; "DUE TO THE LOW DEGREE OF ASTIGMATISM OF THE PATIENT BEFORE SURGERY -- 0.75D, THE PATIENT DID NOT CHOOSE ASTIGMATISM CORRECTION IN THE SURGICAL DESIGN AFTER PREOPERATIVE COMMUNICATION WITH THE PATIENT, AND THE PATIENT SHOWED VISUAL SENSITIVITY AND DISCOMFORT AFTER SURGERY, SO THE PATIENT REPORTED POOR CORRECTED VISION WITHOUT ADDING ASTIGMATISM CORRECTION, AND HIS VISION WAS RESTORED TO THE BEST CORRECTED VISION AFTER REPLACING THE ASTIGMATISM CORRECTION CRYSTAL!". Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281082&pc=MTA | 2023826-2023-00395 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED; "DUE TO THE LOW DEGREE OF ASTIGMATISM OF THE PATIENT BEFORE SURGERY -- 0.75D, THE PATIENT DID NOT CHOOSE ASTIGMATISM CORRECTION IN THE SURGICAL DESIGN AFTER PREOPERATIVE COMMUNICATION WITH THE PATIENT, AND THE PATIENT SHOWED VISUAL SENSITIVITY AND DISCOMFORT AFTER SURGERY, SO THE PATIENT REPORTED POOR CORRECTED VISION WITHOUT ADDING ASTIGMATISM CORRECTION, AND HIS VISION WAS RESTORED TO THE BEST CORRECTED VISION AFTER REPLACING THE ASTIGMATISM CORRECTION CRYSTAL!". Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678113&pc=MTA | 2023826-2023-01214 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -2.75 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON 16 MAY 2023 FOR A SAME MODEL DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892521&pc=MTA | 2023826-2023-01748 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A DEFORMED OPTIC AND THE HAPTICS BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131352&pc=MTA | 2023826-2023-02389 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED REFRACTIVE TREATMENT (LASIK/PRK/ETC). THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "ALL IS FINE AND PATIENT IS HAPPY, LASER TOUCH UP TREATMENT ON (B)(6) 2023, HAS CORRECTED THE POST ICL" THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066600&pc=MTA | 2023826-2023-02225 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -8.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH, BUT DIFFERENT MODEL/POWER WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792960&pc=MTA | 2023826-2023-01479 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BEFORE REPLACING THE FIRST LENS, THE DOCTOR PERFORMED LENS ALIGNMENT AND ROTATION IN THE EYE FOUND THAT THE ASTIGMATIC AXIS OF THE PATIENT'S SUBJECTIVE OPTOMETRY HAD CHANGED. THEREFORE, IT IS SUSPECTED THAT THE FIRST LENS IMPLANTED IS AN TICL." Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678778&pc=MTA | 2023826-2023-01178 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6:EXPLANT DATE (B)(6) 2023. B5-ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: PT INFORMATION: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17696360&pc=MTA | 2023826-2023-03777 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971404&pc=MTA | 2023826-2023-04572 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO REFRACTIVE SUPRISE. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17992929&pc=MTA | 2023826-2023-04742 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748688&pc=MTA | 2023826-2023-01354 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16746899&pc=MTA | 2023826-2023-01350 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS ON (B)(6) 2023 AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761001&pc=MTA | 2023826-2023-04001 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16810401&pc=MTA | 2023826-2023-01482 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS "GOOD". Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. UNABLE TO PERFORM LENS WIDTH DUE TO HAPTIC DAMAGE, LENSES HAVE BEEN SWITCHED IN YELLOW BAGS WHILE RETURNING. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNKNOWN. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16810453&pc=MTA | 2023826-2023-01481 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS "GOOD." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. UNABLE TO PERFORM LENS WIDTH DUE TO HAPTIC DAMAGE, LENSES HAVE BEEN SWITCHED IN YELLOW BAGS WHILE RETURNING. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592020&pc=MTA | 2023826-2023-03620 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF A -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A2- " ON (B)(6) 2023" SHOULD BE REMOVED AND " ON (B)(6) 1989" SHOULD BE ADDED. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954043&pc=MTA | 2023826-2023-04567 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SUPRISE AND REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH THE TORIC MODEL AND STRONGER MYOPIC POWER, SAME LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. 6B: "ON (B)(6) 2023" SHOULD BE REMOVED AND "ON (B)(6) 2023" SHOULD BE ADDED. DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213970&pc=MTA | 2023826-2023-02554 | 2023-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17214401&pc=MTA | 2023826-2023-02555 | 2023-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.50 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991214&pc=MTA | 2023826-2023-01769 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A DIFFERENT POWER TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL . VISUAL INSPECTION FOUND THE OPTIC TORN AND SPLIT AND THE HAPTIC TORN AND BROKEN. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DIMENSIONAL INSPECTION COULD NOT BE PERFORMED DUE TO EXCESSIVE LENS DAMAGE. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410063&pc=MTA | 2023826-2023-03193 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434430&pc=MTA | 2023826-2023-03192 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477568&pc=MTA | 2023826-2023-03350 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PER REPORTER PATIENT WAS DISSATISFIED WITH VA. ON (B)(6) 2023 LENS WAS EXCHANGED FOR TORIC LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17218627&pc=MTA | 2023826-2023-02556 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -13.50/1.5/062 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4) Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840230&pc=MTA | 2023826-2023-04285 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17330141&pc=MTA | 2023826-2023-02928 | 2023-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME LENGTH BUT TORIC LENS/DIFFERENT POWER AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION, LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148654&pc=MTA | 2023826-2023-02406 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -3.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. ON (B)(6)2023 THE LENS WAS EXPLANTED AND LATER ON THIS DATE REPLACED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17818024&pc=MTA | 2023826-2023-04288 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN AND OPTIC SPLIT. UNABLE TO PERFORM DIMENSIONAL/FUNCTIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477662&pc=MTA | 2023826-2023-03352 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BROKEN/BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840052&pc=MTA | 2023826-2023-04284 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18803156&pc=MTA | 2023826-2024-00837 | 2023-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2024, THE LENS WAS EXCHANGED FOR SAME SIZE, DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4: A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: 734061.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633295&pc=MTA | 2023826-2023-03533 | 2023-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO 13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - PRODUCT WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166188&pc=MTA | 2023826-2023-05169 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE IN A PATIENT WHO HAD SURGERY IN LATE MARCH/EARLY (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.5/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED LASIK SURGERY TO REDUCE REMAINING CYLINDER POST ICL AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166188&pc=QCB | 2023826-2023-05169 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THERE IS REFRACTIVE SURPRISE IN A PATIENT WHO HAD SURGERY IN LATE MARCH/EARLY (B)(6) 2023. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.5/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON PERFORMED LASIK SURGERY TO REDUCE REMAINING CYLINDER POST ICL AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862008&pc=MTA | 2023826-2023-04063 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2021. REFRACTIVE CHANGE OVERTIME AND LOW VAULT WERE OBSERVED(ASSESSED ON (B)(6) 2022) . LENS WAS EXCHANGE WITH A DIFFERENT POWER LENS ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455639&pc=MTA | 2023826-2023-03219 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "REPLACE ASTIGMATIC CRYSTALS BASED ON POSTOPERATIVE OPTOMETRY RESULTS." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455641&pc=MTA | 2023826-2023-03220 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "REPLACE ASTIGMATIC CRYSTALS BASED ON POSTOPERATIVE OPTOMETRY RESULTS." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358296&pc=MTA | 2023826-2023-02969 | 2023-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF A -16.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358386&pc=MTA | 2023826-2023-02968 | 2023-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF A -16 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17874320&pc=MTA | 2023826-2023-04417 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM612.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581258&pc=MTA | 2023826-2023-03589 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO TRANSIENT LOSS OF BCVA . THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581262&pc=MTA | 2023826-2023-03588 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO TRANSIENT LOSS OF BCVA . THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17545681&pc=MTA | 2023826-2023-03517 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761235&pc=MTA | 2023826-2023-04075 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395190&pc=MTA | 2023826-2023-05776 | 2023-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18370009&pc=MTA | 2023826-2023-05621 | 2023-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. "DOCTOR CALCULATED LENS WRONGLY." Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375462&pc=MTA | 2023826-2023-05765 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -1.25 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(3) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(3) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560295&pc=MTA | 2023826-2024-00361 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17694988&pc=MTA | 2023826-2023-03947 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT HAS REFRACTIVE SURPRISE. THE SURGEON PERFORMED LRI SURGERY. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS REPORTED AS A USER ERROR WITH NO DEVICE CAUSALITY AND NOTED "SURGEON PLANNED ASTIGMATISM CORRECTION WITH LRI WHICH DOESN'T WORK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H1 - TYPE OF EVENT - SERIOUS CORRECTED TO MALFUNCTION. H4 - MANUFACTURE DATE: 31-MAR-2026 CORRECTED TO 05-APRIL-2023. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17694989&pc=MTA | 2023826-2023-03948 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 20-JUL-2023. IT WAS REPORTED THAT THE PATIENT HAS REFRACTIVE SURPRISE. THE SURGEON PERFORMED LRI SURGERY. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS REPORTED AS A USER ERROR WITH NO DEVICE CAUSALITY AND NOTED "SURGEON PLANNED ASTIGMATISM CORRECTION WITH LRI WHICH DOESN'T WORK". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H1 - TYPE OF EVENT: SERIOUS CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271396&pc=MTA | 2023826-2023-05501 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17891552&pc=MTA | 2023826-2023-04271 | 2023-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. EXCESSIVE VAULT WAS REPORTED AND BLURRED VISION. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18233713&pc=MTA | 2023826-2023-05366 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18057401&pc=MTA | 2023826-2023-04943 | 2023-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2M VICMO13.2; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18066236&pc=MTA | 2023826-2023-04855 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136064&pc=MTA | 2023826-2023-05082 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136489&pc=MTA | 2023826-2023-05058 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18478055&pc=MTA | 2023826-2024-00053 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416892&pc=MTA | 2023826-2023-05827 | 2023-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062883&pc=MTA | 2023826-2023-04964 | 2023-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS USER ERROR. THE REPORTER STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS USER ERROR. THE REPORTER STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4) Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323057&pc=MTA | 2023826-2023-05653 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED DUE TO REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2023 A REPLACEMENT LENS OF DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18478523&pc=MTA | 2023826-2024-00054 | 2023-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH TORIC LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271708&pc=MTA | 2023826-2023-05505 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258189&pc=MTA | 2023826-2023-05312 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258205&pc=MTA | 2023826-2023-05311 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT DIOPTER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE(ASSESSED ON (B)(6) 2023). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196960&pc=MTA | 2023826-2023-05303 | 2023-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 18- SEP-2023. REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18321863&pc=MTA | 2023826-2023-05652 | 2023-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS REPORTED, LENS REMAINS IMPLANTED. CAUSE OF EVEN WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18522597&pc=MTA | 2023826-2024-00230 | 2023-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -5.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11043675&pc=MTA | 2023826-2020-03081 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: ON 20 NOV 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE ENTITLED, 'POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS WITH CENTRAL PORT DESIGN IN 45 TO 55 YEARS OLD PATIENTS: ONE YEAR FOLLOW-UP.' THE STUDY ARTICLE REPORTS LASER TOUCH-UP OF RESIDUAL ERROR IN 5 EYES POST-ICL. ARTICLE STATES "5 EYES REQUIRED LASER TOUCH UP TO CORRECT RESIDUAL REFRACTIVE ERROR, BEING 2 EYES OF THE SAME PATIENT SUBMITTED TO PHOTOREFRACTIVE KERATECTOMY (PRK) AND 3 EYES OF TWO PATIENTS TO FEMTOSECOND LASER IN SITU KERATOMILEUSIS (FEMTOLASIK)." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7783466&pc=MTA | 2023826-2018-01203 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. DATE RECEIVED SHOULD BE CORRECTED TO 19-JUL-2018 IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908155&pc=MTA | 2023826-2018-01426 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.00/+3.0/080 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7908160&pc=MTA | 2023826-2018-01427 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.50/+3.0/X044 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S., (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7916208&pc=MTA | 2023826-2018-01451 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/4.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, AS OF THE DATE OF MDR REPORTING NO INTERVENTION HAS BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928342&pc=MTA | 2023826-2018-01467 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/X092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8143302&pc=MTA | 2023826-2018-01787 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/166 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND THE LENS REMAINS IMPLANTED. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS REMAINS IMPLANTED. THE PATIENTS CURRENT CONDITION IS VISUAL ACUITY IS NOT AS GOOD AS IT SHOULD BE BUT THE PATIENT CONDITION IS OK. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955093&pc=MTA | 2023826-2020-00892 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-04-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER REPORTED REFRACTIVE SURPRISE. "HAVE NOT SEEN THE PATIENT YET AS WE ARE WAITING FOR RESULTS TO STABILIZE." PATIENTS VISIT HAS BEEN POSTPONE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10523407&pc=MTA | 2023826-2020-02024 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENT'S LEFT EYE. PATIENT REPORTS "LEFT EYE IS NOT A CLEAR VISION." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856340&pc=MTA | 2023826-2020-02698 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS RESIDUAL REFRACTIVE ERROR WHICH HAD TO BE CORRECTED WITH LASIK. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10984792&pc=MTA | 2023826-2020-02859 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO PATIENTS LEFT EYE (OS). REFRACTIVE SURPRISE AND LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, "WE ACTUALLY JUST PREFORMED THE REPOSITION TODAY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087624&pc=MTA | 2023826-2020-03176 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11308938&pc=MTA | 2023826-2021-00226 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. ON (B)(6) 2020 REFRACTIVE TREATMENT (LASIK, PRK, ETC.) WAS PERFORMED. REPORTEDLY, "THE AMOUNT OF HIPER CORRECTION IS ONLY XPLAINED BY A MISTAKE IN THE LENS POWER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416119&pc=MTA | 2023826-2021-00565 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. AN "UNEXPECTED REFRACTIVE OUTCOME" AND A 25 DEGREES ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS MDR IS N/A AS IT IS A DUPLICATE CLAIM. REFER TO THE MDR# 2023826-2021-00634 OF THE ACTUAL CLAIM (B)(4). CLAIM # (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11779032&pc=MTA | 2023826-2021-01319 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2021. PATIENT HAS RECENTLY REPORTED CLOUDY VISION AND PRESSURE SENATION. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148271&pc=MTA | 2023826-2021-02474 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, THE LENS WAS "CHANGED" DUE TO REFRACTION ERROR. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840605&pc=MTA | 2023826-2021-04363 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE PATIENT STATES AN COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE PATIENT STATES "POST-ICL VISION LOOKING AROUND THE HOUSE-10, 15, 20, 30-FEET AWAY IS NOT GREAT. IT IS BLURRY MID-RANGE VISION IS ANNOYING." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12850386&pc=MTA | 2023826-2021-04450 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: HYPEROPIC OUTCOME AFTER SPHERICAL LENS IMPLANT WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459217&pc=MTA | 2023826-2022-00392 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+0.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS BLURRED VISION. REPORTEDLY, "PREPARE FOR FOLLOW-UP OBSERVATION AND VISUAL FUNCTION TRAINING." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY:UNK. SEX: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657017&pc=MTA | 2023826-2022-00665 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM6_13.7; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE POST OP REPORT INDICATES "SHE SAYS THAT SEEING IS BETTER OVERALL. ESPECIALLY IN THE DISTANCE. HOWEVER; SHE IS NOW USING READING GLASSES WITH +2.00DPT. I GAVE HER CONTACT LENSES WITH +0.50 DPT FOR THE LEFT EYE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT - "LENS WILL REMAIN INSIDE. THE PATIENT DID NOT SEE BETTER WITH CONTACT LENSES; ONLY IN THE INTERMEDIATE AREA. SO WE'RE GOING TO TELL HER IT'S BEST TO LEAVE IT AS IS". CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13657018&pc=MTA | 2023826-2022-00664 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE POST OP REPORT INDICATES "SHE SAYS THAT SEEING IS BETTER OVERALL. ESPECIALLY IN THE DISTANCE. HOWEVER; SHE IS NOW USING READING GLASSES WITH +2.00DPT. I GAVE HER CONTACT LENSES WITH +0.50 DPT FOR THE LEFT EYE". ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT "LENS WILL REMAIN INSIDE. THE PATIENT DID NOT SEE BETTER WITH CONTACT LENSES; ONLY IN THE INTERMEDIATE AREA. SO WE'RE GOING TO TELL HER IT'S BEST TO LEAVE IT AS IS". CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14048250&pc=MTA | 2023826-2022-01145 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE; HYPEROPIC OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14283152&pc=MTA | 2023826-2022-01402 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE ON (B)(6) 2022; AS A REPLACEMENT LENS TO A PREVIOUS LOW VAULTING AND CATARACT ISSUE. IT WAS MENTIONED THE PROBLEM WAS RESOLVED; HOWEVER NOTED "AFTER LENS REPLACEMENT THERE IS A DECREASE IN VISUAL ACUITY; BUT THE PATIENT IS UNDER OBSERVATION. IT IS EXPECTED THAT VISION WILL RECOVER IN THE FUTURE". LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-01401 FOR INITIAL LENS CLAIM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: (B)(4): LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14456995&pc=MTA | 2023826-2022-01374 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE A PATIENT EXPERIENCED A STREAK IN THE LEFT EYE (OS) AFTER ICL IMPLANTATION. THE STREAK MAY HAVE COME FROM HIS LPI. Manufacturer Narrative: PATIENT IDENTIFIER-RACE:UNK. DATE OF EVENT: UNK. OTHER #: IMPLANT-EXPLANT DATE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14457201&pc=MTA | 2023826-2022-01373 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE A PATIENT EXPERIENCED BLUR IN THE RIGHT EYE (OD) AFTER ICL IMPLANTATION. MONITORING AT THE MOMENT WITH CONSIDERATION OF LASIK. Manufacturer Narrative: PATIENT IDENTIFIER-RACE:UNK. DATE OF EVENT: UNK. OTHER #: IMPLANT DATE-EXPLANT: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14749667&pc=MTA | 2023826-2022-01944 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THERE WAS REFRACTIVE SURPRISE REPORTED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14749668&pc=MTA | 2023826-2022-01943 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2; -6.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS REFRACTIVE SURPRISE REPORTED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282555&pc=MTA | 2023826-2022-02879 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SURPRISE. STATUS OF THE EYE IS REPORTED AS "RESOLVED." CAUSE OF THE EVENT IS USER ERROR. REPORTEDLY, "NEEDED A TORIC RATHER THAN A SPHERICAL LENS." Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: G3: SUPP1 DATE: 10/04/2022. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15440641&pc=MTA | 2023826-2022-03146 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO HAD AN IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A -6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A PATIENT WHO HAD A 13.2MM VTICMO13.2 -16.50/1.5/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2022 EXPERIENCED A 6D MYOPIC SHIFT AT HIS 1 DAY POST-OP VISIT. PER UPDATED INFORMATION THE REPORTER STATES " THE PATIENT RETURNED FOR HIS 1 WK POST OP CHECK TODAY WITH PLANO REFRACTION AND PERFECT VISION. THERE IS NO ISSUE AT THIS TIME. I BELIEVE THE 1 DAY MYOPIC SHIFT SEEN ON EXAM WAS DUE TO USE OF MIOSTAT DURING SURGERY AND THIS EFFECT HAS RESOLVED". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15709881&pc=MTA | 2023826-2022-03909 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15709888&pc=MTA | 2023826-2022-03910 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801769&pc=MTA | 2023826-2022-04131 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -7.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON 12-JUL-2022 AND A LASIK REVISION ON (B)(6) 2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801803&pc=MTA | 2023826-2022-04130 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED INITIALLY ON (B)(6) 2022 AND A LASIK REVISION ON 27-JUL-2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801953&pc=MTA | 2023826-2022-04132 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2024. IT WAS INDICATED THAT THE PATIENT HAD RESIDUAL REFRACTION AFTER IMPLANTATION. LASIK SURGERY WAS PERFORMED ON (B)(6)2022. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15808829&pc=MTA | 2023826-2022-04009 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1MM VICM6 OF -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811139&pc=MTA | 2023826-2022-04010 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1MM VICM6 OF -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15831959&pc=MTA | 2023826-2022-04018 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -4.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED AND INCREASES OF CYL. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15834977&pc=MTA | 2023826-2022-03242 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 12.6 MM VICM612.6, -15.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S RIGHT EYE (OD) (B)(6) 2021. ON (B)(6) 2021, LVC WAS PERFORMED DUE TO RESIDUAL REFRACTION AND POOR UDVA. Manufacturer Narrative: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15835000&pc=MTA | 2023826-2022-03243 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM VICM612.6, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S LEFT EYE (OS) (B)(6) 2021. ON (B)(6) 2021, LVC WAS PERFORMED DUE TO RESIDUAL REFRACTION AND POOR UDVA. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946952&pc=MTA | 2023826-2022-04449 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S EYE ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S EYE ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY, "LENS WAS IMPLANTED UPSIDE DOWN". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND TORN. SUBSEQUENT IMPLANTATION OF NEW LENS WAS INITIATED BUT REPORTEDLY, "LENS GOT UNSTERILE AS IT FELL DOWN ON FLOOR AND THE SURGERY HAS TO BE ABORTED". PROBLEM NOT RESOLVED. WORK ORDER SEARCH FOUND NO SIMILAR EVENTS. PATIENT INJURY NOT REPORTED. D6: EXPLANT DATE (B)(6) 2023. H6: 4627-LENS EXPLANTED. CLAIM#(B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: B5: REPORTEDLY, "PATIENT IS NOT COMING AND IS HAPPY WITH THE VISION BILATERALLY." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE:03-MAY-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC BROKEN AND THE HAPTIC BROKEN. DUE TO SIGNIFICANT LENS DAMAGE DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE NOT POSSIBLE TO BE PERFORMED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010741&pc=MTA | 2023826-2022-04480 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010748&pc=MTA | 2023826-2022-04486 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION AND REFRACTIVE CHANGE WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: CORRECTION LENS DIOPTER IS -11.0/1.0/118 . CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16280256&pc=MTA | 2023826-2023-00396 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS), AND REFRACTIVE MISS WAS OBSERVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16280258&pc=MTA | 2023826-2023-00314 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD), AND REFRACTIVE MISS WAS OBSERVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16517890&pc=MTA | 2023826-2023-00844 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON AN UNKNOWN DATE THE LENS WAS EXPLANTED. D9: RETURN DATE: 30-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL L INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: LENS DIOPTER SHOULD BE CORRECTED TO: -7.0. ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. THE PROBLEM HAS NOT RESOLVED. ADDITIONAL INFORMATION, A RELATIVELY LARGE HYPEROPIC SHIFT WAS FOUND. REPORTEDLY, THE EXCHANGE IS GOING TO BE DONE IN (B)(6) BUT THE EXACT DATE HAS NOT BEEN DECIDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, "I CAN SEE MORE COMFORTABLY THAN IN THE PREVIOUS VISIT AND BEFORE THE EXCHANGE AND WILL WAIT AND SEE...ANOTHER EXCHANGE OR SOME INTERVENTION HAS NOT BEEN PLANNED". PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16517916&pc=MTA | 2023826-2023-00843 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 11-NOV-2022. THE SURGEON REPORTS THAT THERE WAS AN UNDESIRED REFRACTIVE OUTCOME. LENS REMAINS IMPLANTED. THE PROBLEM HAS NOT RESOLVED. REPORTEDLY, THE EXCHANGE IS GOING TO BE DONE IN (B)(6) BUT THE EXACT DATE HAS NOT BEEN DECIDED. CLAIM#(B)(4). Manufacturer Narrative: DIOPTER SHOULD BE CORRECTED TO STATE: -7.50. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON AN UNKNOWN DATE THE LENS WAS EXPLANTED. D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, "I CAN SEE MORE COMFORTABLY THAN IN THE PREVIOUS VISIT AND BEFORE THE EXCHANGE AND WILL WAIT AND SEE...ANOTHER EXCHANGE OR SOME INTERVENTION HAS NOT BEEN PLANNED". PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SIMILAR LENS" SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16563879&pc=MTA | 2023826-2023-00884 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON NOTES RESIDUAL ASTIGMATISM AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16610416&pc=MTA | 2023826-2023-01083 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON NOTES A POSSIBLE LENS EXPLANT AND REPLACEMENT DUE TO A REFRACTIVE MISS AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16646226&pc=MTA | 2023826-2023-01125 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "A PATIENT WHO HAS POST RESIDUAL ASTIGMATISM AFTER TORIC ICL.". Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16651024&pc=MTA | 2023826-2023-01142 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A EVO ICL LENS IMPLANTED INTO HIS RIGHT EYE (OD) AT THE END OF (B)(6) 2023. PATIENT IS DISSATISFIED WITH THE RESULT OF VISUAL ACUITY AND STATES THAT HE HAS READING PROBLEMS. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16651087&pc=MTA | 2023826-2023-01143 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A EVO ICL LENS IMPLANTED INTO HIS LEFT EYE (OS) AT THE END OF (B)(6) 2023. PATIENT IS DISSATISFIED WITH THE RESULT OF VISUAL ACUITY AND STATES HIS VISION IS BLURRY IN DISTANCE. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694625&pc=MTA | 2023826-2023-01185 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "A PATIENT WHO HAS POST RESIDUAL ASTIGMATISM AFTER TORIC ICL.". Manufacturer Narrative: AFTER FURTHER COMMUNICATION WITH REPORTER, INFORMATION PROVIDED WAS PREVIOUSLY REPORTED IN CLAIM# (B)(4). THERE IS NO COMPLAINT. CLAIM# (B)(4) Manufacturer Narrative: PATIENT INFORMATION:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787276&pc=MTA | 2023826-2023-01469 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT HAS REFRACTIVE SURPRISE/REFRACTIVE MISS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16812137&pc=MTA | 2023826-2023-01521 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO HAS HAD A ICL (IMPLANTABLE COLLAMER LENS) IMPLANTED FOR 10 YEARS HAS HAD A 3 DIOPTER SHIFT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16812171&pc=MTA | 2023826-2023-01522 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO HAS HAD A ICL (IMPLANTABLE COLLAMER LENS) IMPLANTED FOR 10 YEARS HAS HAD A 3 DIOPTER SHIFT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850695&pc=MTA | 2023826-2023-01615 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17219513&pc=MTA | 2023826-2023-02562 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17220197&pc=MTA | 2023826-2023-02563 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17607594&pc=MTA | 2023826-2023-03630 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023, DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "BLURRED VISION DUE TO RESIDUAL ASTIGMATISM. SUBJECTIVE REFRACTION AT FINAL VISIT. +0-1.00X130." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17667111&pc=MTA | 2023826-2023-03691 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT AN EVO IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. REFRACTIVE SURPRISE WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831307&pc=MTA | 2023826-2023-04209 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT COMPLAINED OF A BLURRY SPOT IN HER EYE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK, A4: UNK, A5: UNK, A6: UNK, B3: UNK, D4: EXPIRATION DATE UNK, D6A:UNK, D6B: UNK, H4: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18006738&pc=MTA | 2023826-2023-04777 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON AN UNKNOWN DATE.-1 MYOPIC WAS REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187772&pc=MTA | 2023826-2023-04974 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. RESIDUAL REFRACTIVE ERROR IS REPORTED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. - PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. - IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18421048&pc=MTA | 2023826-2023-05810 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -12.5 WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656495&pc=MTA | 2023826-2019-00942 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2019-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), BACK IN 2003/2004. THE PATIENT WAS EXPERIENCING REFRACTIVE CHANGE OVER TIME AND THE SURGEON WAS PLANNING TO EXPLANT AND EXCHANGE THE LENS. BUT THE SURGEON CHANGED HIS MIND AND DECIDED TO MONITOR THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656503&pc=MTA | 2023826-2019-00943 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2019-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), BACK IN 2003/2004. THE PATIENT WAS EXPERIENCING REFRACTIVE CHANGE OVER TIME AND THE SURGEON WAS PLANNING TO EXPLANT AND EXCHANGE THE LENS. BUT THE SURGEON CHANGED HIS MIND AND DECIDED TO MONITOR THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14713784&pc=MTA | 2023826-2022-01864 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.25 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!". ADDITIONAL INFORMATION: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14714037&pc=MTA | 2023826-2022-01865 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.25 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2021. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE!". ADDITIONAL INFORMATION: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15557432&pc=MTA | 2023826-2022-03451 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON OBSERVED REFRACTIVE CHANGE OVER TIME IN THE PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.5 DIOPTER, WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT ON (B)(6) 2018. ON (B)(6) 2022, REFRACTIVE CHANGE OVER TIME WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ASKING FOR EXCHANGE BECUASE OF MYOPICIZATION OF -1.00. SIZING IS OK." PLANNED EXCHANGED FOR (B)(6) 2023. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND WITH ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA SUPP1 G3: (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: B1: PRODUCT PROBLEM. B2: ENTRY SHOULD BE BLANK/VOID. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5 WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 07/NOV/2018. THE PATIENT REFRACTIVE CHANGE OVER TIME. ON 11/JAN/2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, BUT RESIDUE AND DEBRIS ON FOOTPLATE/HAPTIC OF LENS. CLAIM# (B)(4). Manufacturer Narrative: PT. INFO:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337180&pc=MTA | 2023826-2019-00255 | 2018-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +1.5/+4.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EYE HAS TOO MUCH REMAINING ASTIGMATISM. Manufacturer Narrative: ADDITIONAL INFORMATION. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +1.5/+4.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LENS ROTATION AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EYE HAS TOO MUCH REMAINING ASTIGMATISM. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337209&pc=MTA | 2023826-2019-00256 | 2018-12-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +2.5/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EYE HAS TOO MUCH REMAINING ASTIGMATISM. Manufacturer Narrative: ADDITIONAL INFORMATION/ CORRECTED DATA. DESCRIBE EVENT OR PROBLEM: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +2.5/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LENS ROTATION AND THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EYE HAS TOO MUCH REMAINING ASTIGMATISM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920600&pc=MTA | 2023826-2021-04619 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS REFRACTIVE SURPRISE AND LENS DISLOCATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: RECEIVED CQ ON (B)(6) 2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -17.50/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS (13.7MM VTICMO13.7 DIOPTER -17.0/+3.0/105) ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. PROBLEM RESOLVED. REPORTEDLY, "PATIENT IS HAPPY!" THE CAUSE OF THIS EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. CLAIM#: (B)(4) IS REGARDING EXCESSIVE VAULTING WHICH IS UNRELATED TO THIS COMPLAINT. NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO-UNK. DATE OF EVENT-UNK. PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. PRODUCT INFO, IMPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565675&pc=MTA | 2023826-2020-00016 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN 2017. PATIENT HAS COMPLAINED OF BLURRED VISION. LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2019, PATIENT MISSED PLANNED CHECK UP ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS INDICATED BY THE REPORTER THEY BELIEVE THE REASON OF THE BLURRED VISION WAS DUE TO LENS ROTATION. H6 - RESULT CODE: 213 SHOULD BE CORRECTED TO 114. H6 - DEVICE CODE: 2923 CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16834680&pc=MTA | 2023826-2023-01525 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON STATES THAT HE THINKS BUT DOES NOT KNOW IF THE LENS ROTATED 5-10 DEGREES. RESIDUAL ASTIGMATISM WAS ALSO REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372531&pc=MTA | 2023826-2019-00316 | 2018-10-26 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+5.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824424&pc=MTA | 2023826-2018-01262 | 2018-06-11 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.00/+6.0/015 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THERE WAS REDUCTION OF BEST CORRECTED VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4).. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765095&pc=MTA | 2023826-2018-01173 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.0/+2.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941306&pc=MTA | 2023826-2019-01580 | 2019-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941316&pc=MTA | 2023826-2019-01579 | 2019-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932874&pc=MTA | 2023826-2019-01548 | 2019-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THREE MONTHS LATER, THE PATIENT IS NOT SATISFIED WITH HIS VISUAL ACUITY AND COMPLAINS OF BLURRED VISION. THE LENS HAS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAS DECIDED TO WAIT A COUPLE OF MONTHS AND THEN WILL HAVE HIS VISION CHECKED AGAIN. TO DATE THE PATIENT IS FINE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10744678&pc=MTA | 2023826-2020-02533 | 2020-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE AND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10856354&pc=MTA | 2023826-2020-02692 | 2020-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS, THE SURGEON REPORTING EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND PERMANENT LOSS OF BCVA. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "1 PER VAULT, CHANGE ICL WITH 12.6 LENGTH." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343188&pc=MTA | 2023826-2021-00400 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. REPORTEDLY, "ADDITIAONAL REFRACTION INCLUDING CR WILL BE MEASURED. THEN, THE SURGEON WILL CONSIDER A REPLACEMENT SURGERY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11343198&pc=MTA | 2023826-2021-00399 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS WAS EXPLANTED ON (B)(6)2021 DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. REPORTEDLY, "ADDITIAONAL REFRACTION INCLUDING CR WILL BE MEASURED. THEN, THE SURGEON WILL CONSIDER A REPLACEMENT SURGERY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE.VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. THE WIDTH OF THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. UNABLE TO PERFORM DIAMETER VERIFICATION AS THE LENS WAS DAMAGED. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214206&pc=MTA | 2023826-2021-02660 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM613.2, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 . THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, VAULT IS OK NOW AND PATIENT IS MORE HAPPY WITH DISTANCE AND NEAR VISION. PATIENT UNDER OBSERVATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PATIENT IS SEEING MUCH BETTER. HAS TO GET USED SEEING IN DARK AND LIGHT CONDITIONS. REPORTEDLY, "IT SEEMS PROBLEM IS RESOLVED." CLAIM #(B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214232&pc=MTA | 2023826-2021-02659 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM613.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, VAULT IS OK NOW AND PATIENT IS MORE HAPPY WITH DISTANCE AND NEAR VISION. PATIENT UNDER OBSERVATION. Manufacturer Narrative: PATIENT IS SEEING MUCH BETTER. HAS TO GET USED SEEING IN DARK AND LIGHT CONDITIONS. REPORTEDLY, "IT SEEMS PROBLEM IS RESOLVED." CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220828&pc=MTA | 2023826-2022-01102 | 2022-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE REPORTED. LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS USER ERROR - ICL 13.7 INSTEAD OF 12.6. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN,RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS USER ERROR - ICL 13.7 INSTEAD OF 12.6. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: G3: DATE RECEIVED BY MANUFACTURER: 09/08/2022. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14649879&pc=MTA | 2023826-2022-01839 | 2022-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND DIFFERENT MODEL. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED "THE PATIENT'S RIGHT EYE HAD ASTIGMATISM OF -0.75D BEFORE OPERATION. AND ARCH HEIGHT IS LOW. THE ICL WITHOUT ASTIGMATISM WAS SELECTED DURING THE FIRST OPERATION. IT WAS FOUND THAT THE PATIENT HAD ASTIGMATISM OF -1.50D AFTER OPERATION. CONSIDERING THAT IT WAS CAUSED BY THE SURGICAL INCISION; THE ICL WITH ASTIGMATISM WAS SELECTED DURING THE SECOND OPERATION". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS IN A RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED; HAPTIC DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(6). Manufacturer Narrative: CORRECTED DATA: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862479&pc=MTA | 2023826-2022-01906 | 2022-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE REPORTED. REPORTEDLY RESIDUAL ASTIGMATISM WAS CORRECTED BY INCISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, "THE MAIN COMPLAINT OF THIS CASE IS THE LOW VAULT. DUE TO LOW VAULT AFTER THE INITIAL SURGERY, THE SURGEON DECIDED TO EXCHANGE WITH A LARGER SIZE . HE ALSO CONSIDERED CHANGING THE DIOPTER AT THE SAME TIME.' THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1583 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: A2- CORRECTED TO : 25-JAN-1997 IN INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109368&pc=MTA | 2023826-2023-00085 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109425&pc=MTA | 2023826-2023-00086 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS REPLACED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588754&pc=MTA | 2023826-2022-03592 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: UNK. DEVICE EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED AN 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE, IRIDOCORNEAL ANGLES ARE OPEN AND THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT WAS REPORTED AS "UNKNOWN". H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 15-FEB-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588810&pc=MTA | 2023826-2022-03593 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED EXCESSIVE VAULT AND REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE, IRIDOCORNEAL ANGLES ARE OPEN AND THE PATIENT IS HAPPY." THE CAUSE OF THE EVENT WAS REPORTED AS "UNKNOWN". H6: WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15810735&pc=MTA | 2023826-2022-04175 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -03.50 DIOPTER IMPLANTALE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR# 2023826-2022-04176 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15811248&pc=MTA | 2023826-2022-04177 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT NOW HAS EXPERIENCED LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 DIFFERENT MODEL LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865570&pc=MTA | 2023826-2023-01628 | 2023-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT, CLAIM# (B)(4) FOR PATIENT'S FELLOW EYE (OS). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865571&pc=MTA | 2023826-2023-01627 | 2023-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT, CLAIM# (B)(4) FOR PATIENT'S FELLOW EYE (OD). CLAIM (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17193864&pc=MTA | 2023826-2023-02529 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN CONTACT LENS CASE DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302269&pc=MTA | 2023826-2023-02759 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A,12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND BIG PUPIL, PHOTOPHOBIA, AND BLURRED VISION WERE OBSERVED. REPORTEDLY, "NO CORRELATION BETWEEN WTW, SULCUS-SULCUS AND ICL SIZE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER ALSO STATED, "PENDING EXCHANGE ICL." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344449&pc=MTA | 2023826-2023-02859 | 2023-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2023. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344451&pc=MTA | 2023826-2023-02860 | 2023-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2023. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PROBLEM NOT RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17591904&pc=MTA | 2023826-2023-03631 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. IT WAS INDICATED THAT THE PATIENT EXPERIENCED REFRACTIVE CHANGE WHICH WAS ASSESSED ON (B)(6) 2023, THE PATIENT ALSO HAD LOW VAULTING WITH NO ROTATION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT POWER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN WITH NO DEVICE CAUSALITY AND NOTED "PATIENT HAS NOTICED DECREASED VA OS OVER LAST 6 MONTHS SINCE (B)(6) 2023. Manufacturer Narrative: CORRECTED DATA: B5: IMPLANT DATE SHOULD BE CORRECTED TO: (B)(6) 2015. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017240&pc=MTA | 2023826-2023-04814 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986858&pc=MTA | 2023826-2023-04599 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH DIFFERENT POWER LENS (SPHERE/CYLINDER/AXIS) DUE TO REFRACTIVE SURPRISE AND LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11375703&pc=MTA | 2023826-2021-00501 |  | Malfunction | STAAR SUGICAL COMPANY | 2021-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER STATED THAT A 12.6MM, VICM5_12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN (B)(6) 2019, THE SURGEON NOTED REFRACTIVE SURPRISE OF -1.25D ALTHOUGH THEY EXPECTED 0D. IN ADDITION THE LENS WAS SMALL AND THEY NOTICED LOW VAULTING. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8261656&pc=MTA | 2023826-2019-00159 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: (OS) LEFT EYE: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON SEEN PATIENT WHO WAS SEEKING REFRACTIVE SURGERY. UPON EXAMINATION, THE PATIENTS UNCORRECTED DISTANCE VISUAL ACUITY WAS COUNTING FINGERS AT 2 METERS OF CORRECTED 20/22 IN EACH EYE THE PATIENT WAS MYOPIC AND PRESBYOPIA AND WAS THOROUGHLY EXPLAINED TO PATIENT BEFORE SURGERY. ON (B)(6) 2018, THE PATIENT HAD UNDERGONE REFRACTIVE SURGERY BY IMPLANTING A 12.6MM, VICMO12.6, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO HER LEFT EYE (OS). UDVA ON THE FIRST POST-OPERATIVE FOLLOW UP WAS 20/28. ON PATIENTS LAST FOLLOW UP, PATIENT ENDED UP HYPEROPIC WITH UDVA OF 20/50. REFRACTION WAS +0.50 SPH 0.75 CYL X 150. THE VAULT WAS ALSO MEASURED TO BE A BIT HIGH WHICH WAS AROUND 800 MICRONS. THE SURGEON WOULD LIKE TO EXCHANGE THE LENS WITH A SMALLER SIZE (12.1MM) WITH LENS POWER OF -8.00 DIOPTER. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURER DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8261800&pc=MTA | 2023826-2019-00160 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-01-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: (OD) RIGHT EYE: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON SEEN PATIENT WHO WAS SEEKING REFRACTIVE SURGERY. UPON EXAMINATION, THE PATIENTS UNCORRECTED DISTANCE VISUAL ACUITY WAS COUNTING FINGERS AT 2 METERS OF CORRECTED 20/22 IN EACH EYE THE PATIENT WAS MYOPIC AND PRESBYOPIA AND WAS THOROUGHLY EXPLAINED TO PATIENT BEFORE SURGERY. ON (B)(6) 2018, THE PATIENT HAD UNDERGONE REFRACTIVE SURGERY BY IMPLANTING A 12.6MM, VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO HER RIGHT EYE (OD). UDVA ON THE FIRST POET-OPERATIVE FOLLOW UP WAS 20/40. ON PATIENTS LAST FOLLOW UP, PATIENT ENDED UP HYPEROPIC WITH UDVA OF 20/100. REFRACTIVE WAS +1.75 SPH -1.00 CYL X20. THE VAULT ALSO MEASURED TO BE A BIT HIGH WHICH WAS AROUND 800 MICRONS. THE SURGEON WOULD LIKE TO EXCHANGE THE LENS WITH A SMALLER SIZE (12.1MM) WITH LENS POWER OF -8.00 DIOPTER. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE CODE 1583: CODE SHOULD BE ADDED TO ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491987&pc=MTA | 2023826-2020-01957 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN UNEXPECTED REFRACTIVE RESULT WITH ANTERIOR VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, 14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN UNEXPECTED REFRACTIVE RESULT WITH ANTERIOR VAULT. LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10931407&pc=MTA | 2023826-2020-02739 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT REFREACTIVE SURPRISE HAS BEEN OBSERVED IN PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REFRACTIVE SURPRISE AND EXCESSIVE VAULT WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - "EXCESSIVE VAULT" SHOULD BE CORRECTED TO "LOW VAULT" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654121&pc=MTA | 2023826-2021-00952 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS EXCESSIVE VAULT AND 20/40 VA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119243&pc=MTA | 2023826-2021-02332 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS 0 VAULT AND SLIGHTLY BLURRY VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: THE REASON FOR THE REPLACEMENT LENS IS "LOW ARCH AND HIGH". THE PATIENT WAS CONCERNED ABOUT THE RISK OF SURGERY. THE LAST REPLACEMENT LENS WAS AN ICL. THE ASTIGMATISM WAS CORRECTED BY LASER SURGERY. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457498&pc=MTA | 2023826-2021-03551 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S EYE. THE SURGEON REPORTS "300% VAULT" AND +2.00D HYPEROPIC OUTCOME. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE PATIENT CONTINUES TO IMPROVE AND SO THE DATA IS SHIFTING. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE : UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15964379&pc=MTA | 2023826-2022-04176 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -03.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022; AS A REPLACEMENT LENS. IT WAS INDICATED THAT THE PATIENT EXPERIENCED LOW VAULTING WITH RESIDUAL ASTIGMATISM. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: G2: "OTHER: ITALY" SHOULD BE CORRECTED TO "OTHER: FRANCE". CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739724&pc=MTA | 2023826-2023-01368 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THERE IS ZERO VAULTING WITH REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5_13.2 -6.50 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS THERE IS ZERO VAULTING WITH REFRACTIVE CHANGE OVER TIME. ON (B)(6) 2023, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924450&pc=MTA | 2023826-2023-01822 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.; -9.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH, DIFFERENT MODEL/POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 24-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163492&pc=MTA | 2023826-2023-02309 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REFRACTIVE SURPRISE AND EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484564&pc=MTA | 2023826-2023-03356 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PATIENTS VISION AFTER SURGERY WAS 20/50 +1 WITH VAULT AROUND 1000. AFTER 2 WEEKS WITH NO VISUAL IMPROVEMENT, THE LENS WAS EXCHANGED FOR A TORIC LENS OF SHORTER LENGTH AND DIFFERENT POWER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17894899&pc=MTA | 2023826-2023-04407 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT DIOPTER LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981074&pc=MTA | 2023826-2023-04657 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH REFRACTIVE CHANGE OVER TIME WAS REPORTED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.50 DIOPTER WAS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH REFRACTIVE CHANGE OVER TIME WAS REPORTED. ON (B)(6) 2023 THE LENS WAS EXCHANGE FOR A LONGER LENGTH AND THE PROBLEM WAS RESOLVED. D6B- 10-OCT-2023. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15890872&pc=MTA | 2023826-2022-04363 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814889&pc=MTA | 2023826-2023-04282 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862111&pc=MTA | 2023826-2022-01909 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision | N/A | 1 | Event Description: AN ARTICLE WAS RECEIVED ON (B)(6) 2022 ENTITLED 'DETERMINANTS OF SUBJECTIVE QUALITY OF VISION AFTER PHAKIC INTRAOCULAR LENS IMPLANTATION'. THE ARTICLE REPORTS THAT 1 PATIENT UNDERWENT ICL ROTATION (15 DEGREES OF AXIS) IN ONE EYE AFTER ONE MONTH. ANOTHER PATIENT UNDERWENT ICL EXCHANGE IN ONE EYE DUE TO LOW VAULT AND HYPEROPIC REFRACTIVE OUTCOME 16 MONTHS AFTER THE FIRST ICL IMPLANTATION. Manufacturer Narrative: PATIENT ID, AGE OR DOB, SEX, WEIGHT, ETHNICITY, RACE: UNK DATE OF EVENT: UNK MODEL #, OPERATOR OF DEVICE, EXPLANT DATE: UNK. INITIAL REPORTER NAME AND ADDRESS: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273844&pc=MTA | 2023826-2023-00370 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LOW VAULTING; LENS ROTATION AND REFRACTIVE CHANGE OVER TIME. LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - LENS WAS REPLACED ON AN UNKNOWN DATE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412314&pc=MTA | 2023826-2019-00423 | 2018-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -01.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT, ACD<3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8351772&pc=MTA | 2023826-2019-00274 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/4.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SUPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354792&pc=MTA | 2023826-2019-00275 | 2019-01-08 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/4.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. IT WAS REPORTED THAT REFRACTIVE SURPRISE WAS OBSERVED AND SURGEON HAS NOT YET MADE THE DECISION WITH PATIENT. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552452&pc=MTA | 2023826-2019-00649 | 2019-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRY VISION. THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: SERIAL NUMBER, EXP. DATE, UDI: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11352467&pc=MTA | 2023826-2021-00490 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412440&pc=MTA | 2023826-2019-02275 | 2019-10-29 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES LOSS OF BCVA AND REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988414&pc=MTA | 2023826-2020-00949 | 2020-04-09 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED AS AN EXCHANGE LENS. POST-OP, THE NEAR VISION WAS SLIGHTLY IMPROVED BUT STILL NOT IDEAL. THE PATIENT STILL EXPERIENCED A REFRACTIVE SURPRISE AND WILL BE MONITORED. THE LENS REMAINED IMPLANTED. SEE MFR REPORT # 2023826-2020-00948 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155602&pc=MTA | 2023826-2021-02441 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS SPHERE TO BE OUT OF SPECIFICATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15288032&pc=MTA | 2023826-2022-02900 | 2020-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. OFF LABE - ACD <3.0. (B)(4). Manufacturer Narrative: A2: CORRECTED TO - AGE AT TIME OF EVENT, DATE OF BIRTH: (B)(6) 1995 IN THE INITIAL MDR. B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." IN INITIAL MDR. D4: EXPIRATION DATE - 30-JUN-2022 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11453120&pc=MTA | 2023826-2021-00602 | 2020-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH TORIC LENS (-10.50/1.0/171 (SPHERE/CYLINDER/AXIS)) DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11697121&pc=MTA | 2023826-2021-01090 | 2020-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. REPORTEDLY AFTER EXCHANGE THE POSTOP VISUAL ACUITY WAS 1.2 AND NO GHOSTING WAS REPORTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11714206&pc=MTA | 2023826-2021-01112 | 2020-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222296&pc=MTA | 2023826-2021-02819 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -1.75 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. UNABLE TO PERFORM A DIMENSIONAL AND REFRACTIVE VERIFICATION SINCE LENS HAS SIGNIFICANT DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378560&pc=MTA | 2023826-2021-03376 | 2021-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072264&pc=MTA | 2023826-2021-02071 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON REPORTS, "THE PATIENT HAD POOR NEAR VISION AFTER THE FIRST LENS IMPLANTATION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599748&pc=MTA | 2023826-2021-03907 | 2021-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A LIQUID VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN LENS. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED WITHIN THE SAME LOT. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS, INCLUDING THE SUSPECTED DEVICE,HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COALLMER LENS; -15.0 DIOPTER; INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599734&pc=MTA | 2023826-2021-03917 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363349&pc=MTA | 2023826-2021-03286 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT WITH FIBERS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524337&pc=MTA | 2023826-2021-03717 | 2021-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1,-7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12874462&pc=MTA | 2023826-2021-04518 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND POWER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE RECEIVED BY MANUFACTURER: THIS PRODUCT IS NOT MARKETED IN THE US .(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13319427&pc=MTA | 2023826-2022-00244 | 2021-08-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -02.25 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRED VISON AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS, SEE MFR. REPORT # 2023826-2022-00190 FOR THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -02.25 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT COMPLAINED OF BLURRED VISON AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS, SEE MFR. REPORT # 2023826-2022-00190 FOR THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- PER EMAIL COMMUNICATION ON (B)(6), IT WAS REPORTED THAT "THE PATIENT DIDN'T HAVE 'BLURRED VISION' AFTER THE REPLACEMENT LENS WAS IMPLANTED". THIS EVENT WAS REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12903199&pc=MTA | 2023826-2021-04615 | 2021-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIEN'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13581963&pc=MTA | 2023826-2022-00591 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -4.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639725&pc=MTA | 2023826-2022-01509 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LENS OF DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION COULD NOT BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15462009&pc=MTA | 2023826-2022-03136 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE PATIENT IS 43 YEARS OLD. CONSIDERING THE PREBYOPIA PROBLEM, IN ORDER TO DELAY THE TIME OF WEARING READING GLASSES, HE COMMUNICATED REPEATEDLY. FINALLY, HE DECIDED TO RETAIN 50 DEGREES OF MYOPIA DURING THE SURGICAL DESIGN. HOWEVER, THERE IS A BIG DIFFERENCE BETWEEN POSTOPERATIVE CORRECTION OF VISION AND ESTIMATED CORRECTION OF VISION, WHICH CANNOT MEET THE PATIENT'S LONG-TERM NEEDS. AFTER CONSULTATION WITH THE SURGEON, HE DECIDED TO REPLACE IT WITH A CRYSTAL WITH AN UNRETAINED DEGREE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK.ACD<3.0 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15751843&pc=MTA | 2023826-2022-04040 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR TORIC LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H10- CLAIM# (B)(4) SHOULD BE CORRECTED TO CLAIM# (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE: 1494 OFF-LABEL USE (ACD<3.0MM). DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN/BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640034&pc=MTA | 2023826-2022-03702 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6, -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 18-MAY-2022. IT WAS REPORTED THAT LIMBAL RELAXING INCISIONS (LRIS) WERE MADE TO REDUCE ASTIGMATISM BUT IT DID NOT REDUCE THE EXPECTED AMOUNT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC LENS TO ADDRESS THE ASTIGMATISM (HIGH CYLINDER POWER) AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: DEVICE CODE 1494- OFF LABEL USE- PATIENT ACD<3.00MM.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640127&pc=MTA | 2023826-2022-03703 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6, -8.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT LIMBAL RELAXING INCISIONS (LRIS) WERE MADE TO REDUCE ASTIGMATISM BUT IT DID NOT REDUCE THE EXPECTED AMOUNT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC LENS TO ADDRESS THE ASTIGMATISM (HIGH CYLINDER POWER) AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: DEVICE CODE 1494- OFF LABEL USE- PATIENT ACD<3.00MM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993591&pc=MTA | 2023826-2022-04600 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM RESOLVED. THE REPORTER STATED " AFTER THE OPERATION, HER BINOCULAR VISION WAS 1.2, BUT SHE FELT UNCOMFORTABLE NEAR VISION, SO SHE ASKED HER DOMINANT EYE TO BE LOWERED BY 50 DEGREES. AT THE REPEATED REQUEST OF THE PATIENT, THE RIGHT EYE INTRAOCULAR LENS REPLACEMENT WAS PERFORMED ON THE MORNING OF NOVEMBER 5, 2022. THE POSTOPERATIVE VISUAL ACUITY WAS 1.0. THE PATIENT SAID THAT HE WAS VERY SATISFIED WITH THE CURRENT STATE, AND HE FELT COMFORTABLE LOOKING NEAR AND FAR." Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916214&pc=MTA | 2023826-2023-01799 | 2022-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -7.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2022. THE SURGEON REPORTS THERE WAS REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY NOTED "PATIENT HAS SEEN NO PROBLEMS AT CLOSE RANGE WITH SATISFACTORY RESULTS". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16622829&pc=MTA | 2023826-2023-01116 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; 13.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075853&pc=MTA | 2023826-2022-04748 | 2022-11-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM12.0 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 15-FEB-2013. REFRACTIVE CHANGE OVERTIME WAS OBSERVED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. H4: DEVICE MANUFACTURE DATE: 23-OCT-2012. H6: 1494- OFF-LABEL: AGE<21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16542287&pc=MTA | 2023826-2023-00794 | 2022-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE PATIENT'S AGE WAS OLD AND HAD DIFFICULTY SEEING CLOSE AFTER SURGERY. THE DIOPTER OF THE LENS WAS ADJUSTED." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-EHTNICITY :UNK. OFF LABEL - AGE <45. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347483&pc=MTA | 2023826-2023-02986 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347676&pc=MTA | 2023826-2023-02987 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A TORIC LENS OF DIFFERENT POWER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17075871&pc=MTA | 2023826-2023-02258 | 2023-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THE PATIENT HAD 1D ASTIGMATISM BEFORE IMPLANTATION, BUT DUE TO THE PRICE A VICMO MODEL WAS REQUIRED. AFTER IMPLANTATION THE POSTOPERATIVE VISUAL QUALITY WAS LOWER THAN EXPLANTED. ON (B)(6)2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: MODEL# SHOULD BE CORRECTED TO VICMO12.1 CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18212758&pc=MTA | 2023826-2023-05372 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH AND DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THERE WAS NO DEVICE CAUSALITY. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549360&pc=MTA | 2023826-2023-03448 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME SIZE , DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437836&pc=MTA | 2023826-2024-00030 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS DUE TO REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - AGE<21. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778122&pc=MTA | 2023826-2023-04109 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO POOR VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395378&pc=MTA | 2023826-2023-05777 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416672&pc=MTA | 2023826-2023-05825 | 2023-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416770&pc=MTA | 2023826-2023-05826 | 2023-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525783&pc=MTA | 2023826-2024-00175 | 2023-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS REPLACED WITH A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18496445&pc=MTA | 2023826-2024-00117 | 2023-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17393777&pc=MTA | 2023826-2023-03125 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER AS A REPLACEMENT LENS INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PERILIMABAL ON (B)(6) 2023, DUE TO BLURRED VISION AND "OD HORIZONTAL JUXTAPOSED METAMORPHIC EYE." CAUSE OF THE EVENT IS UNKNOWN AND THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "MIOCHOL WAS ON BACK ORDER MIOSTATE USED OU FOR INITIAL SURGERY. PT HAS RETINAL EVAL PENDING ALTHOUGH MOCULAR OCT SHOWS NO ABNORMALITIES WHEN COMPARED TO PREOP MOC OCT. APPT (B)(6) 2023 - RETINA." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PERILIMABAL. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18498823&pc=MTA | 2023826-2024-00209 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICH12.1; +4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE DOCTOR DID NOT CONSIDER THE ALTERATION IN REFRACTION THAT HAPPENED WHEN TRANSITIONING FROM THE ICH MODEL TO THE VICH MODEL. AS A RESULT THE PATIENT NOW HAS A REMAINING REFRACTION OF -0.75 -0.50X85. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8315215&pc=MTA | 2023826-2019-00230 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2019-02-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH12.6, 8.50/3.5/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LEN INTO PATIENTS RIGHT EYE (OD) IN 2013. ON (B)(6) 2019 THE SURGEON REPORTED MYOPIC PROGRESSION. CURRENTLY THE LENS REMAINS IMPLANTED. REQUESTS FOR ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8315217&pc=MTA | 2023826-2019-00231 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2019-02-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICH12.6, 8.50/4.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LEN INTO PATIENTS LEFT EYE (OS) IN 2013. ON (B)(6) 2019 THE SURGEON REPORTED THE PATIENT MYOPIC PROGRESSION. CURRENTLY THE LENS REMAINS IMPLANTED. REQUESTS FOR ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869317&pc=MTA | 2023826-2020-00627 | 2019-12-18 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-23 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Improper or Incorrect Procedure or Method | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; AN ACD VALUE OF 3.2MM WAS USED INSTEAD OF THE ACTUAL ACD VALUE OF 2.53MM IN ORDER TO CALCULATE THE LENS. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721909&pc=MTA | 2023826-2021-01028 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -07.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE ON (B)(6) 2018. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11773680&pc=MTA | 2023826-2021-01383 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -3.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH TORIC LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS WAS FOUND ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11683469&pc=MTA | 2023826-2021-01068 | 2020-10-19 | Injury | STAAR SURGICAL COMPANY | 2021-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH DIFFERENT DIOPTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM, CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190874&pc=MTA | 2023826-2022-01297 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT, BLURRED VISION, AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190879&pc=MTA | 2023826-2022-01296 | 2020-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY FIBRE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009098&pc=MTA | 2023826-2021-01886 | 2021-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -03.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DUE TO REFRACTIVE SURPRISE AND EXCESSIVE VAULT THE LENS WAS REMOVED ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459227&pc=MTA | 2023826-2022-00384 | 2021-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#(B)(4) Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12921591&pc=MTA | 2023826-2021-04632 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND REFRACTIVE CHANGE OVER TIME. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075679&pc=MTA | 2023826-2022-04744 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -8.0 DIOPTER, WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592023&pc=MTA | 2023826-2023-03623 | 2022-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, THE PROBLEM IS NOT RESOLVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075680&pc=MTA | 2023826-2022-04745 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS WITH A -8.0 DIOPTER, WAS IMPLANTED AS A REPLACEMENT INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2022. REPORTEDLY, "LENS TOO BIG. SURGEON PREFERENCE CHANGE." THE PROBLEM WAS NOT RESOLVED. THE REPORTER INDICATED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. "BLURRY VISION" AND "VAULT SMALL" WAS ALSO REPORTED. Manufacturer Narrative: B3 - DATE OF EVENT: 20-MAY-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO "SULCUS TOO SMALL FOR 12.1." AND BLURRY VISION. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE: (B)(6) 2022 SHOULD BE ADDED TO THE INTIAL MDR. H1 - CORRECTED TO: SERIOUS INJURY IN THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: OFF-LABEL: PRE-EXISTING PROGRESSIVE MYOPIA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15810623&pc=MTA | 2023826-2022-03960 | 2022-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A TORIC MODEL, LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. Manufacturer Narrative: (B)(4). OFF LABEL ACD<3.0 <45 YEARS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 20-DEC-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787278&pc=MTA | 2023826-2023-01468 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER EMAIL REPLY, "PATIENT WAS SEEN (B)(6) 2023, EVO ICL NICELY CENTERED. DRYNESS FROM SURGERY HAS RESOLVED AND VISION CONTINUING TO IMPROVED." CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2023. REPORTEDLY, "EYE VERY DRY. STARTED TEARS." DRYNESS FROM SURGERY HAS RESOLVED AND VISION CONTINUIING TO IMPROVED. CLAIM# (B)(4). Manufacturer Narrative: H6: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569168&pc=MTA | 2023826-2023-03513 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SUPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL USE ACD<3.0. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999784&pc=MTA | 2023826-2023-04729 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17845437&pc=MTA | 2023826-2023-04238 | 2022-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE ACD<3.0. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15611640&pc=MTA | 2023826-2022-03644 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -5.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "PATIENT FELT ABNORMAL NEAR VISION AFTER SURGERY AND STRONGLY REQUESTED TO REPLACE THE LENS". Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598566&pc=MTA | 2023826-2022-01744 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Patient-Device Incompatibility | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -1.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE CHANGE OVERTIME AND ASTIGMATISM POST-SURGERY WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A NEW MODEL LENS OF DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178423&pc=MTA | 2023826-2021-02627 | 2021-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Unintended Movement | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING BEEN IMPLANTED UPSIDE DOWN AND PATIENT COMPLAINED OF BLURRED VISON. VISUAL ACUITY WAS POOR. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865569&pc=MTA | 2023826-2023-01626 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM612.6 IMPLANTABLE COLLAMER LENS OF A -6.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND A TORIC LENS WAS IMPLANTED DUE TO REFRACTIVE SUPRISE, REFRACTIVE CHANGE OVER TIME, GLARES/HALOS, AND BLURRED VISION. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. REPORTEDLY, "EVERYTHING IS GOOD, NOTHING TO ADD." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#(B)(4). Manufacturer Narrative: H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135976&pc=MTA | 2023826-2021-00042 | 2020-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT MODEL LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18187759&pc=MTA | 2023826-2023-05236 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. PATIENT IS HYPEROPIC +25 +.50 X086 = 20/15. UNCORRECTED VA IS 20/25. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE DESCRIPTION: THE EVO ICL LENS FEATURES AN OPTIC DIAMETER THAT VARIES WITH THE DIOPTRIC POWER; THE SMALLEST OPTIC DIAMETER BEING 4.9MM AND THE LARGEST 6.1MM. THE EVO ICL LENS IS CAPABLE OF BEING FOLDED AND INSERTED INTO THE POSTERIOR CHAMBER THROUGH AN INCISION OF 3.5MM OR LESS. THE EVO ICL LENS IS INTENDED TO BE PLACED ENTIRELY WITHIN THE POSTERIOR CHAMBER DIRECTLY BEHIND THE IRIS AND IN FRONT OF THE ANTERIOR CAPSULE OF THE HUMAN CRYSTALLINE LENS. WHEN CORRECTLY POSITIONED, THE EVO ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. MODE OF ACTION: THE ICL LENS FUNCTIONS AS A REFRACTIVE ELEMENT TO OPTICALLY REDUCE MODERATE TO HIGH MYOPIA WITH OR WITHOUT ASTIGMATISM. WARNINGS: A PATIENT WITH MESOPIC PUPIL SIZE THAT IS GREATER THAN THE OPTIC DIAMETER OF THE ICL LENS MAY EXPERIENCE SYMPTOMS OF GLARE AND/OR HALOS. PATIENTS SHOULD BE ADVISED ABOUT THIS POTENTIAL RISK PRIOR TO ICL LENS IMPLANTATION. PRECAUTIONS: PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE. - PATIENTS WITH HIGHER DEGREES OF MYOPIA AND/OR MYOPIC ASTIGMATISM EXPERIENCE LOWER EFFICACY AND HIGHER RATES OF ADVERSE EVENTS (AES) AND COMPLICATIONS. - IF A METHOD OF POWER CALCULATION DIFFERENT FROM THAT USED IN THE ICL LENS CLINICAL STUDY (I.E., LENS POWER CALCULATED BY STAAR SURGICAL USING STAAR'S PROPRIETARY SOFTWARE) IS USED, THE EFFECTIVENESS OF THE ICL LENS FOR MYOPIA WITH OR WITHOUT ASTIGMATISM MAY NOT BE CONSISTENT WITH THE RESULTS REPORTED IN THE ICL LENS CLINICAL STUDY RESULTS SECTION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523962&pc=MTA | 2023826-2021-03725 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/501K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16644891&pc=MTA | 2023826-2023-01130 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "I HAD EVO ICL PERFORMED A LITTLE UNDER 6 MONTHS AGO AND I'M UNHAPPY WITH THE RESULTS. NEITHER OF MY EYES ARE 20/20". Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16644892&pc=MTA | 2023826-2023-01129 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "I HAD EVO ICL PERFORMED A LITTLE UNDER 6 MONTHS AGO AND I'M UNHAPPY WITH THE RESULTS. NEITHER OF MY EYES ARE 20/20 AND MY RIGHT EYE HAS A BIT OF A HAZE OVER EVERYTHING WHICH MAY CONTRIBUTE TO NOT BEING ABLE TO READ FROM THE ONE EYE." Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284774&pc=MTA | 2023826-2017-00156 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS,-4.0/+1.5/17 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LEFT EYE WAS IMPLANTED WITH THE CALCULATED RIGHT EYE (OD) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284878&pc=MTA | 2023826-2017-00155 | 2017-01-03 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+1.5/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE RIGHT EYE (OD) WAS IMPLANTED WITH THE CALCULATED LEFT EYE (OS) LENS, RESULTING IN REFRACTIVE SURPRISE. REPOSITIONING SURGERY IS PLANNED FOR (B)(6) 2017. Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4). Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 AND THE OUTCOME WAS GOOD. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2017 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/25. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE CUSTOMER REPORTED THAT THE SURGEON INADVERTENTLY IMPLANTED THE LENS INTERCHANGED BETWEEN EYES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6508705&pc=MTA | 2023826-2017-00682 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem | Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00/+2.00/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE CUSTOMER STATED THAT THE DR. MADE A MISTAKE IN CALCULATION OF THE OPTIC. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON PERFORMED THE LIMBAL INCISION TO REDUCE ASTIGMATISM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560960&pc=MTA | 2023826-2024-00365 | 2023-11-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, TRANSIENT LOSS OF BCVA (BLURRY VISION WITH SHADOW) AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11573132&pc=MTA | 2023826-2021-00765 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND THE PATIENT EXPERIENCED OVER VAULT, NARROW ANGLE, OVERCORRECTION AND LENS ROTATION. IT WAS REPORTED THAT THE LENS WAS EXCHANGED AND NEW VAULT IS GOOD, PATIENT IS SEEING WELL AFTER EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10956213&pc=MTA | 2023826-2020-02850 | 2020-10-22 | Injury | STAAR SURGICAL COMPANY | 2020-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,VICMO 13.2, -17.50 DIOPTER,IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. IT REPORTED THAT THE PROBLEM RESOLVED, PATIENT IS OKAY. THE REPORTER INDICATED " THERE WAS AN IRIS PROLAPS DURING EXPLANTATION OF THE ICL. THE SURGEON DECIDED TO WAIT WITH NEW IMPLANTATION OF ICL AND THE PATIENT WEARS CL." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372191&pc=MTA | 2023826-2021-00442 | 2021-02-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, ELEVATED IOP (26MMHG). CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON(B)(6) 2021, "LENS EXCHANGE IS PLANNED". Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT IMPACT CODE (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939017&pc=MTA | 2023826-2021-01668 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -12.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION, BLURRED VISION AND HYPHEMA, INFLAMMATION. THIS RESOLVED THE PROBLEM. REPORTEDLY, AC BLEEDING WAS REMOVED BY IRRIGATION. INFLAMMATION WAS TREATED BY MEDICATION. A NEW LENS IMPLANT IS PLANNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12537699&pc=MTA | 2023826-2021-03695 | 2021-08-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A SPHERICAL ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS EYE ON (B)(6) 2021. THE SURGEON REPORTED THE PATIENT WAS MIOTIC, WITH A SMALL PUPIL. THE PATIENT COMPLAINED OF BLURRED VISION AND WAS SEEING "YELLOW". THE MIOSIS LASTED FOR AT LEAST 5 DAYS AND WAS DUE TO THE SURGEON USING MIOSTAT AT THE END OF THE SURGERY. THE YELLOW TINT COMPLAINT RESOLVED WITHIN A FEW DAYS AND MAY HAVE SOMEHOW BEEN RELATED TO THE MIOSIS. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. EXPLANT DATE: NA. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944355&pc=MTA | 2023826-2022-01026 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 13.2MM VICM5_13.2; -9.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AS WELL AS HYPEROPIA. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE IT WAS NOTED A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT POWER WAS IMPLANTED. THIS DID NOT RESOLVE THE PROBLEM. SEE MFR# 2023826-2022-01027 FOR REPLACEMENT LENS. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICAITONS. CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17802053&pc=MTA | 2023826-2023-04080 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED UNREACTIVE (FIXED) PUPIL, BLURRED VISION AND IRIS ATROPHY. POST-OP, THE PUPIL IS STILL LARGE. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK . H6: HEALTH IMPACT- CLINICAL CODE: 4581 - UNREACTIVE (FIXED) PUPIL, IRIS ATROPHY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: POST-OP THE PUPIL WAS STILL LARGE AND POORLY CONSTRICTED. AT THE PATIENTS LAST VISIT ON (B)(6) 2023, THE DOCTOR INFORMED PATIENT TO REDUCE VIGAMOX + PREDFORTE ON TILL FINISHED. USE BACK PILOCARPINE TDS ON LE , OPTIVE FUSION WHEN NECESSARY. PATIENT WILL BE REVIEWED IN (B)(6) 2023. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14376862&pc=MTA | 2023826-2022-01317 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE AND DIZZINESS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS; HOWEVER, FUNCTIONAL INSPECTION FOUND THE LENS SPHERE TO BE OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15332386&pc=MTA | 2023826-2022-03005 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM6_12.1; -04.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022 AND (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). EXPIRATION DATE: 09-30-2023. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15332387&pc=MTA | 2023826-2022-03006 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022 AND (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442847&pc=MTA | 2023826-2022-03082 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM6_12.6 OF -06.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2021 AND A REVISION WAS ALSO PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442849&pc=MTA | 2023826-2022-03081 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM OR 12.1 MM VICM6 OF -3.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480051&pc=MTA | 2023826-2022-03260 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480541&pc=MTA | 2023826-2022-03261 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM6; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS INDICATED THAT LASER VISION CORRECTION WAS PERFORMED ON (B)(6) 2022; DUE TO THE PATIENT HAVING RESIDUAL REFRACTION AND POOR UDVA. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16430450&pc=MTA | 2023826-2023-00622 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN SIZE AND MODEL WAS IMPLANTED INTO THE PATIENT'S EYE. PATIENT EXPERIENCED A SIGNIFICANT HYPEROPIC SHIFT 2 WEEKS POST OP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- HYPEROPIC SHIFT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066593&pc=MTA | 2023826-2023-02222 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. SURGEON NOTES THERE IS 1000UM, WITH DIFFICULTY FOCUSING AND STARBURST AROUND LIGHT. IOP IS NORMAL. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - STARBUST, 1000UM. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18561124&pc=MTA | 2023826-2024-00368 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: BLURRY VA IN THE LEFT EYE (OS) WITH REDNESS POST ICL IMPLANTATION WAS REPORTED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581:REDNESS. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587053&pc=MTA | 2023826-2021-03865 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED CQ RECEIVED ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH TORIC LENS (13.2MM VTICMO13.2, DIOPTER -15.5/+1.50/074) ON (B)(6) 2021. PROBLEM RESOLVED. H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16427707&pc=MTA | 2023826-2023-00686 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16427709&pc=MTA | 2023826-2023-00687 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THE EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. Manufacturer Narrative: HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15396076&pc=MTA | 2023826-2022-03059 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT HAD DIFFICULTY LOOKING CLOSE AND EXPERIENCED BLURRED VISION. THE PATIENT STRONGLY REQUESTED TO REMOVE THE LENS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE: 2993 SHOULD BE CORRECTED TO 1401. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457491&pc=MTA | 2023826-2021-03554 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LOW VAULT, A DECREASE IN BCVA, AND A DECREASE IN THE NUMBER OF INDOTHELIAL CELLS. REPORTEDLY, THE SURGEON DOES NOT PLAN TO SEE THE PATIENT AGAIN FOR ONE YEAR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. MODEL#: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. HEALTH EFFECT CLINICAL CODE: (B)(4) (ENDOTHELIAL CELL LOSS). DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 2.8MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17194710&pc=MTA | 2023826-2023-02531 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM), KERATOCONUS. CLAIM# (B)(4). Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND LOSS OF VISUAL ACUITY. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT'S POST-OP REFRACTION ON (B)(6) 2023: PLANO -3.752 X 025 W/ VAULT OF 1070"". H6-HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED AND "4614" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND LOSS OF VISUAL ACUITY. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT'S POST-OP REFRACTION ON (B)(6) 2023: PLANO -3.752 X 025 W/ VAULT OF 1070". CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" SHOULD BE ADDED. H6 - MEDICAL DEVICE PROBLEM CODE: "4614" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH THE SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS OTHER. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. D6B.- "(B)(6) 2023" SHOULD BE CORRECTED TO "(B)(6) 2023" . B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS OTHER. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359537&pc=MTA | 2023826-2021-03269 | 2021-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS BEING IMPLANTED UPSIDE DOWN. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THERE WAS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON DECIDED TO ADD NEW PIS DURING ICL IMPLANTATION. THE SURGEON PLANS TO EXPLANT/EXCHANGE THE LENS BUT CURRENTLY THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND RE-IMPLANTED CORRECTLY. THE PROBLEM WAS RESOLVED AND THE PATIENT IS PERFECT, VISION IS GOOD. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15513701&pc=MTA | 2023826-2022-03377 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAD BLURRED VISION AND NEAR VISION WAS DIFFICULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM AND IT WAS NOTED "THE PATIENT HAD DIFFICULTY LOOKING CLOSE AND STRONGLY REQUESTED TO TAKE IT OUT". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710646&pc=MTA | 2023826-2022-01915 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Unintended Movement | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF REFRACTIVE SURGERY CASE REPORTS IN THE YEAR 2022. ENTITLED "NON-VISCOELASTIC REPOSITIONING OF DISLOCATED TORIC AND NON-TORIC IMPLANTABLE COLLAMER LENS AFTER BLUNT TRAUMA" BY ROBERTO ZALDIVAR, MD; RICCARDO VINCIGUERRA, MD; ROGER ZALDIVAR, MS, MD. TWO PATIENTS WITH HISTORY OF MYOPIA AND ASTIGMATISM WHO PREVIOUSLY UNDERWENT ICL IMPLANTATION WERE REFERRED TO INSTITUTO ZALDIVAR, MENDOZA,ARGENTINA, FOR VISUAL LOSS AFTER A BLUNT TRAUMA. IN ONE PATIENT THE ICL WAS DISLOCATED IN THE ANTERIOR CHAMBER AND PUPILLARY DISLOCATION. IN THE OTHER PATIENT THE LEFT EYE SHOWED A HALF-DISLOCATED TORIC ICL WITH PUPILLARY ENTRAPMENT. IN BOTH PATIENTS THE ICL'S WERE ROTATED AND REPOSITIONED. BOTH CASES HAD A NORMAL FOLLOW-UP WITH NO SIGNIFICANT COMPLICATION. THE CAUSE WAS DUE TO BLUNT FORCE TRAUMA. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- PUPILLARY DISLOCATION/ENTRAPMENT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11372192&pc=MTA | 2023826-2021-00443 | 2021-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-02-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "LENS EXCHANGE IS PLANNED." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, ELEVATED IOP (26MMHG) BUT OPEN ANGLE. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, "LENS EXCHANGE IS PLANNED." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15055918&pc=MTA | 2023826-2022-02273 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO_13.2; -16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND BLURRED VISION WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: ADVERSE EVENT. B2:REQUIRED INTERVENTION. B5: ADDITIONAL INFORMATION: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO_13.2; -16.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND BLURRED VISION WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. H6: HEALTH EFFECT CLINICAL CODE: 2137-BLURRED VISION. 4581- SIGNIFICANT REDUCTION IRIDOCORNEAL ANGLES. H6: HEALTH EFFECT IMPACT CODE: 4629-DEVICE REPLACEMENT. H6: MEDICAL DEVICE PROBLEM CODE: 1401-MISFOCUSING. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591731&pc=MTA | 2023826-2022-03483 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WILL BE REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME, KERATOCONUS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE VISION WAS WORSENING AND CORNEA TOPOGRAPHY SHOWED KERATOCONUS DEVELOPING OU VERSUS 6 MONTHS PRIOR. THE LENS WAS REPLACED WITH A DIFFERENT MODEL (VICM5) BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE VA IS GOOD AND PATIENT IS DOING WELL, HAS SOME NIGHT TIME GLARE CAUSED BY LIGHTS. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: PATIENT SEX, WEIGHT, ETHNICITY AND RACE: UNK. DATE OF EVENT, IMPLANT DATE: UNK. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591791&pc=MTA | 2023826-2022-03484 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WILL BE REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REFRACTIVE CHANGE OVERTIME, KERATOCONUS. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE VISION WAS WORSENING AND CORNEA TOPOGRAPHY SHOWED KERATOCONUS DEVELOPING OU VERSUS 6 MONTHS PRIOR. THE LENS WAS REPLACED WITH A DIFFERENT MODEL (VICM5) BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE VA IS GOOD AND PATIENT IS DOING WELL, HAS SOME NIGHT TIME GLARE CAUSED BY LIGHTS. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8197958&pc=MTA | 2023826-2018-01925 | 2018-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Capsular Bag Tear; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR/ DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 IMPLANTABLE COLLAMRE LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THIS COMPLAINT LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT A 13.2 VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. A VAULT VALUE OF 0-50 M WITH ROTATION AND CAPSULAR TEAR WAS REPORTED. LOSS OF VA WAS ALSO REPORTED BUT PROVIDED PRE-OP AND POST-OP DATA DOES NOT SUPPORT THIS CLAIM. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. THE REPORTER INDICATED THAT THE INITIAL LENS WAS 12.6MM BUT PROVIDED THE SERIAL NUMBER OF A 13.2MM LENS. NO ADDITIONAL INFORMATION HAS BEEN FORTHCOMING. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190415&pc=MTA | 2023826-2022-01286 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A VIVA IMPLANTABLE COLLAMER LENS WAS REMOVED FROM THE PATIENTS LEFT EYE (OS) AND REPLACED WITH AN EVO IMPLANTABLE COLLAMER LENS ON (B)(6) 2022 DUE TO UNCOMFORTABLE VISION "NAUSEA IF PATIENT LOOKS VERY CLOSE (NEAR)." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): NO SERIAL NUMBERS REPORTED. (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION A 13.2MM, VICM6_13.2, -5.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON 14-MAR-2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (TORIC IMPLANTABLE LENS) AND POWER LENS DUE TO PERMANENT LOSS OF BCVA AND THE PROBLEM RESOLVED. "NORMAL" WAS REPORTED FOR STATUS OF THE EYE. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT-RELATED FACTOR, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965166&pc=MTA | 2023826-2020-00897 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD DISCOMFORT DUE TO ASTIGMATISM AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL ICL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965173&pc=MTA | 2023826-2020-00898 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD DISCOMFORT DUE TO ASTIGMATISM AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL ICL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711503&pc=MTA | 2023826-2017-01175 | 2017-04-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, DIOPTER -15.5/+2.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, REFRACTIVE SURPRISE, AND UNCOMFORTABLE VISION. THE LENS WAS REPOSITIONED THREE TIMES ON (B)(6) 2015, (B)(6) 2016, AND (B)(6) 2017 BUT UNSUCCESSFUL, THE LENS WAS ROTATED AGAIN. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696913&pc=MTA | 2023826-2018-01041 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT, AND BLURRED VISION. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE TORN HAPTIC PIECES ARE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696917&pc=MTA | 2023826-2018-01043 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT, AND BLURRED VISION. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE.(B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755755&pc=MTA | 2023826-2021-04210 | 2021-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; Eye Pain | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT HAS A VAGUE OF THE LEFT EYE BRINGING A VERY SERIOUS IMPACT ON THEIR LIVES TO MAKE IT VERY PAINFUL AND ANXIOUS, COMMUNICATED, AND DECIDED TO REMOVE CRYSTALS." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL AND FUNCTIONAL INSPECTION DUE TO SIGNIFICANT DAMAGE OF THE LENS OPTIC PART. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4555190&pc=MTA | 2023826-2015-00193 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO A REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT DIFFERENT DIOPTER POWER (-18.0) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NEW INCISION. REFRACTIVE SURPRISE. GLARE. LENS EXPLANTED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH, WIDTH AND DIOPTER POWER WERE MEASURED AND THE RESULTS OF THE MEASUREMENTS WERE COMPARED AGAINST THE ORIGINAL VALUES AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYES MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, WRONG LENS, UPSIDE-DOWN LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587339&pc=MTA | 2023826-2023-00950 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE PATIENT HAD SOME HALOES/BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED (HALOES/BLURRED VISION HAD RESOLVED). THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567853&pc=MTA | 2023826-2022-03574 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. PATIENT EXPERIENCED HALOS AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057498&pc=MTA | 2023826-2022-04706 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A DIOPTER OF -12.0 WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO GLARE AND HALOS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15331880&pc=MTA | 2023826-2022-02942 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated; Optical Problem | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED HE HAD AN ICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN HIS LEFT EYE (OS). THE PUPIL WAS CONSTRICTED POST-OP BUT HAD RESOLVED AFTER 5-6 DAYS. AT 1 WEEK POST-OP, HIS VISION WAS 20/30 AND THE ICL HAD ROTATED 15°. THE PATIENT REPORTED BLURRED VISION AND HALOES IN LOW LIGHT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK, INITIAL REPORTER OCCUPATION: UNK, DEVICE MANUFACTURE DATE: UNK, (B)(4). Manufacturer Narrative: B5: THE LENS WAS REPOSITIONED AT ONE MONTH POST-OP. THE LENS REMAINED IMPLANTED. THE PATIENT IS SEEING VERY WELL. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078375&pc=MTA | 2023826-2022-04754 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT REPORTED HAS DEVELOPED POSITIVE DYSPHOTOPSIA AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT EXPERIENCED "POSITIVE DYSPHOTOPSIA" (GLARES/ HALOES). CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. HEALTH IMPACT- CLINICAL CODE: 4581 - POSITIVE DYSPHOTOPSIA. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078608&pc=MTA | 2023826-2022-04753 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT REPORTED HAS DEVELOPED POSITIVE DYSPHOTOPSIA AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE PATIENT EXPERIENCED "POSITIVE DYSPHOTOPSIA" (GLARES/ HALOES). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030859&pc=MTA | 2023826-2022-04696 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -1.75 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. STATUS OF THE EYE IS "QUIET.". Manufacturer Narrative: DEVICE EVALUATED BY MFR: PATIENT RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.75 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES; AND GLARES/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "QUIET". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR; AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030862&pc=MTA | 2023826-2022-04697 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -2.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. STATUS OF THE EYE IS "QUIET.". Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES; AND GLARES/HALOES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "QUIET". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR; AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012303&pc=MTA | 2023826-2015-01021 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012304&pc=MTA | 2023826-2015-01022 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154940&pc=MTA | 2023826-2015-01345 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154947&pc=MTA | 2023826-2015-01346 | 2015-03-13 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135234&pc=MTA | 2023826-2017-01923 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -4.5/+2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589778&pc=MTA | 2023826-2017-00880 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS EXPLANT RESOLVED THE PROBLEM. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. CORRECTED DATA: CONCLUSION CODE: AS PER DFU, VICMOS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. DEVICE CODE: (B)(4) PREVIOUS CODE IS NOT APPLICABLE, BLURRED VISION AND GLARE/HALOS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589780&pc=MTA | 2023826-2017-00879 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO GLARE/HALOS AND BLURRED VISION. THE LENS EXPLANT RESOLVED THE PROBLEM. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. CORRECTED DATA: DEVICE CODE: (B)(4) PREVIOUS CODE IS NOT APPLICABLE, BLURRED VISION AND GLARE/HALOES WERE REPORTED. CONCLUSION CODE: AS PER DFU, VICMOS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525079&pc=MTA | 2023826-2016-00333 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/130 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE, BLURRED VISION AND GLARES/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041103&pc=MTA | 2023826-2016-01094 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY 20/25. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE. WORK ORDER SEARCH: (B)(4) SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6041106&pc=MTA | 2023826-2016-01093 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AS OF (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. FOREIGN MATERIAL (DRIED, DISCOLOURED MATERIAL) WAS FOUND ON THE LENS SURFACE. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5797034&pc=MTA | 2023826-2016-00924 | 2016-03-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARES/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A TORIC ICL AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559295&pc=MTA | 2023826-2017-00792 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED AT ONE DAY POST-OP, THE LENS HAD AN EXCESSIVE VAULT, REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT EXPERIENCED GLARE/HALOES. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614664&pc=MTA | 2023826-2015-00346 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); BLURRED VISION; HALO; (GLARE); SIZE INCORRECT FOR PATIENT; (LENS EXPLANTED); (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614665&pc=MTA | 2023826-2015-00347 | 2014-10-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919768&pc=MTA | 2023826-2015-00838 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND EXCESSIVE VAULT WAS OBSERVED. PATIENT EXPERIENCED BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5312180&pc=MTA | 2023826-2015-01675 | 2015-06-18 | Injury | STAAR SURGICAL COMPANY | 2015-12-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/006 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE, HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH PERFORMED AND FOUND NO SIMILAR COMPLAINTS. LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO,LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5685199&pc=MTA | 2023826-2016-00756 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/008 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND A REFRACTIVE SURPRISE. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5831916&pc=MTA | 2023826-2016-00948 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/113 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: (DEVICE FROM SAME LOT EVALUATED); WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6216613&pc=MTA | 2023826-2016-01850 | 2016-10-09 | Injury | STAAR SURGICAL COMPANY | 2016-12-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+5.00/054 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AFTER TWO WEEKS, THE LENS WAS REPOSITIONED. ADDITIONALLY, ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ROTATION, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE: PREVIOUS DATE WAS INCORRECT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4), SECONDARY SURGICAL INTERVENTION, REPOSITIONING, LOSS OF BCVA, LENS EXPLANT. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540150&pc=MTA | 2023826-2016-00375 | 2016-01-15 | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+2.0/076 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS DISLOCATION/SUBLUXATION AND ROTATION. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS AND BLURRED VISION. THE LENS WAS REPOSITIONED AND THE INCISION WAS ENLARGED. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THAT LENS ALSO TILTED AND SUBLUXED - SEE MFR REPORT # 2023826-2016-00376. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC SCRATCHED; DIMENSIONAL INSPECTION FOUND THE LENS IS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047717&pc=MTA | 2023826-2015-01115 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+4.0/94 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON 11/10/2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. LENS EXPLANTED. METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5540385&pc=MTA | 2023826-2016-00376 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS DISLOCATED/SUBLUXED. THE LENS REMAINS IMPLANTED. THIS WAS THE REPLACEMENT LENS - SEE MFR REPORT # 2023826-2016-00375 FOR THE FIRST LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80 AND THE PATIENT EXPERIENCED BLUR/HALOS AND DIPLOPIA. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432678&pc=MTA | 2023826-2016-00150 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/052 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/25). THE LENS WAS EXCHANGED FOR A SHORTER LENS (VICMO12.6, -11.0 DIOPTER) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND SCRATCHES. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5085796&pc=MTA | 2023826-2015-01195 | 2015-07-27 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+5.0/100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY. THE PATIENT ALSO EXPERIENCED PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), BLURRED VISION AND INFLAMMATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: BASED ON THE ABOVE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES CAN BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT DATE OF BIRTH - UNK. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671131&pc=MTA | 2023826-2016-00738 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND LOSS OF BEST-CORRECTED VISUAL ACUITY (POST-OP BCVA - 20/20). THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: THE LENS DESCRIPTION IS -18.0/+2.0/X087 (SPHERE/ CYLINDER/ AXIS). PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/40. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC AND HAPTIC WERE TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5664100&pc=MTA | 2023826-2016-00661 | 2016-04-07 | Injury | STAAR SURGICAL COMPANY | 2016-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED AND IT IS PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THERE WAS FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE DIOPTER MEASUREMENT IS WITHIN SPECIFICATION. CONCLUSION CODE: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT'S AGE OVER 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867161&pc=MTA | 2023826-2016-01048 | 2016-02-12 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Torn Material; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS TORE DURING DELIVERY INTO THE EYE. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION/DISLOCATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE EXCHANGED LENS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CORRECTED DATA: THE CORRECT REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834798&pc=MTA | 2023826-2016-00996 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+5.0/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE - FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. THE PATIENT IS UNDER THE AGE OF 21 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5834802&pc=MTA | 2023826-2016-00997 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+5.0/125 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY AND BLURRED VISION. THE REPORTER INDICATED THE LENS ROTATION WAS DUE TO WEAKNESS OF ZONULES. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS OLD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6001950&pc=MTA | 2023826-2016-01394 | 2016-08-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Use of Device Problem; Insufficient Information | Blurred Vision; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/001 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/SIZE LENS BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, BUT FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT. DATE OF BIRTH AND WEIGHT: - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=MTA | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18430037&pc=QCB | 2023826-2024-00047 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON 11-SEPT-2023. REPORTER STATES THE WRONG LENS WAS IMPLANTED. ON 13-SEPT-2023 REPORTER STATES THE "GOOD LENS IMPLANTED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2024-00046 FOR INITIAL LENS. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216360&pc=MTA | 2023826-2023-02162 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE PATIENT HAD BLURRY VISION AS A RESULT OF THE WRONG IMPLANTED LENS - SHOULD HAVE BEEN INCLUDED IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -06.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON EXCHANGED THE LENS FOR A SAME LENGTH BUT DIFFERENT POWER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H10 - DISREGARD SUPPLEMENTAL #1 MDR AS IT WAS SUBMITTED IN ERROR WITH THE WRONG CLAIM DETAIL CORRECTIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627897&pc=MTA | 2023826-2022-01722 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL/ CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+5.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED RESIDUAL ASTIGMATISM. REPORTEDLY "THE SURGEON IS PLANNING TO EXCHANGE THE LENS WITH SPHERICAL ICL AND PERFORM THE LRI FOR ASTIGMATISM". IT WAS LATER REPORTED "AS I CHECKED WITH THE SURGEON, THE PATIENT STILL HAS THE RESIDUAL ASTIGMATISM (-1.0 CYL). HOWEVER, THE PATIENT DIDN'T FEEL UNCOMFORTABLE AND HASN'T COME BACK TO THE CLINIC FOR FOLLOW-UP. THE SURGEON SAID THAT IT DOESN'T SEEM TO HAVE FURTHER TREATMENT." THE LENS REMAINS IMPLANTED. CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778983&pc=MTA | 2023826-2021-01321 | 2021-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -1.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PRESCRIBED DISTANCE GLASSES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. D6A: IMPLANT DATE CORRECTED TO (B)(6) 2021. H6-HEALTH IMPACT CODE: 4642-DOCTOR PRESCRIBED DISTANT GLASSES. INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. SERIAL # - UNK. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11778984&pc=MTA | 2023826-2021-01320 | 2021-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM6 13.2, -1.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE AND PERMANENT LOSS OF BCVA WAS OBSERVED. PRESCRIBED DISTANCE GLASSES. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. H6-HEALTH IMPACT CODE: 4642-DOCTOR PRESCRIBED DISTANT GLASSES. INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER ADDITIONAL INTERVENTION IS NOT CONSIDERED, AND THERE IS NO NEED TO EXCHANGE. THE LENS REMAINS IMPLANTED. D6A: IMPLANT DATE CORRECTED TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM6 13.2 OF -1.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE, PERMANENT LOSS OF BCVA AND BLURRED VISION WAS REPORTED. THE PATIENT WAS PRESCRIBED DISTANT GLASSES. PER REPORTER PLANNED EXCHANGE IN THE FUTURE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15230266&pc=MTA | 2023826-2022-02685 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -3.00 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODE SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE EXCHANGE LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15554010&pc=MTA | 2023826-2022-03517 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. "THE PATIENT IS (B)(6), PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." Manufacturer Narrative: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136767&pc=MTA | 2023826-2022-02525 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME/ MODEL AND SIZE LENS BUT DIFFERENT DIOPTER DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136769&pc=MTA | 2023826-2022-02526 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A TORIC MODEL LENS OF THE SAME SIZE DUE TO REFRACTIVE SURPRISE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352414&pc=MTA | 2023826-2022-03015 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH A DIFFERENT DIOPTER. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4: DATE OF THIS REPORT-04-AUG-2022 FOR SUPPLEMENTAL #1. G3: DATE RECEIVED BY MANUFACTURER: 19-OCT-2022 FOR SUPPLEMENTAL #1. CLAIM# (B)(4). . Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 31-JAN-2025. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15674658&pc=MTA | 2023826-2022-03576 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -17.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2022 AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF THE SAME MODEL; SIZE AND DIFFERENT DIOPTER WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. " THE DEFECT WAS MAINLY DUE TO THE WRONG REFRACTION SUPPLIED FOR CALCULATION." Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16356424&pc=MTA | 2023826-2023-00476 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS OF A +5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REFRACTIVE SUPRISE WAS OBSERVED. Manufacturer Narrative: B1: CORRECTED TO: PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN THE INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN/BENT TO THE LENS. UNABLE TO PERFORM LENS INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS OF A +5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR SAME MODEL/LENGTH LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16603910&pc=MTA | 2023826-2023-00955 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM6_12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON(B)(6) 2023 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- HEALTH EFFECT - CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTICS BENT, WITH BLOOD STAINS ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTION (FUNCTIONAL) INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669961&pc=MTA | 2023826-2020-02332 | 2020-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION : LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS DUE TO PATIENT REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. HEALTH EFFECT CODE IN INITAL MDR 4582 CORRECTED TO 2137. MEDICAL DEVICE PROBLEM CODE IN INITIAL MDR 1401 CORRECTED TO 2993. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15341923&pc=MTA | 2023826-2022-03007 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT HAS BEEN REPORTED THE LENS HAS ROTATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.50/+2.5/171 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHORT DIAMETER". H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048073&pc=MTA | 2023826-2023-04718 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE WRONG POST-OPERATIVE CORRECTION FOR CALCUATION WAS USE BY SURGEON BY ERROR. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. REPORTEDLY "HE USED THE WRONG PARAMETER FOR CALCULATION." ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS: "SURGEON USED THE WRONG POSTOPERATIVE CORRECTION FOR CALCULATIONS". D6B. - (B)(6) 2023 SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" AMD "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983359&pc=MTA | 2023826-2020-00933 | 2014-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT IS HAPPY. H6- HEALTH EFFECT IMPACT CODE 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. UNABLE TO ENTER CODE IN H6 FIELD. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# : (B)(4). Manufacturer Narrative: CORRRECTED DATA: H6: CLINICAL CODE 4582 IS NOT APPLICABLE IN PREVIOUS MDR. H6: RESULT CODE 114 IS NOT APPLICABLE IN PREVIOUS MDR. H6: CONCLUSION CODE 4310 SHOULD BE IN SUPPLEMENTAL MDR #1. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983362&pc=MTA | 2023826-2020-00932 | 2014-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. PATIENT IS HAPPY. H6- HEALTH EFFECT IMPACT CODE 4629: SECONDARY SURGICAL INTERVENTION; LENS EXCHANGE. UNABLE TO ENTER CODE IN H6 FIELD. CLAIM# (B)(4). Manufacturer Narrative: CORRRECTED DATA: H6-CLINICAL CODE 4582 IS NOT APPLICABLE IN PREVIOUS MDR. H6-RESULT CODE 114 IS NOT APPLICABLE IN PREVIOUS MDR. H6- CONCLUSION CODE 4310 SHOULD BE IN SUPPLEMENTAL MDR #1. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686315&pc=MTA | 2023826-2022-03836 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A ICM120V4 IMPLANTABLE COLLAMER LENS OF A -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2022, LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686372&pc=MTA | 2023826-2022-03837 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A ICM120V4 IMPLANTABLE COLLAMER LENS OF A -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. ON(B)(6)2022 , LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: B5: "ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED." CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT/ ETHINICTY/ RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14059294&pc=MTA | 2023826-2022-01159 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -10.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING AND REFRACTIVE SURPRISE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298674&pc=MTA | 2023826-2022-02908 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT AND REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: 11/30/2015. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE CHANGE OVER TIME" SHOULD BE REMOVED FROM THE ORIGINAL SUBMISSION FOR CORRECTION. H10: "ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHININ THE SAME LOT" SHOULD BE REMOVED AND CORRECTED TO SAY "NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT". Manufacturer Narrative: CORRECTED DATA: H6 HEALTH EFFECT - CLINICAL CODE-2137 SHOULD BE OMITTED AND CODE 4582 SHOULD BE ADDED FOR CORRECTION. DEVICE CODE-1401 SHOULD BE OMITTED FOR CORRECTION. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298880&pc=MTA | 2023826-2022-02909 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT EXPERIENCED LOW VAULT AND REFRACTIVE CHANGE OVERTIME, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 HEALTH EFFECT - CLINICAL CODE-2137 SHOULD BE OMITTED AND CODE 4582 SHOULD BE ADDED FOR CORRECTION. DEVICE CODE: 1401 SHOULD BE OMITTED FOR CORRECTION. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REFRACTIVE CHANGE OVER TIME" SHOULD BE REMOVED FROM THE ORIGINAL SUBMISSION FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: 09/28/2012. THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15432637&pc=MTA | 2023826-2022-03135 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT IS EXPERIENCING HIGH VAULT AND REFRACTIVE SURPRISE AFTER IMPLANTABLE COLLAMER LENS IMPLANTATION IN THE LEFT EYE (OS). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -+6.5/-1.5/84 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY "CONTROL WAS SCHEDULED FOR (B)(6). PATIENT DID NOT COME FOR THE CHECK UP". " CORRECTED DATA: D9- LENS WAS NOT RETURNED TO MANUFACTURER. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE VAULTING WILL BE MEASURED AT THE MONTHLY CHECKUP ON (B)(6). ALL FINE. PATIENT IS HAPPY". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402698&pc=MTA | 2023826-2022-01529 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS AN OVERSIZED ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B3 - DATE OF EVENT (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -17.00/2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE LOW VAULT WITH ROTATION, LENS "DISLOCATION"/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER MARKED THE CAUSE OF THE EVENT AS THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. D6A - IMPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - DEVICE CODE: 1401 -2923 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - TYPE IF INVESTIGATION CODE: 4110-41111 :" LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE ADDED TO INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402699&pc=MTA | 2023826-2022-01530 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS AN OVERSIZED ICL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B3 - DATE OF EVENT (B)(6) 2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS, -17.00/2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE LOW VAULT WITH ROTATION, LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER MARKED THE CAUSE OF THE EVENT AS THE DEVICE." SHOULD BE ADDED TO THE INITIAL MDR. D4 - MODEL # TMICL12.6 SHOULD BE ADDED TO INITIAL MDR. D4 - SERIAL # (B)(6) SHOULD BE ADDED TO INITIAL MDR. D4 - UNIQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE ADDED TO INITIAL MDR. D6A - IMPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - CLINICAL CODE: 2137 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - DEVICE CODE: 1401 -2923 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - TYPE IF INVESTIGATION CODE: 4110-41111 :" LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD BE ADDED TO INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691191&pc=MTA | 2023826-2023-01211 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO_12.6 IMPLANTABLE COLLAMER LENS -15.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2032 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2023-01210 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: PATIENT IDENTIFIERS: (B)(6). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION: LENS DISLOCATION/ SUBLUXATION AND REFRACTIVE SURPRISE. REPORTEDLY "INTRAOCULAR LENS DISPLACEMENT, SEE THE EDGE OF INTRAOCULAR LENS". ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, TWO SIZES LONGER LENGTH, 1.0D STRONGER POWER AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE PATIENT IS IN GOOD CONDITION". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. H6- TYPE OF INVESTIGATION CODE: 4110 - LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6-HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"2993, 1401, 2923" SHOULD BE ADDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17714675&pc=MTA | 2023826-2023-03902 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Blurred Vision; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE PATIENT WAS NOT HAPPY WITH THE RESULTS. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5 - REFRACTIVE SURPRISE SHOULD BE INCLUDED. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12806826&pc=MTA | 2023826-2021-04330 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15821272&pc=MTA | 2023826-2022-03964 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS WITH A -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND REFRACTIVE SUPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "DUE TO THE PATIENT'S DIFFICULTY IN VISUAL PROXIMITY, THE ICL DEGREE WAS ADJUSTED TO -8.5." Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119844&pc=MTA | 2023826-2021-00022 | 2020-12-05 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2021-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2008. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINED IMPLANTED. THE PLAN IS TO EXPLANT AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING, REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS SATISFIED WITH VISUAL OUTCOME. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR (PROGRESSIVE MYOPIA). H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CORRECTED DATA: A2: DATE OF BIRTH (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11119845&pc=MTA | 2023826-2021-00021 | 2020-12-05 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2021-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2008. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS REMAINED IMPLANTED. THE PLAN IS TO EXPLANT AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING, REFRACTIVE CHANGE OVERTIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS SATISFIED WITH VISUAL OUTCOME. THE CAUSE OF THE EVENT WAS PATIENT-RELATED FACTOR (PROGRESSIVE MYOPIA). H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CORRECTED DATA: A2: DATE OF BIRTH - (B)(6) 1989. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15673922&pc=MTA | 2023826-2022-03635 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. OFF LABEL USE: ACD <3.0. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE." DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH "RESIDUE"/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773059&pc=MTA | 2023826-2015-00544 | 2014-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND NOTICED REFRACTIVE SURPRISE. THE PATIENT HAD A PRE-EXISTING DIAGNOSIS OF ACCOMMODATIVE ESOTROPIA. THERE WAS A LARGE DIFFERENCE BETWEEN THE MANIFEST AND CYCLOPLEGIC REFRACTIONS PREOPERATIVELY, +1.50 IN RIGHT EYE AND +1.75 IN THE LEFT EYE. THE ACCOMMODATIVE SPASM WAS ABOUT 1.5 D LEADING TO THE PATIENT'S COMPLAINT OF BLURRED VISION. FIVE MONTHS AFTER THE SURGICAL PROCEDURE THE SURGEON DECIDED TO EXCHANGE THE LENS WITH A DIFFERENT POWER, HOWEVER, THE LENS REMAINS IN THE RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) REFRACTIVE SURPRISE. METHOD: WORK ORDER SEARCH, DEVICE HISTORY REVIEW AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. THE DEVICE HISTORY REVIEW DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. MEDICAL REVIEW - THE ACCOMMODATIVE SPASM WAS ABOUT 1.5 D LEADING TO PATIENT COMPLAINTS OF BLURRED VISION. FIVE MONTHS AFTER THE SURGICAL PROCEDURE, THE SURGEON DECIDED TO EXCHANGE BOTH LENSES WITH NEW ONES WITH DIFFERENT POWER. AT THE TIME OF THIS MEDICAL REVIEW, ONLY THE LENS IN THE LEFT EYE HAD BEEN EXCHANGED, HOWEVER , WE DON'T KNOW THE POWER AND OUTCOME. THE LENS IN THE RIGHT EYE REMAINS IMPLANTED. THE MOST PROBABLE CAUSE OF THIS EVENT IS RESIDUAL MYOPIA SECONDARY TO ACCOMMODATIVE SPASM TO COMPENSATE FOR PRE-EXISTING HYPEROPIA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113371&pc=MTA | 2023826-2018-01753 | 2016-05-02 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -5.00/+5.0/150 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA) AND BLURRED VISION. THE SURGEON HAS NOT DECIDED AT THIS TIME IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THERE WAS NO CATARACT, NO ICL ROTATION AND NO PROGRESSION OF KERATOCONUS. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426150&pc=MTA | 2023826-2017-00460 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/X009 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. AT ONE DAY POST-OP THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED - THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20 AND THE ISSUE WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6796051&pc=MTA | 2023826-2017-01254 | 2017-04-21 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25 AND THE PATIENT IS HAPPY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928358&pc=MTA | 2023826-2018-01437 | 2018-07-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT REPORTED HER NEAR VISION WAS POOR. THE PATIENT COULDN'T SEE CLEARLY FROM CLOSE (40CM), SO SHE REQUESTED TO REMOVE THE ICL LENS AND DIDN'T WANT ANOTHER LENS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254835&pc=MTA | 2023826-2019-00097 |  | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND DIOPTER LENS DUE REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635312&pc=MTA | 2023826-2015-00382 | 2014-11-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THERE WAS AN INCIDENT OF REFRACTIVE SURPRISE AS WELL AS LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT):UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600864&pc=MTA | 2023826-2015-00312 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (REFRACTIVE SURPRISE); (CHOROIDAL EFFUSION-SUSPECT); BLURRED VISION. (LENS IMPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873721&pc=MTA | 2023826-2016-01071 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS, +7.0/+1.0./X025 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4949670&pc=MTA | 2023826-2015-00894 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-07-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT RECOVERED WELL AND IS SATISFIED WITH THE RESULT. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS NOT RETURNED. CLAIM # (B)(4), LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244431&pc=MTA | 2023826-2015-01536 | 2015-07-13 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS NOTED ON (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL WITH DIFFERENT DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND NO DAMAGE; LENS MEASUREMENT WAS WITHIN SPECIFICATION. RESULT (LENS MEASURED WITHIN SPECIFICATION). CONCLUSION (LENS MEASURED WITHIN SPECIFICATION). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -16.00. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089928&pc=MTA | 2023826-2016-01601 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.5/+6.0/086 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5720709&pc=MTA | 2023826-2016-00373 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND WAS REPOSITIONED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED AND IS PLANNED TO EXCHANGE WITH ANOTHER LENS. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5509447&pc=MTA | 2023826-2016-00286 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2016-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. (B)(4). NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF BLURRED VISION AND THE LENS EXCHANGE. (B)(4). THERE WAS NO BREAK IN THE MATERIAL DUE TO MATERIAL INTEGRITY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION NARRATIVE WAS INCORRECT. VISUAL INSPECTION DID NOT FIND THE OPTIC TORN. VISUAL INSPECTION FOUND THE HAPTIC TORN. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7161172&pc=MTA | 2023826-2018-00004 | 2017-11-23 | Injury | STAAR SURGICAL COMPANY | 2018-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND WHITE MATERIAL ON THE LENS. ON (B)(6) 2017, THE PATIENT'S BCVA WAS 20/13. Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: THE DEVICE WAS RETURNED TO THE MANUFACTURER OF 15-DEC-2017. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY, WITH SMALL AMOUNTS OF LIQUID PRESENT, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND SMALL PARTICULATES ON LENS SURFACE. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918200&pc=MTA | 2023826-2017-01534 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 DIOPTER -15.0/+4.0/091 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON NOTED LENS ROTATION, REFRACTIVE CHANGE OVER TIME, AND LENS REPOSITIONING. THE CLAIM REPORT STATES THERE IS A PLAN TO EXCHANGE THE LENS ("RE-IMPLANTATION"), AND THE SAME REPORT STATES THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: LENS REMAINS IMPLANTED; PLAN TO EXCHANGE. LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: PREVIOUS: LENS WAS EXCHANGED IN (B)(6) 2016 (EXACT DATE NOT KNOWN) FOR A LARGER, DIFFERENT DIOPTER LENS. NO ADDITIONAL INFORMATION IS AVAILABLE. - UPDATED: THE LENS WAS EXCHANGED FOR A LARGER, DIFFERENT DIOPTER LENS ON (B)(6) 2016. THE LENS EXCHANGE RESOLVED THE PROBLEM. EXPLANTED DATE: IT IS CORRECTED FROM (B)(6) 2016 TO (B)(6) 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS REPOSITIONING. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609744&pc=MTA | 2023826-2017-00917 | 2016-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/080 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE SURPRISE WERE NOTED. THE LENS WAS REPOSITIONED ON (B)(6) 2016, BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR ANOTHER LENS OF A DIFFERENT MODEL/DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/100 ON (B)(6) 2016. Manufacturer Narrative: THIS DEVICE IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC TORN/BROKEN. (B)(4). CONCLUSION CODE AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE AT THE TIME OF IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6618987&pc=MTA | 2023826-2017-00910 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-06-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/169 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, BLURRED VISION. THE LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2016, BUT UNSUCCESSFUL. ON (B)(6) 2016 THE LENS WAS EXCHANGED. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144234&pc=MTA | 2023826-2017-01947 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS REPORTED AS HAVING ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED POOR VISUAL ACUITY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS SUCCESSFULLY REPOSITIONED AND REMAINS IMPLANTED. AT FIRST DAY POST-OP, THE PATIENT HAD ENLARGED PUPILS BUT IS MUCH HAPPIER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970600&pc=MTA | 2023826-2017-01597 | 2017-07-16 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -13/+4.0/154 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, IT WAS NOTICED THAT THE LENS ROTATED. ON (B)(6) 2017, THE SURGEON REDIALED THE LEFT EYE. THE PATIENT WAS HAPPY FOR A WEEK, THEN THE LENS ROTATED AGAIN ON (B)(6) 2017. AS OF (B)(6) 2017, THE VAULT IS 410 MICRONS, BCVA IS 20/80 AND UCVA IS LESS THAN 20/200. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE SURGEON STATED THAT PATIENT DIDN`T WANT EXCHANGE FOR A NON-TORIC ICL. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183458&pc=MTA | 2023826-2018-00047 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. REFRACTIVE SURPRISE, RESIDUAL ASTIGMATISM, AND LENS ROTATION WAS NOTICED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +1.5/+1.00/8 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE, RESIDUAL ASTIGMATISM, AND LENS ROTATION WAS NOTICED. THE LENS REMAINS IMPLANTED. (B)(4) Manufacturer Narrative: THE LENS WAS REPOSITIONED BUT THE PROBLEM REMAINED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7089423&pc=MTA | 2023826-2017-01829 | 2017-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-12-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.0/+5/086 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND REFRACTIVE SURPRISE WERE NOTICED. THE LENS WAS THEN REPOSITIONED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180692&pc=MTA | 2023826-2018-00039 | 2017-11-30 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/+3.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE, LENS ROTATION, BLURRED VISION, AND "RESIDUAL POWER." Manufacturer Narrative: THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. "PATIENT IS COMFORTABLE WITH 6/6 VISION." (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7180665&pc=MTA | 2023826-2018-00035 | 2017-12-03 | Injury | STAAR SURGICAL COMPANY | 2018-01-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+6.0/099 (SPHERE/ CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS REPOSITIONED. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT IS REPORTED. LENS REPOSITIONING DID NOT RESOLVE THE PROBLEM. SURGEON STATES THAT THE LENS ROTATED POST-OPERATIVELY. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED REFRACTIVE SURPRISE IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871070&pc=MTA | 2023826-2018-01370 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +5.0/+4.0/X070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION, NOT ASSOCIATED TO A LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE REPORTER INDICATED THE LENS WAS INITIALLY RECEIVED TORN BUT THE SURGEON DECIDED TO IMPLANT THE LENS. AT A POST-OP VISIT ON (B)(6) 2018, THE PATIENT REPORTED DIMINISHED VISUAL ACUITY (VA), THE VA WAS 20/200 WITHOUT CORRECTION, 20/30 WITH CORRRECTION. THE LENS WAS EXPLANTED AND THE LENS TEAR WAS CONFIRMED. ANOTHER SAME MODEL/DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN OFF AND RESIDUE ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5041164&pc=MTA | 2023826-2015-01101 | 2015-07-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, +1.0/+5.0/025 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS ROTATED AFTER SURGERY. THE SURGEON DECIDED TO ADJUST THE AXIS AND REPOSITIONED THE LENS ON (B)(6) 2015. DURING FOLLOW-UP VISIT, IT WAS NOTED THE LENS ROTATED AGAIN. THE LENS REMAINS IMPLANTED. A SUPPLEMENTAL WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: REFRACTIVE SURPRISE WAS REPORTED. HEALTH PROFESSIONAL AND DISTRIBUTOR ARE NOT THE REPORT SOURCE. (B)(4). Manufacturer Narrative: (B)(4). METHOD: (EVALUATION METHOD) - LENS WORK ORDER SEARCH. RESULTS: (EVALUATION RESULTS) - A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7136378&pc=MTA | 2023826-2017-01938 | 2016-09-20 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+6.0/006 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, THE LENS HAD ROTATED AND WAS REPOSITIONED AGAIN. THE LENS WAS NOT STABLE AND TO DATE, REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/6.0/006 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED ON (B)(6) 2017. AT A POST-OP VISIT ON (B)(6) 2017, IT WAS REPORTED THAT THE LENS WAS NOT STABLE AND "ROTATED AGAIN BUT LESS". THE LENS REMAINS IMPLANTED IN THE PATIENT'S EYE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701661&pc=MTA | 2023826-2017-01156 | 2012-01-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2, DIOPTER -12.0/+1.0/91 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME. THE CUSTOMER STATED CURRENTLY THE LENS REMAINS IMPLANTED, AND THE DR. PLANS TO EXCHANGE THE LENS WITH HIGHER REFRACTIVE POWER TO RESOLVED THE PROBLEM FOR THE PATIENT. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.0/91 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR SAME MODEL DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: SHOULD HAVE BEEN RECORDED. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: FUNCTIONAL INSPECTION WAS PERFORMED BUT DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6307761&pc=MTA | 2023826-2017-00166 | 2012-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/1.5/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT HAS EXPERIENCED REFRACTIVE CHANGE OVERTIME. AS OF THE DATE OF THIS MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR ANOTHER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6465367&pc=MTA | 2023826-2017-00586 | 2015-08-13 | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. ON (B)(6) 2015, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS OPTIC TORN. ONE HAPTIC WAS TORN AND BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428720&pc=MTA | 2023826-2017-00464 | 2015-11-11 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/105 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE AND LENS DISLOCATION/SUBLUXATION. THE SURGEON ATTEMPTED TO REPOSITION THE LENS BUT THE RESULT WAS NOT STABLE. THE LENS WAS THEN EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE SURGEON IS PLANNING TO PERFORM CORNEAL REFRACTIVE SURGERY TO REDUCE THE ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED WAS RETURNED IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, LENS HAPTIC BROKEN, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DRIED, DISCOLORED MATERIAL). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7205788&pc=MTA | 2023826-2018-00104 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS REMAINS IMPLANTED IN THE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS PLEASED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6420636&pc=MTA | 2023826-2017-00449 | 2016-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +3 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH THE SAME SIZE AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS, CLEAR AND THICK SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6442677&pc=MTA | 2023826-2017-00516 | 2016-05-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/79 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR ANOTHER LENS OF A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY ON (B)(6) 2017 WAS 20/25. DEVICE EVALUATION: THE DEVICE WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197059&pc=MTA | 2023826-2016-01802 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LENS THE SAME SIZE, MODEL, BUT A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THAT THE LENS WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6393825&pc=MTA | 2023826-2017-00364 | 2016-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-03-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+0.5/86 DIOPTER IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR THE SAME LENS MODEL WITH A DIOPTER -7/ +0.5/84 DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE NOT APPLICABLE, PATIENT'S ACD IS ABOVE 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT'S ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6429971&pc=MTA | 2023826-2017-00470 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+1.5/31 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND ENLARGEMENT OF INCISION SURGERY. ON (B)(6) 2016, THE LENS WAS EXPLANTED. THE LENS WAS THEN EXCHANGED WITH THE SAME SIZE, AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED.Event Description: THERE WAS NO ENLARGEMENT OF THE INCISION ON (B)(6) 2016. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THERE WAS NO ENLARGEMENT OF INCISION ON 08/08/2016 AS PREVIOUSLY REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION: OFF LABEL USE (MORE THAN 45 YEARS OF AGE AT THE TIME OF LENS IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6381058&pc=MTA | 2023826-2017-00320 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, DIOPTER -11.50/+3.00/0.87 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ALSO ON (B)(6) 2016, THE SURGEON RE-POSITIONED THE LENS. ON (B)(6) 2016 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (PREVIOUS): REPOSITIONING WAS PERFORMED ON (B)(6) 2016. (CORRECTED) REPOSITIONING WAS PERFORMED ON (B)(6) 2016. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993869&pc=MTA | 2023826-2017-01661 | 2016-11-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: REPORTEDLY, THE SURGEON IMPLANTED A 13.2MM VICMO13.2 DIOPTER -15.50 DIOPTER INTO THE PATIENT'S EYE ON (B)(6) 2016. THE SURGEON NOTES A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED, PLANS TO BE EXCHANGED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7200518&pc=MTA | 2023826-2018-00100 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -01.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON NOTED A REFRACTIVE SURPRISE AND PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -1.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED AND EXCHANGED DUE TO REFRACTIVE SURPRISE ON (B)(6) 2018. THE PROBLEM WAS NOT RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6414543&pc=MTA | 2023826-2017-00433 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE DOCTOR ALSO STATED THE PATIENT EXPERIENCED RESIDUAL MIXED ASTIGMATISM, AND BELIVES THAT THE LENS MAY BE OVERROTATED. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE (B)(4) NO CODE AVAILABLE (RESIDUAL MIXED ASTIGMATISM). RESULT CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801614&pc=MTA | 2023826-2017-01300 | 2017-01-18 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6)2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, SO THE LENS WAS EXPLANTED ON (B)(6)2017 AND WAS REPLACED WITH THE SAME MODEL LENS BUT WITH A -8.5 DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION- DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918306&pc=MTA | 2023826-2017-01535 | 2017-01-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 DIOPTER -11.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE SURGEON EXPLANTED THE LENS DUE TO REFRACTIVE SURPRISE. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ON (B)(6) 2017, THE CLAIMANT DISCLOSED THE SERIAL NUMBER FOR THE ASSOCIATED LENS. ON THIS DATE, IT WAS DETERMINED THAT THIS CASE IS A DUPLICATE OF 2023826-2017-01630 (CLAIM# (B)(4)). PLEASE DISREGARD THE INFORMATION IN THE CURRENT CLAIM AND INSTEAD REFERENCE 2023826-2017-01630. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE INFORMATION: UNK, AS SERIAL NUMBER IS UNK. PATIENT CODE (B)(4) (NO CODE AVAILABLE): SECONDARY SURGERY, LENS EXPLANT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7398191&pc=MTA | 2023826-2018-01499 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-04-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED. THE PROBLEM WAS RESOLVED AND "THE PATIENT IS REALLY HAPPY". (B)(4) Manufacturer Narrative: DEVICE EVALUATIONS - LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND TWO FEET WERE BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS FOR A 13.7MM LENS. FUNCTIONAL INSPECTION FOUND THE POWER OF THE RETURNED LENS TO BE -8.6 D INSTEAD OF THE LABELLED POWER OF -120 D. WORK ORDER SEARCH - A WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATIONS - THE CORRECT LABELLED POWER SHOULD BE -12.0 D NOT -120 D. (B)(4). Manufacturer Narrative: CORRECTIVE DATA: EXPLANT DATE: CHANGED FROM UNKNOWN TO(B)(6) 2018. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT IS IS AMBLYOPIC IN THE FELLOW EYE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6648555&pc=MTA | 2023826-2017-00993 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT DIOPTER LENS, DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: PREVIOUS: STATED LENS HAD NOT YET BEEN EVALUATED. UPDATED: LENS HAS BEEN EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (SAME PATIENT, FELLOW EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6651664&pc=MTA | 2023826-2017-00994 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT DIOPTER LENS, DUE TO REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT (SAME PATIENT, FELLOW EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6388842&pc=MTA | 2023826-2017-00317 | 2017-02-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, THE LENS WAS EXCHANGED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.25 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THAT THE FIRST LENS WAS ORDERED BY ERROR 13.2 -2.25 INSTEAD OF 13.2 -18.0. ON (B)(6) 2017, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, THE LENS WAS EXCHANGED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6395339&pc=MTA | 2023826-2017-00362 | 2017-02-06 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.0/76 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND SIMILAR DIOPTER DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND THICK, STICKY, CLEAR RESIDUE ON PRODUCT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6402737&pc=MTA | 2023826-2017-00376 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, WAS EXPLANTED AND REPLACED WITH A LESSER POWER ICL DUE TO HYPEROPIC RESULT . THERE WAS NO PATIENT INJURY AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE UCVA AFTER ICL EXCHANGE WAS 20/20-2. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6717761&pc=MTA | 2023826-2017-01197 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR SIZE LENS OF A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/13 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE HAPTIC SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721360&pc=MTA | 2023826-2017-01191 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND THE LENS WAS EXCHANGED FOR A SIMILAR SIZE LENS OF A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017 THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE LENS WAS EXCHANGED FOR A SAME SIZE LENS OF A DIFFERENT DIOPTER. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC AND THE LENS WAS RETURNED WITH SOME MOISTURE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: PATIENT AGE >(B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487511&pc=MTA | 2023826-2017-00629 | 2017-02-28 | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+0.5/X013 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR SAME SIZE, SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT YET BEEN RETURNED. Manufacturer Narrative: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS.(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6473540&pc=MTA | 2023826-2017-00593 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-04-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, DIOPTER -11.50/+3.0/X085 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6808731&pc=MTA | 2023826-2017-01352 | 2017-04-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT, REFRACTIVE SURPRISE, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRY VISION WAS NOTICED. THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN THE OPERATIVE EYE WITH ANY CATARACT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6667615&pc=MTA | 2023826-2017-01034 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, DIOPTER -2.0/+3.5/005 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. UPDATED BASED ON PRODUCT EVALUATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595444&pc=MTA | 2023826-2017-00893 | 2017-04-26 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -6.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CONCLUSION: NO FAILURE DETECTED; DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670093&pc=MTA | 2023826-2017-01029 | 2017-05-05 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIFFERENT POWER, THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6670109&pc=MTA | 2023826-2017-01030 | 2017-05-05 | Injury | STAAR SURGICAL COMPANY | 2017-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIFFERENT POWER LENS, THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701659&pc=MTA | 2023826-2017-01166 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/83 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THIS CASE IS UNDER MEDICAL CONSULT. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183462&pc=MTA | 2023826-2018-00048 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON 26 APR 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS THEN EXCHANGED FOR A DIFFERENT LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF CLEAR AND WHITE RESIDUE ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE PERFORMED. CORRECTED DATA: DATE OF THIS REPORT: IT IS CORRECTED FROM "18-AUG-2017" TO " 15-DEC-2017". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A NEW LENS WAS IMPLANTED LATER AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213916&pc=MTA | 2023826-2018-00123 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME AND THE PATIENT WANTED THE LENS REMOVED DUE TO DECREASED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WAS NO PATIENT INJURY WHEN LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND SCRATCHES ON THE OPTIC. THERE WAS PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213918&pc=MTA | 2023826-2018-00122 | 2017-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME AND THE PATIENT WANTED THE LENS REMOVED DUE TO DECREASED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WAS NO PATIENT INJURY WHEN LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6691787&pc=MTA | 2023826-2017-01123 | 2017-05-22 | Injury | STAAR SURGICAL COMPANY | 2017-07-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, DIOPTER +6.0/+4.0/150 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. THE PHYSICIAN REPOSITIONED THE LENS ON (B)(6) 2017 AND PERFORMED A YAG, AND THE LENS REMAINS STABLE. HOWEVER, THERE IS NO CHANGE IN REFRACTION AND VISUAL ACUIT IS < 6/60 AND DOES NOT IMPROVE BEYOND 6/36 WITH GLASSES. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED AND THE SURGEON HAS NOT DECIDED WHAT THE NEXT COURSE OF ACTION WILL BE. Manufacturer Narrative: THE SURGEON HAS RE-ORIENTED THE LENS PER ADVICE FROM MEDICAL. THE PATIENT'S VISION IS NOW AT 6/12 (EXPECTED OUTCOME) AND VAULT IS AT 800. ALFAGAN DOSE HAS BEEN REDUCED TO 1 TIME PER DAY. THE PATIENT IS SATISFIED WITH THE RESULTS, PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: \PREVIOUS: INITIAL MDR STATED "THE PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE" _CORRECTED: "A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON." PREVIOUS: THE SUPPLEMENTAL INFORMATION STATED: "THE SURGEON HAS RE-ORIENTED THE LENS PER ADVICE FROM MEDICAL." _CORRECTED: "THE SURGEON HAS RE-ORIENTED THE LENS." PREVIOUS: THE SUPPLEMENTAL INFORMATION STATED "ALFAGAN" DOSE, CORRECT SPELLING IS PRESUMED TO BE "ALPHAGAN" DOSE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653553&pc=MTA | 2023826-2017-01005 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS FOUND RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6657557&pc=MTA | 2023826-2017-01018 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.5/20 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE AND DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674235&pc=MTA | 2023826-2017-01075 | 2017-05-24 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/080 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT ALSO EXPERIENCES HYPEROPIA. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725514&pc=MTA | 2023826-2017-01204 | 2017-06-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZED, SIMILAR DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PREVIOUS: THE LENS WAS EXCHANGED FOR SAME SIZE, SIMILAR DIOPTER. _CORRECTED: THE LENS WAS EXCHANGED FOR A SAME SIZE , LOWER DIOPTER. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED FOUND THAT THE PRODUCT IS WITHIN THE VERIFICATION. (B)(4). Manufacturer Narrative: PREVIOUSLY: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED FOUND THAT THE PRODUCT IS WITHIN THE VERIFICATION. _CORRECTED: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS REMEASURED AND FOUND THAT THE PRODUCT IS WITHIN THE SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993839&pc=MTA | 2023826-2017-01659 | 2017-06-19 | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 DIOPTER -11.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT DIOPTER LENS ON (B)(6) 2017, THE SURGEON REPORTING A REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725922&pc=MTA | 2023826-2017-01216 | 2017-06-22 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -7.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER AND TYPE LENS DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN LNES CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725634&pc=MTA | 2023826-2017-01205 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.00/+3.0/X096 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON PLANS TO PERFORM AN ADDITIONAL LENS ROTATION. Manufacturer Narrative: STATES THAT AN ADDITIONAL LENS ROTATION IS PLANNED. UPDATED/CORRECTED: REPORTEDLY, THE LENS WAS REPOSITIONED ON (B)(6) 2017. AS OF (B)(6) 2017 THE "PATIENT IS HAPPY AFTER REPOSITIONING." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725726&pc=MTA | 2023826-2017-01206 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -13.50/+3.50/X192 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. SURGEON WANTS TO KNOW THE AMOUNT OF REQUIRED ROTATION FOR PLANO REFRACTION. Manufacturer Narrative: THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE LENS ROTATED BACK AGAIN AFTER REPOSITIONING, SO THE SURGEON PLANS TO MAKE "BIOPTIC LVC" TO "CORRECT RX." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711486&pc=MTA | 2023826-2017-01182 | 2017-06-27 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -0.5/+6/109 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, MISALIGNMENT. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. ATTEMPTS TO CONTACT TO THE CUSTOMER FOR ADDITIONAL INFORMATION HAVE BEEN ON GOING. Manufacturer Narrative: THE IMPLANTED VTICM5_12.6 LENS WAS REPOSITIONED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE SURGEON DECIDED NOT TO DO A SECOND REPOSITIONING, THE PATIENT IS HAPPY WITH THE SLIDE CYLINDER AT THE MOMENT. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/22. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874037&pc=MTA | 2023826-2017-01445 | 2017-06-30 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND FAR SIGHT TENDENCY. THE LENS WAS EXCHANGED FOR ANOTHER LENS, SAME MODEL BUT DIFFERENT DIOPTER, AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S HYPEROPIA IMPROVED, BUT NOT COMPLETELY. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN, AND FOREIGN MATERIALS ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF 45 YEARS OLD AND IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6970595&pc=MTA | 2023826-2017-01603 | 2017-07-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -1.5/1.5/175 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS THEN EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS PRESENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BROKEN. METHOD CODE: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -1.5/1.5/175 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. A SIMILAR LENS WAS THEN IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. THE DATE OF EXPLANT IS CORRECTED FROM (B)(6) 2017 TO (B)(6) 2017. THE DEVICE MANUFACTURED DATE IS CORRECTED FROM 25-JUN-2015 TO 26-JUN-2015. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6865125&pc=MTA | 2023826-2017-01452 | 2017-07-21 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICH12.1, DIOPTER +4.0/+1.5/071 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON. REPORTEDLY, 1% ATROPINE WAS USED IN THE LEFT EYE. SURGEON IS AWAITING CONSULTATION TO DETERMINE THE NEXT COURSE OF ACTION. Manufacturer Narrative: LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SAME MODEL DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050900&pc=MTA | 2023826-2017-01707 | 2017-07-27 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.50/+5.0/070 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335039&pc=MTA | 2023826-2018-00340 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -9.50/+4.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6) 2018, THE LENS WAS REMOVED DUE PATIENT'S DISSATISFACTION WITH HIS VISUAL ACUITY WHICH WAS 20/30 AFTER THE SURGERY. HIS PRE-OP VISUAL ACUITY IS 20/20. THE LENS WAS EXCHANGED FOR LONGER DIFFERENT LENS. THE PATIENT CURRENT VISUAL ACUITY IS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND HAD PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952444&pc=MTA | 2023826-2017-01564 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+1.0/087 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SIMILAR LENS. THE PATIENT'S UCVA IS 20/30 AS OF (B)(6) 2017 AND MANIFEST REFRACTION IS +0.50/-1.25/30 AND 20/20 AS OF (B)(6) 2017. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855452&pc=MTA | 2023826-2017-01420 | 2017-08-14 | Malfunction | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -8.50/4.0/91 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ADDITIONAL REFRACTIVE TREATMENT (ENCHANCEMENT) WAS DONE. THE SURGEON DECIDED TO WAIT A FEW WEEKS AFTER THE SURGERY TO ALLOW THE EYE AND CORNEA TO HEAL AND GET STABLE. THE SURGEON DECIDED THAT NO EXCHANGE IS NEEDED, BUT WILL PERFORM A LASER TOUCH UP IF NECESSARY. Manufacturer Narrative: ADDITIONAL REFRACTIVE TREATMENT (ENHANCEMENT) WAS NOT PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6912388&pc=MTA | 2023826-2017-01496 | 2017-08-30 | Injury | STAAR SURGICAL COMPANY | 2017-10-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -7.0/+1.5/111 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: LENS EXCHANGE WAS PERFORMED ON (B)(6) 2017. PATIENT VISUAL ACUITY WAS NOT CORRECTED COMPLETELY AFTER EXCHANGE. THE DOCTOR HAS DECIDED TO PERFORM LASIK FOR ADDITIONAL CORRECTION. PATIENT POST-OP, PRE-LASIK, MR IS SPH. -0.1 AND CYL. -1.75. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908143&pc=MTA | 2023826-2017-01500 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 16 NOV 2017 AND EXCHANGED WITH A SAME MODEL BUT LOWER DIOPTER LENS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND THAT LENS HAD A VERY YELLOW VISUAL ASPECT. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS MEASURED TO BE -9.31D AT THE TIME OF RELEASE. WORK ORDER SEARCH FOUND NO OTHER SIMILAR DEFECT WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. CLAIM# (B)(4). Manufacturer Narrative: IMPLANTED MAY 2017. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021358&pc=MTA | 2023826-2017-01670 | 2017-09-05 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOSS OF BCVA. BCVA WENT FROM 20/25 TO 20/35. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BCVA ON (B)(6) 2017 WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7146627&pc=MTA | 2023826-2017-01944 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE LENS WAS +3.0D FROM THE TARGET. THE LENS WAS EXCHANGED FOR A SIMILAR LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION WERE PERFORMED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6941354&pc=MTA | 2023826-2017-01559 | 2017-09-15 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/079 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. A REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON, INCIDENT DATE REPORTED AS (B)(6) 2017. THE SURGEON IS CONSIDERING LENS EXCHANGE OR ROTATION. Manufacturer Narrative: UPDATED INFORMATION: LENS REPOSITIONING PERFORMED. SURGEON STATES, "VISUS IS NOW 1.25." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). [ANTERIOR ENDOTHELIUM CHAMBER DEPTH LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183428&pc=MTA | 2023826-2018-00050 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND IT RESOLVED THE PROBLEM. THERE IS NO PLANS IMPLANT ANOTHER LENS AS THE PATIENT WAS DISSATISFIED WITH THE BLURRED VISION. ON (B)(6) 2017, THE PATIENT'S BCVA IS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE ITS HAPTIC AND PRESENCE OF CLEAR SURGICAL RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7006958&pc=MTA | 2023826-2017-01687 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 A REFRACTIVE SURPRISE WAS REPORTED BY THE SURGEON. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 REFRACTIVE SURPRISE WAS NOTED BY THE SURGEON. THE LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7140439&pc=MTA | 2023826-2017-01950 | 2017-10-16 | Malfunction | STAAR SURGICAL COMPANY AG | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2005. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME. IT WAS STATED THAT THE LENS REMAINS IMPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS TO RESOLVE THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7140505&pc=MTA | 2023826-2017-01951 | 2017-10-16 | Malfunction | STAAR SURGICAL COMPANY AG | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2005. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVERTIME. IT WAS STATED THAT THE LENS REMAINS IMPLANTED, AND THE SURGEON PLANS TO EXCHANGE THE LENS TO RESOLVE THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135226&pc=MTA | 2023826-2017-01920 | 2017-10-19 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/180 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT CURRENTLY IS NOT PLANNED TO HAVE ANY ADDITIONAL SURGERY. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050082&pc=MTA | 2023826-2017-01757 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -0.50/+5.5/089 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". THE DIOPTER VALUE OF THE INTIAL LENS IMPLANTED IS CORRECTED FROM "-0.50/+5.5/089 SPHERE/ CYLINDER/ AXIS" TO " -0.5/+5.5/084 SPHERE/ CYLINDER/ AXIS". DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017". (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050099&pc=MTA | 2023826-2017-01756 | 2017-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 DIOPTER -1.50/+4.5/089 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, AND LENS RE-ALIGNMENT. REPORTEDLY, THE LENS WAS RE-ALIGNED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ANY FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED FROM " 26-OCT-2017" TO "25-OCT-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "26-OCT-2017" TO " 25-OCT-2017" IN THE INITIAL MDR. (B)(4) Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) IS CORRECTED FROM "-1.50/ +4.5/ 089 SPHERE/ CYLINDER/ AXIS" TO "-2.0/ +4.5/ 072 SPHERE/ CYLINDER/ AXIS". (B)(4) Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213955&pc=MTA | 2023826-2018-00120 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213961&pc=MTA | 2023826-2018-00121 | 2017-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISUAL DAMAGE AND RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE . WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156010&pc=MTA | 2023826-2017-01970 | 2017-12-05 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.0/+4.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343058&pc=MTA | 2023826-2018-00427 | 2017-12-27 | Injury | STAAR SURGICAL COMPANY | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, -2.75 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED IN THE EYE. IT WAS REPORTED THAT THIS EYE HAD A LVC (LASER VISION CORRECTION) 15 YEARS AGO. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR 12.6 VTICM5_12.6, -2.00/+1.00/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. DOCTOR AND PATIENT ARE HAPPY AND SATISFIED WITH THE RESULTS. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484262&pc=MTA | 2023826-2018-00640 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+6.0/152 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), DECREASED VISUAL ACUITY, BLURRED VISION, ENLARGEMENT OF PI OR ADDITION OF NEW PI PERFORMED VIA YAG, AND ADMINISTRATION OF TOPICAL 2% PILOCARPINE. THE STATUS OF THE EYE IS LISTED AS SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER. ADDITIONAL INFORMATION-FIRST DAY POST-OP THE VISUAL ACUITY WITHOUT CORRECTION WAS "ALMOST 20/20," BUT "A FEW DAYS AFTER" IT WAS 20/50. POST-YAG LASER AND TOPICAL PILOCARPINE THE VA WAS 20/30 BUT "TODAY" IS 20/80 WITHOUT CORRECTION. THE LAST VISIT DATE IS LISTED AS (B)(6) 2018. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON ALSO NOTICED LOSS OF BCVA. "PELLUCID DEGENERATION, PREVIOUS CORNEAL SURGERY (NON- REFRACTIVE)" IS CORRECTED TO "PELLUCID DEGENERATION". (B)(4) Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT: ON (B)(4) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS. THE PROBLEM IS RESOLVED. (SECONDARY SURGERY, LENS EXCHANGE). CLAIM # (B)(4). Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: ON (B)(6) 2019 THE SURGEON REPORTS THE PATIENT'S BCVA IS 20/30 IN HIS RIGHT (OD) EYE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA, ENLARGEMENT OR ADDITION OF PI-YAG, ADMINISTRATION OF TOPICAL 2% PILOCARPINE, SUPERIOR IRIDECTOMY, NARROW ANTERIOR CHAMBER). CONCLUSION CODE: OFF-LABEL USE [PREVIOUS HISTORY: PELLUCID DEGENERATION, CORNEAL SURGERY (NON-REFRACTIVE)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7403701&pc=MTA | 2023826-2018-00514 | 2018-03-12 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/050 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2018 WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY THE LENS REMAINS IMPLANTED AND "VISION IS STILL BLURRY." Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION: THE LENS WAS REPOSITIONED AGAIN ON (B)(6) 2018. AFTER REPOSITIONING, THE PATIENT IS "VERY HAPPY WITH THE OUTCOME" AND "HE WAS READING WITH NO EFFORT AT ALL." THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE : NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7660091&pc=MTA | 2023826-2018-00997 | 2018-03-21 | Injury | STAAR SURGICAL COMPANY | 2018-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499550&pc=MTA | 2023826-2018-00709 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT LOWER POWER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7499555&pc=MTA | 2023826-2018-00710 | 2018-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -10.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 REFRACTIVE SURPRISE WAS NOTICED. THE LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7571599&pc=MTA | 2023826-2018-00845 | 2018-05-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER LENS, THE SURGEON NOTING REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594283&pc=MTA | 2023826-2018-00894 | 2018-05-21 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/+5.0/036 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. REPORTEDLY, THE SURGEON PERFORMED THE "WRONG CALCULATION." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). CONCLUSION CODE : USER ERROR (SURGEON PERFORMED WRONG CALCULATIONS) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6314688&pc=MTA | 2023826-2017-00158 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/102 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AFTER THE SURGERY WITH AN UNCORRECTED VISUAL ACUITY (UCVA) OF 6/36 AND A VAULT VALUE OF 250 MICROMETERS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. SECONDARY SURGICAL INTERVENTION, EXPLANT, LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673303&pc=MTA | 2023826-2017-01052 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -17.5/+2.0/X174 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ASTIGMATISM WAS NOTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4). Manufacturer Narrative: CORRECTED DATA: TYPE OF REPORTABLE EVENT: "MALFUNCTION" INSTEAD OF "SERIOUS INJURY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6674075&pc=MTA | 2023826-2017-01054 |  | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.0/+6.0/96 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BLURRY VISION WAS REPORTED. THE LENS WAS THEN ROTATED 10 DEGREES ANTI CLOACKWISE ON (B)(6) 2017 WHICH IMPROVED UNAIDED VISION FROM 0.7 TO 0.9. Manufacturer Narrative: CORRECTED DATA: PREVIOUS REPORT STATEMENT "SINGLE-USE DEVICE WAS REPROCESSED AND REUSED ON A PATIENT" IS INCORRECT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6755173&pc=MTA | 2023826-2017-01247 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE BENT AND BROKEN. THERE WERE ALSO FIBERS PRESENT ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762514&pc=MTA | 2023826-2017-01264 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER REACHED OUT TO US ASKING FOR ASSISTANCE WITH CALCULATIONS FOR A TORIC- ICL CASE IN WHICH THE PATIENT HAD SURGERY IN (B)(6) ON (B)(6) 2016. THE PATIENTS POST-OP IN THE (OD) RIGHT EYE CHECKED OUT PERFECT HOWEVER THE (OS) LEFT EYE SHOWED THE LENS IN GOOD POSITION AND THERE WAS GOOD VAULT BUT THERE WAS A LOT OF RESIDUAL ASTIGMATISM PRESENT. THE REPORTER STATED THAT THE PATIENT'S OUTCOME IN THE (OS) LEFT EYE WAS -2.25 ASTIGMATISM. THE REPORTER STATED THAT THEY NEEDED US TO CALCULATE IF THEY COULD ROTATE THE ICL TO HELP ALLEVIATE SOME OF THE RESIDUAL ASTIGMATISM AS LASER VISION CORRECTION WAS NOT A GREAT OPTION DUE TO PATIENT'S CORNEAL CONDITION. A ROTATION WAS PROVIDED TO THE REPORTER AND THEY WERE ABLE TO PERFORM SURGERY ON (B)(6) 2017. THE REPORTER INDICATED THAT THEY WOULD LET US KNOW ON THE PATIENTS ONE WEEK FOLLOW UP WHAT THE OUTCOMES OF THE PROCEDURE ARE. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CORRECTED DATA: HOWEVER, THE LEFT (OS) EYE SHOWED THE LENS IN GOOD POSITION AND WITH GOOD VAULT; BUT THE SURGEON NOTED REFRACTIVE SURPRISE (LOT OF RESIDUAL ASTIGMATISM). (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795668&pc=MTA | 2023826-2017-01318 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6)2016 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS, AND THE PROBLEM WAS RESOLVED. THE DR. STATED THAT WHEN HE ORDERING THE PRIMARY LENS, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS PRODUCT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PMA 510(K):_THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE SURGEON STATED THAT WHEN ORDERING PRIMARY ICL, THE RIGHT EYE ASTIGMATISM READINGS HAVE BEEN MIXED WITH LEFT EYE CYLINDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201657&pc=MTA | 2023826-2018-00093 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): -THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201662&pc=MTA | 2023826-2018-00095 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5/+4.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE SURGEON HAS DECIDED TO TREAT THE PATIENT USING PRK-REFRACTIVE LASER TREATMENT. THE LENS WILL REMAIN IMPLANTED. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419742&pc=MTA | 2023826-2018-00526 |  | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -3.50/+6.0/172 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. (B)(4). AGE IS (B)(6) YEARS, NOT (B)(6) YEARS AS PREVIOUSLY REPORTED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE IN SPECIFICATION. CONCLUSION: - THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF KERATOCONUS AND PREVIOUS CORNEAL SURGERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557296&pc=MTA | 2023826-2018-00803 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED PATIENT HAD A 12.5MM, TICM125V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT RECENTLY CAME BACK WITH REFRACTIVE SURPRISE, LENS REMAINS IMPLANTED. IF ADD'L INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE IMPLANTED DATE IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2007". ADDITIONAL DATA: THE REPORTER INDICATED THAT THE PATIENT HAD A 12.5MM TICM145V4, -16.0/+2.5/019 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IMPLANTED INTO THE LEFT EYE (OS) ON (B)(6) 2007. THE PATIENT RECENTLY HAD REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. SECONDARY SURGICAL INTERVENTION, EXPLANT AND NEW LENS IMPLANTED LATER. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639948&pc=MTA | 2023826-2018-00956 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, +2.5/+3.0/030 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE IMPLANTABLE COLLAMER LENS (VTICH12.6) WAS REMOVED AND EXCHANGED ON (B)(6) 2018. THE EXCHANGED RESOLVED THE PROBLEM. (B)(4) Manufacturer Narrative: THE SURGEON IMPLANTED A 12.6MM VTICH12.6 +4.50/+3.5/031 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THAT LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4903447&pc=MTA | 2023826-2015-00807 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-07-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 AND REFRACTIVE CHANGE WAS NOTED ON (B)(6) 2015, DUE TO MYOPIA PROGRESSION. THE LENS REMAINS IMPLANTED AND THE SURGEON PLANS TO EXCHANGE FOR A DIFFERENT POWER. THE PATIENT'S LAST BCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) REFRACTIVE SURPRISE. CONCLUSIONS: BASED ON THE COMPLAINT INFORMATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT YET EXCHANGED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6755542&pc=MTA | 2023826-2017-01248 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-31 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE AND CLOSE TO LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS EXCHANGED FOR A VICM5_13.2, -6.50 DIOPTER, AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE HAPTICS WERE SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6117561&pc=MTA | 2023826-2016-01637 | 2016-10-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/046 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGE FOR THE SAME MODEL, SIZE BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED FOR THE SAME MODEL, SIZE, DIFFERENT DIOPTER LENS. (B)(4): BLURRED VISION/REFRACTIVE SURPRISE. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6682110&pc=MTA | 2023826-2017-01090 | 2017-05-30 | Injury | STAAR SURGICAL COMPANY | 2017-06-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -1.5/+3.0/090 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. APPROXIMATELY TWO MONTHS AFTER RECOVERY, PATIENT WAS HIT IN THE EYE BY A CHILD. THE PATIENT'S VISION WORSENED AND PATIENT BEGAN TO EXPERIENCE REFRACTIVE SURPRISE. PHYSICIAN REPOSITIONED THE LENS TWICE WHICH DID NOT RESOLVE THE ISSUE. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. SURGEON PLANS TO DO ANOTHER REPOSITION AFTER CONSULTING WITH THE MEDICAL AFFAIRS TEM Manufacturer Narrative: ON (B)(6) 2017 THE LENS WAS REPOSITIONED ACCORDING TO RE-CALCULATION RESULTS. THE PATIENT IS HAPPY, THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4) CONCLUSION CODE:OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.82MM)]. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711472&pc=MTA | 2023826-2017-01181 | 2017-06-23 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -15.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON ASSUMED THAT THE LENS IS UPSIDE DOWN, AND PLANNED TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2017 LENS WAS EXCHANGED FOR SIMILAR SIZE AND DIOPTER LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5750951&pc=MTA | 2023826-2016-00803 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AN ICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE POST-OP. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED AS A RESULT OF A CONTROL VISIT (6 MONTHS), THE PATIENT'S VISION WAS TO CLOSE TO EXPECTATION. THE EVENT WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4773094&pc=MTA | 2023826-2015-00545 | 2014-03-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-14 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND NOTICED A REFRACTIVE SURPRISE. THE PATIENT HAD A PRE-EXISTING DIAGNOSIS OF ACCOMMODATIVE ESOTROPIA. THERE WAS A LARGE DIFFERENCE BETWEEN THE MANIFEST AND CYCLOPLEGIC REFRACTIONS PREOPERATIVELY, +1.50 IN RIGHT EYE AND +1.75 IN THE LEFT EYE. THE ACCOMMODATIVE SPASM WAS ABOUT 1.5 D LEADING TO THE PATIENT'S COMPLAINT OF BLURRED VISION. FIVE MONTHS AFTER THE SURGICAL PROCEDURE THE SURGEON DECIDED TO EXCHANGE THE LENS IN THE LEFT EYE WITH A DIFFERENT POWER. THE POWER AND THE OUTCOME IS UNKNOWN. SEE MFR # 2023826-2015-00544 FOR THE RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH, DEVICE HISTORY REVIEW AND MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. THE DEVICE HISTORY REVIEW DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. MEDICAL REVIEW - THE ACCOMMODATIVE SPASM WAS ABOUT 1.5 D LEADING TO PATIENT COMPLAINTS OF BLURRED VISION. FIVE MONTHS AFTER THE SURGICAL PROCEDURE, THE SURGEON DECIDED TO EXCHANGE BOTH LENSES WITH NEW ONES WITH DIFFERENT POWER. AT THE TIME OF THIS MEDICAL REVIEW, ONLY THE LENS IN THE LEFT EYE HAD BEEN EXCHANGED, HOWEVER, WE DON'T KNOW THE POWER AND OUTCOME. THE MOST PROBABLE CAUSE OF THIS EVENT IS RESIDUAL MYOPIA SECONDARY TO ACCOMMODATIVE SPASM TO COMPENSATE FOR PRE-EXISTING HYPEROPIA. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5244802&pc=MTA | 2023826-2015-01534 | 2015-01-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. REFRACTIVE SURPRISE WAS NOTED IN (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL BUT DIFFERENT DIOPTER SIZE. THE RESOLVED THE PROBLEM. , POST-OP, UCVA WAS 20/20 SEE MFR. #2023826-2015-01535 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5250502&pc=MTA | 2023826-2015-01535 | 2015-01-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/X075 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. REFRACTIVE SURPRISE WAS NOTED IN (B)(6) 2015. THE ICL WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SIMILAR LENS MODEL.THIS RESOLVED THE PROBLEM. POST-OP, UCVA WAS 20/20 SEE MFR. #2023826-2015-01534 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). CORRECTED DATA: BEST ESTIMATE IS (B)(6) 2015 ((B)(6) 2015). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -11.00/1.0/X075 (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5325949&pc=MTA | 2023826-2015-01712 | 2015-05-19 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/063 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4) . DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK EVENT PROBLEM AND EVALUATION CODES: PATIENT: NEW INCISION, SECONDARY SURGERY, REFRACTIVE SURPRISE. DEVICE: EXPLANTED. METHOD: WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074358&pc=MTA | 2023826-2016-00927 | 2015-11-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +5.50/+3.0/060 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL INFORMATION RECEIVED: THE REPORTER STATED THAT THE LENS WAS EXPLANTED ON (B)(6) 2016. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH KERATOCONUS IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (DISCOLORED). FUNCTIONAL AND DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5900669&pc=MTA | 2023826-2016-01136 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO 12.6 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654713&pc=MTA | 2023826-2016-00656 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/122 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS IS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE DIOPTER MEASUREMENT IS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596424&pc=MTA | 2023826-2016-00489 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.5/+5.0/076 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT HAD LENS EXCHANGE, AND PROBLEM WAS RESOLVED. PATIENT IS SLIGHTLY MONOVISION. PREVIOUS REPORT SOURCE, USER FACILITY, WAS INCORRECT. CORRECT INFORMATION IS DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596425&pc=MTA | 2023826-2016-00490 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -15.5/+5.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS HAD A REFRACTIVE CHANGE OVERTIME AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR ANOTHER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4). DEVICE DISCARDED BY USER. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917266&pc=MTA | 2023826-2016-01182 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/118 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/14. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: DATE ON FOLLOW-UP #1 IS INCORRECT. CORRECT DATE WAS 08/29/2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5324056&pc=MTA | 2023826-2015-01717 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+3.5/026 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT'S VISION WAS BLURRED, WITH SHADOW. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH ANOTHER LENS (ICL), SAME MODEL, LENGTH AND -7.5 DIOPTER, THIS RESOLVED THE PROBLEM. THE PATIENT WAS DOING WELL WITH BEST CORRECTED VISUAL ACUITY (BCVA) 6/6. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (HAIR,FIBERS) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694868&pc=MTA | 2023826-2016-00800 | 2016-02-17 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION CONCLUDES THE LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: PATIENT CODE - (B)(4) (EXPLANTED) - SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. PATIENT CODE - (B)(4) (NEW INCISION) - CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. DEVICE CODE - (B)(4) (EXPLANTED) - SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089924&pc=MTA | 2023826-2016-01600 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.5/+6.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S LAST VISIT WAS ON (B)(6) 2016 WITH UNCORRECTED VISUAL ACUITY (UCVA) 20/16. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5730070&pc=MTA | 2023826-2016-00818 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+4.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: BLURRED VISION WAS REPORTED. LENS REPOSITIONING WAS PERFORMED AND THE PROBLEM WAS RESOLVED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613052&pc=MTA | 2023826-2016-00537 | 2016-05-17 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTED A REFRACTIVE SURPRISE AND DECIDED TO EXCHANGE THE LENS WITH ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/20. THE DEVICE WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. (B)(4). INITIAL REPORT IS INCORRECT; REPORT SOURCE WAS A DISTRIBUTOR, NOT A USER FACILITY (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. (B)(4) REFRACTIVE SURPRISE. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074344&pc=MTA | 2023826-2016-00908 | 2016-06-03 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+5.0/113 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: ON (B)(6) 2016, THE REPORTER STATED THAT THE PATIENT EXPERIENCED REFRACTIVE SURPRISE ASSOCIATED WITH LENS MISALIGNMENT. FIVE MONTHS AFTER SURGERY, THE SURGEON DECIDED TO PERFORM A SECONDARY INTERVENTION TO REPOSITION THE LENS. TICL WAS STILL FOUND TO BE OFF-AXIS AFTER REPOSITIONING AND THE SURGEON DECIDED TO REMOVE THE LENS ON (B)(6) 2016 AND EXCHANGE THE LENS FOR LONGER ONE WITH THE SAME POWER. THIS SOLVED THE PROBLEM. DEVICE EVALUATION - RETURNED PRODUCT EVALUATION FOUND THE LENS WAS RETURNED DRY AND BENT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). REPORT SOURCE: "USER FACILITY" TO "DISTRIBUTOR". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5799981&pc=MTA | 2023826-2016-00931 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.0/042 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826220&pc=MTA | 2023826-2016-00988 | 2016-07-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, +0.50/+1.0/001 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT IS ABOVE THE AGE OF 45 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR PATIENTS ABOVE THE AGE OF 45 YEARS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074347&pc=MTA | 2023826-2016-00910 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER ON FOLLOW-UP REPORT #2 IS INCORRECT; CORRECT DATE IS (B)(6) 2017. LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS WITHIN SPECIFICATIONS. LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND EXCHANGED ON THE SAME DAY WITH VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER. THIS RESOLVED THE PROBLEM AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15 AND UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT, IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972072&pc=MTA | 2023826-2016-01271 | 2016-08-21 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972074&pc=MTA | 2023826-2016-01270 | 2016-08-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/177 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059254&pc=MTA | 2023826-2016-01537 | 2016-09-19 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+6.0/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR ANOTHER LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6223834&pc=MTA | 2023826-2017-00013 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/089 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2010. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EXPLANT DATE: N/A. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5867090&pc=MTA | 2023826-2016-01047 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A REFRACDTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639951&pc=MTA | 2023826-2018-00967 | 2018-01-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON REPORTS MILD REFRACTIVE ERROR (MYOPIA, ASTIGMATISM) AND WHITE FOREIGN MATTER ADHERED TO THE SURFACE OF THE LENS, WHICH HAS INCREASED FROM THE FIRST EXAM TO THE FOLLOW-UP EXAM. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS UNDERGOING STEROID TREATMENT AND IS UNDER FOLLOW-UP. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. CORRECTED DATA: DATE OF EVENT: IT IS CORRECTED TO (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: EXACT DATE OF EVENT IS UNKNOWN; (B)(6) 2018 IS REPORTED. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639952&pc=MTA | 2023826-2018-00966 | 2018-01-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE SURGEON REPORTS MILD REFRACTIVE ERROR (MYOPIA, ASTIGMATISM) AND WHITE FOREIGN MATTER ADHERED TO THE SURFACE OF THE LENS, WHICH HAS INCREASED FROM THE FIRST EXAM TO THE FOLLOW-UP EXAM. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS UNDERGOING STEROID TREATMENT AND IS UNDER FOLLOW-UP. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: EXACT DATE OF EVENT IS UNKNOWN; (B)(6) 2018 IS REPORTED. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978094&pc=MTA | 2023826-2017-01614 | 2017-02-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/X063 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS NOTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. (B)(4). ADDITIONAL CONCLUSION CODE FOR OFF-LABEL USE ADDED (PREVIOUS CORNEAL/REFRACTIVE SURGERY). (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A". (B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491165&pc=MTA | 2023826-2017-00636 | 2017-03-16 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 6.5/+2.5/91 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, AND LENS ROTATION. ON (B)(6) 2017 THE LENS WAS REPOSITIONED, THE SURGEON COMMENTED THE LENS OUTCOME WAS GOOD AFTER THE SURGERY UNTIL (B)(6) 2017, THE LENS DID ROTATE AGAIN. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THE CUSTOMER ALSO MENTIONED THAT THE DOCTOR PLANNED TO EXPLANT, AND EXCHANGE THE LENS WITH ANOTHER MANUFACTURER'S LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6491169&pc=MTA | 2023826-2017-00635 | 2017-03-16 | Injury | STAAR SURGICAL COMPANY | 2017-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 6.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE SURGEON STATED THAT THE LENS OUTCOME ONE DAY WAS GOOD, BUT IT DID ROTATE OVER A FEW WEEKS. AS OF THE DAY OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. THE CUSTOMER ALSO MENTIONED THAT THE DOCTOR PLANNED TO EXPLANT, AND EXCHANGE THE LENS WITH ANOTHER MANUFACTURER'S LENS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 6.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). THE SURGEON STATED THAT THE LENS OUTCOME ONE DAY WAS GOOD, BUT IT DID ROTATE OVER A FEW WEEKS. THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4) . CLAIM # (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE ALSO PERFORMED. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 6.5/+2.5/76 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND LENS ROTATION (NOT ASSOCIATED TO A LOW VAULT). THE SURGEON STATED THAT THE LENS OUTCOME ONE DAY WAS GOOD, BUT IT DID ROTATE OVER A FEW WEEKS. THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4) . CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6927297&pc=MTA | 2023826-2017-01530 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2017-10-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/175 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED WITH LOW VAULTING), THERE WAS LOSS OF BEST-CORRECTED VISUAL ACUITY AND THE SURGEON NOTED A REFRACTIVE SURPRISE. THE LENS WAS ALSO NOTED TO HAVE TORN DURING INSERTION. THE LENS WAS REPOSITIONED ON (B)(6) 2017 BUT HAD ROTATED AGAIN, WITH LOSS OF BEST-CORRECTED VISUAL ACUITY AFTER REPOSITIONING INTERVENTION. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/60. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DOCTOR HAS REMOVED THE LENS AFTER MAKING SURE THAT THE TORN HAPTIC AFFECTS THE STABILITY OF LENS IN SULCUS. THE DOCTOR WILL RE-EVALUATE IN A MONTH AFTER REMOVAL. WORK ORDER SEARCH: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE LENS WAS REMOVED ON (B)(6) 2017. THE REPORTER STATED THAT THE DOCTOR IS EXPECTED TO RE-ASSESS THE PRE-OP REFRACTION AND MEASUREMENTS. THIS WAS SCHEDULED TO HAPPEN IN (B)(6) BUT MAYBE PATIENT DID NOT COME BACK YET. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6940698&pc=MTA | 2023826-2017-01493 | 2017-08-28 | Injury | STAAR SURGICAL COMPANY | 2017-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -7.0/+2.5/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS NOTICED. ON (B)(6) 2017, THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955554&pc=MTA | 2023826-2017-01578 | 2017-09-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/X135 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTED A REFRACTIVE SURPRISE. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: PATIENT IDENTIFIER IS CORRECTED FROM "(B)(6)" TO "(B)(6)" IN THE INITIAL. THE SURGEON NOTED A REFRACTIVE SURPRISE CAUSED DUE TO "WRONG ROTATION OF LENS". (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: "CURRENTLY, THE LENS REMAINS IMPLANTED." ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR POWER LENS. REPORTEDLY, "THE OUTCOME IS VERY GOOD. UCVA 20/20." (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7487562&pc=MTA | 2023826-2018-00673 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD ROTATED (NOT ASSOCIATED TO A LOW VAULT) AND THERE WAS A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED 3 TIMES BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT HAD LASER TOUCH UP ON (B)(6) 2017. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR A LONGER LENS. Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXCHANGED ON (B)(6) 2018 FOR A LONGER LENS. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED INFORMATION: INJECTION MODEL SYSTEM: FTP; INJECTOR MODEL: MSI-PF; CARTRIDGE MODEL: SFC-45 NOT IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594290&pc=MTA | 2023826-2018-00890 | 2018-05-10 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+2.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ADDITIONALLY, THE SURGEON STATES THE LENS WAS REPOSITIONED ON (B)(6) 2018; EACH TIME THE LENS WAS REPOSITIONED IT ROTATED AT A 90 DEGREE ANGLE AGAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978090&pc=MTA | 2023826-2017-01613 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0/+2.5/X096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS REPOSITIONED. A REFRACTIVE SURPRISE FORM WAS COMPLETED BY THE SURGEON. SURGEON STATES PATIENT IS "NOT HAPPY" AND HIS BCVA IS "IMPROVING WITH REFRACTION." SURGEON IS INQUIRING AS TO WHETHER REALIGNMENT OR LENS EXCHANGE SHOULD BE PERFORMED. Manufacturer Narrative: ADDITIONAL ITEMS REPORTED BY SURGEON: LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THE DIOPTER VALUE OF THE INITIAL LENS, VTICMO13.2 IMPLANTED IS CORRECTED FROM -5.0/+2.5/096 TO -6.0/+2.5/096 DIOPTER. (B)(4). Manufacturer Narrative: UPDATED INFORMATION: ACCORDING TO CLAIMANT THE LENS WAS REPOSITIONED AN ADDITIONAL TIME ON (B)(6) 2018. AT A FOLLOW-UP VISIT THE SURGEON NOTES THE PATIENT'S VISION IS NOT IMPROVING. THE SURGEON STATES IT'S UNKNOWN IF THE LENS HAS ROTATED AGAIN BUT IT'S NOT GIVING A GOOD REFRACTIVE OUTCOME. THE SURGEON BELIEVES THE ICL MAY BE SMALL. REPORTEDLY, THE PATIENT IS UNHAPPY AND HAS DECIDED TO REMOVE THE LENS. (B)(4) Manufacturer Narrative: EXPLANT DATE: "UNK" IS CORRECTED TO "N/A".(B)(4) Manufacturer Narrative: UPDATED INFORMATION: LENS WAS EXPLANTED AND A NEW LENS WAS IMPLANTED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S.. SECONDARY SURGERY, LENS REPOSITIONING. KERATOCONUS EYE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213549&pc=MTA | 2023826-2016-01833 | 2015-11-19 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | PHAKIC INTRAOCULAR LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: CONCLUSION: PER DFU FOR THIS LENS MODEL, VICMOS ARE CONTRAINDICATED FOR PATIENTS ABOVE THE AGE OF 45 YEARS OLD. THIS PATIENT WAS ABOVE 45 AT THE AGE OF IMPLANTATION. PER DFU FOR THIS LENS MODEL, PRE-EXISTING DIAGNOSIS OF KERATOCONUS IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. DATE RECEIVED BY MANUFACTURER ON INITIAL REPORT IS 11/30/2016. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6210714&pc=MTA | 2023826-2016-01830 | 2016-09-13 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT.Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE DATE FOR THE INITIAL MDR IS (B)(6) 2016. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.50/+3.0/16 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED." (B)(4). Manufacturer Narrative: DEVICE MANUFACTURE DATE: PREVIOUS DATE INCORRECT. (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6227597&pc=MTA | 2023826-2017-00015 | 2016-10-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-05 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE . AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR PRODUCT EVALUATION. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXCHANGED DUE TO THE PATIENT EXPERIENCING REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. METHOD: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DATE RECEIVED BY MANUFACTURER FOR INITIAL REPORT WAS 12/09/2016. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED, BUT NOT SOLD IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179489&pc=MTA | 2023826-2016-01765 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. AT THE TIME OF REPORTING IT IS UNKNOWN IF THE LENS WILL BE EXCHANGED. Manufacturer Narrative: LENS DIOPTER DESCRIPTION IS -9.50/+2.0/90 (SPHERE/CYLINDER/AXIS). DATE RECEIVED BY THE MANUFACTURER FOR THE INITIAL MDR FOR THIS CLAIM IS 11/16/2016. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) TO OTHER SERIOUS (IMPORTANT MEDICAL EVENTS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT AVAILABLE TO RETURN.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978337&pc=MTA | 2023826-2017-01630 | 2017-01-19 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE/REFRACTIVE CHANGE OVERTIME. THE POSTOPERATIVE REFRACTION WAS NOT PREDICTED (HYPEROPIA-MANIFEST REFRACTION, CYCLOPLEGIC REFRACTION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE CORRECTED DATE IS 07-SEPT-2017. THE COMPLAINT WAS FIRST RECEIVED ON CLAIM # (B)(4), MFR REPORT #: 2023826-2017-01535. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020606&pc=MTA | 2023826-2017-01695 | 2017-02-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/002 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7020855&pc=MTA | 2023826-2017-01696 | 2017-02-11 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/180 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL LENS (12.1 MM) BUT DIFFERENT DIOPTER (-7.0/+1.0/012) AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144255&pc=MTA | 2023826-2017-01912 | 2017-07-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+5.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.6MM) BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE FOR SAME MODEL/DIFFERENT DIOPTER LENS. PATIENT CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5521849&pc=MTA | 2023826-2016-00325 | 2016-02-13 | Injury | STAAR SURGICAL COMPANY | 2016-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS ROTATED AND HAD A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS ALSO HAD REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED FOR A LONGER LENS, SAME MODEL AND SIMILAR DIOPTER. THE PATIENT IS DOING FINE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. (B)(4). THE DATES ON THE INITIAL WERE INCORRECTLY REPORTED AS (B)(6) 2016. THE CORRECT DATE IS (B)(6) 2016. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438806&pc=MTA | 2023826-2018-00564 | 2010-11-24 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED-ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM IS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438807&pc=MTA | 2023826-2018-00566 | 2010-11-24 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. CURRENTLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. CORRECTED INFORMATION-ON 10-JUL-2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED.DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS.(B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832781&pc=MTA | 2023826-2017-01389 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -9.5/+1.5/102 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY AND ALL IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832815&pc=MTA | 2023826-2017-01390 | 2013-10-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -13.5/+2.5/078 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE CHANGE OVER TIME. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY AND EVERYTHING IS FINE. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365387&pc=MTA | 2023826-2017-00289 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/2.0/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING WITH REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND OTHER UNKNOWN ADVERSE EVENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A CENTRIFUGE TUBE, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645753&pc=MTA | 2023826-2017-00988 | 2016-02-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -11.5/+2/107 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, GLARES/HALOES, DECENTRALIZATION, AND "FUNNY VISION." THE LENS WAS REPOSITIONED.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. THE CURRENT LENS HAS BEEN EXCHANGED FOR A LONGER, SIMILAR DIOPTER LENS ON (B)(6) 2017. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A TEAR IN THE LENS OPTIC AND PIECE OF HAPTIC MISSING FROM THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703130&pc=MTA | 2023826-2017-01160 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.0/+3.0/110 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6703148&pc=MTA | 2023826-2017-01161 | 2016-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -10.5/+2.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN EXPERIENCING LOW VAULT (NO ROTATION) AND REFRACTIVE SURPRISE. THE LENS IS PLANNED TO BE EXCHANGED BUT CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED AND "ALL ARE HAPPY". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7599401&pc=MTA | 2023826-2018-00908 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. LOW VAULT,REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED AND THE SURGEON REPORTED THAT HE PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE LENS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. THIS EXCHANGE RESOLVED THE PROBLEM AND "PATIENT IS HAPPY". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOR DIOPTER FOUND LENS WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7601418&pc=MTA | 2023826-2018-00907 | 2016-04-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. LOW VAULT,REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED AND THE SURGEON REPORTED THAT HE PLANS TO EXCHANGE THE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. 'PATIENT IS HAPPY', PROBLEM RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438818&pc=MTA | 2023826-2018-00562 | 2016-06-15 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+1.5/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME AND LOW VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED STATEMENT ON 12-SEP-2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443489&pc=MTA | 2023826-2017-00512 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/X089 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.8MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6443524&pc=MTA | 2023826-2017-00513 | 2016-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/X087 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS AND SIMILAR DIOPTER DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SHALLOW ANTERIOR CHAMBER. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6199353&pc=MTA | 2023826-2016-01798 | 2016-10-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT REPORTED A LOW VAULT AND REFRACTIVE SURPRISE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: ON (B)(6) 2016, THE LENS WAS EXCHANGED WITH LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE AND FOREIGN MATERIAL - FIBERS PRESENT ON LENS SURFACE. (B)(4). - WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6842513&pc=MTA | 2023826-2017-01337 | 2016-10-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12/1/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, REFRACTIVE CHANGE OVER TIME, AND HYPEROPIA SIDE-EFFECT AFTER A WEEK. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP STATED THAT THE BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/13 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/16. Manufacturer Narrative: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US BUT NOT MARKETED IN THE US. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373632&pc=MTA | 2023826-2017-00315 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (CORRECTED DATA): LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017. (PREVIOUS DATA): STATED LENS WAS EXCHANGED FOR A SHORTER LENS. (B)(4). (CORRECTED DATA): LENS EXCHANGED FOR A LONGER LENS. (PREVIOUS DATA)-LENS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SCRATCH ON THE LENS OPTIC. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) AND THE LENS HAVING A SCRATCH ON OPTIC. (B)(4). Manufacturer Narrative: DEVICE EXPLANTED IN 2017. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404238&pc=MTA | 2023826-2017-00392 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18/5/009 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, REFRACTIVE SURPRISE, AND BLURRED VISION. AS THE DATE MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPDATED/ADDITIONAL INFORMATION-THE PATIENT COMPLAINED OF BLURRED VISION, "AS IF SEEING EVERYTHING THROUGH WATER." THE SURGEON SUSPECTED ISSUES MAY BE RELATED TO A VERY DIFFICULT PERSONAL PERIOD THE PATIENT IS GOING THROUGH. THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6415314&pc=MTA | 2023826-2017-00434 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/144 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, FLUCTUATIONS IN VISUAL ACUITY, AND ACCOMMODATIVE SPASM. ON (B)(6) 2017 A SECONDARY SURGERY FOR IRIDECTOMY WAS PERFORMED, AND THE PATIENT STARTED PILOCARPINE TREATMENT. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING THE PATIENT'S STATUS. AT THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. CONCLUSION: OFF LABEL USE (KERATOCONUS EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6388872&pc=MTA | 2023826-2017-00330 | 2017-02-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE PATIENT ALSO EXPERIENCED REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: ON (B)(6) 2017, THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE), AND THE LENS MISSING A PIECE OF THE HAPTIC. INCORRECT DIOPTER SIZE -17.0, SHOULD BE -15.5. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812399&pc=MTA | 2023826-2017-01361 | 2017-04-05 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND CLEAR AND RED RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE (PRE-EXISTING CONDITION OF CATARACT IN THE OPERATIVE EYE OR NON-TRAUMATIC CATARACT IN THE FELLOW EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565214&pc=MTA | 2023826-2017-00790 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY NEAR VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SMALL CONTAINER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6563963&pc=MTA | 2023826-2017-00803 | 2017-04-13 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT WAS EXPERIENCING EXCESSIVE VAULT AND MYDRIASIS. AS OF (B)(6) 2017, THE PROBLEM IS RESOLVED. REPORTER STATED THAT THE SYMPTOMS RESOLVED THEMSELVES. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL SYMPTOMS: REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL INFORMATION: REPORTER STATES THE HIGH VAULT WAS PROBABLY CAUSED BY TOO MUCH VISCOELASTIC LEFT IN THE EYE. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL SYMPTOMS: REFRACTIVE SURPRISE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ADDITIONAL INFORMATION: REPORTER STATES THE HIGH VAULT WAS PROBABLY CAUSED BY TOO MUCH VISCOELASTIC LEFT IN THE EYE. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832836&pc=MTA | 2023826-2017-01391 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -8.0/+1.0/087 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION WAS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -8.0/+1.0/087 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES); WAS NOT SELECTED IN PREVIOUS MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -8.0/+1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON 03/06/2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832866&pc=MTA | 2023826-2017-01392 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -7.5/+1.0/079 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2017, THE SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. Manufacturer Narrative: CORRECTED DATA: THE DATE OF IMPLANTATION IS CORRECTED FROM (B)(6) 2015 TO (B)(6) 2015. CLAIM# (B)(4). Manufacturer Narrative: PLANNED EXCHANGE IS CANCELLED AS THE PATIENT IS PREGNANT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -7.5/+1.0/079 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE SURGEON NOTICED LOW VAULT AND REFRACTIVE CHANGE OVER TIME ON (B)(6) 2017. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 OF -7.5/+1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(4) 2015. SURGEON REPORTED LOW VAULT AND REFRACTIVE CHANGE OVER TIME. SURGERY WAS CANCELED IN 2017 DUE TO PATIENT BEING PREGNANT. ON (B)(6)2019 THE LENS WAS EXCHANGED FOR A LARGER LENGTH LENS AND PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6847031&pc=MTA | 2023826-2017-01408 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -13.0/2.0/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS EXCHANGE FOR A SHORTER LENS RESOLVED THE PROBLEM. CORRECTED DATA: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, AND SHALLOW AC. THE LENS WAS EXCHANGED ON (B)(6) 2017 FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801617&pc=MTA | 2023826-2017-01279 | 2017-06-11 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: BCVA WAS 20/25 AT LAST EXAM (B)(6) 2017). IT IS UNKNOWN WHY BCVA DECREASED COMPARED TO BASELINE (20/20). (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6908164&pc=MTA | 2023826-2017-01489 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT/DEFORMED AND THE HAPTIC BENT/DEFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890299&pc=MTA | 2023826-2017-01470 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -3.75 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON IT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS FOUND TO BE 20/24. THIS INFORMATION SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139572&pc=MTA | 2023826-2017-01930 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/87 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH - NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7139744&pc=MTA | 2023826-2017-01929 | 2017-09-27 | Injury | STAAR SURGICAL COMPANY | 2017-12-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/91 (SPHERE/CYLINDER/AXIS) , IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCE LOW VAULT AND REFRACTIVE SURPRISE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON 11 APR 2018 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050172&pc=MTA | 2023826-2017-01749 | 2017-10-23 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS A REFRACTIVE SURPRISE. THE DOCTOR PLANS TO EXPLANT THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20 AND DOCTOR NOTED THAT EVERYTHING IS FINE NOW. (B)(4). REPORT SOURCE: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152456&pc=MTA | 2023826-2017-01961 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE SURGEON ALSO NOTED THE FOLLOWING: LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS NOTED AS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152530&pc=MTA | 2023826-2017-01960 | 2017-11-06 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING. THE SURGEON ALSO NOTED THE FOLLOWING: LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS NOTED AS RESOLVED AFTER THE LENS EXCHANGE. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213972&pc=MTA | 2023826-2018-00139 | 2017-12-05 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, DIOPTER -13.0 INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SMALLER LENS DUE TO BLURRED VISION AND EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL, AND IN AN ENVELOPE. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT, MISSING A PIECE OF HAPTIC, AND WHITE RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7308043&pc=MTA | 2023826-2018-00308 | 2018-01-15 | Injury | STAAR SURGICAL COMPANY | 2018-03-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP STATUS OF THE EYE APPEARS NORMAL, GOOD OUTCOME AND PATIENT IS PLEASED. (B)(4). Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LEN- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7359037&pc=MTA | 2023826-2018-00338 | 2018-02-07 | Injury | STAAR SURGICAL COMPANY | 2018-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, -12.00/+1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE WAS NOTICED. THE LENS REMAINS IMPLANTED. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS . VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: CORRECTED DATA: "ADVERSE EVENT" IS CORRECTED TO "PRODUCT PROBLEM". "OTHER SERIOUS" IS REMOVED FROM THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639247&pc=MTA | 2023826-2018-00978 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1, -11.0/+2.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597242&pc=MTA | 2023826-2018-00900 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PROGRESSIVE MYOPIA WAS REPORTED AS A PRE-EXISTING CONDITIONS. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULTING, REFRACTIVE SURPRISE, AND BLURRED VISION. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A TORN HAPTIC (TORN PIECE MISSING) AND DEBRIS ON THE LENS SURFACE. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7597627&pc=MTA | 2023826-2018-00899 | 2018-05-15 | Injury | STAAR SURGICAL COMPANY | 2018-06-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH AND DIFFERENT DIOPTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PROGRESSIVE MYOPIA WAS REPORTED AS A PRE-EXISTING CONDITIONS. Manufacturer Narrative: REPORT SOURCE SHOULD NOT BE "UNKNOWN", REPORT SOURCE IS IN THE UNITED STATES. REPORT SOURCE: HEALTH PROFESSIONAL AND USER FACILITY. SHOULD HAVE BEEN RECORDED IN INITIAL MDR. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696918&pc=MTA | 2023826-2018-01031 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/066 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE CAUSE IS LISTED AS A PATIENT-RELATED FACTOR AND STATES THE LENS WAS DIFFICULT TO MANIPULATE INTO THE CORRECT POSITION BECAUSE THE "PATIENT WAS VERY NERVOUS AND HECTIC." THE SURGEON ALSO STATES THE SULCUS SIZE DOES NOT CORRELATE WITH THE WHITE TO WHITE MEASUREMENT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED/CORRECTED INFORMATION: ON (B)(4) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM IS RESOLVED AND THE "PATIENT IS HAPPY WITH 20/20 VISION." (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002182&pc=MTA | 2023826-2018-01540 | 2018-10-04 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS "EXCHANGED" FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988487&pc=MTA | 2023826-2020-00984 | 2020-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 20 MAR 2013. LOW VAULT AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A VICMO13.2, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT MODE/LENGTH LENS DUE TO LOW VAULT AND REFRACTIVE CHANGE OVERTIME. THE EXCHANGE RESOLVED THE PROBLEM. D7-EXPLANT DATE: (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6060685&pc=MTA | 2023826-2016-01533 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND UNREACTIVE PUPIL. THE LENS WAS REPOSITIONED, AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2 VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, UNREACTIVE PUPIL, AND IRIS PUSHED TOWARDS THE CORNEA. THE LENS WAS REPOSITIONED AND THE ANTERIOR CHAMBER WAS WASHED AND IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. (B)(4). CORRECTED DATA: EXPLANTED DATE: N/A. (B)(6). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6094948&pc=MTA | 2023826-2016-01635 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON EXPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, INTERMITTENT ANGLE CLOSURE, BLURRY VISION AND DILATED EYE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PATIENT IS DOING WELL. THE REPORTER STATED THE LENS WAS IMPLANTED BY ANOTHER SURGEON IN (B)(6). THE CUSTOMER IN (B)(6) STATED THE LENS WAS IMPLANTED ON (B)(6) 2016. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6126126&pc=MTA | 2023826-2016-01685 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 0.0/+4.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE LENS WAS EXPLANTED BUT THE EXACT DATE IS UNKNOWN AS OF 03/13/2017. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, 0.0/+4.0/083 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRED VISION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15. THE LENS WAS EXPLANTED ON (B)(6) 2017. IT HAS SINCE BEEN EXCHANGED AND THE PATIENT IS HAPPY. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). DATE RECEIVED BY MANUFACTURER FOR FOLLOW-UP #1 WAS 11/23/2016. DATE RECEIVED BY MANUFACTURER FOR FOLLOW-UP #2 WAS 02/16/2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244089&pc=MTA | 2023826-2017-00054 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO MYOPIC REFRACTION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.1MM) BUT DIFFERENT DIOPTER POWER (-8.0). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL WITH LIQUID. A VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A SMALL TEAR AND A SCRATCH ON THE HAPTIC. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6244134&pc=MTA | 2023826-2017-00055 |  | Injury | STAAR SURGICAL COMPANY | 2017-01-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO MYOPIC REFRACTION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL (12.1MM) BUT DIFFERENT DIOPTER POWER (-7.5). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LENS VIAL WITH LIQUID. A VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR ON THE HAPTIC AND A SMALL PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6993874&pc=MTA | 2023826-2017-01660 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. DATE IS UNKNOWN. LENS WAS IMPLANTED IN (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199927&pc=MTA | 2023826-2015-01454 | 2014-07-05 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -5.5/+2.0/X081 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT, LENS ROTATION, BLURRED VISION AND REFRACTIVE CHANGE WAS NOTED (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND WAS EXCHANGED FOR A LONGER LENS LENGTH. THE PROBLEM WAS RESOLVED. PATIENT'S LAST VISIT ON (B)(6) 2014 - UCVA 20/32. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -5.5/+2.0/X081. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4806370&pc=MTA | 2023826-2015-00588 | 2014-08-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND NOTED EXCESSIVE VAULT, WITH A REFRACTIVE SURPRISE/REFRACTIVE CHANGE OVERTIME, IN THE POST-OPERATIVE PERIOD. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740720&pc=MTA | 2023826-2015-00489 | 2014-10-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WERE NOTED, PATIENT EXPERIENCED A REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH A DIFFERENT DIOPTER ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542595&pc=MTA | 2023826-2015-00175 | 2014-12-01 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.0MM ICM130V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME (OVERCORRECTION) AND BLURRED VISION. THE ICL WAS EXCHANGED FOR ANOTHER SAME MODEL LENS BUT A DIFFERENT DIOPTER POWER (-15.0), WHICH RESOLVED THE PROBLEM. THE PATIENT IS HEALING WELL OU. Manufacturer Narrative: (B)(4). NEW INCISION. REFRACTIVE CHANGE OVERTIME/OVERCORRECTION. EXPLANTED. VAULTING. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS TORN. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED BUT WITH THE RE-MEASUREMENT OF THE LENS, IT CAN BE JUSTIFIED THAT THE COMPLAINT WAS NOT PRODUCT RELATED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161953&pc=MTA | 2023826-2015-01270 | 2015-04-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. LENS ROTATION WITH LOW VAULT WAS OBSERVED ON (B)(6) 2015. PATIENT HAD BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS EXCHANGED FOR A LONGER LENS WITH A SIMILAR DIOPTER. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193772&pc=MTA | 2023826-2015-01444 | 2015-10-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/112 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE. THE SURGEON IS PLANNING ON EXPLANTING THE LENS. THE LENS REMAINS IMPLANTED.Event Description: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A LONGER LENS WITH A DIFFERENT DIOPTER. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.80MM AT THE TIME OF IMPLANTATION. CLAIM #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4) METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4) LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5370894&pc=MTA | 2023826-2016-00057 | 2015-12-08 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO MECHANICAL COMPLICATION OF ICL INITIAL ENCOUNTER. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED MILD BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND STABLE. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. RETURNED LENS WAS FOUND WITHIN DIMENSIONAL SPECIFICATIONS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED THE EXCESSIVE VAULT OF THE IMPLANTED LENS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER LENGTH MICL (12.6 MM) FOUR DAYS AFTER IMPLANTATION TO ADDRESS EXCESSIVE VAULTING AND SUBSEQUENT NARROWING OF THE ANGLE WITHOUT IOP ELEVATION. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2016, THE CAUSE OF THE EVENT WAS OVERSIZED LENS (NOT PRODUCT RELATED) AND VAULTING WAS ASSESSED ON THE FIRST POSTOP DAY. THE SAME REPORT STATED THAT PATIENT WAS STABLE FOLLOWING ICL EXCHANGE AND BCVA WAS 20/20. GIVEN THE INFORMATION THAT AFTER IMPLANTATION OF A SHORTER LENGTH ICL THE VAULTING RESOLVED, IT IS VERY LIKELY THAT PREOPERATIVE BIOCALCULATION ERROR HAS CAUSED THE SELECTION OF WRONG ICL LENGTH (OVERSIZED). CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, IT IS VERY LIKELY THAT PREOPERATIVE BIOCALCULATION ERROR HAS CAUSED THE SELECTION OF WRONG ICL LENGTH (OVERSIZED). (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5464177&pc=MTA | 2023826-2016-00200 | 2016-01-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+0.5/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT AND A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923619&pc=MTA | 2023826-2016-01221 | 2016-05-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 1999. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.2MM) AND THE PROBLEM WAS NOT RESOLVED. THE VAULT IS LOW AND THE PLAN IS TO EXCHANGE FOR A 13.7MM. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 1999. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.2) OF A DIFFERENT MODEL. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. NOT APPLICABLE. THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. IMPLANT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. MFR DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6089943&pc=MTA | 2023826-2016-01603 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/093 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23.Event Description: THE PROBLEM WAS RESOLVED AFTER LENS EXCHANGE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CONCLUSION CODE: (OFF-LABEL USE): PATIENT ACD WAS BELOW 3.00 MM. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC BROKEN; DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND WITHIN SPECIFICATIONS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247436&pc=MTA | 2023826-2017-00065 | 2016-08-03 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0/+2.5/098, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT REPORTING LOW VAULT, ROTATION OF THE LENS, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE REFRACTIVE CHANGE OVER TIME IS DUE TO THE LOW VAULT AND ICL ROTATION. THE LENS HAS BEEN RETURNED, BUT HAS NOT YET BEEN EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED AND HAS BEEN EVALUATED. (B)(4). THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (HAIRS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245220&pc=MTA | 2023826-2018-00205 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7245224&pc=MTA | 2023826-2018-00207 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-02-06 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE CHANGE OVER TIME, LOW VAULT, AND LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. CORRECTED/UPDATED INFORMATION: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH WHITE LIQUID ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268474&pc=MTA | 2023826-2018-00194 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -12.5/+1.5/65 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT, LENS ROTATION, AND REFRACTIVE CHANGE OVER TIME WAS NOTICED. A SURGERY FOR A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 AND A NEW LONGER LENGTH LENS WAS IMPLANTED LATER ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. EXPLANT DATE: (B)(6) 2018. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM" IN THE INITIAL MDR. "EXCESSIVE VAULT" IS CORRECTED TO "LOW VAULT" IN THE INITIAL MDR. TYPE OF REPORTABLE EVENT: IT IS CORRECTED FROM "SERIOUS INJURY" TO "MALFUNCTION" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6673242&pc=MTA | 2023826-2017-01044 | 2017-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE DR. MENTIONED THAT THE LENS IS DECENTERED IN PUPIL TEMPORALLY, HE WONDERED IF A TEMPORAL HAPTIC COULD BE DAMAGED/FOLDED. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. ACCORDING TO THE DR., THE PATIENT IS MANAGING THE PROBLEM WITH CONTACT LENSES, AND THE DR. ALSO PLANNED TO EXCHANGE THE LENS. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6595410&pc=MTA | 2023826-2017-00892 | 2017-05-04 | Injury | STAAR SURGICAL COMPANY | 2017-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -7.00/+6.0/118 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND LENS ROTATION. THE LENS WAS REPOSITIONED (B)(6) 2017.TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUS: LENS REMAINS IMPLANTED. CURRENT: ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A LONGER LENS. VAULT IS NORMAL, PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6865845&pc=MTA | 2023826-2017-01453 | 2017-05-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, DIOPTER -9.0/+3.5/095 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED BECAUSE LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE WERE NOTED BY THE SURGEON. THE REPLACEMENT LENS WAS OF LONGER LENGTH AND SIMILAR DIOPTER. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC BENT AND DEFORMED. (B)(4). DEVICE MANUFACTURER DATE: IT IS CORRECTED TO "27-MAY-2016". (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169880&pc=MTA | 2023826-2018-00012 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT; THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF RESIDUE AND DEBRIS ON SURFACE OF THE LENS. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169893&pc=MTA | 2023826-2018-00013 | 2017-11-25 | Injury | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS ROTATION. THE PATIENT EXPERIENCED BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE VAULT IS ADEQUATE AND PATIENT IS EXTREMELY HAPPY. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481819&pc=MTA | 2023826-2018-00628 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/2.0/082 (SPHERE/AXIS/SPHERE), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED ON (B)(46 2018 DUE TO LOW VAULTING, LENS ROTATION, AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT ALONG WITH PRESENCE OF RESIDUE ON THE LENS SURFACE. THE LENS WAS ALSO MISSING THE TORN PIECE OF ITS HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7639140&pc=MTA | 2023826-2018-00977 | 2018-04-03 | Injury | STAAR SURGICAL COMPANY | 2018-06-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/+2.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589138&pc=MTA | 2023826-2017-00867 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | INPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: ON (B)(6) 2017 THE REPORTER STATED THAT HE NEEDED TO RETURN A LENS THAT WAS "REMOVED FROM THE EYE DUE TO SIZING ISSUES." ATTEMPTS TO RETRIEVE ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED/CORRECTED INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, DIOPTER -16.0/+4.5/115 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN EXPERIENCING LOW VAULT, LENS ROTATION, BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM IS RESOLVED. ADDITIONAL INFORMATION ADDED. LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS: 26-APR-2017. _CORRECTED: 25-APR-2017. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114673&pc=MTA | 2023826-2017-01860 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.00/+1.50/89 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING THAT CAUSE LENS ROTATION TWICE AND REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6769268&pc=MTA | 2023826-2017-01284 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INPLANTED AN ICL IMPLANTABLE COLLAMER LENS, AND THERE WAS A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/180 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, LENS ROTATION AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT STILL HAS SOME REFRACTION LEFT, BUT THE SURGEON DECIDED NOT TO ROTATE THE LENS. THE SURGEON IS PLANNING TO DO A LASER TOUCH-UP. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826240&pc=MTA | 2023826-2016-00255 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, LENS DISLOCATION/SUBLUXATION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.6. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC BENT AND FIBERS ON LENS SURFACE (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594279&pc=MTA | 2023826-2018-00891 | 2013-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -2.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS LOW VAULT, REFRACTIVE CHANGE OVER TIME AND REFRACTIVE SURPRISE. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULTING, REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVERTIME. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172583&pc=MTA | 2023826-2016-01747 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 20106 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (SPOT, DRIED DISCOLORED MATERIAL) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE: AT THE TIME OF IMPLANTATION, THE PATIENT HAD CATARACT IN THE OPERATIVE EYE OR NON-TRAUMATIC CATARACT IN THE FELLOW EYE, PREVIOUS CORNEAL OR REFRACTIVE SURGERY, LOW ENDOTHELIAL CELL COUNT, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PREVIOUS HISTORY OF IRITIS, SYNECHIA, PIGMENT DISPERSION, PSEUDOEXFOLIATION, ALL OF WHICH ARE CONTRAINDICATED AS PER THE DFU OF THIS LENS MODEL. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7268468&pc=MTA | 2023826-2018-00181 | 2016-11-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND REFRACTIVE CHANGE OVER TIME. THE SURGEON IS MAKING PLANS TO REPLACE THE LENS. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. SECONDARY SURGERY, LENS EXPLANT. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A KERATOCONUS EYE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675761&pc=MTA | 2023826-2017-01083 | 2016-12-26 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50/+4.5/062 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND REFRACTIVE SURPRISE. THE SURGEON STATED THAT THE LENS SEEMS NOT TO BE ROTATED AT SLIT LAMP, AND REFRACTION WAS CHANGING THE FIRST WEEKS BUT IT IS STABLE SINCE 2 MONTHS. THE DOCTOR PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: THE DOCTOR PLANS TO EXCHANGE THE LENS. UPDATED ON 21-AUG-2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS, SIMILAR DIOPTER. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL AND HAD CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF DEBRIS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS DEVICE IS NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- NO CODE AVAILABLE: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6376873&pc=MTA | 2023826-2017-00305 | 2016-12-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, DIOPTER -5.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCES LOW VAULT, REFRACTIVE SURPRISE, AND RESIDUAL MYOPIA. AS OF THE DATE OF MDR REPORTING THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7118994&pc=MTA | 2023826-2017-01865 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO DECREASE VISUAL ACUITY. AT THE TIME OF SUBMISSION, ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED TAPED TO AN EMPTY CARTRIDGE CASE. VISUAL INSPECTION FOUND DEBRIS ON LENS AND THE HAPTIC MISSING A PIECE. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION CODE: - PER DFU FOR THIS LENS MODEL, IT IS RECOMMENDED ONLY TO USE MICROSTAAR INJECTORS, MODEL MSI-TF OR MSI-PF WITH SFC-45 OR SFC-45 FP CARTRIDGE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7119003&pc=MTA | 2023826-2017-01866 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO DECREASE VISUAL ACUITY. AT THE TIME OF SUBMISSION, ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: DESCRIBE EVENT OR PROBLEM - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED TAPED TO AN EMPTY CARTRIDGE CASE. VISUAL INSPECTION FOUND DEBRIS AND CLEAR RESIDUE ON LENS AND NO VISIBLE DAMAGE TO LENS. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR SHORTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, IT IS RECOMMENDED ONLY TO USE MICROSTAAR INJECTORS, MODEL MSI-TF OR MSI-PF WITH SFC-45 OR SFC-45 FP CARTRIDGE. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281454&pc=MTA | 2023826-2018-00259 | 2017-10-11 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. REPORTEDLY, THE LENS WAS LATER EXPLANTED DUE TO THE SIZE BEING "TOO SMALL" AND REFRACTIVE SURPRISE. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7200513&pc=MTA | 2023826-2018-00099 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2006. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS A REFRACTIVE SURPRISE. THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS BUT TO DATE THE LENS REMAINS IMPLANTED. THE PATIENT EXPERIENCED DECREASED VISUAL ACUITY. LASEK WAS PERFORMED ON THE PATIENT ON (B)(6) 2007 FOR CORRECTING RESIDUAL REFRACTION AFTER ICL IMPLANTATION. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2018. THE PROBLEM WAS RESOLVED AND THE PATIENT WAS SATISFIED WITH HIS VISUAL ACUITY. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488976&pc=MTA | 2023826-2018-00678 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR ANOTHER LENS. THE REPORTER INDICATED THEY FELT THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS HAPPY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND PIECE MISSING, FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7527890&pc=MTA | 2023826-2018-00736 | 2018-04-20 | Injury | STAAR SURGICAL COMPANY | 2018-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TO BE TORN AND THE LENS MISSING THE TORN PIECE OF HAPTIC. THERE WAS PRESENCE OF RESIDUE ON THE PRODUCT. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, IMPLANTABLE COLLAMER LENS OF, -18.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS OF DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS. PATIENT CODE: (B)(4)- SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4621347&pc=MTA | 2023826-2015-00356 | 2014-11-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE WITH AN INCIDENT OF REFRACTIVE SURPRISE. AN EXCHANGE OF THE LENS IS PLANNED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD: WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE LENS WAS EXPLANTED 5 MONTHS AFTER THE INITIAL SURGERY. THE LENS WAS EXCHANGED FOR A SMALLER LENS WITH DIFFERENT POWER TO ADJUST THE ERROR AND THE PROBLEM WAS RESOLVED. EXPLANTED: (B)(6) 2015. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5865428&pc=MTA | 2023826-2016-01033 | 2015-12-10 | Injury | STAAR SURGICAL COMPANY | 2016-08-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5838759&pc=MTA | 2023826-2016-00993 | 2016-06-09 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL WITH PARTICULATES ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PREVIOUSLY REPORTED EVENT DESCRIPTION IS INCORRECT. CORRECT EVENT DESCRIPTION IS THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -5.5/3.0/014 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXCHANGED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975336&pc=MTA | 2023826-2016-01313 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, REFRACTIVE SURPRISE AND RETINAL REPAIR. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT LENS DIOPTER MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT (VAULT VALUE: 1360 UM) AND REFRACTIVE SURPRISE. RETINAL REPAIR WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER SAME MODEL LENS WITH DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING FINE WITH VAULT VALUE 500 UM AND AV 9/10; THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY (UCVA) IS 0.9 BETTER THAN THE PREOPERATIVE BEST CORRECTED VISUAL ACUITY (BCVA) = 0.6. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135217&pc=MTA | 2023826-2017-01919 | 2012-02-21 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/097 (SPHERE/CYLINDAR/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. REPORTEDLY, THE SURGEON NOTES EXCESSIVE VAULT AND REFRACTIVE CHANGE OVER TIME. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: REPORTEDLY, WHEN THE SURGEON EXAMINED THE EYE PRIOR TO PLANNED LENS EXCHANGE SURGERY, IT WAS DISCOVERED THAT THE LENS HAD ROTATED. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594258&pc=MTA | 2023826-2018-00888 | 2018-03-20 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.5/+1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2018. THE EXCHANGE WAS SUCCESSFUL. PROBLEM RESOLVED. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND VISUAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: CORRECTED DATA: IT IS CORRECTED TO (B)(6) 2018 IN THE SUPPLEMENTAL MDR#1. DEVICE EVALUATION CORRECTED TO: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4635334&pc=MTA | 2023826-2015-00381 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THE LENS DESCRIPTION IS -10.0/+3.0/063 (SPHERE/CYLINDER/AXIS). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CONCLUSION: AT THE TIME OF IMPLANTATION, PATIENT WAS (B)(6) (I.E. UNDER THE AGE OF 21 YEARS) AND HAD PRE-EXISTING PROGRESSIVE MYOPIA OR HYPEROPIA, BOTH OF WHICH ARE CONTRAINDICATED AS PER THE DFU OF THIS LENS MODEL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5678984&pc=MTA | 2023826-2016-00743 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS HAD A LOW VAULT, ROTATION, AND A REFRACTIVE SURPRISE, AND WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BROKEN AND FOREIGN MATERIAL (PARTICULATES) ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.84MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4634764&pc=MTA | 2023826-2015-00380 | 2014-11-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-26 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WAS NOTED AS WELL AS LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME, AND BLURRED VISION. THE LENS WAS REPOSITIONED. THE LENS WAS EXPLANTED ON (B)(6) 2014. THE PATIENT'S LAST BVCA WAS 20/60 AS OF (B)(6) 2014. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). NOT RETURNED. Manufacturer Narrative: PER MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826245&pc=MTA | 2023826-2016-00256 | 2015-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, LENS DISLOCATION/SUBLUXATION AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16.6. Manufacturer Narrative: (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION CONCLUDED LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5968182&pc=MTA | 2023826-2016-01309 | 2016-07-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/110 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ROTATION AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (SMALL FIBERS). (B)(4). METHOD CODE: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5018700&pc=MTA | 2023826-2015-01065 | 2015-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-08-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/+1.50X45 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT COMPLAINT OF BLURRED VISION. ON EXAMINATION, THE PATIENT'S UCVA: 6/15 WITH -0.50/+1.75X115. A REFRACTIVE SURPRISE WAS REPORTED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. EVALUATION CONCLUSION CODE(S) BASED ON THE COMPLAINT HISTORY, DEVICE HISTORY RECORD REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4)- NO CODE AVAILABLE (REFRACTIVE SURPRISE), UNLABELED DEVICE CODE: (B)(4)- NO INFORMATION. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED.(B)(4). OTHER TEXT : LENS REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191011&pc=MTA | 2023826-2016-01785 | 2016-01-30 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THIS WAS PERFORMED ON (B)(6) 2016. AFTER IMPLANTATION, REFRACTIVE SURPRISE AND A LOSS OF BCVA WAS OBSERVED (20/40 PRE-OPERATIVE TO 20/70 POST-OPERATIVE). THE PATIENT COMPLAINED OF BLURRED VISION IMMEDIATELY AFTER SURGERY. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: DATE OF BIRTH IS UNKNOWN; BIRTH YEAR IS (B)(6). (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900450&pc=MTA | 2023826-2017-01487 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO SURGEON NOTED REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH (13.2) BUT DIFFERENT DIOPTER (-12.0), AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/16. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817519&pc=MTA | 2023826-2017-01367 | 2017-07-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/170 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS REPOSITIONED ON (B)(6) 2017. THE REPORTER INDICATED THE REASON FOR THE PROBLEM WAS NOT A ROTATION OF THE LENS CAUSED THROUGH A SIZING PROBLEM, IT WAS THE SURGEON WHO IMPLANTED THE LENS 5º CW (CLOCKWISE) INSTEAD OF 5º CCW(COUNTER CLOCKWISE). THE PATIENT'S CURRENT CONDITION IS UCVA - 20/20, AND PATIENT IS HAPPY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6962643&pc=MTA | 2023826-2017-01592 | 2017-08-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO BLURRED VISION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL (VTICMO12.6) AND DIOPTER (-5.0/+1.5/098) LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS ON LENS). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589758&pc=MTA | 2023826-2017-00773 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A PHAKIC IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE HAPTIC BENT. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL AND FUNCTIONAL TESTING WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 +1, THE EYE IS QUIET. DUANE'S RETRACTION SYNDROME. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451075&pc=MTA | 2023826-2020-01854 | 2016-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, -03.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451111&pc=MTA | 2023826-2020-01853 | 2016-12-28 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, -04.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8055981&pc=MTA | 2023826-2018-01622 | 2017-09-29 | Injury | STAAR SURGICAL COMPANY | 2018-11-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -3.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164819&pc=MTA | 2023826-2018-01818 | 2017-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/1.0/058 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENTS IS RPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTION INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773081&pc=MTA | 2023826-2018-01192 | 2017-11-24 | Injury | STAAR SURGICAL COMPANY AG | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE PATIENT DECIDED TO PROCEED WITH LASEK, WHICH WAS PERFORMED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773096&pc=MTA | 2023826-2018-01191 | 2017-11-24 | Injury | STAAR SURGICAL COMPANY AG | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM, ICM120V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE REPORTER INDICATED THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE PATIENT DECIDED TO PROCEED WITH LASEK, WHICH WAS PERFORMED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: DATE OF REPORT SHOULD BE CORRECTED TO 16-JUL-2018 FROM 17-JUL-2018 IN ORIGINAL MDR. DATE RECEIVED SHOULD BE CORRECTED TO 16-JUL-2018 FROM 17-JUL-2018 IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135976&pc=MTA | 2023826-2020-01174 | 2018-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672610&pc=MTA | 2023826-2019-01040 | 2018-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH BLACK STAINS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538838&pc=MTA | 2023826-2020-00001 | 2018-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE OVERTIME ON (B)(6) 2019. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME MODEL, SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- DATE OF EVENT IN INITIAL MDR IS CORRECTED TO (B)(6) 2018. B5-THE REPORTER INDICATED THAT A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME MODEL, SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE CHANGE OVER TIME. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7700430&pc=MTA | 2023826-2018-01066 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14/1.5/24 (SPHERE/CYLIBDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC TORN/BROKEN/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824451&pc=MTA | 2023826-2018-01256 | 2018-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.5/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCING A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS REPORT IS BEING SUBMITTED PER FDA REQUEST. THIS REPORT WAS PREVIOUSLY FILED WITH INCORRECT MFR NUMBER. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7696916&pc=MTA | 2023826-2018-01035 | 2018-06-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY THE LENS REMAINS IMPLANTED. THE SURGEON STATES THERE IS A "SIGN ERROR ON ASTIGMATISM" THAT OCCURRED DURING LENS CALCULATION. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED STATEMENT: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR THE SAME SIZE, DIFFERENT SPHERE/CYLINDER/AXIS LENS. THE PROBLEM IS RESOLVED. THE STATUS OF THE EYE IS REPORTED AS "PERFECT". (B)(4). Manufacturer Narrative: PREVIOUS MDR SHOULD INCLUDE THE FOLLOWING STATEMENT: "USE/HANDLING AND ADDITIONAL PROCESSING IMPACTS LENS DIMENSIONS/FUNCTIONS; RETURNED LENS MAY NOT MEET ORIGINAL SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE CYLINDER AND AXIS TO BE WITHIN SPECIFICATIONS BUT FOUND THE SPHERE TO BE 0.07D OFF SPECIFICATIONS. FUNCTIONAL INSPECTION AT THE TIME OF FINAL INSPECTION PRIOR TO PRODUCT RELEASE FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765576&pc=MTA | 2023826-2018-01175 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.00/4.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7961068&pc=MTA | 2023826-2018-01509 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-10-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO12.6,-18.00/3.0/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 WITH AN OBSERVATION OF REFRACTIVE SURPRISE. REPOSITIONING OF THE LENS WAS CONDUCTED AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THAT 'THE PATIENT IS ABSOLUTELY FINE' AND THAT 'VN IS 6/9'. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS DUE TO A USER ERROR, SURGEON ADMITTED TO IMPLEMENTING INCORRECT DATA IN OCOS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721256&pc=MTA | 2023826-2018-01075 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -0.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS A >/= 2-LINE DROP IN BCDVA AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE DECREASE IN BCDVA REFLECTS THE SUBJECT'S CHALLENGE WITH NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PREVIOUS STATEMENT: THE LENS REMAINS IMPLANTED. CORRECTED AND ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2018. THE VISION LOSS THAT WAS INITIALLY REPORTED HAS BEEN RESOLVED. THE CAUSE IS LISTED AS DEVICE. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721259&pc=MTA | 2023826-2018-01073 | 2018-06-27 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, -0.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS A >/= 2-LINE DROP IN BCDVA AT THE PATIENT'S 1-MONTH POSTOPERATIVE VISIT. THE LENS REMAINS IMPLANTED AND THE INVESTIGATION IS ONGOING. Manufacturer Narrative: ADDITIONAL INFORMATION: IT IS THE OPINION OF THE MEDICAL MONITOR THAT THE DECREASE IN BCDVA REFLECTS THE SUBJECT'S CHALLENGE WITH NEUROADAPTATION TO THE EDOF OPTIC. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. CORRECTED AND ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2018. THE CAUSE IS LISTED AS DEVICE. THE VISION LOSS INITIALLY REPORTED HAS BEEN RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT AND CURVED. (B)(4). Manufacturer Narrative: UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7914678&pc=MTA | 2023826-2018-01450 | 2018-06-28 | Injury | STAAR SURGICAL COMPANY | 2018-09-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/+3.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10432773&pc=MTA | 2023826-2020-01790 | 2018-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6,-10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHAGED FOR A TORIC LENS DUE TO TRANSIENT LOSS OF BCVA. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746178&pc=MTA | 2023826-2019-01140 | 2018-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR DIFFERENT MODEL AND DIOPTER LENS DUE TO OBSERVATION OF REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906882&pc=MTA | 2023826-2018-01446 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906886&pc=MTA | 2023826-2018-01443 | 2018-07-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, IMPLANTABLE COLLAMER LENS, -6.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN VIAL WITH CLEAR SURGICAL DEBRIS/RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272659&pc=MTA | 2023826-2019-00145 | 2018-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: G4- "COMBINATION PRODUCT" SHOUDL BE REMOVED FROM PREVIOUS MDR. H6- INVESTIGATION CONCLUSION CODE: 4310 SHOULD BE CORRECTED TO 4315 IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860735&pc=MTA | 2023826-2018-01248 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: DIOPTER CORRECTED TO "-9.00/+1.0/167" IN THE INITIAL MDR. THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/+1.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860748&pc=MTA | 2023826-2018-01249 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/+10/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED WITH EXCHANGE PLANNED. PATIENT IS FINE. Manufacturer Narrative: CORRECTED DATA: DIOPTER CORRECTED TO: "-8.00/0.5/002" IN THE INITIAL MDR. EXPIRATION DATE CORRECTED TO: 30-APR-2021 IN THE INITIAL MDR. MANUFACTURE DATE CORRECTED TO: 11-MAY-2018 IN THE INITIAL MDR. ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -8.00/0.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THE REPORTER STATED THAT "THE EXCHANGE WENT WELL AND THE PATIENT IS HAPPY". THE PROBLEM WAS RESOLVED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021132&pc=MTA | 2023826-2018-01579 | 2018-08-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS NOTED AS UNKNOWN. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS.THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113360&pc=MTA | 2023826-2018-01756 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+1.5/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC BENT AND DEFORMED. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. DIMENSIONAL AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227438&pc=MTA | 2023826-2019-00059 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM IS RESOLVED WITH THE EXCHANGE; "ALL IS FINE!". (B)(4). Manufacturer Narrative: ADDED APPLICABLE OUTCOME/ PREVIOUSLY STATED THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL, SPHERE, CYLINDER, AXIS LENS. REPORTEDLY, POST-OP RESULTS ARE NOT CURRENTLY AVAILABLE. DISTRIBUTOR CHECKED IN ERROR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2013. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME AND "VISUAL POWER PROBLEM ON CLOSE DISTANCES." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- (B)(4). CLAIM# COMPLAINT-(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975584&pc=MTA | 2023826-2018-01531 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -8.50/1.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7975592&pc=MTA | 2023826-2018-01532 | 2018-09-19 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.7, -9.00/0.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. A REPLACEMENT LENS (DIFFERENT LENGTH AND DIOPTER) WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR, THE SURGEON USED THE WRONG RX DATA FOR ICL POWER CALCULATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8229422&pc=MTA | 2023826-2019-00023 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8268673&pc=MTA | 2023826-2019-00018 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270834&pc=MTA | 2023826-2019-00017 | 2018-10-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE , "SIMILARE" (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. WITHIN THE SAME SURGERY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8089154&pc=MTA | 2023826-2018-01716 | 2018-10-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.50/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON PLANS TO DO A PHOTOREFRACTIVE KERATECTOMY (PRK) TO RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: IT WAS REPORTED THAT "AFTER PRK REFRACTION NOW IS -0,25/-0,25 WITH A BCVA OF 20/20". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8276381&pc=MTA | 2023826-2019-00156 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS BUT DIFFERENT SPHERE/CYLINDER/AXIS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149315&pc=MTA | 2023826-2018-01806 | 2018-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS.THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT AND CLEAR SURGICAL RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8166385&pc=MTA | 2023826-2018-01798 | 2018-11-03 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/X166 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR (I.T. ERROR AT THE SURGEON'S OFFICE. THE PATIENTS REFRACTION DATA WAS INPUT INCORRECTLY AS PLUS CYLINDER INSTEAD OF A MINUS CYLINDER). THE PATIENT HAD RESIDUAL ASTIGMATISM POST-OPERATIVE AND A REFRACTIVE TREATMENT (LASIK) IS PLANNED TO CORRECT PROCEDURE. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT HAD LASIK SURGERY ON (B)(6) 2018 FOR CORRECTION AND THE PATIENT IS HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8149333&pc=MTA | 2023826-2018-01801 | 2018-11-08 | Injury | STAAR SURGICAL COMPANY | 2018-12-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILAR DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS NOT RESOLVED. REPORTEDLY, THE "PATIENT IS DISSATISFIED WITH THE MYOPIA REMAINING." THE CAUSE IS LISTED AS UNKNOWN. SEE MFR REPORT# 2023826-2018-01804 FOR CASE CONTINUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451066&pc=MTA | 2023826-2020-01852 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451118&pc=MTA | 2023826-2020-01851 | 2018-11-15 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHTT EYE (OD) ON (B)(6) 2015. REFRACTIVE CHANGE OVERTIME AND RESIDUAL MYOPIA WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10141191&pc=MTA | 2023826-2020-01229 | 2018-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT.. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8333964&pc=MTA | 2023826-2019-00237 | 2019-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBER ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8404547&pc=MTA | 2023826-2019-00357 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-03-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED FOR A LONGER LENS ON (B)(6) 2019. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE LENS WAS DISCARDED. (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140603&pc=MTA | 2023826-2020-01211 | 2019-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440522&pc=MTA | 2023826-2019-02332 | 2019-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS EXCHANGED FOR A SAME SIZE BUT DIFFERENT MODEL (TORIC) LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440662&pc=MTA | 2023826-2019-02333 | 2019-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -4.00 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS EXCHANGED FOR A SAME SIZE BUT DIFFERENT MODEL (TORIC) LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574170&pc=MTA | 2023826-2019-00636 | 2019-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5 13.2, -09.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAIN IMPLANTED. REPORTER INDICATED EXCHANGE PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795449&pc=MTA | 2023826-2019-01240 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836679&pc=MTA | 2023826-2019-01348 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1,-11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836776&pc=MTA | 2023826-2019-01349 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM, VICM5_12.1,-11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836981&pc=MTA | 2023826-2019-01350 | 2019-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICM5_12.6,-4.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CONCLUSION CODE: 4310 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. RESULT CODE: 114 SHOULD BE REMOVED FROM SUPPLEMENTAL #1 MEDWATCH REPORT AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY WITH SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC BENT, AND FIBERS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: "DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." SHOULD HAVE ALSO BEEN INCLUDED IN SUPPLEMENTAL #1 MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9339100&pc=MTA | 2023826-2019-02181 | 2019-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE(OD) ON (B)(6)2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR DIFFERENT DIOPTER LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUAITION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9010196&pc=MTA | 2023826-2019-01654 | 2019-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND SAME SIZE LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076363&pc=MTA | 2023826-2019-01717 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.75 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852361&pc=MTA | 2023826-2019-01384 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE MANUFACTURE DATE: CORRECTED TO 25-APR-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - ON (B)(6) 2019 IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED AND PATIENT HAS ISSUES WITH HER BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILLIAR BODY. SEE MFR# 2023826-2019-02272 FOR EXCHANGED REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852363&pc=MTA | 2023826-2019-01383 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5 12.1, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - ON (B)(6) 2019 IT WAS REPORTED THAT THE PROBLEM WAS NOT RESOLVED AND PATIENT HAS ISSUES WITH HER BINOCULAR VISION, NASAL PAIN, AND A CONSTANTLY CONTRACTION OF THE CILLIAR BODY. REPORTER PROVIDED PATIENT STATUS, AND STATED "WITH WEARING CONTACT LENSES THERE ARE NO ISSUES, OS IS HER DOMINANT EYE, MAYBE THE HIGHER RX ON OS CAUSE THE PROBLEMS. THE PATIENT ALSO TAKES PSYCHOTROPIC DRUGS. SHE WILL COME BACK FOR EXAM IN 1 MONTH". SEE MFR# 2023826-2019-02271 FOR EXCHANGED REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588684&pc=MTA | 2023826-2020-00072 | 2019-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. G4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703107&pc=MTA | 2023826-2020-00263 | 2019-07-16 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350976&pc=MTA | 2023826-2019-02184 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTICS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WIDTH TO BE WITHIN SPECIFICATION. HOWEVER, OVERALL LENGTH VERIFICATION WAS NOT ABLE TO BE PERFORMED DUE TO THE RETURNED CONDITION OF THE LENS. FUNCTIONAL VERIFICATION WAS FOUND TO BE OUT OF SPECIFICATIONS DUE TO THE RETURNED CONDITION OF THE LENS (FOOTPLATE BROKEN/MISSING), LENS NOT SITING PROPERLY ON LENS VERIFICATION HOLDER. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - DATE OF EVENT: (B)(6) 2019. H10 - ADDITIONAL MANUFACTURER NARRATIVE: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299869&pc=MTA | 2023826-2019-02082 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9299940&pc=MTA | 2023826-2019-02083 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND A HAPTIC TEAR. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372949&pc=MTA | 2023826-2019-02244 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND BLACKISH RESIDUE AND THE HAPTIC SCRATCHED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "PROBLEM WAS RESOLVED." SHOULD BE REMOVED AND REPLACED WITH "PROBLEM WAS NOT RESOLVED." IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9513144&pc=MTA | 2023826-2019-02481 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND STAINS ON THE HAPTIC/OPTIC. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9506333&pc=MTA | 2023826-2019-02479 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -7.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION FINDINGS 213 SHOULD BE ADDEX TO THE PREVIOUS MDR. H6- INVESTIGATION CONCLUSION 4310 SHOUDL BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403872&pc=MTA | 2023826-2020-01738 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -1.75 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE PROBLEM WAS RESOLVED. CORRECTED STATEMENT: PROBLEM WAS NOT RESOLVED. REFERENCE FILE# (B)(4) FOR REPLACEMENT LENS. H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS FIBER ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC AND OPTIC WAS DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10- REFERENCE MFR# 2023826-2020-01755 (CLAIM# (B)(4)). CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857171&pc=MTA | 2023826-2020-00580 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE LENS MAY BE EXCHANGED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5 13.2 OF -8.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A TORIC LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE: (B)(6)2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857186&pc=MTA | 2023826-2020-00581 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. THE LENS MAY BE EXCHANGED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VICM5 13.2 OF -8.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A TORIC LENS AND THE PROBLEM WAS RESOLVED. D7- EXPLANT DATE: (B)(6) 2020. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824585&pc=MTA | 2023826-2020-00491 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VICMO 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE TORIC LENS ON (B)(6) 2019 DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION CODE 10 SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913389&pc=MTA | 2023826-2020-00740 | 2019-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH BUT DIFFERENT MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED. THE DIMENSIONAL VERIFICATION FOUND THE LENS IN SPECIFICATION BUT REFRACTIVE (FUNCTIONAL) VERIFICATION FAILED. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516586&pc=MTA | 2023826-2019-02472 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND PUPIL HYPER REACTION WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "DUE TO THE PUPIL HYPER REACTION, PATIENT WAS COMPLAINING ABOUT QUALITY OF VISION." LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE 2137 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988391&pc=MTA | 2023826-2020-00954 | 2020-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND IT WAS UNKNOWN IF THE PROBLEM WAS RESOLVED, SEE MFR REPORT # 2023826-2020-00955 FOR THE EXCHANGED LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335265&pc=MTA | 2023826-2020-01562 | 2020-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED BY THE REPORTER THE LENS WAS REMOVED AND EXCHANGED ON (B)(6)2020. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025652&pc=MTA | 2023826-2020-01005 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "INDUCED ASTIGMATISM." THE SURGEON BELIEVES THE ASTIGMATISM WAS CAUSED NOT BY THE LENS BUT THE "LOAD" ON THE WOUND AND CORNEA DURING SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH FOREIGN BLOOD STAINS/FIBERS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION FINDINGS CODE 213 SHOULD BE INCLUDED IN SUPPLEMENTAL MDR#1. H6 - INVESTIGATION CONCLUSION CODE 4310 SHOULD BE INCLUDED IN SUPPLEMENTAL MDR#1. CLAIM# : (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476385&pc=MTA | 2023826-2020-01929 | 2020-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H3-DEVICE EVALUATION: FUNCTIONAL VERIFICATION WAS PERFORMED AND LENS WAS FOUND NOT TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135977&pc=MTA | 2023826-2020-01175 | 2020-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075976&pc=MTA | 2023826-2020-01088 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFO: THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. THE PROBLEM IS RESOLVED. H6-PATIENT CODE 3191 (LENS EXCHANGE, SECONDARY SURGERY). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5 - DATE OF LENS EXCHANGE: (B)(6) 2020, SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308537&pc=MTA | 2023826-2020-01498 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. ON (B)(6) 2020 A DIFFERENT DIOPTER LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7740000&pc=MTA | 2023826-2018-01148 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-01 | MTA | UNK | Misfocusing | Blurred Vision; No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN HINDAWI JOURNAL OF OPHTHALMOLOGY WAS RECEIVED ON 06-JUL-2018 THAT CITED A CLINICAL STUDY, ARTICLE TITLED 'TWO-YEAR OUTCOMES OF VISION IMPLANTABLE COLLAMER LENS WITH A CENTRAL HOLE FOR CORRECTING HIGH MYOPIA.' THE ARTICLE REPORTS 1 EYE LOST ONE LINE OF CDVA. REPORTEDLY, THE LOSS OF CDVA WAS DUE TO AGE-RELATED CATARACT PROGRESSION AND WAS NOT DUE TO THE ICL. THE ICL WAS EXPLANTED, PHACOEMULSIFICATION AND IOL WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (SECONDARY SURGERY, LENS EXCHANGE, PHACOEMULSIFICATION). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7891566&pc=MTA | 2023826-2018-01420 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/+1.0/69 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER DUE TO OBSERVATION OF REFRACTIVE SURPRISE AND REFRACTIVE CHANGE OVERTIME. IT WAS REPORTED THAT THE DESIRED DEGREE OF MYOPIA WAS NEVER ACHEIVED AND OVER TIME IT DRIFTED EVEN LESS MYOPIC. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION WAS RECEIVED. REPORTER INDICATED THAT THE EXCHANGED LENS RESOLVED THE PROBLEM. LAST REPORT STATES THAT VISION IS IMPROVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. THE LENS WAS RETURNED IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7918080&pc=MTA | 2023826-2018-01453 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VTICMO13.2, -8.50/5.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). REFRACTIVE SURPRISE WAS OBSERVED AND ROTATION/REPOSITIONING WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE DOCTOR DID THE REPOSITIONING LAST WEEK AND HE IS WAITING FOR PATIENT'S FOLLOW UP TO SEE THE RESULT". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION: " THE ICL WAS ROTATED ANTI-CLOCKWISE TO THE DESIRED AXIS. REFRACTION IN HER LAST FOLLOW UP VISIT (3 WEEKS FOLLOWING THE PROCEDURE) IS AS FOLLOWS: -0.25 -0.75 AXIS 70 WITH UCVA OF 20/25. PATIENT IS HAPPY". (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10404437&pc=MTA | 2023826-2020-01755 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -02.25 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE , SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7928374&pc=MTA | 2023826-2018-01397 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE PATIENT REPORTED HER NEAR VISION WAS POOR. THE PATIENT COULDN'T SEE CLEARLY FROM CLOSE (40CM), SO SHE REQUESTED TO REMOVE THE ICL LENS AND DIDN'T WANT ANOTHER LENS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT; LENS WAS CAUGHT IN VIAL CAP. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8198020&pc=MTA | 2023826-2018-01926 | 2018-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A VICMO13.2, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A VAULT VALUE OF 300, LENS TEAR BEFORE/DURING LOADING, LENS DISLOCATION/SUBLUXATION AND BLURRED VISION WAS REPORTED. REPOSITIONING WAS REPORTED AS A SECONDARY INTERVENTION. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTION: EVENT - THE REPORTER INDICATED THAT A VICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT THE LENS WAS FOUND TORN/BROKEN BEFORE OPENING THE VIAL. LENS DISLOCATION/ SUBLUXATION ALONG WITH BLURRED VISION AND LENS REPOSITIONING WERE ALSO REPORTED. THE VAULT VALUE WAS REPORTED AS 300 M. ON (B)(6) 2018 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION THE DEVICE WAS THE CAUSE OF THE EVENT. DEVICE CODE 2993: PREVIOUS CODE NOT APPLICABLE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7927069&pc=MTA | 2023826-2018-01471 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE PATIENT STATES THAT THE SURGEON PERFORMED THE "SIGHT TEST INCORRECTLY, RESULTING IN THE AXIS OF MY ASTIGMATISM BEING AT LEAST 4-10 DEGREES OUT." REPORTEDLY, "INSTEAD OF ORDERING NEW LENSES TO CORRECT THE PRESCRIPTION," THE SURGEON FIT THE LENS "AT AN ANGLE." PATIENT REPORTS POOR AND BLURRED VISION, VISION BEING "MUCH POORER THAN I WOULD'VE EXPECTED," ESPECIALLY IN LOWER LIGHT CONDITIONS. ACCORDING TO THE PATIENT, "IT'S A STRUGGLE TO PERFORM MY DAILY ACTIVITIES." STATES PREVIOUSLY WAS ABLE TO ACHIEVE CLOSE TO 20/20 VISION WITH GLASSES. THIS CLAIM IS FOR THE LEFT (OS) EYE, REGARDING EVO VISIAN ICLS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7927076&pc=MTA | 2023826-2018-01473 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-02 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE PATIENT STATES THAT THE SURGEON PERFORMED THE "SIGHT TEST INCORRECTLY, RESULTING IN THE AXIS OF MY ASTIGMATISM BEING AT LEAST 4-10 DEGREES OUT." REPORTEDLY, "INSTEAD OF ORDERING NEW LENSES TO CORRECT THE PRESCRIPTION," THE SURGEON FIT THE LENS "AT AN ANGLE." PATIENT REPORTS POOR AND BLURRED VISION, VISION BEING "MUCH POORER THAN I WOULD'VE EXPECTED," ESPECIALLY IN LOWER LIGHT CONDITIONS. ACCORDING TO THE PATIENT, "IT'S A STRUGGLE TO PERFORM MY DAILY ACTIVITIES." STATES PREVIOUSLY WAS ABLE TO ACHIEVE CLOSE TO 20/20 VISION WITH GLASSES. THIS CLAIM IS FOR THE RIGHT (OD) EYE, REGARDING EVO VISIANICLS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8052105&pc=MTA | 2023826-2018-01626 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/1.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SUPRISE WAS OBSERVED, THE SURGEON DECIDED TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). POST SURGERY RESULTS ARE PENDING AT THE TIME OF THIS REPORT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATENT IS FINE AT THE MOMENT, STRUGGLING WITH RX DUE TO REFRACTIVE SUPRISE. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS USER ERROR. SURGEON USED THE WRONG AXIS OF THE RX. Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -7.50/1.5/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) -2018. REFRACTIVE SURPRISE WAS OBSERVED. IT WAS REPORTED THAT THE SURGEON USED THE WRONG AXIS OF THE RX. ON (B)(6) 2018, REPORTER STATED: "PATIENT CONDITIONS ARE NOT BAD, BUT THE UCVA IS ONLY 20/25 BECAUSE OF THE REFRACTIVE SURPRISE". ON (B)(6) 2018 IT WAS STATED: "PATIENT IS FINE-AT THE MOMENT, HE'S ONLY A BIT STRUGGLING WITH HIS RX ON OS (BECAUSE OF REFRACTIVE SURPRISE)". THE SURGEON PLANS TO DO A REPOSITIONING OF THE LENS (19 DEGREES CCW). IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). "THE ROTATION HAS NOT BEEN DONE DUE TIME PROBLEMS OF THE PATIENT, YET IT IS PLANNED FOR MID OF DECEMBER". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE LENS WAS REPOSITIONED ON (B)(6) 2019, "STABLE REFRACTION AFTER REPOSITIONING." AND THE PROBLEM WAS RESOLVED. PATIENT CODE ADDITIONAL INFORMATION: 3191 - SECONDARY SURGICAL INTERVENTION, REPOSITIONING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974926&pc=MTA | 2023826-2018-01528 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/+2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7974933&pc=MTA | 2023826-2018-01527 | 2018-04-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 AND (B)(6) 2018 THE LENS WAS RE-POSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATES "SUSPECTED ZONULAR WEAKNESS THAT COULD CAUSED LENS ROTATION MULTIPLE TIMES." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7710867&pc=MTA | 2023826-2018-01071 | 2018-05-18 | Injury | STAAR SURGICAL COMPANY | 2018-07-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+6.0/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS REPORTED AS HAVING A REFRACTIVE SURPRISE AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER STATED THAT THE "DOCTOR DECIDED TO ROTATE THE LENS BUT AFTER SIX DAYS THE LENS RETURNED TO ITS INITIAL POSITION OBTAINING THE SAME REFRACTIVE SURPRISE. THE DOCTOR WILL ORDER THE SAME LENS WITH THE SAME POWER BUT A BIGGER SIZE BECAUSE THE VAULT AFTER THESE DAYS HAS DECREASED." (B)(4). Manufacturer Narrative: ADVERSE EVENT" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. 'REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICE)" NOT INCLUDED IN THE SUPPLEMENTAL MDR#2. (B)(4). Manufacturer Narrative: LENS WAS EXCHANGED FOR A SIMILAR POWER AND A LARGER SIZE ON (B)(6) 2018. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7851176&pc=MTA | 2023826-2018-01292 | 2018-07-24 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICM5_12.1, -13.00+/4.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018, LENS REMAINS IMPLANTED, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182343&pc=MTA | 2023826-2018-01887 | 2018-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -0.50/+6.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE "LENS WAS NOT STABLE" ON (B)(6) 2019. ON (B)(6) 2019 THE REPORTER STATED, "THE SURGEON TRIED TO CORRECT THE REFRACTIVE SURPRISE WITH A REPOSITIONING OF BOTH LENSES. FOR THIS HE ROTATED THE MISALIGNED TICLS TO HE AXIS OF THE IODS. AFTER THAT (1DAY POSTOP) FIRST THE REFRACTIONS WERE GOOD AND EVERYTHING WAS FINE. AT THE NEXT EXAMINATION MID OF (B)(6) THE TICLS HAVE ROTATED AGAIN TO THE SAME MISALIGNED AXIS AS BEFORE. IN THIS POSITION THE LENSES WERE NOW STABLE AND SO THE SURGEON DECIDED TO MAKE A LASER TOUCH UP TO CORRECT THE RESIDUAL REFRACTION ON BOTH EYES". PATIENT CODE: 3191 SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: (B)(4). NOT INCLUDED IN THE SUPPLEMENTAL MDR#2 CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT HAD DECIDED NOT TO HAVE A LVC PROCEDURE TO CORRECT CURRENT REFRACTIVE ERROR AFTER TICL IMPLANTATION. THE PLAN IS TO EXPLANT THE LENS. CORRECTED DATA: PATIENT CODE: 3191 - (SECONDARY SURGICAL INTERVENTION, LASER TOUCH UP) - PREVIOUS CODE NOT APPLICABLE, PATIENT DECIDED NOT TO HAVE LASER TOUCH UP (LVC). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272716&pc=MTA | 2023826-2019-00152 |  | Injury | STAAR SURGICAL COMPANY | 2019-01-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+2.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD ROTATED AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, THE LENS ROTATED AGAIN. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO LENS ROTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048919&pc=MTA | 2023826-2016-01514 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS - +8.00/+6.00/93 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. CORRECTION DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6725096&pc=MTA | 2023826-2017-01217 | 2017-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: CORRECTED DATA: CONCLUSION: PREVIOUS CODE IS NOT APPLICABLE, THIS LENS MODEL IS CONTRAINDICATED IN THE PRESENCE OF ANY OF THE FOLLOWING CIRCUMSTANCES AND/OR CONDITIONS SUCH AS FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: IT WAS REPORTED ON (B)(6) 2019 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050057&pc=MTA | 2023826-2017-01758 | 2017-09-08 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -12.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REPORTEDLY, ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO LOSS OF BCVA. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND MOISTURE ON THE LENS BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589617&pc=MTA | 2023826-2018-00878 | 2018-04-25 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) INSPECTION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A SAME MODEL/DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048973&pc=MTA | 2023826-2020-01038 | 2020-03-25 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; CALCULATIONS WERE PERFORMED INCORRECTLY BY THE USER. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048977&pc=MTA | 2023826-2020-01037 | 2020-03-25 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE AND MODEL, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; CALCULATIONS WERE PERFORMED INCORRECTLY BY THE USER. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227439&pc=MTA | 2023826-2019-00056 | 2009-03-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/029 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY AND EXCHANGE WENT WELL!" (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227424&pc=MTA | 2023826-2019-00053 | 2009-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -13.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE SURGEON REPORTS LOW VAULT AND REFRACTIVE CHANGE OVER TIME. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS REMAINS IMPLANTED. UPDATED INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED. "PATIENT IS HAPPY AND EXCHANGE WENT WELL!" PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (LENS EXCHANGE, SECONDARY SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107521&pc=MTA | 2023826-2018-01749 | 2016-02-27 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN ICM120V4, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. "DUE TO OBSERVATION OF LOW VAULT" SHOULD BE CORRECTED TO "DUE TO OBSERVATION OF LOW VAULT AND REFRACTIVE CHANGE OVERTIME." CONCOMITANT MEDICAL PRODUCT: CARTRIDGE MODEL AND LOT NOT APPLICABLE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8108029&pc=MTA | 2023826-2018-01748 | 2016-02-27 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN ICM120V4, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. "DUE TO OBSERVATION OF LOW VAULT" SHOULD BE CORRECTED TO "DUE TO OBSERVATION OF LOW VAULT AND REFRACTIVE CHANGE OVERTIME." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7773875&pc=MTA | 2023826-2018-01187 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICMO13.2, -13.0/+5.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 REPOSITIONING WAS PERFORMED DUE TO MISALIGNMENT. ONE DAY AFTER REPOSITIONING UCVA/BCVA WAS 20/40 AND 20/20, IOP WAS 15 MMHG, VAULT- 902 MICRONS. RX WAS THE SAME AS BEFORE MANIPULATION. IT WAS REPORTED THAT THE SURGEON PLANS TO REPLACE THE ICL DUE TO THE RESIDUAL REFRACTIVE ERROR AND HIGH VAULT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107279&pc=MTA | 2023826-2018-01671 | 2018-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A LONGER LENS DUE TO BLURRED VISION AND LOW VAULT. THE REPORTER STATES, "EXCHANGE DIDN'T RESOLVE LOW VAULT, PATIENT COMPLAINT ABOUT BLURRED VISION BUT UCVA ACHIEVED PRE-OP BCVA." REFERENCE MFR REPORT# 2023826-2018-01665 FOR CASE CONTINUATION. Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021583&pc=MTA | 2023826-2018-01572 | 2018-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+1.0/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2018. THE "VAULT WAS GREATLY IMPROVED" AND THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND BLUE RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624221&pc=MTA | 2023826-2020-00147 | 2018-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE AND DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7790898&pc=MTA | 2023826-2018-01237 | 2018-07-09 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6,-5.5/+1.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS REMAINS IMPLANTED AND THE SURGEON IS NOT SURE IF THE LENS WILL BE EXPLANTED OR EXCHANGED, THE SURGEON WILL BE MONITORING OVER THE NEXT MONTHS. PATIENT'S CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824085&pc=MTA | 2023826-2018-01269 | 2018-07-09 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS REFRACTIVE SURPRISE. THE CAUSE OF THE EVENT IS NOTED AS "MISMATCH OF LENS SIZE." ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER, DIFFERENT POWER AND MODEL LENS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765084&pc=MTA | 2023826-2018-01171 | 2018-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.0/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT AND REFRACTIVE SURPRISE. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL DRY WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS REMOVED ON AN UNKNOWN DATE. ON (B)(6) 2019 A LONGER LENGRTH REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. H6- PATIENT CODE 3191: SECONDARY SURGERY; LENS EXPLANT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8164913&pc=MTA | 2023826-2018-01819 | 2018-10-14 | Injury | STAAR SURGICAL COMPANY | 2018-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, REFRACTIVE SURPRISE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED WITH A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS "RPORTED" AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). "WORD" ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8241995&pc=MTA | 2023826-2019-00087 | 2018-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION, AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE STATUS OF THE EYE IS NOTED AS "GOOD". CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843596&pc=MTA | 2023826-2019-01339 | 2019-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE SURGEON REPORTS REFRACTIVE CHANGE OVER TIME. THE LENS WAS EXCHANGED WITH A LONGER, DIFFERENT MODEL LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SPHERICAL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY AND THE POST OP UCVA IS 20/20. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D7-EXPLANT DATE : (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SPHERICAL AND LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY AND THE POST OP UCVA IS 20/20. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D7- EXPLANT DATE (B)(6) 2019 IS CORRECTED TO (B)(6) 2019. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894316&pc=MTA | 2023826-2019-01430 | 2019-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL WITHIN A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894317&pc=MTA | 2023826-2019-01429 | 2019-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON B)(6) 2019. ON B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL WITHIN A ZIPLOC BAG. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979191&pc=MTA | 2023826-2020-00927 | 2019-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, SURGEON "USED LRI TECHNIQUE MAKING THE MAIN INCISION SUPERIORLY TO REDUCE WTR CYLINDER AND WAS ABLE TO REDUCE BUT NOT FULLY DUE TO HIGH AMOUNT OF CYLINDER POWER. SURGEON FELT CURRENT VAULTING WAS TOO LOW(230R/262L) AND NEEDS TO CHANGE THE LENS WITH DIFFERENT POWER AND LENGTH" WHILE "RE-ENTERING MAIN INCISION TO DECREASE ASTIGMATISM FURTHER AND INCREASE -0.50 MORE MYOPIC LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN SUPPLEMENTAL MDR #1 B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMAPAIRMENT IN SUPPLEMENTAL MDR #1 H6 - MEDICAL DEVICE PROBLEM CODE 1583 SHOULD BE INCLUDED IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979197&pc=MTA | 2023826-2020-00926 | 2019-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS REMAINS IMPLANTED. PER THE REPORTER ON (B)(6) 2020, SURGEON "USED LRI TECHNIQUE MAKING THE MAIN INCISION SUPERIORLY TO REDUCE WTR CYLINDER AND WAS ABLE TO REDUCE BUT NOT FULLY DUE TO HIGH AMOUNT OF CYLINDER POWER. SURGEON FELT CURRENT VAULTING WAS TOO LOW (230R/262L) AND NEEDS TO CHANGE THE LENS WITH DIFFERENT POWER AND LENGTH" WHILE "RE-ENTERING MAIN INCISION TO DECREASE ASTIGMATISM FURTHER AND INCREASE -0.50 MORE MYOPIC LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS EXCHANGED ON (B)(6) 2020 FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM." H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION- LENS EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN SUPPLEMENTAL MDR #1. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMAPAIRMENT IN SUPPLEMENTAL MDR #1. H6 - MEDICAL DEVICE PROBLEM CODE 1583 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960165&pc=MTA | 2023826-2020-00866 | 2019-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416976&pc=MTA | 2023826-2019-02201 | 2019-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782973&pc=MTA | 2023826-2020-00407 | 2020-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/40 (PRE-OP BCVA WAS 20/20) SO CLAIM # (B)(4) (MFR REPORT # 2023826-2020-00408) WAS OPENED. TO DATE, THE BCVA HAS NOT INCREASED BUT VAULT IS BETTER. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT SHOULD HAVE BEEN SPECIFIED THAT THE LENS WAS EXPLANTED AND REPLACED IN A SEPERATE SURGERY ON THE SAME DATE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8547398&pc=MTA | 2023826-2019-00644 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +04.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711588&pc=MTA | 2023826-2017-01176 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -15.5/+2.0/65 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS WAS REPOSITIONED ON(B)(6)2015. THE PROBLEM WAS RESOLVED, THE LENS IS IN POSITION AT DESIRED AXIS. AS OF THE DATE OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, AND WILL NOT RETURN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2 -15.5/+2.0/065 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS REPOSITIONED ON (B)(6) 2015 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2018 IT WAS REPORTED THAT REFRACTIVE SURPRISE AND LOW VAULT WITH ROTATION WERE OBSERVED IN THE PATIENT; THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2018 WHICH RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BENT AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194494&pc=MTA | 2023826-2018-01903 | 2018-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS REPOSITIONED. THE SURGEON REPORTS LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: EVENT- LENS WAS EXPLANTED. EXPLANT DAT- UNK. PATIENT CODE 3191: NO CODE AVAILABLE - SECONDARY SURGERY, EXPLANT. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #: (B)(4). Manufacturer Narrative: DESCRIBE EVENT OR PROBLEM: ADDITIONAL INFORMATION: ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM IS RESOLVED, THE PATIENT IS "VERY HAPPY!" DISTRIBUTOR PREVIOUSLY CHECKED IN ERROR. PATIENT CODE 3191 PREVIOUSLY STATED LENS EXPLANT UPDATED, 3191 SHOULD STATE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8227441&pc=MTA | 2023826-2019-00057 | 2018-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." Manufacturer Narrative: ADDITIONAL INFORMATION: REFER TO MFR REPORT # 2023826-2019-01063 FOR CLAIM# (B)(4). FOR FURTHER INVESTIGATION DETAILS. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL INFORMATION: PROBLEM IS NOT RESOLVED DUE TO THERE STILL BEING REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: PREVIOUSLY STATED LENS WAS EXPLANTED AT A LATER DATE. UPDATED INFORMATION: LENS WAS EXCHANGED ON 15-JAN-2019 WITH A LONGER, DIFFERENT MODEL LENS. PROBLEM IS RESOLVED. DISTRIBUTOR WAS CHECKED IN ERROR. PREVIOUSLY STATED LENS EXPLANT. UPDATED-LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 -11.0/+5.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES LOW VAULT, LENS ROTATION, AND REFRACTIVE SURPRISE DUE TO THE ROTATION OF THE LENS. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS RE-ROTATED AFTER REPOSITIONING, THE SURGEON STATING "NON-STABLE SULCUS POSITION (TOO SHORT ICL)." THE LENS WAS EXPLANTED ON A LATER DATE. ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS AND IF PROBLEM WAS RESOLVED HAS BEEN REQUEST. IF ADDITIONAL DETAILS ARE AVAILABLE A SUPPLEMENTAL MDR WILL BE SUBMITTED. DEVICE EVALUATION- THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND RESIDUE ON LENS SURFACE. PATIENT CODE- 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172154&pc=MTA | 2023826-2018-01843 | 2018-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.5/+2.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTES REFRACTIVE SURPRISE AND LENS DISLOCATION. ON (B)(6) 2018 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. SURGEON ALSO STATES THE VAULT IS 200UM AND PLANS TO IMPLANT A LONGER LENS. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019, THE LENS WAS REPOSITIONED A 2ND TIME. REPORTEDLY, THE "PATIENT AND SURGEON ARE HAPPY WITH THE RESULTS." THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172170&pc=MTA | 2023826-2018-01851 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -18.0/+3.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS ROTATION, PIGMENT DISPERSION, AND BLURRED VISION. REPORTEDLY, THE ICL HAD ROTATED TO THE VERTICAL POSITION. IT IS NOTED THAT THE PATIENT DOES NOT WANT TO UNDERGO ANOTHER OPERATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE STATUS OF THE EYE IS NOTED AS "CALM" AFTER THE EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7508629&pc=MTA | 2023826-2018-00723 | 2010-06-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -08.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2010. THE REPORTER INDICATED THE LENS HAD A LOW VAULT AND THERE WAS REFRACTIVE CHANGE OVERTIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. THE LENS WAS EXCHANGED FOR A LONGER LENS (13.7) BUT DIFFERENT MODEL AND DIOPTER - VTICM5_13.7, -6.00/+1.0/X086 (SPHERE/CYLINDER/AXIS). THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT. MALFUNCTION TO SERIOUS INJURY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488984&pc=MTA | 2023826-2018-00677 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND REFRACTTIVE SURPRISE. THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE FOR ANOTHER LENS. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018 AND EXCHANGED FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. (B)(4). CORRECTED DATA: PRODUCT PROBLEM TO ADVERSE EVENT MALFUNCTION TO SERIOUS INJURY (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL VERIFICATION WAS PERFORMED ON THE LENS. INVESTIGATION FOUND LENS WAS WITHIN SPECIFICATION AT THE TIME OF RELEASE. RESULTS OF RETURNED LENS FUNCTIONAL EVALUATION ARE INCONCLUSIVE; LENS MAY HAVE BEEN MISSHAPEN DURING PERIOD LENS REMAINED IMPLANTED OR DUE TO HANDLING DURING EXPLANT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887952&pc=MTA | 2023826-2019-01353 | 2019-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. BLURRED VISION, INTERMITTENT ANGLE CLOSURE WITHOUT ELEVATED IOP, AND EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. PER THE REPORTER, "DOCTOR THOUGHT THERE WAS NO RELATED FACTOR WITH PRODUCT." Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND DEBRIS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8693088&pc=MTA | 2023826-2019-01082 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR THE SAME MODEL AND DIOPTER LENS DUE TO BLURRED VISION AND ACCRETION ON THE LENS, LENS OPACITY. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "THERE WERE WHITE MATERIALS ADHERED ON THE LENS AND THE LEVEL DID NOT GET WELL BUT WORSE, THEY DECIDED TO REMOVE THE LENS". Manufacturer Narrative: H6 - DEVICE CODE 1426 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - FOREIGN BODY ANALYSIS: LIGHT-COLORED DEBRIS ON THE LENS WAS ANALYZED USING FOURIER TRANSFORM INFRARED SPECTROSCOPY (FTIR) AND WAS FOUND TO BE A MIXTURE OF PROTEIN AND CARBOHYDRATE, CONSISTENT WITH A BIOLOGICAL MATERIAL. H6 - RESULT CODE 213 CORRECTED TO RESULT CODE 114 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS, NO WHITE PARTICULATE. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965322&pc=MTA | 2023826-2019-01611 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; No Apparent Adverse Event | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. WHITE ACCRETION ON THE LENS AND LENS OPACITY WERE OBSERVED. RINDERON EYE DROPS INITIALLY. CURRENT TREATMENT FLUOROMETHOLONE EYE DROPS (PATIENT IS AN ATHLETE, CANNOT USE STEROIDS FOR LONG). THE LENS REMAINS IMPLANTED. PER THE REPORTER ON PATIENT'S (B)(6) 2019 VISIT, "BLURRED VISION HAS BEEN GENERATED FROM THREE WEEKS BEFORE THIS DAY." Manufacturer Narrative: RESULT CODE: 213 SHOULD BE CORRECTED TO 114, AS IT IS NOT APPLICABLE. DEVICE CODE: 1401 SHOULD HAVE BEEN INCLUDED. PATIENT CODE 3191: NO CODE AVAILABLE (PATIENT WAS TREATED WITH MEDICATIONS, RINDERON AND FLUOROMETHLONE) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1426. ADVERSE EVENT SHOULD HAVE BEEN SELECTED IN INITIAL MEDWATCH REPORT. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN SELECTED IN INITIAL MEDWATCH REPORT. AT A FOLLOW UP VISIT IT WAS REPORTED THAT THE "WHITE ACCRETIONS OF THE PATIENT'S GOT SMALLER WITHOUT SPECIFIC MEDICAL TREATMENT. DOCTOR THINKS THE RISK OF RECURRENCE WILL BE LOW. REPORTED LENS OPACITY MEANS THAT ICL LENS OPACITY WAS DUE TO WHITE ACCRETIONS AND NOT PATIENT'S CATARACT." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136933&pc=MTA | 2023826-2019-01806 | 2018-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138555&pc=MTA | 2023826-2020-01118 | 2018-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -05.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED WITH A SAME SIZE TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - PATIENT CODE: 2137 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574235&pc=MTA | 2023826-2019-00674 | 2019-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM512.1, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED, THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE CODE: 1494 - OFF-LABEL USE SHOULD BE REMOVED FROM PREVIOUS MEDWATCH REPORT, AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: AT THE LATEST DISCUSSION WITH THE SURGEON, THE SURGEON BELIEVES THAT IT IS POSSIBLE HE MISTAKINGLY IMPLANTED THE LENS MEANT FOR THE PATIENT'S OTHER EYE. THE SURGEON ISN'T CONSIDERING EXCHANGING THE LENS. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE SPHERE TO BE OUT OF SPECIFICATIONS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552454&pc=MTA | 2023826-2019-00645 | 2019-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS REFRACTIVE SURPRISE, BLURRY VISION, NARROW ANGLES, AND "1+ PIGMENTED CELL". THE LENS REMAINS IMPLANTED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: SERIAL NUMBER, EXP. DATE, UDI: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8617956&pc=MTA | 2023826-2019-00819 | 2019-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED ALONG WITH THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION OF THE LENS SPHERE WAS FOUND TO BE 0.45D OUT OF SPECIFICATION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMAITON: THE REPORTER INDICATED THAT THAT A 13.7MM, VICM5_13.7, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR, PATIENT HAD SOME PREBYOPIA DUE TO AGE. ALTHOUGH THE PATIENT HAD GOOD VISION AFTER INITIAL LENS IMPLANT, THE PATIENT WAS DISSATISFIED WITH NEAR VISION. THE INITIAL LENS WAS EXCHANGED BECAUSE THE PATIENT WANTED BETTER NEAR VISION. THERE ARE CURRENTLY NO ADVERSE EVENTS IN THE PATIENT AND NO COMPLAINTS FROM THE PATIENT. CORRECTED DATA DATE OF BIRTH: CORRECTED FROM (B)(6) 2019 TO (B)(6) 1968. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: PREVIOUS MDR SHOULD INCLUDE "USE/HANDLING AND ADDITIONAL PROCESSING IMPACTS LENS DIMENSIONS/FUNCTIONS; RETURNED LENS MAY NOT MEET ORIGINAL SPECIFICATIONS." DATE RECEIVED BY MANUFACTURER: DATE IN INITIAL MDR SHOULD BE CORRECTED TO 18-APR-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8618086&pc=MTA | 2023826-2019-00820 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: THE REPORTER INDICATED THAT THAT A 13.7MM, VICM5_13.7, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR, PATIENT HAD SOME PREBYOPIA DUE TO AGE. ALTHOUGH THE PATIENT HAD GOOD VISION AFTER INITIAL LENS IMPLANT, THE PATIENT WAS DISSATISFIED WITH NEAR VISION. THE INITIAL LENS WAS EXCHANGED BECAUSE THE PATIENT WANTED BETTER NEAR VISION. THERE ARE CURRENTLY NO ADVERSE EVENTS IN THE PATIENT AND NO COMPLAINTS FROM THE PATIENT. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTICS WERE BROKEN, BENT AND DEFORMED. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). PATIENT DATE OF BIRTH: CORRECTED FROM (B)(6) 2019 TO (B)(6) 1968. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF REPORT: "17-APR-2019" SHOULD BE CORRECTED TO "18-APR-2019" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: "17-APR-2019" SHOULD BE CORRECTED TO "18-APR-2019" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9513282&pc=MTA | 2023826-2019-02482 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VICM5_12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND STAINS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. D4- EXPIRATION DATE: 31-DEC-2021. H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). B5 - THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043596&pc=MTA | 2023826-2020-01031 | 2019-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -03.50 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND LENS WAS EXCHANGED ON (B)(6) 2019 WITH A TORIC LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED MOISTURIZED IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9379814&pc=MTA | 2023826-2019-02245 | 2019-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS/ CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF-LABEL USE; ACD < 3.00MM SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412469&pc=MTA | 2023826-2019-02274 | 2019-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -16.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES LOSS OF BCVA AND REFRACTIVE SURPRISE. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXCHANGED FOR A VTICMO12.6, -17.50/0.5/088 (SPHERE/CYLINDER/AXIS), LENS ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION: LENS EXCHANGE) CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988419&pc=MTA | 2023826-2020-00948 | 2019-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS BUT THE NEAR VISION WAS SLIGHTLY IMPROVED BUT STILL NOT IDEAL. PATIENT WILL BE MONITORED. SEE MFR REPORT # 2023826-2020-00949 FOR EXCHANGED LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398560&pc=MTA | 2023826-2020-01750 | 2020-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D11- NOT APPLICABLE IN THE INITAL MDR. ADDITIONAL INFORMATION: H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAMEE LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774158&pc=MTA | 2023826-2019-01195 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT HAS AN IMPLANTABLE COLLAMER LENS IMPLANTED. THE SURGEON REPORTS A CASE OF LARGE REFRACTIVE SURPRISE (HIGH CYLINDER) AFTER IMPLANTING IT IN PIGGYBACK SCENARIO. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: UPDATED INFORMATION RECEIVED: THE SURGEON PERFORMED A ROTATION OF THE ICL AND THIS SOLVED THE PROBLEM. VISION IS IMPROVED AND CYLINDER DECREASED. PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, SURGEON ROTATED LENS). (B)(4). Manufacturer Narrative: UNK-ESUBMITTER DOES NOT ALLOW FOR A BLANK OR "UNK" ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814242&pc=MTA | 2023826-2020-00473 | 2019-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8850039&pc=MTA | 2023826-2019-01369 | 2019-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 12.6MM, MICL 12.6, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE REPORTER STATED, "BLURRY AND DOUBLE VISION POST OP, NO PATIENT INJURY. PROBLEM RESOLVED AFTER EXPLANTING THE LENS". PATIENT IS STABLE, VISION WAS BETTER AND CLOUDINESS WAS GONE. NO PLANS TO IMPLANT ANOTHER LENS, PATIENT DID NOT WANT THE LENS REPLACED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE; ACD < 3.00MM. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440448&pc=MTA | 2023826-2020-01894 | 2020-01-02 05:00 | Injury | STAAR SUGICAL COMPANY | 2020-08-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1, -6.5 DIOPTER, INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND DIFFERENT MODEL LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717957&pc=MTA | 2023826-2020-00334 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE ISSUE. Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE: 3001 SHOULD BE CORRECTED TO 1401. CLAIM#:(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628975&pc=MTA | 2023826-2019-00891 | 2019-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 17 FEB 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL/FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 2682 PATIENT-DEVICE INCOMPATIBILITY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639513&pc=MTA | 2023826-2020-00188 | 2019-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. REPORTEDLY, THE LENS WAS EXCHANGED ON (B)(6) 2019 WITH A SAME SIZE, DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - DEVICE CODE: 2682 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416217&pc=MTA | 2023826-2020-01720 | 2020-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER (-06.50) LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8066004&pc=MTA | 2023826-2018-01665 | 2018-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO BLURRED VISION AND "ICL TILTED." THE REPORTER STATES, "AFTER FIRST EXCHANGE, ICL LENS TILTED, SURGEON DIDN'T WANT TO EXCHANGE ANOTHER ICL. PATIENT IS OK." REFERENCE MFR REPORT# 2023826-2018-01671 FOR INFORMATION REGARDING FIRST LENS IMPLANT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: METHOD CODE 4110: LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS IN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729355&pc=MTA | 2023826-2019-01113 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENTS NEXT VISIT IS IN (B)(6), NO TREATMENT HAS BEEN PERFORMED AND THE PATIENTS CURRENT CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -4.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, DOCTOR ADMITTED THAT THE LENS WAS UPSIDE DOWN. HE REMOVED THE LENS AND RE-IMPLANTED THE LENS. IT WAS REPORTED THAT THE PATIENT IS DOING VERY WELL NOW. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REMOVED AND RE-IMPLANTED LENS. CLAIM#: (B)(4). Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD HAVE BEEN INDICATED IN SUPPLEMENTAL MDR#1 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8729476&pc=MTA | 2023826-2019-01114 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENTS NEXT VISIT IS IN (B)(6), NO TREATMENT HAS BEEN PERFORMED AND THE PATIENTS CURRENT CONDITION REMAINS THE SAME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATES THAT A 12.1MM, VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. REFRACTIVE SURPRISE AND LOSS OF BCVA WAS OBSERVED, DOCTOR ADMITTED THAT THE LENS WAS UPSIDE DOWN. HE REMOVED THE LENS AND RE-IMPLANTED THE LENS. IT WAS REPORTED THAT THE PATIENT IS DOING VERY WELL NOW. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, REMOVED AND RE-IMPLANTED LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8051091&pc=MTA | 2023826-2018-01593 | 2018-06-28 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2018-11-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM12.6, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. REFRACTIVE SUPRISE WAS OBSERVED, LENS REMAINS IMPLANTED. REMOVAL AND EXCHANGED IS CURRENTLY BEING PLANED TO RESOLVE THE ISSUE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THE ORIGINAL MDR STATED "LENS REMAINS IMPLANTED" WHICH SHOULD BE UPDATED TO "LENS REMOVED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADVERSE EVENT SHOULD HAVE BEEN ADDED IN SUPPLEMENTAL MEDWATCH REPORT #2. REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #2. UNK SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #2. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: 12-SEP-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747635&pc=MTA | 2023826-2020-00314 | 2019-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9747651&pc=MTA | 2023826-2020-00315 | 2019-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6111670&pc=MTA | 2023826-2016-01655 | 2016-09-29 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Positioning Failure; Inadequacy of Device Shape and/or Size | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, AND THE LENS WENT INTO THE PATIENT'S EYE UPSIDE DOWN. THE REPORTER INDICATED THERE WAS A REFRACTIVE SURPRISE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION TO DATE OF REPORT: 09/30/2016. ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. (B)(4): SECONDARY SURGERY, LENS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, SAME DIOPTER AND SIZE. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201647&pc=MTA | 2023826-2018-00091 | 2017-07-01 | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Shelf Life Exceeded; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -4.50/+1.5/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A LONGER LENS, THE SURGEON NOTING LOW VAULT, LENS ROTATION, AND LOSS OF BCVA. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5523847&pc=MTA | 2023826-2016-00327 | 2015-09-14 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND BLURRED VISION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT HAPTIC WAS TORN AND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. (B)(4). AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN A PATIENT UNDER THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7169824&pc=MTA | 2023826-2018-00053 | 2017-12-06 | Malfunction | STAAR SURGICAL COMPANY | 2018-01-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; No Consequences Or Impact To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS IMPLANTED AND REFRACTIVE SURPRISE WAS OBSERVED AFTER SURGERY. THE REPORTED STATED THAT TIME CORRECTED THE PROBLEM AND THAT THE LENS REMAINS IMPLANTED. LENS REMOVAL WAS NOT NECESSARY, PATIENT IS DOING GREAT AND IS VERY HAPPY. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5654727&pc=MTA | 2023826-2016-00659 | 2010-12-28 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/092 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND ROTATION AND REMAINS IMPLANTED. IT IS PLANNED TO EXPLANT THE LENS. Manufacturer Narrative: THE SURGEON PERFORMED A SECONDARY INTERVENTION TO REPOSITION THE LENS TO THE ORIGINAL AXIS. AN EXCHANGE IS PLANNED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. OUTCOME ATTRIBUTED TO ADVERSE EVENT: UNK. EXPLANT DATE: N/A. DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6808530&pc=MTA | 2023826-2017-01329 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE. LENS ROTATION WAS NOTICED FOUR MONTHS AFTER THE IMPLANTATION. THE DOCTOR STATED THAT THERE HAS BEEN NO INCIDENCE OF TRAUMA TO THE PATIENT. Manufacturer Narrative: PATIENT EXPEREINCED LOSS OF BCVA AND REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECT DATE FOR FOLLOW-UP #1 IS 09/13/2017. FOLLOW-UP #1 SHOULD BE "ADDITIONAL INFORMATION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6808548&pc=MTA | 2023826-2017-01330 |  | Malfunction | STAAR SURGICAL COMPANY | 2017-08-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). LENS ROTATION WAS NOTICED IN THE EYE AND CAUSING REFRACTIVE CHANGE. THE DOCTOR STATED THAT THERE HAS BEEN NO INCIDENCE OF TRAUMA TO THE PATIENT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECT DATE FOR FOLLOW-UP #1 IS 09/13/2017. FOLLOW-UP #1 SHOULD BE "ADDITIONAL INFORMATION". CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365365&pc=MTA | 2023826-2017-00279 | 2017-01-13 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. Manufacturer Narrative: DEVICE EVALUATION-LENS RETURNED DRY IN SMALL CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DRY RESIDUE AND DEBRIS ON LENS AND PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9717743&pc=MTA | 2023826-2020-00331 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S EYE IN (B)(6) 2019. REFRACTIVE SURPRISE WAS OBSERVED. REPORTEDLY"PATIENT IS CONSIDERING ICL EXCHANGE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2137. H6 - RESULT CODE: 213 SHOULD BE CORRECTED TO 114. (B)(4). Manufacturer Narrative: IMPLANT DATE: (B)(6) 2019. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8965298&pc=MTA | 2023826-2019-01562 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; No Apparent Adverse Event; Appropriate Term/Code Not Available | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE IN 2019. WHITE ACCRETION ON THE LENS, BLURRED VISION, AND LENS OPACITY WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY THERE IS A POSSIBILITY THE LENS MAY BE EXPLANTED IN THE FUTURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE: 1426 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2137. DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1401. DEVICE CODE: 3191 NO CODE AVAILABLE (OPACITY ON THE LENS) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: 2019. IMPLANTED: 2019. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW; BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764121&pc=MTA | 2023826-2019-01156 | 2019-02-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | Blurred Vision; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -05.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REPORTEDLY "POST-OP REFRACTION, SHOWED HYPEROPIC SHIFT (+1.0D)" IN PATIENT EYE. PER REPORTER ON (B)(6) 2019, "PATIENT'S UCVA WAS 20/13 AND POST-OP REFREATION SHOWED +1.0/-0.25/109 ON THE LAST VISIT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT CODE: 2692 SHOULD BE CORRECTED TO 2137 IN INITIAL MEDWATCH REPORT. DEVICE CODE: 3189 SHOULD BE CORRECTED TO 1401 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8219629&pc=MTA | 2023826-2019-00045 | 2018-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available; Missing Value Reason | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_13.2, -9.5/+1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE CAUSE OF EVENT WAS A PATIENT RELATED FACTOR. SURGEON PLANS TO EXCHANGE THE LENS IN (B)(6); IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: ADVERSE EVENT SHOULD BE MARKED IN PREVIOUS MDR. REQUIRED INTERVENTION SHOULD BE MARKED IN PREVIOUS MDR. REFRACTIVE SURPRISE WAS ALSO REPORTED. THE LENS WAS EXCHANGED FOR A LONGER LENGTH, STRONGER POWER LENS. SHOULD BE ADDED TO THE INITIAL MDR. DATE RECEIVED BY MFR: SHOULD BE CORRECTED TO "OCTOBER 14,2019" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: "LENS REMAINS IMPLANTED" SHOULD BE CORRECTED TO "LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. "THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR"SHOULD BE CORRECTED TO " THE CAUSE OF THE EVENT WAS USER ERROR." IN THE INITIAL MDR. ADDITIONAL INFORMATION: SECONDARY SURGICAL INTERVENTION - EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARK. ETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361178&pc=MTA | 2023826-2017-00262 | 2017-01-02 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Blurred Vision; Overwear Syndrome | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15/0.5/149 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED EXCESSIVE VAULT, GLARE/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. THE LENS REMAINS IMPLANTED AS OF (B)(6) 2017. Manufacturer Narrative: CORRECTED DATA: EXCESSIVE VAULT, GLARES/HALOES, AND PHOTOPHOBIA WAS OBSERVED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCHES: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039570&pc=MTA | 2023826-2023-01884 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. TOXIC ANTERIOR SEGMENT SYNDROME(TASS) AND BLURRED VISION WAS OBSERVED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITION OF ORAL STEROIDS AND OINTMENTS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WASREPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: H6: HEALTH IMPACT CODE: "4625 - ADDITIONAL SURGERY: ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "2137" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163013&pc=MTA | 2023826-2020-01257 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -09.0, IMPLANTABLE COLLAMER LENS INTO PATIENT'S LEFT EYE (OS) IN 2009. REPORTEDLY, LENS "HAS ROTATED." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6 - SHOULD BE CORRECTED TO "H6 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT", IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVER TIME ("VISION STARTED CHANGING"). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163014&pc=MTA | 2023826-2020-01256 |  | Malfunction | STAAR SURGICAL COMPANY AG | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4, -09.0, IMPLANTABLE COLLAMER LENS INTO PATIENT'S RIGHT EYE (OD) IN 2009. REPORTEDLY, LENS "HAS ROTATED." LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THE PATIENT HAD EXPERIENCED REFRACTIVE CHANGE OVER TIME (PATIENT'S VISION STARTED TO CHANGE). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - SHOULD BE CORRECTED TO "H6 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT", IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: IMPLANTED DATE: 2009. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595254&pc=MTA | 2023826-2023-00997 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A DIOPTER OF -9.0 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL DUE TO REFRACTIVE SURPRISE AND BLURRED VISON. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "SURGICAL INCISION INDUCED MORE ASTIMATISM THAN EXPECTED," AND "REFRACTIVE MISS - PATIENT HAD 1.D OF ASTIGMASTISM AFTER SPHERICAL ICL PLACEMENT, DESPITE HAVING ONLY 0.5D PRE-OP SO I EXCHANGED THE SPHERICAL ICL FOR A TORIC." THE STATUS OF THE EYE IS REPORTED AS, "HEATHY." Manufacturer Narrative: WEIGHT-RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- RETURNED TO MANUFACTURER: "13-MAR-2023" SHOULD BE CORRECTED TO "10-MAR-2023." B5- "CAUSE IS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED." DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00D INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND BLURRED VISION. REPORTEDLY "1D ASTIGMATISM AFTER SPHERICAL ICL PLACEMENT, DESPITE HAVING ONLY 0.50D PRE-OP". ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE DIFFERENT MODEL/ SAME LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS NOTED AS "HEALTHY". THE CAUSE FOR THIS EVENT WAS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED.- "SURGICAL INCISION INDUCED MORE ASTIGMATISM THAN EXPECTED". H6- MEDICAL DEVICE PROBLEM CODE: "2017" SHOULD BE REMOVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987063&pc=MTA | 2023826-2018-01508 | 2018-09-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2018-10-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED LENS MICL12.1, -9.5 DIOPTER, IMPLANTED IN PATIENTS RIGHT EYE ON (B)(6) 2018 IMPROVED VAULT BUT VAULT CONTINUES TO BE HIGH. AT 3 MONTH EXAM, PATIENT WAS 20/20 UDVA, WITH A +0.25 SPH (OD), AND A 2+ VAULT. PATIENT COMPLAIN IS WITH INTERMEDIATE AND NEAR VISION AFTER LONG DAYS. PATIENT HAS CHOSEN NOT TO EXPLANT LENS. ADDITIONAL INFORMATION, PILO 1% AND LIMIFY HAVE BENN OF MINIMAL HELP. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7995562&pc=MTA | 2023826-2018-01555 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A MICL12.1, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018 AS A REPLACEMENT LENS. AT 3 MONTH POST-OP EXAM, THE PATIENT WAS 20/20 UDVA, WITH A PI +0.25X075 (OS), AND A 2+ VAULT. THE PUPIL WAS MINIMALLY REACTIVE AT 6-7MM. PATIENT'S COMPLAIN IS WITH INTERMEDIATE AND NEAR VISION AFTER LONG DAYS. PATIENT HAS CHOSEN NOT TO EXPLANT THE LENS. PILO 1% AND LUMIFY OF MINIMAL HELP. Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794738&pc=MTA | 2023826-2022-01891 | 2021-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REPORTEDLY GLARES/HALOES AND BLURRED VISION OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2021 WHICH RESOLVED THE PROBLEM. CAUSE IS DUE TO THE DEVICE. STATUS OF EYE IS WITHIN NORMAL LIMITS. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880650&pc=MTA | 2023826-2021-04519 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL13.7, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT GLARE/HALOES BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ABERATIONS IMPROVED BUT STILL PRESENT. PER UPDATED INFORMATION THE REPORTER REPORTED THAT THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS RESIDUAL REFRACTIVE ERROR AND GLARE. LASIK FOR RESIDUAL ASTIGMATISM IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880665&pc=MTA | 2023826-2021-04520 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL13.7, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT GLARE/ HALOES BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED "GLARE/ ABERATIONS IMPROVED BUT STILL PRESENT. PER UPDATED INFORMATION THE REPORTER REPORTED THAT THE LENS REMAINS IMPLANTED BUT THE PATIENT HAS RESIDUAL REFRACTIVE ERROR AND GLARE. LASIK FOR RESIDUAL ASTIGMATISM IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072260&pc=MTA | 2023826-2021-02079 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, DIOPTER +7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), GLARE/HALOS, AND BLURRED VISION. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: PREOP BCVA REPORTED AS 20/200, POST EXPLANT BCVA 20/40. ON (B)(6) 2021 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE IS REPORTED AS "INTERNAL ANATOMY UNPREDICTED." CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6329973&pc=MTA | 2023826-2017-00236 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-02-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.50 INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2016 (EXACT DATE UNKNOWN). AFTER IMPLANTATION, THE PATIENT COMPLAINED OF SEEING GLARES, HALOES, AND BLURRED VISION. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AS OF DATE OF SUPPLEMENTAL MDR SUBMISSION, THE LENS HAS BEEN PLANNED TO EXCHANGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE IMPLANT DATE REPORTED WAS (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED DUE TO GLARE/HALOES, WHICH RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. CORRECTIONAL DATA: THE CORRECT IMPLANT DATE REPORTED WAS (B)(6) 2017. (B)(4). Manufacturer Narrative: EXACT IMPLANT DATE NOT KNOWN; IMPLANT IN (B)(6) 2016. THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153640&pc=MTA | 2023826-2023-00127 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D9: RETURN DATE (B)(6) 2023. G2: FOREIGN; OTHER: (B)(6). H3: DEVICE EVALUATED: NO. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153746&pc=MTA | 2023826-2023-00128 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED WITH GLARE AND HALOES AND BLURRED VISION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: D9: RETURN DATE 06-FEB-2023. G2: FOREIGN OTHER: MALAYSIA H3: DEVICE EVALUATED: NO. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483348&pc=MTA | 2023826-2020-01951 | 2018-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2016. LENS OPACITY (ASC) WAS OBSERVED ON 15/JAN/2019. GLARE/HALOES AND BLURRED VISION WAS ALSO OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HALOS SINCE SURGERY, AT NIGHT. ALPHAGAN HELPED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: B3-DATE OF EVENT IN INITIAL MDR SHOULD BE REPORTED AS (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483350&pc=MTA | 2023826-2020-01952 | 2018-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2016. LENS OPACITY (ASC) WAS OBSERVED ON (B)(6) 2019. GLARE/HALOES AND BLURRED VISION WAS ALSO OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, "HALOS SINCE SURGERY, AT NIGHT. ALPHAGAN HELPED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION- ANALYSIS OF PRODUCTION RECORDS 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: B3-DATE OF EVENT IN INITIAL MDR SHOULD BE REPORTED AS (B)(6) 2018. CLAIM# : (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10930997&pc=MTA | 2023826-2020-02815 | 2020-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO THE LENS WAS STUCK IN CARTRIDGE/INJECTION SYSTEM AND TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE PATIENT COMPLAINED OF GLARE/HALOS AND BLURRED VISION. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE PATIENT IS STABLE WITH NO DIPLOPIA. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052011&pc=MTA | 2023826-2023-02086 | 2023-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052113&pc=MTA | 2023826-2023-02087 | 2023-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE PATIENT COMPLAINED OF GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364576&pc=MTA | 2023826-2019-02196 | 2019-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. GLARE/HALOES AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT COMPLAINED ABOUT GLARE, BLURRED VISION, AND THE IMPLANTED LENS. BUT HE STILL COMPLAINED ABOUT VISUAL IMPAIRMENT AFTER REMOVING THE LENS. PER THE REPORTER ON (B)(6) 2019, "CURRENTLY THE PATIENT IS UNDER CONTINUOUS OBSERVATION AND NO TREATMENT." Manufacturer Narrative: B5: IT WAS LATER MENTIONED THAT THE PATIENT'S VISUAL IMPAIRMENT HAS INCLUDED TO HAVE FLOATERS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED BY THE PATIENT THAT THEY HAD TRIED OUT RGP, SCLERALS LENSES BUT THIS STILL DID NOT RESOLVE THE PROBLEM AND VISION IS STILL MUCH WORSE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16375402&pc=MTA | 2023826-2023-00543 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. THE PROBLEM WAS RESOLVED. THE PATIENT COULD NOT TOLERATE THE HALO. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991203&pc=MTA | 2023826-2023-01764 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -3.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT PREFERRED LENS EXPLANT AND USE OF GLASSES. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991204&pc=MTA | 2023826-2023-01765 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION ON (B)(6)2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT PREFERRED LENS EXPLANT AND USE OF GLASSES. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021086&pc=MTA | 2023826-2018-01581 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8021095&pc=MTA | 2023826-2018-01582 |  | Malfunction | STAAR SURGICAL COMPANY | 2018-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICMO13.2, -16.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE ON (B)(6) 2018. REPORTER STATES "MY ISSUE IS THAT I CAN SEE GREAT OUTDOORS BUT SUFFER FROM PATCHES INDOORS OR UNDER NON/NATURAL LIGHT. I HAVE BEEN TO MY CHECKUPS AND THEY SAY ALL IS OK. BUT THE PATCHES ARE MAKING MY VISION INDOORS EXHAUSTIVE. WE HAVE HAD A SECOND OPINION FROM ANOTHER OPTICIAN AND THE GRAYISH AND BLURRED IMAGES IS CAUSED DUE TO THE LIGHT REFRACTION ON THE LENS ESPECIALLY IN LOW LIGHT WHEN THE PUPIL IS AT ITS BIGGEST. WE BELIEVE MY WIFE'S ISSUE IS BECAUSE SHE STILL HAS ASTIGMATISM. THIS IS THE CULPRIT OF THE BLURRY IMAGES ESPECIALLY IN LOW LIGHT CONDITIONS." TO DATE THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS UNDER TREATMENT WITH BRIMONIDINE (TO REDUCE HALOES) AND WILL BE CLOSELY MONITORED. EVENT: "VTICMO13.2, -16.00/2.0/90" SHOULD BE CORRECTED TO "VTICMO13.2, -16.00/2.0/94" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11661950&pc=MTA | 2023826-2021-01047 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. PATIENT IS EXPERIENCING GLARE/HALOES, BLURRED VISION, CIRCLES IN VISION AND RAINBOW HALOES. LENS REMAINS IMPLANTED. THE REPORTER STATED " I BELIEVE HE IS GOING TO WAIT A COUPLE OF MONTHS AND SEE IF HE CAN ADAPT TO THE LENSES." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION " THE PATIENT IS NOTICING AN IMPROVEMENT IN THEIR SYMPTOMS AND IS HOLDING OFF ON MAKING A DECISION UNTIL WINTER." CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11662621&pc=MTA | 2023826-2021-01048 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE REPORTER STATED " HE IS COMPLAINING OF HALOS/GLARES AND THEN A REFRACTIVE SURPRISE." IT WAS LATER REPORTED THAT THE PATIENT IS ALSO EXPERIENCING BLURRED VISION, CIRCLES IN VISION AND RAINBOW HALOES. LENS REMAINS IMPLANTED. THE REPORTER STATED " I BELIEVE HE IS GOING TO WAIT A COUPLE OF MONTHS AND SEE IF HE CAN ADAPT TO THE LENSES." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION " THE PATIENT IS NOTICING AN IMPROVEMENT IN THEIR SYMPTOMS AND IS HOLDING OFF ON MAKING A DECISION UNTIL WINTER." CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (OHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279827&pc=MTA | 2023826-2020-01502 | 2020-02-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE PATIENT COMPLAINED SHE COULD NOT SEE WELL, WAS SEEING HALOS, LINES AND VIBRATION OF THE LENS. HER NEAR VISION WAS BLURRED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 STATEMENT - "PATIENT IS DEPRESSED" SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H1: TYPE OF REPORTABLE EVENT IN INITIAL MDR SHOULD BE CORRECTED TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391264&pc=MTA | 2023826-2023-03041 | 2022-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON 08/SEP/2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND BLURRED VISION. PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391269&pc=MTA | 2023826-2023-03040 | 2022-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Patient-Device Incompatibility; Optical Problem | Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6, 16.0 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS AND BLURRED VISION. PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8412661&pc=MTA | 2023826-2019-00418 | 2018-08-14 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICMO 13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AND THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. PATIENT CURRENT CONDITION IS GOOD. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414779&pc=MTA | 2023826-2019-00419 | 2018-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATD THAT THE SURGEON IMPLANTED A 13.2MM, VICMO13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO GLARE/HALOES AND BLURRED VISION. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER SHOULD BE 14-FEB-2019 IN INITIAL MDR (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16210270&pc=MTA | 2023826-2023-00247 | 2021-11-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.50/+3.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 06-SEP-2021. THE PATIENT EXPERIENCED LENS ROTATION; UNREACTIVE (FIXED) PUPIL; AND IMMEDIATELY AFTER SURGERY; GLARE/HALOES AND BLURRED VISION. LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNREACTIVE (FIXED) PUPIL. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473809&pc=MTA | 2023826-2024-00133 | 2023-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: EXCESSIVE VAULT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES, PIGMENT DISPERSION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16329543&pc=MTA | 2023826-2023-00464 | 2022-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Blurred Vision; Visual Disturbances; Halo; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOES AND DISCOMFORT DUE TO OVER CORRECTION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- OVER CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8379600&pc=MTA | 2023826-2019-00351 | 2018-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Blurred Vision; Visual Disturbances; Halo; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/039 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, WITH GLARE/HALOS AND BLURRED VISION. THE PATIENT HAD DISCOMFORT WITH THEIR VISION IN COMPARISON WITH THE ICL LENS IN THE RIGHT EYE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. METHOD CODE 4117 NO LONGER APPLICABLE. INITIAL MDR SHOULD HAVE INCLUDED DEVICE CODE 1494: OFF-LABEL USE, PRE-EXISTING HISTORY OF KERATOCONUS AND CORNEAL CROSS-LINKING. INITIAL MDR SHOULD HAVE INCLUDED METHOD CODE 4111. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE PATIENT IS STABLE AND HAS RETURNED TO PREOPERATIVE REFRACTION. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2019. D7 - UPDATED TO (B)(6) 2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S., BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876813&pc=MTA | 2023826-2017-01461 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -14.5/+2.0/071 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT BEGAN EXPERIENCING GLARE/HALOES AND BLURRED VISION. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER AND SIMILAR DIOPTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DATE OF THIS REPORT: IT IS CORRECTED TO "(B)(6) 2017" FROM "(B)(6) 2017" IN THE INITIAL. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191998&pc=MTA | 2023826-2016-01811 | 2016-11-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND GLARE/HALOES. THE LENS REMAINS IMPLANTED IN THE PATIENTS EYE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL BUT DIFFERENT DIOPTER AND THE PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). METHOD: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752257&pc=MTA | 2023826-2017-01144 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -7.5/+1.0/92 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6752269&pc=MTA | 2023826-2017-01143 | 2016-12-24 | Injury | STAAR SURGICAL COMPANY | 2017-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -9/+2/88 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO GLARE/HALOES, BLURRED VISION, AND FOREIGN BODY SENSATION. THERE IS NO PLANS TO EXCHANGE FOR A NEW LENS. ON (B)(6) 2017, PATIENT'S BCVA WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7135224&pc=MTA | 2023826-2017-01924 | 2017-10-26 | Injury | STAAR SURGICAL COMPANY | 2017-12-20 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/+4.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTED REFRACTIVE SURPRISE, BLURRED VISION, GLARE/HALOES. REPORTEDLY, THE LENS WAS REPOSITIONED. NO ADDITIONAL INFORMATION AVAILABLE AT THE TIME OF MDR SUBMISSION. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THE REPORTER STATES THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2018. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY, LENS REPOSITIONING H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438817&pc=MTA | 2023826-2018-00563 | 2013-01-22 | Injury | STAAR SURGICAL COMPANY AG | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE SURGEON REPORTS LOW VAULT, REFRACTIVE CHANGE OVER TIME, GLARE/HALOS, AND BLURRED VISION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684682&pc=MTA | 2023826-2017-01095 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BCVA, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION, GLARE/HALOES AND PUPIL MODERATELY DILATED. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH VICMO12.6 AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND UNCORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). SECONDARY SURGERY INTERVENTION, EXCHANGE, PIGMENT DISPERSION, SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES,PUPIL MODERATELY DILATED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7074520&pc=MTA | 2023826-2017-01826 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LOW VAULT, REFRACTIVE CHANGE OVERTIME, READING PROBLEMS, BLURRED VISION, AND GLARE/HALOS WAS NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7075520&pc=MTA | 2023826-2017-01827 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LOW VAULT, REFRACTIVE CHANGE OVERTIME, READING PROBLEMS, BLURRED VISION, AND GLARE/HALOS WAS NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: METHOD: NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540586&pc=MTA | 2023826-2017-00742 | 2016-11-29 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT STATED THAT HE EXPERIENCED GLARE/HALOES AT NIGHT. THE LENS REMAINS IMPLANTED, THE PATIENT FEELS THE SIDE EFFECT ARE NOT SIGNIFICANT. Manufacturer Narrative: PREVIOUS: (B)(6) 2017, CORRECTED: (B)(6) 2016. ADDITIONAL INFORMATION: PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, AND IRIS ATROPHY. PREVIOUS: (B)(4; CORRECTED: (B)(4). (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT STATED THAT HE EXPERIENCED GLARE/HALOES AT NIGHT, EXCESSIVE VAULTING, LOSS OF BCVA AND IRIS ATROPHY. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906082&pc=MTA | 2023826-2018-01439 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. GLARE/HALOES AND BLURRED VISION WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7918895&pc=MTA | 2023826-2018-01454 | 2018-06-03 | Injury | STAAR SURGICAL COMPANY | 2018-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO13.7, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND BLURRED VISION. ON (B)(6) 2018 A SHORTER LENGTH LENS WAS IMPLANTED. THIS REPLACEMENT LENS RESOLVED THE PROBLEM. IT WAS REPORTED THAT THE MAIN REASON OF REMOVAL WAS THE HIGH VAULT THAT RESULTED IN A DILATED UNREACTIVE PUPIL. PATIENT COMPLAINED MAINLY FROM HALOES AND DISRUPTED IMAGES. Manufacturer Narrative: (B)(6). (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "30-AUG-2018" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269224&pc=MTA | 2023826-2019-00146 | 2018-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARE/HALOS AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT STILL COMPLAINS OF GLARE/HALOS AND BLURRED VISION. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. POST-OP PATIENT CONDITION - DRY CORNEAL EPITHELIUM. SEE MFR REPORT # 2023826-2019-00147 FOR EXCHANGED LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634216&pc=MTA | 2023826-2020-00202 | 2018-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. PRK, REFRACTIVE TREATMENT WAS ALSO REPORTED AS A SECONDARY INTERVENTION. IT WAS REPORTED THAT THE EXCHANGE LENS RESOLVED THE PROBLEM. AT PATIENTS POST OP FOLLOW UP VISIT " SOME GLARE AT NIGHT " AND " S/P PRK DOING BETTER, MUCH BETTER VAULT" WAS NOTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS ICL (HIGH VAULT CAUSING SECONDARY ABERRATIONS, REFRACTORY TO CONTACT LENS CORRECTION). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN SPECIMEN CONTAINER WITH RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9634632&pc=MTA | 2023826-2020-00203 | 2018-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, GLARE/HALOES AND BLURRED VISION. PRK, REFRACTIVE TREATMENT WAS ALSO REPORTED AS A SECONDARY INTERVENTION. IT WAS REPORTED THAT THE EXCHANGE LENS RESOLVED THE PROBLEM. AT PATIENTS POST OP FOLLOW UP VISIT " GLARE AT NIGHT "AND " S/P PRK DOING BETTER, MUCH BETTER ICL VAULT" WAS NOTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS ICL (HIGH VAULT CAUSING SECONDARY ABERRATIONS, REFRACTORY TO CONTACT LENS CORRECTION). Manufacturer Narrative: CORRECTED DATA: D6: SHOULD BE "(B)(6) 2018" IN THE INITIAL MDR. ADDITIONAL INFORMATION: H6- WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN SMALL CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7911897&pc=MTA | 2023826-2018-01448 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER EXPLANTING THE LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7912042&pc=MTA | 2023826-2018-01449 | 2017-07-20 | Injury | STAAR SURGICAL COMPANY | 2018-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATINETS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND BLURRED VISION. IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED AFTER EXPLANTING THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190948&pc=MTA | 2023826-2020-01349 | 2017-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: (B)(6) 2017. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189334&pc=MTA | 2023826-2020-01281 | 2017-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATIONS. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906085&pc=MTA | 2023826-2018-01440 | 2018-08-23 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. GLARE/HALOES AND BLURRED VISION WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A MICRO-CENTRIFGUE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269212&pc=MTA | 2023826-2019-00147 | 2018-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT COMPLAINED OF GLARE/HALOS AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THIS IS THE EXCHANGED LENS FOR MFR REPORT # 2023826-2019-00146. Manufacturer Narrative: CHANGED FROM "ADVERSE EVENT" TO "PRODUCT PROBLEM". CLAIM#: (B)(4). Manufacturer Narrative: DATE OF REPORT, DATE RECEIVED BY MANUFACTURER CORRECTED TO 24-DEC-2018 FOR INITIAL MDR. CONCOMITANT MEDICAL PRODUCTS: PREVIOUS CARTRIDGE MODEL LOT NUMBER AND FOAM TIP PLUNGER MODEL LOT NUMBER NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2019 THE LENS WAS EXPLANTED AND POSTOPERATIVE THE EYE IS "QUIET" AND THE "PATIENT WAS FEELING FINE". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364572&pc=MTA | 2023826-2019-02197 | 2019-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. GLARE AND BLURRED VISION WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT COMPLAINED ABOUT GLARE, BLURRED VISION, AND THE IMPLANTED LENS. BUT HE STILL COMPLAINED ABOUT VISUAL IMPAIRMENT AFTER REMOVING THE LENS. PER THE REPORTER ON (B)(6) 2019, "CURRENTLY THE PATIENT IS UNDER CONTINUOUS OBSERVATION AND NO TREATMENT." Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 3191 SHOULD BE IN THE INITIAL MDR. H6- PATIENT CODE 2227 SHOULD BE ADDED IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS LATER INDICATED BY THE PATIENT THAT THEY HAD TRIED RGP/SCLERALS LENSES BUT THIS DID NOT RESOLVE THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516581&pc=MTA | 2023826-2019-02471 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL HYPER REACTION, GLARE/HALOES, AND BLURRED VISION WAS OBSERVED. LENS WAS REMOVED ON (B)(6) 2019. REPORTEDLY "DUE TO THE PUPIL HYPER REACTION. PATIENT WAS COMPLAINING ABOUT QUALITY OF VISION." LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE 3191 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR H6 - PATIENT CODE 3191 - 'SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE' CORRECTED TO 'SECONDARY SURGICAL INTERVENTION, LENS EXPLANT' IN INITIAL MDR H6 - PATIENT CODE 3191 - NO CODE AVAILABLE, PUPIL HYPER REACTION SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320740&pc=MTA | 2023826-2018-00322 | 2016-12-01 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+1.5/180 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2016. ON (B)(6) 2017 THE LENS WAS EXCHANGED. REPORTEDLY, THE PATIENT EXPERIENCED "PROBLEMS READING AND DOING COMPUTER WORK." THE SURGEON STATES THE EYE IS "CALM AND QUIET" AFTER THE EXCHANGE AND THAT THE EYE IS 0.5 MORE MYOPIC THAN EXPECTED. THE SURGEON ALSO REPORTS THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE AND "PROBLEMS READING AND DOING COMPUTER WORK." (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6396523&pc=MTA | 2023826-2017-00367 | 2017-02-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE LENS HAS YET TO BE EXPLANTED AS OF (B)(6) 2017 AND THE PATIENT EXPERIENCED GLARE/HALOES AND BLURRED VISION. Manufacturer Narrative: THE PATIENT EXPERIENCED GLARE AND HALOS DUE TO CENTRAL AQUAPORT HOLE. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS. THE PROBLEM IS RESOLVED, SYMPTOMS DISAPPEARED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213739&pc=MTA | 2023826-2016-01837 | 2014-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE LENS WAS EXCHANGED DUE TO LOW VAULTING AND 'REFRACTIVE CHANGE OVER TIME'. NO REPORT OF REFRACTIVE SURPRISE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6520730&pc=MTA | 2023826-2017-00707 | 2015-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -7.0/+1.5/089 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT, REFRACTIVE SURPRISE, AND REFRACTIVE CHANGE OVER TIME. THE SURGEON STATED THAT THE ICL IS ON THE AXIS. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: STATED LENS REMAINS IMPLANTED. LENS WAS EXCHANGED ON (B)(6) 2017 FOR A LARGER, DIFFERENT DIOPTER LENS. THE PROBLEM WAS RESOLVED. FINAL POST-OP RESULTS: VAULT 500UM, UCVA 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND RED RESIDUE), AND A PIECE OF HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6461375&pc=MTA | 2023826-2017-00573 | 2015-10-24 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, DIOPTER -16.5/+5.5/088 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT AND LENS ROTATION. IN ADDITION, THE PATIENT SUFFERED SIGNIFICANT VISION LOSS AND GLARE WHILE DRIVING AT NIGHT THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727937&pc=MTA | 2023826-2015-00470 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Blurred Vision; Visual Impairment | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT WAS HAVING VISUAL ACUITY PROBLEMS. THE PATIENT COMPLAINED OF GHOSTING/BLURRING VISION. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE SURGEON IS GOING TO WAIT UNTIL THE EYE HEALS BEFORE IMPLANTING ANOTHER LENS. SUTURES WERE NOT REQUIRED AND THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: PT WEIGHT: UNK. DATE OF EVENT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED BY FACILITY. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - MANY FACTORS COULD CAUSE VISUAL ACUITY CHANGE SUCH AS UNDER CORRECTED REFRACTION AND PATHOPHYSIOLOGICAL CHANGE OF VISUAL SYSTEM. GIVEN THE REPORTED INFORMATION, THE CAUSE OF THE EVENT IS UNKNOWN. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY, ICL (12.6 MM;-8.5 D) WAS EXPLANTED SIX WEEKS POSTOPERATIVELY TO ADDRESS PATIENT VISUAL DISTURBANCES ("GHOSTING"/"BLURRY VISION"). ACCORDING TO THE REPORT IMPLANTATION OF ANOTHER LENS WAS NOT SCHEDULED YET. THE PATIENT WAS DOING WELL. THE SAME REPORT STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. VISUAL DISTURBANCES HAVE BEEN IDENTIFIED AS POTENTIAL COMPLICATIONS AFTER ICL IMPLANTATION. THE OPTICAL DIAMETER OF THE ICL RANGES FROM 4.9 MM TO 5.8 MM AND THE PRECAUTIONS SECTION OF THE DFU INSTRUCTS PHYSICIANS THAT "PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE". IT FURTHER PRECAUTIONS THAT "THE EFFECT OF PUPIL SIZE ON VISUAL SYMPTOMS IS NOT KNOWN". THE DFU INDICATES THE "SMALLER THE OPTIC DIAMETER, THE GREATER THE INCIDENCE OF SUBJECTIVE PATIENT SYMPTOMS". (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727929&pc=MTA | 2023826-2015-00469 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT WAS HAVING VISUAL ACUITY PROBLEMS. THE PATIENT COMPLAINED OF GHOSTING/BLURRING VISION. THE INCISION WAS ENLARGED TO REMOVE THE LENS. THE SURGEON IS GOING TO WAIT UNTIL THE EYE HEALS BEFORE IMPLANTING ANOTHER LENS. SUTURES WERE NOT REQUIRED AND THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED BY FACILITY. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY FACTORS COULD CAUSE VISUAL ACUITY CHANGE SUCH AS UNDER CORRECTED REFRACTION AND PATHOPHYSIOLOGICAL CHANGE OF VISUAL SYSTEM. GIVEN THE REPORTED INFORMATION, THE CAUSE OF THE EVENT IS UNKNOWN. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6146503&pc=MTA | 2023826-2016-01717 | 2015-02-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Impairment | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -4.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE CUSTOMER EXPERIENCED REFRACTIVE SURPRISE AND LOSS OF BEST CORRECTED VISUAL ACUITY (PRE-OP 20/20 / POST-OP 20/25). CUSTOMER COMPLAINED FROM FIRST POST-OP DAY, BUT A FEW MONTHS LATER BECAME PREGNANT. THE LENS REMAINS IMPLANTED. PATIENT TO SEE SURGEON FOR CORRECTION AFTER DELIVERY OF BABY. Manufacturer Narrative: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5240530&pc=MTA | 2023826-2015-01521 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | NA | 1 | Event Description: THE PATIENT REPORTED THROUGH SOCIAL MEDIA HAD A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IMPLANTED IN THEIR LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT REPORTED FOUR DAYS AFTER SURGERY, WAS UNABLE TO READ ANYTHING/BLURRY. TO DATE THE PATIENT'S VISION HAS NOT IMPROVED. THE PATIENT WAS PRESCRIBED EYE DROPS (NEVANAC AND GENTEAL GEL) FOR A MONTH. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572294&pc=MTA | 2023826-2016-00409 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE CHANGE OVER TIME. THE REPORTER INDICATED THE PLAN IS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542548&pc=MTA | 2023826-2015-00121 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Blurred Vision; Visual Impairment; Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD) AND IS NOW EXPERIENCING GHOST/BLURRY IMAGES OF OBJECTS IN LOW LIGHT AND DAY LIGHT. THE PATIENT ALSO HAS GLARES/HALOS/FOG DURING THE DAY AND AT NIGHT THE PUPIL DILATES MORE THAN THE DIAMETER OF THE LENS CAUSING BLURRINESS. THE SURGEON PRESCRIBED ALPHAGAN DROPS WHICH MAKES HIS EYES RED, TEARY AND CRUSTY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE: UNK. PT WEIGHT: UNK. EVENT DATE: UNK. EXPIRATION DATE - UNK. IMPLANT AND EXPLANT DATES: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: LENS EDGE EFFECTS MAY BE A SOURCE OF OPTICAL ABERRATIONS-SUCH AS GLARE AND HALOS. DYSPHOTOPSIAS MAY OCCUR WHEN PUPILS DILATE LARGER THAN THE OPTICAL DIAMETER OF THE ICLS. BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542549&pc=MTA | 2023826-2015-00199 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Blurred Vision; Visual Impairment; Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS) AND IS NOW EXPERIENCING GHOST/BLURRY IMAGES OF OBJECTS IN LOW LIGHT AND DAY LIGHT. THE PATIENT ALSO HAS GLARES/HALOS/FOG DURING THE DAY AND AT NIGHT THE PUPIL DILATES MORE THAN THE DIAMETER OF THE LENS CAUSING BLURRINESS. THE SURGEON PRESCRIBED ALPHAGAN DROPS WHICH MAKES HIS EYES RED, TEARY AND CRUSTY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. EXPIRATION DATE - UNK. IMPLANT AND EXPLANT DATES: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: LENS EDGE EFFECTS ARE A SOURCE OF OPTICAL ABERRATIONS-SUCH AS GLARE AND HALOS. DYSPHOTOPSIAS MAY RESULT WHEN PUPILS DILATE LARGER THAN THE OPTICAL DIAMETER OF THE ICLS. BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6211496&pc=MTA | 2023826-2016-01841 | 2016-10-06 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Blurred Vision; Visual Impairment; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12/03/78 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND LOSS OF BCVA (BEST CORRECTED VISUAL ACUITY). THE LENS WAS EXCHANGED FOR ANOTHER LENS VTICM 13.2 MM -12/03/79 DIOPTER. THE PROBLEM WAS RESOLVED AFTER THE LENS WAS EXCHANGED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION AND FUNCTIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THE LENS WAS EXCHANGED FOR A 13.2 MM VTICM5_13.2, NOT A VTICM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5593263&pc=MTA | 2023826-2016-00410 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-04-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +03.50 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). THE LENS WAS REMOVED ON (B)(6) 2016 AND EXCHANGED WITH A LENS OF DIFFERENT DIOPTER, WHICH RESOLVED THE PROBLEM. (B)(4). CONCLUSION CODE: PER DFU OF THIS LENS MODEL, VICL IS CONTRAINDICATED IN PATIENTS WITH PREVIOUS CORNEAL OR REFRACTIVE SURGERY. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBERS ON LENS. DIMENSIONAL INSPECTION AND FUNCTIONAL INSPECTION CONCLUDED THE LENS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919703&pc=MTA | 2023826-2016-01192 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Torn Material; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED THE LENS TORE DURING DELIVERY INTO THE EYE AND SUBSEQUENTLY THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629094&pc=MTA | 2023826-2015-00368 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THERE WAS A REFRACTIVE SURPRISE AND THE VISUAL ACUITY WAS NOT GOOD ENOUGH. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS IMPLANTED ON (B)(6) 2014. SEVEN MONTHS AFTER THE INITIAL SURGERY, THE SURGEON DECIDED TO REPOSITION THE LENS 30 CW TO THE TARGET POSITION OF 3 CCW. HOWEVER, THE LENS ROTATED AGAIN TO THE SAME POSITION. THE SURGEON IS PLANNING TO REMOVE AND REPLACE THE LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: LOW VAULT WAS REPORTED. LENS WAS EXPLANTED SOME TIME AFTER REPOSITIONING THE LENS. DEVICE EVALUATION: AN EVALUATION WAS PERFORMED; ADDITIONAL INVESTIGATION FOUND THE DEVICE WAS IN SPECIFICATION AT TIME OF RELEASE. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AS WELL AS PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, RUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED. Manufacturer Narrative: EVALUATION METHOD: DEVICE HISTORY RECORD REVIEW RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4820232&pc=MTA | 2023826-2015-00166 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-04 04:00 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing; Insufficient Information | Blurred Vision; Visual Impairment; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THROUGH SOCIAL MEDIA THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD). THE PATIENT REPORTED THREE DAYS POST-OP, HE HAD BLURRY VISION AND COULDN'T READ SMALL PRINT. THE PATIENT REPORTED LARGE LETTERS ARE OK BUT STILL FUZZY AND NOT AS CRISP LIKE WHEN HE WORE CONTACT LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5495703&pc=MTA | 2023826-2016-00263 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/118 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5496053&pc=MTA | 2023826-2016-00264 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Blurred Vision; Visual Impairment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE LENS ROTATING AND THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY. THE LENS WAS EXCHANGED FOR A LONGER SPHERICAL LENS WITH ANOTHER DIOPTER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791519&pc=MTA | 2023826-2018-01239 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Blurred Vision; Vitrectomy; No Code Available | UNK | 1 | Event Description: ARTICLE STATED THAT THE PATIENT'S PAST OCULAR HISTORY WAS SIGNIFICANT FOR PLACEMENT OF VISIAN ICL SEVERAL YEARS EARLIER FOR THE TREATMENT OF HIGH MYOPIA. MOST RECENTLY, THE LEFT EYE (OS) ICL WAS REMOVED SECONDARY TO DISLOCATION. CATARACT SURGERY WAS PERFORMED AND A CT LUCIA IOL WAS IMPLANTED. IT WAS REPORTED THAT THE PATIENT TOLERATED THE PROCEDURE WELL. THE POSTOPERATIVE COURSE WAS UNCOMPLICATED, AND THE PATIENT RETAINED EXCELLENT DISTANCE UCVA. Manufacturer Narrative: CORRECTED DATA: AN ARTICLE STATED THAT THE PATIENT'S PAST OCULAR HISTORY WAS SIGNIFICANT FOR PLACEMENT OF VISIAN ICL SEVERAL YEARS EARLIER FOR THE TREATMENT OF HIGH MYOPIA. MOST RECENTLY THE LEFT EYE (OS) ICL WAS REMOVED SECONDARY TO DISLOCATION. FOR UNKNOWN PERIOD OF TIME AFTER THE ICL REMOVAL THE PATIENT'S LEFT EYE EXPERIENCED LOSS OF VISION. EXTREME ZONULAR INSTABILITY WAS REPORTED. CATARACT SURGERY AND VITRECTOMY WAS PERFORMED WITH A REPLACEMENT CT LUCIA IOL; AND SUTURES WERE USED TO CLOSE THE WOUND. IT WAS REPORTED THAT THE PATIENT TOLERATED THE PROCEDURE WELL. THE POSTOPERATIVE COURSE WAS UNCOMPLICATED AND THE PATIENT RETAINED EXCELLENT DISTANCE UCVA. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: "EXTREME ZONULAR INSTABILITY AFTER LENS REMOVAL" NOT INCLUDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND CATARACT SURGERY WITH IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973350&pc=MTA | 2023826-2021-04716 | 2021-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | N/A | 1 | Event Description: REPORTEDLY A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS), HOWEVER THE SURGEON REPORTS A CAPSULAR TEAR AND PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. ATTEMPTS TO OBTAIN CLARIFICATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: PER RESPONSE-THERE WAS NO PATIENT CONTACT WITH THE LENS. THE EVENT WAS CAUSED BY A SURGICAL BLADE THAT WASN'T SHARP ENOUGH. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4960647&pc=MTA | 2023826-2015-01008 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Capsular Bag Tear | NA | 1 | Event Description: THE CUSTOMER REPORTED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE SURGEON INADVERTENTLY HIT THE CAPSULE WHEN WIDENING THE INCISION. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2014. THE SURGEON WAS IN TRAINING AND EVENT WAS DUE TO A USER'S ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM) : BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. AS REPORTED BY THE SURGERY FACILITY, THE ROOT CAUSE OF THE COMPLAINT IS USER ERROR AS SURGEON WAS UNDER TRAINING. THE PRODUCT WAS NOT RETURNED TO STAAR FOR AN EVALUATION. (B)(4). Manufacturer Narrative: PATIENT'S VISION IS 6/5^-1 DESPITE MILD NOTICEABLE LENS CHANGES. REPORT SOURCE IS DISTRIBUTOR, NOT USER FACILITY. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481825&pc=MTA | 2023826-2018-00636 | 2018-02-23 | Injury | STAAR SURGICAL COMPANY | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Capsular Bag Tear; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 12.6 MM VICMO12.6, -13 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO CAPSULAR TEAR. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN; FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS); AND THE HAPTIC TORN PIECE IS MISSING. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266114&pc=MTA | 2023826-2015-01601 | 2014-06-04 | Injury | STAAR SURGICAL COMPANY | 2015-12-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Capsular Contracture; Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -12.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. A ANTERIOR SUBCAPSULAR CATARACT OPACITY WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: DIOPTER: -12.00. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513652&pc=MTA | 2023826-2015-00150 | 2014-12-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2. IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2014 AND EARLY LENS OPACITY (ASCO) WAS OBSERVED. THE SURGEON PLANS TO EXPLANT, HOWEVER, THIS LENS REMAINS IMPLANTED. SEE MFR# 2023826-2015-00149 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). THE LENS REMAINS IMPLANTED CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158221&pc=MTA | 2023826-2020-01234 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY ON (B)(6) 2020. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256324&pc=MTA | 2023826-2022-00148 | 2020-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC/PSC. PHACO-EMULSIFICATION (CCATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -18.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY-ASC/PSC. PHACO-EMULSIFICATION (CCATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION : UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION : UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980707&pc=MTA | 2023826-2020-02950 | 2020-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN 2003. ASC LEN OPACITY WAS OBSERVED ON (B)(6) 2020, 17 YEARS AFTER ICL IMPLANTATION. THE REPORTER STATED " NO FURTHER STEPS ARE NEEDED WITH OD, BECAUSE THEY INTERNALLY SWITCHED OD <-> OS. SO WITH OD ALL IS FINE." . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER REPORTER, THERE IS NO COMPLAINT ASSOCIATED TO THE RIGHT EYE (LENS SERIAL # (B)(6). CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN 2003. ASC LENS OPACITY WAS OBSERVED ON (B)(6) 2020, 17 YEARS AFTER ICL IMPLANTATION. LENS REMAINS IMPLANTED. PER NEW INFORMATION REPORTER MEANT TO SAY THAT THERE IS A MILD ASC LENS OPACITY BUT THE REFRACTIVE SURPRISE IS ONLY AT LEFT EYE (OS) AND NOT AT THE REPORTED RIGHT EYE (OD). FOR RIGHT EYE (OD) NO LENS EXCHANGE IS PLANNED RIGHT NOW. CLAIM # (B)(4). Manufacturer Narrative: 2003. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366535&pc=MTA | 2023826-2021-00476 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE IN (B)(6) 2007. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-UPDATED INFO: THE PATIENT DEVELOPED A CATARACT ASC IN THE LEFT (OS) EYE 13-14 YEARS AFTER ICL IMPLANTATION. THE ICL AND CATARACT WERE REMOVED AND A MULTIFOCAL IOL WAS IMPLANTED WITHOUT COMPLICATION. D6B-UNK. H6-HEALTH IMPACT CODE UPDATED. (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11366550&pc=MTA | 2023826-2021-00475 | 2021-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, INTO THE PATIENT'S EYE IN (B)(6) 2007. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-UPDATED INFO: IN (B)(6) 2021 THE LENS WAS EXPLANTED FROM THE RIGHT (OD) EYE DUE TO LENS OPACITY ASC. CATARACT WAS REMOVED AND A MULTIFOCAL LENS WAS IMPLANTED WITH NO COMPLICATIONS. H6-HEALTH IMPACT CODE UPDATED. (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072276&pc=MTA | 2023826-2021-02070 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: REPORTEDLY AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO "POST-VISAN CATARACTS." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2016. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS OPACITY PSC. THE REPORTER STATES, "POSTERIOR SUBCAPSULAR POLAR AGE-RELATED CATARACTS." PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523570&pc=MTA | 2023826-2021-03689 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER STATES A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO A CATARACT FORMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN GAUZE. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS THROUGHOUT THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872783&pc=MTA | 2023826-2021-04489 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A NS (NUCLEAR SCLEROTIC) CATARACT WAS OBSERVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12875660&pc=MTA | 2023826-2021-04490 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO A NS (NUCLEAR SCLEROTIC) CATARACT WAS OBSERVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783787&pc=MTA | 2023826-2021-04216 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR CATARACT) WAS OBSERVED. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13022168&pc=MTA | 2023826-2021-04813 | 2021-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC) WAS OBSERVED. PHACO-EMULSIFICATION (CATARACT) SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233024&pc=MTA | 2023826-2022-00060 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, ON (B)(6) 2021. THE REPORTER INDICATED IT APPEARED THAT A CRYSTALLINE LENS OPACITY WAS OBSERVED POST-OP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. D4: EXPIRATION DATE UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722040&pc=MTA | 2023826-2022-01933 | 2022-05-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT HAS LENS OPACITY (CATARACT ASC). IT WAS ALSO NOTED "THE INVESTIGATOR DISCUSSED THE DEVELOPMENT OF ANTERIOR SUBCAPSULAR OPACITIES WITH THE PATIENT AND THE DECISION WAS MADE TO MONITOR THE VISION FOR NOW". THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14881811&pc=MTA | 2023826-2022-02099 | 2022-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -14.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITIES) WHICH WAS OBSERVED ON (B)(6) 2022 AND NOTICED IN THE LENS PERIPHERY AT APPROXIMATELY 4 O'CLOCK WHICH WAS DISCOVERED ON SLIT LAMP EXAMINATION. IT WAS NOTED THE ASC WAS DISCUSSED THE WITH PATIENT AND A DECISION WAS MADE TO MONITOR THE VISION CLOSELY. LENS REMAINS IMPLANTED. IT IS REPORTED THE PROBLEM IS NOT RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. - SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637143&pc=MTA | 2023826-2023-03752 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2015. IT WAS INDICATED THE PATIENT DEVELOPED A CATARACT AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048071&pc=MTA | 2023826-2023-04719 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. CATARACTS WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ON (B)(6) 2023 LENS OPACITY(CATARACT) PSC, NS, WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. CAUSE OF EVENT PATIENT RELATED FACTOR: ADL COMPLAINTS, DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ENTERED INTO FORM CORRECTION: G2 - REPORTED AS FOREIGN - CORRECT TO REMOVE FOREIGN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223150&pc=MTA | 2023826-2023-05380 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -3.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT HAD LENS OPACITY, ASC/NS WHICH WAS OBSERVED ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED AND CATARACT WITH A POSTERIOR CHAMBER LENS WAS PERFORMED WAS ADDITIONALLY REPORTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176027&pc=MTA | 2023826-2016-01771 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED A CATARACT HAD DEVELOPED BUT THERE WAS NO VISION LOSS. THE ICL REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: A2: (B)(6) 1980. FACILITY CONFIRMED PATIENT REPORT AND THE DATE OF LAST VISIT WAS (B)(6) 2016. THE PATIENT WILL BE MONITORED AND VA POST-OP WAS 20/20. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414854&pc=MTA | 2023826-2019-00394 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED SHE HAS HAD VISION IMPLANTABLE COLLAMER LENSES IMPLANTED IN BOTH EYES FOR TEN YEARS AND NOW CATARACTS HAVE DEVELOPED. THE LENSES REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. SEE MFR. # 2023826-2019-00393 FOR FELLOW EYE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414858&pc=MTA | 2023826-2019-00393 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED SHE HAS HAD VISION IMPLANTABLE COLLAMER LENSES IMPLANTED IN BOTH EYES FOR TEN YEARS AND NOW CATARACTS HAVE DEVELOPED. THE LENSES REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. SEE MFR. # 2023826-2019-00394 FOR FELLOW EYE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736588&pc=MTA | 2023826-2019-01120 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-26 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A PATIENT WHO'D HAD AN ICL IMPLANTED PREVIOUSLY, CAME FOR A CONSULTATION AT ANOTHER CLINIC. THE SURGEON REPORTS A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S., CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736581&pc=MTA | 2023826-2019-01121 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN GRAEFES ARCHIVE FOR CLINICAL AND EXPERIMENTAL OPHTHALMOLOGY WAS RECEIVED ON 03-JUN-2019, PUBLISHED ONLINE 27-MAY-2019, ENTITLED 'USE OF LOW-VAULT POSTERIOR CHAMBER COLLAGEN COPOLYMER PHAKIC INTRAOCULAR LENSES FOR THE CORRECTION OF MYOPIA: A 3-YEAR FOLLOW UP.' PATIENTS HAD PIOL IMPLANTATION PERFORMED AT BINA EYE HOSPITAL IN TEHRAN, IRAN FROM AUG. 2010 - MAY 2015. THE ARTICLE AIMED TO EXAMINE CATARACT INCIDENCE AFTER IMPLANTING DIFFERENT IMPLANTABLE COLLAMER LENSES IN EYES WITH POOR VAULT. OF THE 311 PATIENTS UNDERGOING SURGERY, 14 EYES FROM 14 PATIENTS HAD LOW VAULT. THE STUDY FOUND THAT LENS OPACITY OCCURRED IN 5 OUT OF THE 316 EYES INVOLVED IN THE STUDY, HOWEVER, NONE OF THESE WERE THE EYES WITH THE ASSOCIATED LOW VAULT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK. LOT #, IMPLANT DATE: UNK. PMA/510K #: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848511&pc=MTA | 2023826-2019-01373 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: ON (B)(6) 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES LENS OPACITIES WITHOUT LOSS OF BCVA. "WE OBSERVED ANTERIOR SUBCAPSULAR OPACITIES WITHOUT EFFECT ON VISION AND WITHOUT PROGRESSION OVER TIME IN 1 MYOPIC EYE AND IN 6 HYPERMETROPIC EYES, INCIPIENT CORTICONUCLEAR CATARACT WITHOUT EFFECT ON VISION IN 1 MYOPIC EYE". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441180&pc=MTA | 2023826-2019-02346 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. IT IS UNKNOWN AS TO WHETHER THE LENS REMAINS IMPLANTED; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441184&pc=MTA | 2023826-2019-02345 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. IT IS UNKNOWN AS TO WHETHER THE LENS REMAINS IMPLANTED; ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158238&pc=MTA | 2023826-2020-01236 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY 2 - 3 WEEKS AFTER SURGERY. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G4 - DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 14-MAY-2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158240&pc=MTA | 2023826-2020-01237 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS PERIPHERY 2 - 3 WEEKS AFTER INITIAL SURGERY. THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND THE PATIENT IS HAPPY. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G4 - DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429095&pc=MTA | 2023826-2020-01826 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY RECENTLY AFTER IMPLANT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNKNOWN ENTRY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429103&pc=MTA | 2023826-2020-01825 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY RECENTLY AFTER IMPLANT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNKNOWN ENTRY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10635882&pc=MTA | 2023826-2020-02347 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | UNK | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. PER REPORTER, "5 YEAR POST-OP CATARACT" IS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10779842&pc=MTA | 2023826-2020-02541 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE EYE IN ORDER TO PERFORM CATARACT SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: PER REPORTER THE LENS THAT WAS EXPLANTED WAS IMPLANTED BY ANOTHER SURGEON APPROXIMATELY 15 YEARS AGO. THE PATIENT IS AN OLDER PATIENT WITH HIGH MYOPIA WHICH ARE RISK FACTORS FOR CATARACT FORMATION. HE IS UNABLE TO SAY WITH CERTAINTY THAT THE ICL CONTRIBUTED TO THE CATARACT FORMATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER-UNK. DATE OF EVENT-UNK. DEVICE MANUFACTURE DATE -UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900101&pc=MTA | 2023826-2020-02751 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: AN ARTICLE FROM JCRS WAS RECEIVED ON (B)(6) 2020 THAT SITED A STUDY TITLED "LONG-TERM ASSESSMENT OF CRYSTALLINE LENS TRANSPARENCY IN EYES IMPLANTED WITH A CENTRAL-HOLE PHAKIC COLLAMER LENS DEVELOPING LOW POSTOPERATIVE VAULT". THE STUDY FOUND OUT OF 24 EYES, 1 EYE HAD DEVELOPED ANTERIOR SUBCAPSULAR OPACITIES. THE PATIENT HAD A 12.6MM IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER IMPLANTED IN HER LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969935&pc=MTA | 2023826-2020-02888 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: RECEIVED A COMPLAINT FROM (B)(6), THE PATIENT REPORTED A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969950&pc=MTA | 2023826-2020-02887 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: RECEIVED A COMPLAINT FROM (B)(6), THE PATIENT REPORTED A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER (PATIENT) INDICATED THE IMPLANTING SURGEON WAS DR. (B)(6) AND THE LENS WAS IMPLANTED IN (B)(6) 2019. THE SURGEON'S OFFICE REPORTED AS PER THEIR RECORDS ALL PARAMETERS WERE ALL RIGHT TILL FIRST MONTH OF POST OP. THEY REQUESTED THE PATIENT TO RETURN FOR FOLLOW UP EVERY 6 MONTHS BUT THE PATIENT DID NOT RETURN. THEY ARE ATTEMPTING TO CONTACT THE PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11310526&pc=MTA | 2023826-2021-00357 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN IMPLANTABLE COLLAMER LENS, ON (B)(6) 2021, DUE TO A CATARACT HAD DEVELOPED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE CATARACT WAS A NUCLEAR SCLEROTIC (NS) AND A CORTICAL CATARACT. THE CAUSE OF THE EVENT WAS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) WAS PERFORMED TO REMOVE THE CATARACT AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL AFTER CATARACT SURGERY, 20/20, WITH NO COMPLICATIONS. H6: HEALTH EFFECT IMPACT CODE: 4625 - PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12019227&pc=MTA | 2023826-2021-01904 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: THE PATIENT REPORTER THAT A 11.0MM, ICH110V3, +20.0 DIOPTER, IMPLANTABLE COLLAMER LENS (ICL) WAS IMPLANTED INTO HER LEFT EYE (OS) ON (B)(6) 2000. THE PATIENT REPORTED "MASSIVE CATARACTS IN MY LEFT EYE". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- WORK ORDER SEARCH: ALL LENSES MANUFACTURED UNDER THE SAME LOT WERE INVESTIGATED AND THERE WAS NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN THE ASSOCIATED LOT. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT - RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112661&pc=MTA | 2023826-2021-02209 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER STATED " I HAVE A PATIENT WHO WAS IN THE HYPEROPIC ICL STUDY IN THE YEAR 2000 . SHE NOW HAS A CATARACT AND IS NEEDING SURGERY".ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER - RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. EXPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602578&pc=MTA | 2023826-2021-03940 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER IS REQUESTING INFORMATION ON HOW TO PROCEED WITH CATARACT SURGERY IN REGARDS TO MEASUREMENTS, ADJUSTMENTS IN IOL SELECTION AND SURGICAL REMOVAL OF THE DEVICE IN A PATIENT THAT HAD ICL IMPLANTATION 20+ YEARS AGO. Manufacturer Narrative: B5- INITIAL MFR REPORT IS NOT APPLICABLE, PER UPDATED INFORMATION THERE IS NO COMPLAINT FOR THIS PATIENT. SURGEON WAS JUST REQUESTING INFORMATION. THE REPORTER STATED, "THERE IS ABSOLUTELY NOTHING WRONG WITH THE PATIENTS ICL." CLAIM # (B)(4). Manufacturer Narrative: AGE OR DOB, SEX, WEIGHT, ETHNICITY, & RACE: UNK. EVENT DATE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637223&pc=MTA | 2023826-2021-04011 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER SUSPECTS THAT A POST-PI PATIENT NOW HAS A CATARACT CAUSED BY THE PI. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFOR:UNK. DATE OF EVENT:UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12641181&pc=MTA | 2023826-2021-04028 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD). POST-OPERATIVE OBSERVATION SHOWED THE DEVELOPMENT OF CATARACT. THE REFRACTIVE RESULT IS STABLE. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-03029. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EVENT DATE: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973360&pc=MTA | 2023826-2021-04706 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | N/A | 1 | Event Description: AN ARTICLE PUBLISHED IN GRAEFE'S ARCHIVE FOR CLINICAL AND EXPERIMENTAL OPHTHALMOLOGY WAS RECEIVED ON (B)(6) 2021 ENTITLED 'INITIAL CLINICAL OUTCOMES OF TWO DIFFERENT PHAKIC POSTERIOR CHAMBER IOLS FOR THE CORRECTION OF MYOPIA AND MYOPIC ASTIGMATISM.' THE ARTICLE STATES 1 EYE WAS SCHEDULED FOR RE-ROTATION BUT THE PATIENT DID NOT SHOW UP; 1 EYE WAS RE-ROTATED; CATARACT EXTRACTION WAS PERFORMED AFTER 4 ICL IMPLANTATIONS-NONE BECAUSE OF ANTERIOR SUBCAPSULAR OPACIFICATION; THERE WAS NO CONTACT OF THE PHAKIC IOL AND THE CRYSTALLINE LENS IN ANY OF THE COMPLICATED CASES; FEMTO-LASIK SURGERY WAS PERFORMED IN 2 EYES BECAUSE A ROTATION WOULD NOT HAVE FULLY CORRECTED THE RESIDUAL REFRACTION FULLY. PLEASE NOTE-LENS RE-ROTATION REPORTED IN MFR REP# (B)(4). 2 CASES OF ICL EXPLANT AND CATARACT REPORTED IN MFR REP# (B)(4) AND (B)(4). 2 CASES OF FEMTO-LASIK SURGERY REPORTED IN MFR REP# (B)(4). AND (B)(4). Manufacturer Narrative: B5-RECEIVED ADDITIONAL INFORMATION INCLUDING ASSOCIATED SERIAL NUMBER REGARDING THE 1 CASE OF LENS ROTATION AND 1 EYE WITH CATARACT DEVELOPMENT. SEE MFR REP# 2023826-2022-00147 AND 2023826-2022-00148 RESPECTIVELY FOR ASSOCIATED REPORTS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: DATE OF EVENT: UNK. PROCODE: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY MODEL #, IMPLANT DATE, EXPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13544065&pc=MTA | 2023826-2022-00528 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE 15 YEARS AGO. THE LENS WAS EXPLANTED WITH SUBSEQUENT FEMTOSECOND CATARACT SURGERY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14121142&pc=MTA | 2023826-2022-01174 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, -21.00 DIOPTER IN THE PATIENTS LEFT EYE (OS). AT ONE YEAR POST-OP, THE EYE DEVELOPED A VISUALLY SIGNIFICANT ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT PHACO ASPIRATION OF THE CATARACT WITH ICL IN SITU AND A STANDARD POWER PCIOL WAS IMPLANTED. ONE YEAR AFTER THE LEFT EYE CATARACT SURGERY, THE PATIENT HAD UDVA OF 20/120, CDVA OF 20/40 AND SUBJECTIVE REFRACTION OF +3.00 DS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138558&pc=MTA | 2023826-2022-01175 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, -21.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). EIGHT YEARS POST-OP, THE EYE DEVELOPED AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT PHACOASPIRATION OF THE CATARACT WITH ICL IN SITU AND A SINGLE PIECE HYDROPHOBIC PCIOL WAS IMPLANTED. ONE YEAR AFTER THE RIGHT EYE CATARACT SURGERY, THE PATIENT HAD CDVA OF 20/40 AND SUBJECTIVE REFRACTION OF -0.75 DC@10 DEGREES. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760487&pc=MTA | 2023826-2022-01965 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: IT WAS REPORTED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED ANTERIOR SUBCAPSULAR OPACITY (CATARACT). IT WAS NOTED PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H10 - IT WAS FOUND CLAIM WAS A DUPLICATE OF CLAIM# (B)(4), MFR# 2023826-2022-01933. DISREGARD INITIAL MEDWATCH REPORT. DETAILS ARE N/A. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621351&pc=MTA | 2023826-2022-03696 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621428&pc=MTA | 2023826-2022-03697 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT LEFT (OS) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7; -15.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) ON (B)(6) 2021. THE SURGEON REPORTS THE PATIENT HAS DEVELOPED POSTERIOR CAPSULAR CATARACTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039552&pc=MTA | 2023826-2023-01649 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: A PRESENTATION TITLED "INFLUENCE OF EXPERIENCE AND THE SURGICAL LEARNING CURVE ON OUTCOMES OF IMPLANTABLE COLLAMER LENS SURGERY",IN A MAJOR REVIEW OF POTENTIAL COMPLICATION IN 2592 EYES, CATARACT WAS THE MAJOR POSTOPERATIVE COMPLICATION REPORTED AS 5.2%. OF THOSE 15.4% WERE REPORTED AS SURGICALLY INDUCED. 33.8% EYES HAD POOR VAULT (<200 MICRONS)." THE PRIMARY OUTCOME MEASURE IS THE DEVELOPMENT OF CATARACT OR ENDOTHELIAL CELL LOSS BETWEEN THOSE GROUPS. CONCLUSION, ICL IMPLANTATION PROVIDED GOOD REFRACTIVE OUTCOMES AND STABILITY IN THE LONG TERM .THERE WAS NO SIGNIFICANT DIFFERENCE IN POST-OPERATIVE ECC LOSS , CATARACT DEVELOPMENT OR COMPLICATIONS WITH DIFFERENT LEVEL OF EXPERTISE. REPORTEDLY, QUESTIONS TO THE PRESENTER FROM THE MODERATOR DISAGREED WITH FINDINGS OF THE PRESENTER AS IN THEIR OWN EXPERIENCE THE INCIDENCE OF CATARACTS IS VERY LOW IF ANY WITH EVO OR V4C. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495098&pc=MTA | 2023826-2023-03442 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED A PROSPECTIVE COMPARATIVE STUDY BETWEEN IMPLANTABLE PHAKIC INTRAOCULAR CONTACT LENS AND IMPLANTABLE COLLAMER LENS IN TREATMENT OF MYOPIA IN ADULTS INDICATED THAT IN A STUDY OF 60 EYES OF 60 PATIENTS , THERE WAS ONE CASE OF LENS OPACITY WHICH DID NOT AFFECT THE FINAL UCDVA. Manufacturer Narrative: CORRECTED DATA: B4- ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. THE COMPLAINT OF LENS OPACITY BELONGED TO A LENS NOT MANUFACTURED BY STAAR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674548&pc=MTA | 2023826-2023-03841 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -15.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO THE PATIENT DEVELOPING A CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346342&pc=MTA | 2023826-2023-05690 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE 29-30 YEARS AGO. IT WAS INDICATED THE LENS WAS NOW EXPLANTED DUE A NATURAL CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346569&pc=MTA | 2023826-2023-05626 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S ON (B)(6) 2023. CATARACT OBSERVED. THE LENS WAS EXCHANGED AFTER CATARACT REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY) CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18428669&pc=MTA | 2023826-2024-00059 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS WAS EXPLANTED AFTER 30 YEARS DUE THE PATIENT HAD DEVELOPED A NATURAL CATARACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10899883&pc=MTA | 2023826-2020-02746 |  | Injury | STAAR SURGICAL COMPANY AG | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINED IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE REPORTER INDICATED THE CATARACT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6), BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601473&pc=MTA | 2023826-2015-00297 | 2013-11-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) AND WAS NOT REPLACED. SEE MFR REPORT #2023826-2015-00296 FOR THE OTHER EYE. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW-REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601598&pc=MTA | 2023826-2015-00296 | 2013-11-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) AND WAS NOT REPLACED. Manufacturer Narrative: (B)(4) - CATARACT, INDUCED; (LENS EXPLANTED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4900125&pc=MTA | 2023826-2015-00771 | 2015-04-30 | Injury | STAAR SURGICAL COMPANY | 2015-07-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) AND AN EARLY ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED. THE LENS REMAINS IMPLANTED. SURGEON PLANS TO REMOVE THE LENS, EXTRACT THE CATARACT AND IMPLANT AN IOL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598784&pc=MTA | 2023826-2022-01733 | 2022-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LENS OPACITY (CATARACT); ANTERIOR SUBCAPSULAR CATARACTS (ASC). THE PROBLEM WAS RESOLVED. PATIENT HAD A CATARACT SURGERY WITH PSEUDOPHAKIC IOL IMPLANTATION. CAUSE REPORTED AS USER ERROR. Manufacturer Narrative: WEIGHT-RACE: UNK. (B)(4). LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14165251&pc=MTA | 2023826-2022-01220 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Cataract | NA | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN J KOREA OPHTHALMOL SOC TITLED, "LONG-TERM CLINICAL OUTCOME OF IMPLANTABLE COLLAMER LENS." THE ARTICLE STATED THAT THIS RETROSPECTIVE STUDY INCLUDED 129 EYES OF 68 PATIENTS WHO UNDERWENT IMPLANTABLE COLLAMER LENS (ICL) IMPLANTATION OF MYOPIA WITH A 10 YEAR FOLLOW UP. "EIGHT EYES (6.2%) DEVELOPED CATARACT DURING FOLLOW-UP, WHICH WAS SYMPTOMATIC IN THREE EYES (2.3%) AND TREATED WITH ICL EXPLANTATION AND PHACOEMULSIFICATION. RHEGMATOGENOUS RETINAL DETACHMENT OCCURRED IN ONE EYE (0.8%) AND WAS TREATED WITH VITRECTOMY." Manufacturer Narrative: WEIGHT - RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6303780&pc=MTA | 2023826-2017-00133 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY ASCO (ANTERIOR SUBCAPSULAR OPACITY), WHICH WAS OBSERVED ON (B)(6) 2016. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/30 WHICH IS THE SAME AS PREOPERATIVE DATA AS OF (B)(6) 2015. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648707&pc=MTA | 2023826-2018-00976 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON WILL BE REMOVING A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS BECAUSE THE PATIENT HAS A CATARACT. ON (B)(6) 2018 THE REPORTER STATED THAT THE LENS REMAINS IMPLANTED, LENS REMOVAL HAS NOT BEEN PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155681&pc=MTA | 2023826-2017-01983 |  | Injury | Unknown Manufacturer | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract | NA | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT DEVELOPED A CATARACT AFTER ICL IMPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6522593&pc=MTA | 2023826-2017-00710 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -11.0/1.0/091 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(6). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6534124&pc=MTA | 2023826-2017-00711 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-05-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, DIOPTER -10.0/2.0/093 IMPLANTABLE COLLAMER LENS (ICL), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT BEGAN TO EXPERIENCE LOW VAULT AND LENS OPACITY. TO DATE, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ACD < 3.0MM (2.9MM)]. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12213000&pc=MTA | 2023826-2021-02697 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED LOW VAULTING AND LENS OPACITY (ASC). THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT". Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE OVERALL LENGTH WITHIN SPECIFICATIONS, WIDE INSPECTION COULD NOT BE PERFORMED DUE TO HAPTIC SIGNIFICANTLY DAMAGED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14283191&pc=MTA | 2023826-2022-01401 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING AND LENS OPACITY (ASC). ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS BUT IT WAS NOTED "AFTER LENS REPLACEMENT; THERE IS A DECREASE IN VISUAL ACUITY; BUT THE PATIENT IS UNDER OBSERVATION. IT IS EXPECTED THAT VISION WILL RECOVER IN THE FUTURE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-01402 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189499&pc=MTA | 2023826-2022-01284 | 2022-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT NO INTERVETNION WAS PERFORMED FOR THE LENS OPACITY AND THE PATIENT IS UNDER FOLLOW-UP OBSERVATION. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS. REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189674&pc=MTA | 2023826-2022-01285 | 2022-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT AND LENS OPACITY WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT NO INTERVENTION WAS PERFORMED FOR THE LENS OPACITY AND THE PATIENT IS UNDER FOLLOW-UP OBSERVATION. Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649701&pc=MTA | 2023826-2023-01087 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, A CENTRAL ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED (VA WAS 20/25). THE ICL VAULT WAS HIGH (300%). THE LENS REMAINED IMPLANTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "REPORTEDLY, ENLARGEMENT OF INCISION AND ADDITION OF SUTURE PERFORMED" SHOULD BE ADDED FOR CORRECTION. H6: 4625-ADDITIONAL SURGERY: ADDITIONAL SUTURE SHOULD BE ADDED FOR ENLARGEMENT OF INCISION SHOULD BE ADDED FOR CORRECTION. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986369&pc=MTA | 2023826-2021-01873 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | N/A | 1 | Event Description: VIDEO COMPLAINT: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM A PATIENT'S EYE DUE TO A CATARACT. ANOTHER LENS WAS EXPLANTED DUE TO HIGH VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT - UNK. PRODUCT CODE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE -UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190574&pc=MTA | 2023826-2022-01287 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT AND ANTERIOR CAPSULE CATARACT FORMATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14190575&pc=MTA | 2023826-2022-01288 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, MICL12.1, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT AND ANTERIOR CAPSULE CATARACT FORMATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: WEIGHT-RACE: UNK. DATE OF EVENT- UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14656722&pc=MTA | 2023826-2022-01787 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS PATIENT-RELATED FACTOR. Manufacturer Narrative: AGE-RACE:UNK. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16154867&pc=MTA | 2023826-2023-00132 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. LOW VAULT AND ANTERIOR SUBCAPSULAR CATARACT FOLLOWING 11 YEARS OF ICL IMPLANTATION WAS REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT ADDED AND PRODUCT PROBLEM OMITTED FOR CORRECTION. B2: OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS ADDED FOR CORRECTION. H1: SERIOUS INJURY ADDED AND MALFUNCTION OMITTED FOR CORRECTION. H6: 2199- SHOULD BE OMITTED FOR CORRECTION. CLAIM#(B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16584913&pc=MTA | 2023826-2023-00976 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2 -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT AND IT WAS REPORTED THAT ON (B)(6) 2023 LENS OPACITY (ASC) WAS OBSERVED. LENS REMAINS IMPLANTED. "OTHER/ LOW VAULT" WAS REPORTED FOR CAUSE OF EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904523&pc=MTA | 2023826-2023-01688 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V4 ICL IMPLANTABLE COLLAMER LENS IN THE PATIENTS EYE. THE PATIENT DEVELOPED A MILD ANTERIOR SUBCAPSULAR OPACITY (CATARACT) AND THE VAULT WAS SMALLER THAN EXPECTED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6539252&pc=MTA | 2023826-2017-00718 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO DEVELOPMENT OF A CATARACT. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY CONFIRMED THE PATIENT REPORT OF A CATARACT IN THE LEFT EYE AND WILL MONITOR THE PATIENT FOR ANY CHANGES. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED LOW ICL VAULTING WAS OBSERVED, AS WELL AS BEGINNING OF ANTERIOR SUBCAPSULAR CATARACT (ASC). UCVA WAS 20/25-. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385985&pc=MTA | 2023826-2022-00346 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract | N/A | 1 | Event Description: AN ARTICLE PUBLISHED IN FRONTIERS IN MEDICAL WAS RECEIVED ON 06/01/2022 ENTITLED 'EIGHT-YEAR OUTCOMES OF IMPLANTATION OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS WITH A CENTRAL PORT FOR MODERATE TO HIGH AMETROPIA.' THE ARTICLE STATES TWO EYES REQUIRED ICL REPOSITIONING AT 1 WEEK POSTOPERATIVELY (A 90 DEGREE ROTATION TO VERTICAL ORIENTATION DUE TO A HIGH NON-TORIC ICL VAULT AND A RE-ALIGNMENT DUE TO A TORIC ICL ROTATION). TWO EYES REQUIRED SIMULTANEOUS ICL EXTRACTION AND CATARACT SURGERY AT 2 AND 3 YEARS POSTOPERATIVELY DUE TO ANTERIOR SUBCAPSULAR CATARACT FORMATION. ONE EYE REQUIRED ICL SIZE CHANGE AT 7 YEARS POSTOPERATIVELY TO PREVENT PERIPHERAL ATTACHMENT BETWEEN THE ICL AND THE CRYSTALLINE LENS AND THE SUBSEQUENT PROGRESSION OF A PERIPHERAL ANTERIOR SUBCAPSULAR CATARACT. ALL THREE EYES THAT DEVELOPED CATARACTS HAD A SLIGHT PRE-EXISTING PERIPHERAL ANTERIOR SUBCAPSULAR CATARACT FORMATION PREOPERATIVELY, AND THE UDVA AT THE FINAL FOLLOW-UP WAS 20/20 OR MORE IN THESE EYES. TWO EYES REQUIRED PHOTOREFRACTIVE KERATECTOMY AT 8 YEARS POSTOPERATIVELY DUE TO MYOPIC REGRESSION. Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY EXPIRATION DATE, IMPLANT DATE, EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513713&pc=MTA | 2023826-2015-00141 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract | P030016 | 1 | Event Description: ADDITIONAL INFORMATION RECEIVED - THE SURGEON INDICATED THE CATARACT WAS RELATED TO THE DEVICE AND WAS EITHER A SURGICAL TRAUMA OR PROGRESSIVE CATARACT. THE PATIENT'S VISUAL ACUITY WAS 20/20. Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.7MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED A CATARACT WAS FORMING ON UPPER LENS. THE ICL WAS IMPLANTED ON (B)(6) 2010 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) - CATARACT, INDUCED. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - SEVERAL FACTORS COULD CAUSE THE LATE ONSET OF CATARACT WHICH INCLUDES AGE, CONCOMMITANT DISEASES SUCH AS DIABETES, TRAUMA, SMOKING, ALCOHOL USE OR SYSTEMIC DRUGS SUCH AS STEROIDS. AGING IS A MAJOR CONTRIBUTING FACTOR FOR PROGRESSIVE CATARACT WHICH COULD NOT BE RULED OUT ESPECIALLY IN THIS CASE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6542762&pc=MTA | 2023826-2017-00754 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S EYE AND A CORTICAL CATARACT HAD DEVELOPED. THE PLAN IS TO EXPLANT THE LENS LATER THIS MONTH. THE REPORTER INDICATED THIS WAS THE SECOND LENS USED FOR THIS PATIENT - SEE MFR REPORT # 2023826-2017-00765. THE REPORTER INDICATED A POSSIBLE CAUSE OF THE EVENT WAS MANIPULATION IN THE PATIENT'S EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15059955&pc=MTA | 2023826-2022-02274 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A ICL OR EVO LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. "PATIENT HAS DEVELOPED CATARACT 5 YEARS AFTER SURGERY." PER THE REPORTER "THE PATIENT IS A JAPANESE ASTRONAUT AND IS BEING EVALUATED BY THE SPACE PROGRAM MEDICAL STAFF, INCLUDING (B)(6) FOR POSSIBLE PUBLICATION AS A CASE STUDY." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262921&pc=MTA | 2023826-2023-02665 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF A -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, ASC LENS OPACITY WAS OBSERVED AND THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283881&pc=MTA | 2023826-2023-02672 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract | P880091 | 1 | Event Description: THE REPORTER, INDICATED THAT A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS OF -23.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. ON (B)(6) 2023, THE LENS WAS REMOVED, DUE TO SECONDARY CATARACT AND SIGNIFICANT VISIONS DECREASE. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000125&pc=MTA | 2023826-2022-04492 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract |  | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "IN 2020, I RECEIVED VISIAN ICL SURGERY IN BOTH MY EYES. TWO YEARS LATER AT 30 YEARS OLD, I AM NEEDING THEM REMOVED AND CATARACT SURGERY TO TREAT CATARACTS IN BOTH MY EYES. MY SURGEON SAID THAT THE CATARACTS WERE CAUSED BY A BUILDUP OF FLUID BETWEEN THE ICL LENS AND MY NATURAL LENS DUE TO A 'KINK' IN THE LENS," AND "IT'S GETTING HARD TO SEE." Manufacturer Narrative: REPORTED THROUGH MEDWATCH. NO FOLLOW UPS CAN BE PERFORMED AS THE PATIENT ELECTED TO NOT BE IDENTIFIED AND CONTACT INFORMATION WAS OMITTED FROM THE REPORT. CLAIM# (B)(4). Manufacturer Narrative: B1: CORRECTED TO: PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN THE INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/ DAMAGE SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000158&pc=MTA | 2023826-2022-04493 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract |  | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "IN 2020, I RECEIVED VISIAN ICL SURGERY IN BOTH MY EYES. TWO YEARS LATER AT 30 YEARS OLD, I AM NEEDING THEM REMOVED AND CATARACT SURGERY TO TREAT CATARACTS IN BOTH MY EYES. MY SURGEON SAID THAT THE CATARACTS WERE CAUSED BY A BUILDUP OF FLUID BETWEEN THE ICL LENS AND MY NATURAL LENS DUE TO A 'KINK' IN THE LENS," AND "IT'S GETTING HARD TO SEE." Manufacturer Narrative: B1: CORRECTED TO: PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN THE INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/ DAMAGE SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: REPORTED THROUGH MEDWATCH. NO FOLLOW UPS CAN BE PERFORMED AS THE PATIENT ELECTED TO NOT BE INDEMNIFIED AND CONTACT INFORMATION WAS OMITTED FROM THE REPORT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10877667&pc=MTA | 2023826-2020-02718 | 2020-01-21 | Malfunction | STAAR SURGICAL COMPANY | 2020-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Cataract | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020 AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS OPACITY (ASC). CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED WITHIN THE SAME SURGERY AS REMOVAL OF INITIAL LENS. PROBLEM WAS RESOLVED. THE REPORTER STATED " WHEN THE ICL LENS LOADING WAS FINISHED, THE OPERATING ROOM NURSE INFORMED DOCTOR THAT IT WOULD TAKE ANOTHER 10 MINUTES TO DISINFECT THE SURGICAL INSTRUMENTS. WHEN THE LENS WAS IMPLANTED THE EYE, IT TOUCHED THE ANTERIOR CAPSULE OF THE CRYSTAL, CAUSING CATARACCT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714835&pc=MTA | 2023826-2022-03903 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Cataract |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. CATARACT FORMATION WAS OBSERVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17657366&pc=MTA | 2023826-2023-03780 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Cataract | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE APPROXIMATELY AROUND 2001. POSSIBLE CATARACT SURGERY MAY BE PLANNED. Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B1: PRODUCT PROBLEM CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: MALFUNCTION CORRECTED TO SERIOUS IN INITIAL MDR. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511797&pc=MTA | 2023826-2020-01990 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY WAS OBSERVED ON (B)(6) 2020. CATARACT SURGERY WAS NOT PERFORMED AND A DIFFERENT TYPE OF LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS DEBRIS ON LENS. CLAIM # (B)(6). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511799&pc=MTA | 2023826-2020-01989 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO AN ANTERIOR SUBCAPSULAR OPACITY WAS OBSERVED ON (B)(6) 2020. CATARACT SURGERY WAS NOT PERFORMED AND A DIFFERENT TYPE OF LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS . CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302644&pc=MTA | 2023826-2023-02839 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. LENS OPACITY WAS OBSERVED ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302731&pc=MTA | 2023826-2023-02840 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2 -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS OPACITY WAS OBSERVED ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998309&pc=MTA | 2023826-2023-01941 | 2023-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A CATARACT WAS OBSERVED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED ON 04-APR-2023 AND THE PROBLEM WAS RESOLVED. CATARACT SURGERY WAS PERFORMED AND AN IOL LENS WAS IMPLANTED DURING THE SAME PROCEDURE AS ICL EXPLANT. THE PATIENT IS PSEUDOPHAKIC. THE CAUSE OF THE EVENT WAS DUE TO DEVICE AND PATIENT-RELATED FACTOR (SHALLOW CHAMBER, AGE AND HIGH MYOPIA). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725719&pc=MTA | 2023826-2023-01328 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VISIAN ICL IMPLANTABLE COLLAMER LENS, INTO HER LEFT EYE (OS), IN 2011. THE REPORTER INDICATED A CATARACT HAD DEVELOPED AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT AGE/ DOB, WEIGHT, ETHNICITY & RACE UNKNOWN. EVENT DATE: UNKNOWN. DEVICE EXPIRATION DATE UNKNOWN. DEVICE EXPLANT DATE: UNKNOWN. INITIAL REPORTER OCCUPATION: UNKNOWN. DEVICE MANUFACTURE DATE: UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -04.50 DIOPTER, INTO HER LEFT EYE (OS) ON (B)(6) 2009. THE REPORTER INDICATED A CATARACT HAD DEVELOPED AND THE LENS REMAINED IMPLANTED. THE SURGEON INDICATED THE PATIENT WAS LAST SEEN ON (B)(6) 2013. D6A: (B)(6) 2009. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687731&pc=MTA | 2023826-2020-02429 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICH12.1 IMPLANTABLE COLLAMER LENS, DIOPTER +7.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. LENS OPACITY ASC CORTICAL AND INCISION ENLARGEMENT WAS ALSO REPORTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687743&pc=MTA | 2023826-2020-02385 | 2020-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICH12.1 IMPLANTABLE COLLAMER LENS, DIOPTER +5.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. LENS OPACITY ASC CORTICAL AND INCISION ENLARGEMENT WAS ALSO REPORTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587376&pc=MTA | 2023826-2023-00949 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2013. THE LENS WAS REPORTED AS LOW VAULTING AND AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587382&pc=MTA | 2023826-2023-00948 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2013. THE LENS WAS REPORTED AS LOW VAULTING AND AN ANTERIOR SUBCAPSULAR OPACITY (CATARACT) WAS OBSERVED ON (B)(6) 2023. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16891772&pc=MTA | 2023826-2023-01749 | 2023-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2; -5.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS LOW VAULTING WITHOUT ROTATION AND LENS CATARACT (CORTICAL OBSERVED ON (B)(6) 2023). LENS REMAINS IMPLANTED. PROBLEM IS NOT RESOLVED. IT WAS ADDITIONALLY NOTED "GOOD. NO COMPLAINTS PER PATIENT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15891259&pc=MTA | 2023826-2022-04271 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM MICL 12.1 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. NARROW VAULT AND PATIENT NEEDED CE (CATARACT EXTRACTION) OD DUE TO CONTACT WITH ICL. FORMING CATARACT. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: HEALTH IMPACT:(B)(4) - CATARACT EXTRACTION INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM MICL 12.1-7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. NARROW VAULT AND PATIENT NEEDED CE (CATARACT EXTRACTION) OD DUE TO CONTACT WITH ICL FORMING CATARACTS. GLARES AND HALOS AND BLURRED VISION REPORTED. THE LENS WAS REMOVED ON 14-NOV-2022. PROBLEM RESOLVED. CAUSE WAS REPORTED AS DEVICE. PANOPTIX AND IOL IMPLANTATION IN PLACE. YAG CAP SCHEDULED. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: D6B: LENS EXPLANTED (B)(6) 2022. H6: 1494- ACD<3.00, AGE >45 AT TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10687738&pc=MTA | 2023826-2020-02387 | 2020-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXPLANTED DUE TO LENS OPACITY ASC AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. REPORTEDLY, THE TIP OF THE KNIFE TOUCHED THE ANTERIOR LENS CAPSULE. Manufacturer Narrative: ADDITIONAL INFORMATION; H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED INTRA-OPERATIVELY WITHIN INITIAL SURGERY; DUE TO LENS OPACITY ASC. PHACOEMULSIFICATION & IOL IMPLANTATION ALSO REPORTED. THE REPORTER STATES THE PROBLEM WAS RESOLVED. H6 - HEALTH IMPACT CODE 4625-ADDITIONAL SURGERY (PHACOEMULSIFICATION & IOL IMPLANTATION). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15770939&pc=MTA | MW5113203 | 2022-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | MTA | VISIAN IMPLANTABLE CONTACT LENS | Unexpected Therapeutic Results; Material Twisted/Bent | Cataract |  | 1 | Event Description: IN (B)(6) 2020, I RECEIVED VISIAN ICL SURGERY IN BOTH MY EYES. TWO YEARS LATER AT 30 YEARS OLD, I AM NEEDING THEM REMOVED AND CATARACT SURGERY TO TREAT CATARACTS IN BOTH MY EYES. MY SURGEON SAID THAT THE CATARACTS WERE CAUSED BY A BUILDUP OF FLUID BETWEEN THE ICL LENS AND MY NATURAL LENS DUE TO A "KINK" IN THE LENS. I STILL HAD NOT PAID OFF MY ICL LENSES AT THIS TIME, IT WAS A DEVASTATING BLOW. I HAVE NOT HAD THE SURGERY YET BUT IT'S GETTING HARD TO SEE. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5453137&pc=MTA | 2023826-2016-00159 | 2015-03-30 | Injury | STAAR SURGICAL COMPANY | 2016-02-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD) ON (B)(6) 2010. THE PATIENT REPORTED A CATARACT WAS FORMING OUTSIDE OF THE VISUAL RANGE. THE FACILITY INDICATED THE DATE OF THE CATARACT DIAGNOSIS WAS (B)(6) 2015. THE CATARACT WAS AN ANTERIOR SUBCAPSULAR AND HAS PROGRESSED. THE FACILITY REPORTED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE AND PATIENT'S VA POST-OP WAS 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PER MEDICAL REVIEW: THE PATIENT COMPLETED 60-MONTH FOLLOW-UP FORM AND REPORTED THAT CATARACT WAS FORMING OUT OF VISUAL AXIS. NO REPORT OF SECONDARY SURGICALLY OR MEDICALLY INTERVENTION PERFORMED. ACCORDING TO DATA CLARIFICATION REQUEST, DATED ON 2/5/16, ANTERIOR SUBCAPSULAR CATARACT WAS NOTED A YEAR AGO AND HAS PROGRESSED, CAUSING THE DECREASE IN VA (20/20-1) SINCE ICL IMPLANTATION. THE SAME REPORT STATED THAT CATARACT WAS NOT DEVICE RELATED. LENS REMAINS IMPLANTED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983994&pc=MTA | 2023826-2016-01330 | 2016-06-06 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract, Induced; Blurred Vision | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM THE FDA - # MW5064163. THE PATIENT REPORTED HAD 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES, IMPLANTED BILATERALLY ON (B)(6) 2012 FOR HIGH MYOPIA. THE PATIENT REPORTED HAD BLURRY VISION IN THE LEFT EYE FOR TWO WEEKS. UPON EXAMINATION AN ANTERIOR SUBCAPSULAR CATARACT WAS FOUND. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5984003&pc=MTA | 2023826-2016-01331 | 2016-06-06 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract, Induced; Blurred Vision | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM THE FDA - # MW5064163. THE PATIENT REPORTED HAD 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES, IMPLANTED BILATERALLY ON (B)(6) 2012 FOR HIGH MYOPIA. THE PATIENT REPORTED HAD BLURRY VISION IN THE LEFT EYE FOR TWO WEEKS. UPON EXAMINATION AN ANTERIOR SUBCAPSULAR CATARACT WAS FOUND. Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5274151&pc=MTA | MW5058376 |  | Injury | VISIO | 2015-12-03 | MTA | ICLS | Insufficient Information | Cataract, Induced; Blurred Vision |  | 1 | Event Description: I HAD IMPLANTABLE CONTACT LENSES IMPLANTED IN 2012. THREE YEARS AFTER SURGERY I NOTICED FUZZY VISION. I WENT TO THE OPHTHALMOLOGIST AND LEARNED I HAVE CATARACTS IN BOTH EYES. I AM (B)(6) AND HAD NO EVIDENCE OF CATARACTS BEFORE. THESE LENSES WERE SUPPOSED TO LAST A DECADE OR MORE, AND NOW I MUST HAVE AN IOL IMPLANTED AND THE ICLS REMOVED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873717&pc=MTA | 2023826-2016-01069 | 2015-07-12 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.50/+1.50/172 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR/POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/50 +1. Manufacturer Narrative: CONCLUSION: THE PRE-EXISTING "GRAINY PERIPHERAL CORTICAL CLOUDING" WAS LISTED ON THE FORM AND THAT CONDITION MIGHT HAVE CAUSED THE CATARACT DEVELOPMENT SINCE IT IS WELL KNOWN THAT OPACITIES START AS VACUOLES AND CLEFTS BETWEEN THE LENS FIBERS AND PROGRESS TOWARDS ANTERIOR, POSTERIOR OR EQUATORIAL CORTEX. IT SHOULD BE NOTED THAT AT THE TIME OF SURGERY THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. GIVEN ALL THIS INFORMATION, OBSERVED LENS OPACITY AND SUBSEQUENT CATARACT DEVELOPMENT WAS AGE RELATED (PATIENT FACTOR) AND NOT CAUSED BY THE DEVICE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC BROKEN OFF. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873711&pc=MTA | 2023826-2016-01070 | 2004-05-19 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -14.50/+3.75/167 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2004. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (POSTERIOR SUBCAPSULAR AND NUCLEAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRED VISION, GLARE AND HALOS. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/30 -2. Manufacturer Narrative: CORRECTED DATA: (B)(4). ADDITIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934940&pc=MTA | 2023826-2016-01143 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934943&pc=MTA | 2023826-2016-01144 | 2016-06-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; Loss of Vision; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012 WITH EXCELLENT RESULTS. THE PATIENT REPORTED DECREASED VISON, WITH GLARE, BLURRY VISION AND HALOS FOR 2 WEEKS. AN EXAM ON (B)(6) 2016 SHOWED AN ANTERIOR SUBCAPSULAR CATARACT . THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/12.5. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955552&pc=MTA | 2023826-2017-01577 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Activation Failure | Cataract, Induced; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT AFTER THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5/+2.0/097 INTO THE PATIENT'S LEFT EYE (OS), THE LENS WOULD NOT PROPERLY UNFOLD. THE SURGEON STATES THAT, "AFTER ALL POSSIBLE MANEUVERS" THE LENS DIDN'T UNFOLD PROPERLY AND THE LENS WAS EXPLANTED. DURING EXPLANT THE LENS WAS SCRATCHED. THIS OCCURRED ON (B)(6) 2017. REPORTEDLY, THE SURGEON THEN PEFORMED PHACO-EMULSIFICATION ON THE EYE AND IMPLANTED AN EXCHANGE IOL ON (B)(6) 2017 "WITHOUT COMPLICATIONS." Manufacturer Narrative: BEST CORRECTED VISUAL ACUITY (BCVA), ON (B)(6) 2017, WAS 20/30 WHICH WAS A DECREASE FROM BASELINE BCVA (20/25). (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, PHACO-EMULSIFICATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6176008&pc=MTA | 2023826-2016-01763 | 2014-09-17 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract, Induced; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.0/+1.50/042 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2014. THE PATIENT EXPERIENCED BLURRED VISION AND CLOUDY VISUAL ACUITY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/25 +2. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE PRODUCT. THE LENS WAS REMEASURED AND WAS FOUND TO BE IN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5462169&pc=MTA | 2023826-2016-00198 | 2012-05-15 | Injury | STAAR SURGICAL COMPANY AG | 2016-02-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR) WHICH WAS NOTED ON (B)(6) 2012. THE PATIENT EXPERIENCED BLURRED VISION AND STARBURSTING OF LIGHT AT NIGHT. CATARACT SURGERY WAS PERFORMED AND A TORIC IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645517&pc=MTA | 2023826-2017-00956 | 2017-03-22 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE REPORTER INDICATED LENS OPACITY (ANTERIOR) WAS FIRST OBSERVED ON (B)(6) 2017 AND THE LENS WAS EXPLANTED DUE TO LOW VAULTING ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT HAD ADDITIONAL SYMPTOMS - MILD BLURRING. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652095&pc=MTA | 2023826-2016-00641 | 2013-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW/NO VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LOW VAULT OF THE LENS WAS FIRST OBSERVED ON (B)(6) 2013. LENS OPACITY WAS OBSERVED (1 TR ASC) AND PATIENT EXPERIENCED BLURRY VISION. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2, WITH NO DISCOMFORT OR COMPLAINTS. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796304&pc=MTA | 2023826-2016-00935 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRY VISION. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/15 -3. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS RE-HYDRATED IN BSS FOR REMEASUREMENT AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796310&pc=MTA | 2023826-2016-00905 | 2016-05-09 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2015. THE PATIENT EXPERIENCED BLURRY VISION. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION - PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054105&pc=MTA | 2023826-2016-01523 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Corneal Edema; Hyphema; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 LENS, -3.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO EXCESSIVE VAULTING, RESULTING IN ELEVATED INTRAOCULAR PRESSURE (IOP), SHALLOW ANTERIOR CHAMBER, CORNEAL EDEMA, HYPHEMA, AND DEVELOPMENT OF A CATARACT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND A SMALL TEAR ON THE HAPTIC. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909198&pc=MTA | 2023826-2016-01140 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Cataract, Induced; Corneal Edema; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE PATIENT HAD BILATERAL ICL IMPLANTABLE COLLAMER LENSES IMPLANTED ON (B)(6) 2015. AT DAY 1 POST-OP, THE PATIENT HAD +4 CORNEAL EDEMA WITH DENSE ENDOTHELIAL FOLDS IN HER RIGHT EYE (OD). TWO WEEKS AFTER THE SURGERY, THE VISION IN THE RIGHT EYE (OD) WAS CLOUDY. THE PATIENT'S VISION GRADUALLY IMPROVED AND WAS TREATED WITH LASIK. THE PATIENT ENDED UP WITH A CYCLOPLEGIC REFRACTION OF -0.25 +0.25 X 095. THE PATIENT HAD 20/20 DISTANCE VISION BUT REPORTS SEEING POORLY AT NEAR (J-1). THE PATIENT WENT TO A DIFFERENT SURGEON ON (B)(6) 2016 FOR A SECOND OPINION. THE ICL WAS FOUND TO BE TOUCHING THE PHYSIOLOGIC LENS AND AN ANTERIOR SUBCAPSULAR CATARACT HAD DEVELOPED. ULTRASOUND SHOWED THE ICL WAS UPSIDE DOWN WITH A REVERSE VAULT. THE PLAN IS TO EXPLANT AND REPLACE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE WAS A LOT OF CORNEAL EDEMA AND CLOUDINESS IN THE EYE. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336760&pc=MTA | 2023826-2019-02174 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method; Optical Problem | Cataract, Induced; Glaucoma; Intraocular Pressure Increased; Uveitis; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN MEDICAL SCIENCE WAS RECEIVED ON 22-OCT-2019, ENTITLED 'A COMPARATIVE STUDY BETWEEN FOLDABLE IRIS CLAW LENS AND ICL (IMPLANTABLE COLLAMER LENS) PHAKIC IOLS IN CORRECTING MODERATE AND HIGH MYOPIA.' THE ARTICLE STATES ONE EYE EXPERIENCED UVEITIS AND HIGH IOP. THIS WAS CONTROLLED AFTER ONE MONTH OF SURGERY. THE REFRACTIVE RESULTS WERE STABLE AND THE BCVA IMPROVED ONE TO TWO LINES FROM PREOPERATIVE VALUES. IN TWO EYES GLARE AND HALOS WERE REPORTED. THIS WAS RELATED TO A SLIGHTLY CENTRAL PI IN ONE EYE AND PUPIL DILATION AFTER AN ACUTE ATTACK OF GLAUCOMA IN THE OTHER. BY THE END OF THE 1-YEAR FOLLOW UP PERIOD GLARE AND HALOS WERE NO LONGER REPORTED. ONE CASE OF ACUTE PUPILLARY BLOCK GLAUCOMA WITH AN IOP SPIKE TO 40MM FIRST DAY POSTOP. THE PI WAS SMALL, AND THE PATIENT UNDERWENT ANOTHER SURGICAL IRIDECTOMY 2 DAYS AFTER SURGERY AND IOP WAS NORMALIZED. ONE EYE DEVELOPED ANTERIOR SUB CAPSULAR CATARACT 3-MONTHS POSTOP DUE TO LOW VAULT. THIS OCCURRED DUE TO IMPROPER MEASURING OF WHITE-TO-WHITE DISTANCE, MEASURED USING A CALIPER WITH INDISTINCT APPEARANCE OF THE LIMBUS DUE TO CONTACT LENS WEAR. Manufacturer Narrative: DEVICE CODE 3001 SHOULD HAVE BEEN INCLUED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT: UNK. PRODUCT CODE-UNK. MODEL #, EXPLANTED: UNK. PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. MANUFACTURE DATE: UNK. (B)(4). IMPROPER OR INCORRECT PROCEDURE OR METHOD (IMPROPER MEASUREMENT OF WHITE-TO-WHITE DISTANCE) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6608173&pc=MTA | 2023826-2017-00900 | 2017-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract, Induced; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/043 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, UNREACTIVE PUPIL (FIXED), ANGLE CLOSURE, AND IRIS ATROPHY. ON (B)(6) 2017 THE PATIENT ALSO EXPERIENCED LENS OPACITY ASCO. THE SURGEON PLANS TO EXCHANGE TO A SHORTER LENS FOR THE PATIENT. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. (B)(4). Manufacturer Narrative: (B)(6). THE LENS WAS EXCHANGED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). H6_WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218705&pc=MTA | 2023826-2015-01478 | 2015-08-10 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF AN ANTERIOR SUBCAPSULAR OPACITY, PUPIL BLOCK, ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS OCCURRED WITH EXPLANT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE: FIBERS. (B)(4): ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF REMAINING VISCO/ FLUIDS. CONCLUSION: THE PATIENT IS UNDER THE AGE OF 21 YEARS. THA PATIENT'S EYE HAS ANTERIOR CHAMBER DEPTH (ACD), LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328083&pc=MTA | 2023826-2015-01750 |  | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HIS VISION HAS CHANGED TO -0.5. THE FACILITY REPORTED THE PATIENT HAS AN EARLY CATARACT, DUE TO THE AGING PROCESS AND NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH RESULTS: PER MEDICAL REVIEW: ON (B)(6) 2015 THE PATIENT COMPLETED 36 MONTH POST-TREATMENT PATIENT FOLLOW-UP FORM AND REPORTED DECREASE OF VISUAL ACUITY FOLLOWING ICL IMPLANTATION. ACCORDING TO DCR FORM, RECEIVED ON 11/19/15, THE CAUSE OF THE EVENT WAS A CATARACT DEVELOPMENT DUE TO AGING PROCESS AND NOT RELATED TO THE DEVICE. ICL REMAINED IMPLANTED AND UCVA WAS 20/25 WITH GOOD PROGNOSIS. SERIOUSNESS IS BASED ON NO REPORTED LOSS OF BCVA AND NO REPORT OF ANY SECONDARY INTERVENTION PLANNED OR SCHEDULED. IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE", THEREFORE THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5452843&pc=MTA | 2023826-2016-00160 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS) ON (B)(6) 2010. THE PATIENT REPORTED HAS LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT REPORTED HIS VISION HAS DECREASED AND IS DETERIORATING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVALUATION RESULTS: PER MEDICAL REVIEW: THE PATIENT COMPLETED 60-MONTH FOLLOW-UP FORM AND REPORTED THAT HIS VISION HAD DECREASED ("DETERIORATING"). HE WAS TOLD THAT EARLY ONSET OF CATARACT WAS OBSERVED. NO SECONDARY SURGICALLY OR MEDICALLY INTERVENTION WAS PERFORMED OR PLANNED AND NO REPORT OF BCVA LOSS. NO INFORMATION/CONFIRMATION WAS RECEIVED TO DATE FROM THE IMPLANTING SURGEON. IT SHOULD BE NOTED THAT LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION, PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). CONCLUSION - (UNABLE TO CONFIRM COMPLAINT); BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5452852&pc=MTA | 2023826-2016-00161 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD) ON (B)(6) 2010. THE PATIENT REPORTED HAS LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT REPORTED HIS VISION HAS DECREASED AND IS DETERIORATING. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536971&pc=MTA | 2023826-2016-00299 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. PROCESS EVALUATION): WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4760961&pc=MTA | 2023826-2015-00534 | 2014-02-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device remains implanted | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: IT WAS REPORTED PATIENT HAD A MICL 13.6 MM IMPLANTABLE COLLAMER LENS IMPLANTED IN THE PATIENT'S LEFT EYE ON (B)(6) 2012. AS PART OF THE POST MARKET STUDY, THE PATIENT REPORTED EXPERIENCING HAVING LOST ACUITY IN HIS NEAR VISION AFTER THE SURGERY AND HAVE DIFFICULTY FOCUSING WITHOUT THE USE OF CONTACTS OR GLASSES TO BALANCE HIS EYESIGHT. ALSO BEING TOLD BY HIS OPTOMETRIST THAT SHE CAN SEE THE BEGINNINGS OF CATARACTS. THE PHYSICIAN'S OFFICE WAS CONTACTED FOR FURTHER INFORMATION. THE PATIENT'S LAST VISIT WAS ON (B)(6) 2014. THE PATIENT IS NOW HYPEROPIC. STATED ADVERSE EVENT IS RELATED TO THE DEVICE. PROGNOSIS IS GOOD. VISUAL ACUITY POST-TREATMENT BCDVA 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). ICL REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536601&pc=MTA | 2023826-2016-00303 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER RIGHT EYE (OD) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY CONFIRMED THE PATIENT REPORT AND INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015. THE TYPE OF CATARACT WAS CORTICAL PERIPHERAL. THE LENS HAD A LOW VAULT AND THIS PROBABLY WAS THE CAUSE OF THE SUBCAPSULAR HAZE. THE LENS REMAINS IMPLANTED AND THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5328002&pc=MTA | 2023826-2015-01751 | 2014-04-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR AND POST SUBCAPSULAR). THE LENS WAS EXPLANTED ON (B)(6) 2014, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S VA WAS 20/25 AND THE PROGNOSIS IS GOOD. Manufacturer Narrative: PER MEDICAL REVIEW: REPORTEDLY, ICL WAS EXPLANTED FROM A (B)(6) - YEAR OLD PATIENT, TWO YEARS POSTOPERATIVELY, TO ADDRESS CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST (DCR), RECEIVED ON 11/19/15, CATARACT WAS EXTRACTED AND IOL WAS IMPLANTED AT THE SAME SURGERY DATE. THE SAME REPORTED ATTRIBUTED CATARACT (ANTERIOR SUBCAPSULAR COMBINED WITH POSTERIOR SUBCAPSULAR) TO THE DEVICE. THE PROGNOSIS AFTER SECONDARY SURGICAL INTERVENTION WAS GOOD (VA 20/25). A WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084245&pc=MTA | 2023826-2015-01194 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOSS OF BEST-CORRECTED VISUAL ACUITY AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THERE IS STILL SEDIMENT (VISCOELASTIC) ON THE CRYSTALLINE LENS CENTER. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CORRECTION: IMPLANTED DATE - (B)(6) 2015. ADDITIONAL INFORMATION - INVESTIGATION SUMMARY: THE MANUFACTURER'S DEVICE ANALYSIS RESULTS - THE EVENTS REPORTED BY PROF. (B)(6) ARE LIMITED TO A SINGLE SITE BY A SINGLE SURGEON IMPLANTING IMPLANTABLE COLLAMER LENSES (LCLS) BETWEEN THE DATES OF MARCH 2015 AND OCTOBER 2015. TEN PATIENTS WERE REPORTED TO HAVE MILKY WHITE PRECIPITATES FORM ON THE SURFACE OF THE ICL AND ON THE ANTERIOR LENS CAPSULE OF THE PHYSIOLOGIC CRYSTALLINE LENS. NO ADDITIONAL REPORTS OF PRECIPITATES HAVE BEEN RECEIVED FROM ANY OTHER SURGEON OR SITE WITH THE SAME LOT NUMBERS AS THE REPORTED EVENTS FROM PROF. (B)(6). ADDITIONALLY, IT IS IMPORTANT TO NOTE THAT IN THE DIRECTIONS FOR USE (DFU) OF THE LCL, THE COMPLICATIONS AND ADVERSE REACTIONS SECTION IDENTIFIES THE POTENTIAL RISK FOR "SECONDARY SURGICAL INTERVENTION TO REMOVE/REPLACE THE LENS." THE PRECAUTIONS SECTION ALSO SUGGESTS "THE LENS SHOULD NOT BE EXPOSED TO ANY SOLUTIONS OTHER THAN THE NORMALLY USED INTRAOCULAR IRRIGATING SOLUTIONS (E.G. ISOTONIC SALINE, BSS, VISCOELASTIC. ETC.)." DURING THE IN-DEPTH INVESTIGATION WITH PROF. (B)(6) TO IDENTIFY ROOT CAUSE, THE INTRAOPERATIVE REGIMEN REVEALED THAT INTRACAMERAL 0.1% CEFUROXIME ANTIBIOTIC FOR PROPHYLAXIS WAS USED IN THE 10 REPORTED CASES, ALONG WITH RINGER'S SOLUTION. LN AN ARTICLE PUBLISHED BY VASAVADA ET AL. (COMPARISON BETWEEN RINGER'S LACTATE AND BALANCED SALT SOLUTION ON POSTOPERATIVE OUTCOMES AFTER PHACOEMULSIFICATION: A RANDOMIZED CLINICAL TRIAL, LNDIAN J OPHTHALMOL. 2009 MAY-JUN; 57(3): 191-195), THE AUTHORS CONCLUDE THAT RINGER'S LACTATE IS HYPOTONIC AND SLIGHTLY ACIDIC (OSMOLALITY 280MMOL, PH 6.0) AS COMPARED TO BSS (OSMOLALITY 302 MMOL, PH 7.4) AND AQUEOUS (OSMOLALITY 302 MMOL, PH 7.4). THEY FURTHER CONCLUDED THAT A STATISTICALLY SIGNIFICANT LARGER NUMBER OF PATIENTS WITH FLARE AND CELLS MAY BE DUE TO THE SLIGHTLY ACIDIC FLUCTUATING PH AND LACK OF BUFFER SYSTEM IN RINGER'S LACTATE, WHICH COULD AFFECT THE BLOOD AQUEOUS BARRIER STABILITY. ON THE OTHER HAND, 855 HAS AN ACETATE CITRATE BUFFER SYSTEM AND THE PH, THOUGH SLIGHTLY ALKALINE, IS STABLE. SUBSEQUENT TO THE SERIES OF REPORTED EVENTS, ADDITIONAL PATIENTS HAVE BEEN TREATED BY PROF. (B)(6) WITH AN ICL USING ONLY COMMERCIALLY AVAILABLE BALANCED SALINE SOLUTION ((B)(4)) AND OCUCOAT ((B)(4)) INTRAOPERATIVE, WITHOUT THE USE OF 0.1% INTRACAMERAL CEFUROXIME ANTIBIOTIC. THERE WERE NO OBSERVED INCIDENTS OF DEPOSITS FOLLOWING THIS PROTOCOL REGIMEN. UPON REVIEW OF THE DEVICE HISTORY RECORDS (DHR), ALL PRODUCTS REPORTED IN THE INCIDENCES, BASED UPON LOT AND SERIAL NUMBER REVIEW, INDICATED SPECIFICATIONS WERE MET DURING MANUFACTURING PROCESS. ALL PRODUCTS PASSED QUALITY INSPECTIONS AND WERE RELEASED PER ACCEPTANCE CRITERIA. FINAL COMMENTS FROM THE MANUFACTURER: BASED UPON A THOROUGH REVIEW OF ALL AVAILABLE DATA AND ON CLINICAL PERSPECTIVE AS A PRACTICING OPHTHALMIC SURGEON WITH EXPERTISE IN THE LCL, DR. (B)(6) (MEDICAL CONSULTANT) AND STAAR CONCLUDED THAT THE OBSERVED PRECIPITATES WERE MOST LIKELY THE RESULT OF INTERACTION BETWEEN THE INTRAOCULAR MEDICATIONS USED AT THE TIME OF SURGERY. (B)(4). Manufacturer Narrative: THE CUSTOMER HAS REPORTED SINCE USING THE NEW SURGICAL PROTOCOL (USING BSS AND OCUCOAT ONLY) NO NEW COMPLAINT CASES OF LOSS OF BCVA, WITH PRECIPITATE ON THE CRYSTALLINE LENS, HAVE BEEN RECEIVED SINCE THE BEGINNING OF 2016. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5471961&pc=MTA | 2023826-2016-00204 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN HER RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5471962&pc=MTA | 2023826-2016-00205 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536606&pc=MTA | 2023826-2016-00302 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT REPORTED THE EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE SURGEON REPORTED THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PROGNOSIS IS EXCELLENT. THE VA POST-OP WAS 20/20. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536972&pc=MTA | 2023826-2016-00300 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS) ON (B)(6) 2011. THE PATIENT REPORTED LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS REMOVED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2015. THE REPORTER INDICATED THE CATARACT WAS A CORTICAL CENTRAL AND THE EVENT WAS NOT DEVICE RELATED. AN IOL WAS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327964&pc=MTA | 2023826-2015-01752 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE LENS WAS EXPLANTED ON (B)(6) 2015, THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: METHOD: MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: PER MEDICAL REVIEW: REPORTEDLY ICL WAS EXPLANTED 3 YEARS POSTOPERATIVELY TO ADDRESS LOW VAULT AND A CATARACT DEVELOPMENT. ACCORDING TO COMPLETED DATA CLARIFICATION REQUEST, DATED ON (B)(6) 2015, THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR AND WAS ATTRIBUTED TO THE DEVICE. FOLLOWING ICL REMOVAL, CATARACT WAS EXTRACTED, AND IOL WAS SUCCESSFULLY IMPLANTED. THE SAME REPORT STATED THAT THE PATIENT PROGNOSIS WAS GOOD (VA:20/20). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILIAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5139024&pc=MTA | 2023826-2015-01342 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-10-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO A LOW VAULT, THE DEVELOPMENT OF A CATARACT AND DECREASED VISUAL ACUITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER MEDICAL REVIEW - CENTRAL AND/OR PERIPHERAL CONTACT BETWEEN THE ICL AND THE CRYSTALLINE LENS IS ASSOCIATED WITH ANTERIOR SUBCAPSULAR CATARACT FORMATION. HOWEVER, THIS REPORT PROVIDES NO INFORMATION ON THE DEGREE OF VAULTING OR THE POSITION OF THE LENS OPACITIES. ADDITIONALLY, THERE IS NO INFORMATION ABOUT THE CRYSTALLINE LENS PRIOR TO THE INITIAL SURGERY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284955&pc=MTA | 2023826-2017-00114 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(6) 1966. NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284956&pc=MTA | 2023826-2017-00115 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE FACILITY REPORTED THE CATARACT (ANTERIOR SUBCAPSULAR) WAS DIAGNOSED ON (B)(6) 2015. THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VISUAL ACUITY WAS 20/15. Manufacturer Narrative: (B)(4). DATE OF BIRTH - (B)(6) 1966. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7277678&pc=MTA | 2023826-2018-00229 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V4 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED 5 YEARS AFTER BEING IMPLANTED IN THE PATIENTS EYE, DUE TO SILICONE- OIL INDUCED CATARACT. PHACOEMULSIFICATION SURGERY WITH IOL (INTRAOCULAR LENS) IMPLANTATION WAS PERFORMED. SILICONE-OIL REMOVAL WAS REPORTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047775&pc=MTA | 2023826-2015-01116 | 2015-02-13 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (CORTICAL) AND A CAPSULAR TEAR. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293145&pc=MTA | 2023826-2015-01653 | 2015-08-18 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -22.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2008. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/40. Manufacturer Narrative: AGE AT TIME OF EVENT IS (B)(6) WHICH IS INCORRECT, CORRECTED AGE AT TIME OF EVENT IS (B)(6). (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THERE WAS NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: DIOPTER: -22.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5710805&pc=MTA | 2023826-2016-00808 | 2016-05-11 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN MICL IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. CATARACT SURGERY WAS PERFORMED. THE REPORTER INDICATED IT WAS A NUCLEAR CATARACT AND WAS UNRELATED TO THE ICL. THE LENS WAS ORIGINALLY IMPLANTED BY ANOTHER SURGEON AT A DIFFERENT FACILITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826213&pc=MTA | 2023826-2016-00983 | 2016-06-08 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. THE PATIENT'S NATURAL LENS WAS ALSO REMOVED AND REPLACED WITH A DIFFERENT INTRAOCULAR LENS (IOL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. EXPLANT DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5437516&pc=MTA | 2023826-2016-00165 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF A CATARACT. AN INTRAOCULAR LENS WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5437517&pc=MTA | 2023826-2016-00162 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. THE CATARACT WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, MICL WAS EXPLANTED APPROXIMATELY TWO AND A HALF YEARS POSTOPERATIVELY TO ADDRESS DEVELOPMENT OF A CATARACT. CATARACT WAS REMOVED AS WELL. NO REPORTED POSTOPERATIVE SEQUELAE. NO FURTHER INFORMATION WAS RECEIVED TO DATE DESPITE MULTIPLE ATTEMPTS. IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) YEARS OLD AND PER DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE", THEREFORE THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: CONCLUSION: (OFF-LABEL, UNAPPROVED, OR CONTRAINDICATED USE.) (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE OTHER HAPTIC WAS FOUND TO BE CRACKED, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION - BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680358&pc=MTA | 2023826-2016-00764 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER STATED THE PATIENT IS DOING WELL AND THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS DISCARDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680364&pc=MTA | 2023826-2016-00765 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. CATARACT SURGERY WAS PERFORMED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER STATED THE PATIENT IS DOING WELL AND THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6684194&pc=MTA | 2023826-2017-01086 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-07-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/93 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ACCORDING TO THE DR. THAT BECAUSE THE SURGERY OPERATING TIME WAS TOO LONG, THE SUSPENSORY LIGAMENT NEAR THE LENS CAPSULE MEMBRANE DOT LIGHT LENS OPACITY. THE LENS WAS EXPLANTED IN THE SAME SURGERY, AND THE PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED, NO IOL LENS WAS IMPLANTED. THE CUSTOMER STATED THAT THE SITUATION IS STABLE, AND OBSERVATION ON REGULAR BASIC. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE : OFF LABEL USE (ACD<3.0MM) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533731&pc=MTA | 2023826-2016-00354 | 2014-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -06.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2014. THE SURGEON PERFORMED CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293227&pc=MTA | 2023826-2015-01652 | 2015-05-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -17.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2006. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015.THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015 . PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: DIOPTER: -17.50. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553204&pc=MTA | 2023826-2016-00408 | 2015-11-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -09.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE; MEASUREMENT WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5711853&pc=MTA | 2023826-2016-00720 | 2016-05-08 | Injury | STAAR SURGICAL COMPANY | 2016-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+0.5/151 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY - ASCO) ON (B)(6) 2016 AND UNREACTIVE PUPIL (FIXED). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENTS: UNK. EXPLANT DATE: N/A. (B)(4) LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873542&pc=MTA | 2023826-2016-01065 | 2016-06-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT AND REFRACTIVE SURPRISE. THE LENS WAS REMOVED, CATARACT SURGERY WAS PERFORMED AND A TORIC IOL WAS IMPLANTED WITH GOOD RESULTS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE DIOPTER WAS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS REPORT SOURCE 'USER FACILITY' IS INCORRECT. THE CORRECT REPORT SOURCE IS 'DISTRIBUTOR'. DEVICE CODE: (B)(4) (EXPLANTED) TO (B)(4)- PREVIOUS CODE NOT APPLICABLE UNDER DEVICE CODE; IT SHOULD BE UNDER PATIENT CODE INSTEAD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533727&pc=MTA | 2023826-2016-00353 | 2012-05-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2015. THE SURGEON PERFORMED CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4). ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5084270&pc=MTA | 2023826-2015-01197 | 2015-05-15 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -20.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975337&pc=MTA | 2023826-2016-01317 | 2016-08-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: LENS REPOSITIONING WAS PERFORMED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSION OF INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7427120&pc=MTA | 2023826-2018-00550 |  | Injury | STAAR SURGICAL COMPANY | 2018-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract, Induced; No Known Impact Or Consequence To Patient | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM MICL12.1, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2017. THE REPORTER STATED THAT THE SURGEON WILL BE REMOVING THE LENS BECAUSE IT IS TOO FLAT AND IT IS CAUSING A CATARACT. A LONGER LENS IN THE SAME POWER WAS ORDERED TO REPLACE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE SURGEON HAS DECIDED NOT TO IMPLANT A LARGER LENS SINCE "IS LIKELY TO CAUSE NARROW ANGLES" AND WILL MONITOR AND CONSIDER CATARACT SURGERY AS NEEDED IN THE FUTURE. (B)(4) Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5327956&pc=MTA | 2023826-2015-01753 | 2014-07-21 | Injury | STAAR SURGICAL COMPANY | 2015-12-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract, Induced; Retinal Tear; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED SHE HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT (ANTERIOR SUBCAPSULAR). THE LENS HAD A LOW VAULT. THE PATIENT HAD A SMALL RETINAL TEAR AND IT WAS REPAIRED. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, (B)(6) YEAR OLD PATIENT RECEIVED ICL IN 2012 AND WAS COMPLAINING ABOUT DECREASE IN VA DUE TO CATARACT DEVELOPMENT BY COMPLETING 36 MONTH FOLLOW-UP FORM. THE PATIENT ALSO REPORTED SMALL RETINAL TEAR THAT WAS REPAIRED. DATA CLARIFICATION REQUEST FORM WAS RECEIVED ON 11/19/2015, STATING THAT ANTERIOR SUBCAPSULAR CATARACT WAS CAUSED BY LOW VAULT OF ICL DECREASING THE VA FROM 20/15 TO 20/20. ICL REMAINED IMPLANTED AND NO SECONDARY OR SURGICALLY INTERVENTION WAS PERFORMED. HOWEVER, THE SMALL RETINAL TEAR, THAT WAS TREATED AND RESOLVED BY CRYOTHERAPY, WAS NOT CAUSED BY DEVICE BUT BY PATIENT RELATED FACTOR (VITREOUS DEGENERATION). IT SHOULD BE NOTED THAT AT THE TIME OF THE ICL IMPLANTATION PATIENT WAS (B)(6) YEARS OLD AND PERF DFU INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE "THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS AND THAT WAS REFLECTED IN DEVICE CAUSALITY. FURTHERMORE, LITERATURE REPORTS THAT ONLY 1-2% OF PATIENTS DEVELOP CLINICALLY SIGNIFICANT CATARACT FOLLOWING ICL IMPLANTATION PROBABLY DUE TO PATIENT RELATED FACTORS (ESPECIALLY HIGH MYOPES AND OLDER PATIENTS). A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4) LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064234&pc=MTA | 2023826-2021-02063 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED AT THE BEGINNING THE EYE DEVELOPED AN ABNORMAL DILATION OF THE PUPIL AND THEN A QUICK CORTICAL TOTAL CATARACT (IOP WITHIN THE NORMAL AT EACH EXAMINATION, NO SIGNS OF UVEITIS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12449121&pc=MTA | 2023826-2021-03482 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT #2023826-2021-03481 AND THE PROBLEM WAS NOT RESOLVED. THE PATIENT NOW HAS A SMALL ANTERIOR SUBCAPSULAR CATARACT AND VISION IS WORSE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807726&pc=MTA | 2023826-2021-04259 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | UNK | 1 | Event Description: AN ARTICLE PUBLISHED IN (B)(6) WAS RECEIVED ON 20-OCT-2021, ENTITLED 'LASER CATARCT SURGERY ON AN EYE WITH A PHAKIC IOL.' THE ARTICLE DESCRIBES TWO CASE EXAMPLES OF (B)(6) WOMEN WHO HAD PREVIOUSLY BEEN IMPLANTED WITH VISIAN ICLS IN THE LEFT (OS) EYE. BOTH DEVELOPED 2+ TO 3+ ANTERIOR SUBCAPSULAR HAZE, ONE ALSO WITH 2+ TO 3+ NUCLEAR SCLEROSIS AND BOTH SUBSEQUENTLY UNDERWENT ICL EXPLANT AND CATARACT SURGERY. Manufacturer Narrative: PATIENT IDENTIFIER, AGE OR DATE OF BIRTH, WEIGHT, RACE:UNK. DATE OF EVENT: UNK. PRODUCT CODE: UNK; IMPLANT DATE, EXPLANT DATE:UNK. DEVICE MANUFACTURE DATE:UNK. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12978932&pc=MTA | 2023826-2021-04659 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: A JOURNAL ARTICLE WAS RECEIVED FOR "FIVE-YEAR OUTCOMES OF EVO IMPLANTABLE COLLAMER LENS IMPLANTATION FOR THE CORRECTION OF HIGH MYOPIA AND SUPER HIGH MYOPIA". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342653&pc=MTA | 2023826-2022-01456 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | UNK | 1 | Event Description: AN ARTICLE PUBLISHED IN THE JOURNAL OF CATARACT AND REFRACTIVE SURGERY WAS RECEIVED ON 11/04/2022 ENTITLED 'PHAKIC INTRAOCULAR LENS EXPLANTATION: A SERIES OF 175 SURGERIES.' THE ARTICLE STATES THAT 15 EYES HAD VISIAN V4 IMPLANTABLE LENSES EXPLANTED, 86% DUE TO CATARACT FORMATION AND 14% DUE TO ENDOTHELIAL CELL LOSS. Manufacturer Narrative: DATE OF EVENT-UNK, ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629536&pc=MTA | 2023826-2022-03707 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THERE SEEMS TO BE ANTERIOR CAPSULAR OPACITY THAT HAS DEVELOPED POSSIBLY FROM THE IMPLANTATION PROCESS; ADDITIONALLY IT RESULTED IN SLIGHTLY DISTORTED CENTRAL VISION THAT IS INFLICTING THE PATIENT. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10900067&pc=MTA | 2023826-2020-02747 |  | Injury | STAAR SURGICAL COMPANY AG | 2020-11-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE REPORTER INDICATED THE CATARACT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: IT WAS ALSO REPORTED THAT PHACOEMULSIFICATION WAS PERFORMED ON THE PATIENT. SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - PHACOEMULSIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848263&pc=MTA | 2023826-2022-04295 | 2022-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -14.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS LENS OPACITY OF ASC AND PSC WHICH WAS OBSERVED ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. THE PROBLEM AS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "PATIENT REPORTED BEING HIT HARD IN THE LEFT EYE AND HER VISION BEGAN TO BLUR/GET HAZY". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11830093&pc=MTA | 2023826-2021-01442 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision | NA | 1 | Event Description: ON 2021/04/16 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN HINDAWI, CASE REPORT IN OPHTHALMOLOGICAL MEDICINE, ENTITLED, 'A CASE OF IMPLANTABLE COLLAMER LENS (ICL) WITH REVERSE ORIENTATION FOR 10 YEARS'. THE STUDY ARTICLE REPORTS A CASE OF INVERTED ICL THAT HAS UNDERGONE EXPLANTATION AND CATARACT SURGERY DUE TO DECREASED VISUAL ACUITY, ANTERIOR SUBCAPSULAR CATARACT, AND LOW ICL VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO, EVENT DATE: UNK. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547267&pc=MTA | 2023826-2023-03522 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE AND LENS OPACITY. IT WAS REPORTED THAT REFRACTIVE TREATMENT AND PHACO-EMULSIFICATION WITH IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM RESOLVED. THE REPORTER STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED, FOR CAUSE DETAILS THE REPORTER STATED " THE REFRACTIVE SURPRISE RESULTED IN A HYPEROPIC LVC (LASER VISION CORRECTION) AND UNSATISFACTORY VISION. THE ASC WAS OF LESSER ISSUE". Manufacturer Narrative: H6- HEALTH EFFECT- IMPACTL CODE 4625: REFRACTIVE TREATMENT AND CATARACT SURGERY WITH IOL IMPLANTATION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH FOREIGN MATERIAL ON LENS SURFACE SPECIFICALLY RESIDUE THROUGHOUT THE LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158233&pc=MTA | 2023826-2020-01235 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON NOTED AN ANTERIOR SUBCAPSULAR OPACITY IN THE PATIENTS EYE ON (B)(6) 2020. THE PATIENT COMPLAINED OF BLURRED VISION. EXPLANTATION OF THE LENS AND CLEAR LENS SURGERY IS PLANNED FOR (B)(6) 2020. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4 - CORRECTED TO (B)(6) 2020. B5 - "IT WAS REPORTED THAT CLEAR LENS SURGERY AND IMPLANTATION OF NON-STAAR LENS WAS PERFORMED, WHICH RESOLVED THE ISSUE." SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CORRECTED DATA: H6 - DEVICE CODE:1401 SHOULD HAVE BEEN INCLUDED. B3 - CORRECT DATE SHOULD BE (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3 - CORRECT DATE SHOULD BE REVERTED BACK TO (B)(6) 2020. H6 - MEDICAL DEVICE PROBLEM CODE: 1401 SHOULD BE REMOVED AS IT IS N/A. H6 - PATIENT CODE: 2137 SHOULD BE REMOVED AS IT IS N/A. B5 - "THE PATIENT COMPLAINED OF BLURRED VISION", SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT AS IT IS N/A. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16474967&pc=MTA | 2023826-2023-00725 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.0/1.0/99 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT FOR A PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THIS WAS DUE TO A EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 12.6, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/99 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE PATIENT EXPERIENCED EARLY LENS OPACITY. REPORTEDLY "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 13.2, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS BECAME CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN). THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846153&pc=MTA | 2023826-2019-01293 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INDIAN JOURNAL OF OPHTHALMOLOGY IN (B)(6) 2019 TITLED, 'COMPARATIVE ANALYSIS OF CLINICAL OUTCOMES BETWEEN TWO TYPES OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENSES FOR CORRECTION OF MYOPIA AND MYOPIC ASTIGMATISM.' THE ARTICLE STATES THAT "ONE EYE (0.49%) IN THE ICL GROUP DEVELOPED VISUALLY SIGNIFICANT CATARACT. IN ALL EYES, THE CATARACT DEVELOPED AFTER A MINIMUM DURATION OF 6 MONTHS FOLLOWING IMPLANTATION. NONE OF THE EYES DEMONSTRATED A LOW VAULT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: "FOURTEEN CASES OF 203 EYES WERE ALSO REPORTED TO HAVE LOST ONE OR MORE LINES OF CDVA AT 1 YEAR POST-OPERATIVELY." SHOULD BE ADDED IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF REPORTABLE EVENT: "MALFUNCTION" SHOULD BE CORRECTED TO "SERIOUS" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL #: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16772345&pc=MTA | 2023826-2023-01383 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | P030016 | 1 | Event Description: A COPY OF AN ARTICLE WAS FOUND FROM IJO, DATED FEBRUARY 2023. THE ARTICLE WAS IN REGARD TO "APPLICATION OF FEMTOSECOND LASER IN ASSISTED CATARACT AFTER ICL IMPLANTATION WITH DIFFERENT VAULT". THE ARTICLE INVOLVED 2 DIFFERENT PATIENTS - CASE 1 - A 54-YEAR-OLD FEMALE, WITH PROGRESSIVE BLURRED VISION IN LEFT EYE FOR 2 YEARS, WHO UNDERWENT ICL IMPLANTATION 11 YEARS AGO. THE LENS WAS REMOVED DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT UNDERWENT FLACS (FEMTOSECOND LASER-ASSISTED CATARACT SURGERY) AND A MONOFOCAL IOL WAS IMPLANTED. AT ONE-WEEK POST-OP THE UDVA WAS 20/20. CASE 2 - A 51-YEAR-OLD FEMALE WITH PROGRESSIVE BLURRING OF VISION IN RIGHT EYE FOR ONE YEAR. AN ICL WAS IMPLANTED 11 YEARS AGO DUE TO HIGH MYOPIA. THE LENS WAS REMOVED DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT UNDERWENT FLACS (FEMTOSECOND LASER-ASSISTED CATARACT SURGERY) AND A TRIFOCAL IOL WAS IMPLANTED. AT ONE-WEEK POST-OP THE UDVA WAS 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - FEMTOSECOND LASER-ASSISTED CATARACT SURGERY (FLACS), IOL IMPLANTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914025&pc=MTA | 2023826-2021-04571 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision | P030016 | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN HIS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. INITIAL REPORTER OCCUPATION: UNK. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11159193&pc=MTA | 2023826-2021-00122 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Improper or Incorrect Procedure or Method | Cataract; Blurred Vision | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN SCIENTIFIC REPORTS WAS RECEIVED ON 18-DEC-2020, ENTITLED 'ETIOLOGY AND OUTCOMES OF CURRENT POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS EXTRACTION.' THE ARTICLE STATES THAT ICL EXTRACTION WAS REQUIRED IN A TOTAL OF 8 OUT OF 770 EYES. FIVE OF THE EYES EXPERIENCED PROGRESSION OF PRE-EXISTING CATARACT FORMATION; INCLUDING NUCLEAR CATARACT (2 EYES); ANTERIOR SUBCAPSULAR CATARACT IN THE PERIPHERAL AREA (1 EYE) AND PARACENTRAL AREA (2 EYES). THREE EYES HAD RESIDUAL REFRACTIVE ERRORS IMMEDIATELY AFTER INITIAL SURGERY DUE TO INACCURATE ICL POWER CALCULATION. IN THE CASES OF CATARACT FORMATION PHACOEMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED. IN THE CASES OF RESIDUAL REFRACTIVE ERRORS THE EXPLANTED LENS WAS REPLACED WITH A REPLACEMENT ICL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14252557&pc=MTA | 2023826-2022-01313 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022 WITH PAINLESS PROGRESSIVE VISION LOSS IN EYE FOR HALF A YEAR. LATER, ON (B)(6) 2022, THE PATIENT RECEIVED ICL IOL EXTRACTION FOR PHACOEMULSIFICATION EXTRACTION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14253045&pc=MTA | 2023826-2022-01314 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO12.1 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS OPACITY WAS OBSERVED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT HAD PAINLESS PROGRESSIVE VISION LOSS, AND WAS ADMITTED TO THE CATARACT DEPARTMENT OF OUR HOSPITAL ON (B)(6) 2022. ON (B)(6) 2022, THE PATIENT UNDERWENT ICL IOL EXTRACTION, PHACOEMULSIFICATION AND IOL IMPLANTATION IN OUR HOSPITAL." Manufacturer Narrative: WEIGHT-RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734972&pc=MTA | 2023826-2021-01214 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Cataract; Blurred Vision | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY OF "OUTCOMES AND COMPLICATIONS OF IMPLANTABLE COLLAMER LENS FOR MILD TO ADVANCE KERATOCONUS". THE STUDY WAS TO EVALUATE THE OUTCOME AND SAFETY OF THE IMPLANTABLE COLLAMER LENS (ICL) IMPLANTED IN 32 EYES OF PATIENTS WITH MYOPIC AND MYOPIC ASTIGMATISM WITH MILD TO ADVANCE KERATOCONUS. OF THE 32 EYES, THE LENSES IN 3 EYES HAD ROTATED AND HAD TO BE REPOSITIONED. ONE PATIENT HAD NON-VISUALLY SIGNIFICANT ANTERIOR SUBCAPSULAR CATARACT. SIX PATIENTS WERE UNSATISFIED WITH THE QUALLTY OF THEIR VISION AND ONE OF THOSE PATIENTS HAD THE LENS EXPLANTED DUE TO RESIDUAL REFRACTIVE ERROR AND UNSATISFACTORY VISION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13374207&pc=MTA | 2023826-2022-00351 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Blurred Vision |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). IT WAS INDICATED THAT THE LENS WAS TOO FLAT AND RESULTED IN MINIMAL CLEARANCE OVER THE NATURAL LENS, AND THERE WAS A SLIGHT CATARACT WITH A "SOMEWHAT" HYPEROPIC OUTCOME. THE LENS WAS EXPLANTED ON (B)(6) 2021. A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS IMPLANTED AND THE PATIENT IS DOING QUITE WELL. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12904968&pc=MTA | 2023826-2021-04570 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; Halo | P030016 | 1 | Event Description: THE PATIENT INDICATED THE SURGEON IMPLANTED A (B)(4) IMPLANTABLE (B)(6) LENS, -12.00 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT REPORTED EXPERIENCING BLURRY/CLOUDY VISION ONE WEEK AFTER ICL IMPLANTATION, REPORTEDLY DUE TO PREDNISOLONE, WHICH WAS PRESCRIBED AFTER IMPLANTATION. THE PREDNISOLONE WAS DISCONTINUED. CATARACTS WERE OBSERVED. THE PATIENT SEES HALOES IN HIS RIGHT EYE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, EXPLANT DATE: UNKNOWN, INITIAL REPORTER OCCUPATION: UNKNOWN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191994&pc=MTA | 2023826-2016-01814 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE DEVELOPMENT OF A CATARACT. THE PATIENT COMPLAINED OF BLURRY VISION AT NIGHT. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING FINE. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER?: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED IN PIECES AND UNABLE TO EVALUATE. THE LENS WAS RETURNED DRY. METHOD CODE(S): LENS WORK ORDER SEARCH. RESULTS CODE(S): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4959171&pc=MTA | 2023826-2015-00960 | 2014-12-08 | Injury | STAAR SURGICAL COMPANY | 2015-08-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY) AND BLURRED VISION. CATARACT SURGERY WAS PERFORMED BY PHACOEMULSIFICATION AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140284&pc=MTA | 2023826-2016-01711 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. AT 3 MONTHS POST-OP, PATIENT WAS 20/25 UNCORRECTED AND THE PRESSURES WERE NORMAL. THE REPORTER INDICATED THE BEGINNING OF AN ANTERIOR SUBCAPSULAR CATARACT WAS NOTED AT THAT VISIT. THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140273&pc=MTA | 2023826-2016-01712 | 2015-08-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED A CATARACT HAD DEVELOPED. THE LENS REMAINS IMPLANTED BUT MAY BE REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE OPACITY (ANTERIOR) WAS OBSERVED ON (B)(6) 2015 AND THE CAUSE WAS UNKNOWN. THE LENS DIOPTER WAS -13.0. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4775745&pc=MTA | 2023826-2015-00562 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY REPORT NO: MW5039118. THE PATIENT REPORTED HAD MICL IMPLANTABLE COLLAMER LENS IMPLANTED IN BOTH EYES IN MID 2012. THE PATIENT WAS HAVING PROBLEMS WITH HER VISION (CLOUDY/BLURRY) AND WAS TOLD CATARACTS WERE DEVELOPING IN BOTH EYES. THE PATIENT EVENTUALLY HAD BOTH LENSES EXPLANTED, ONE LENS IN (B)(6), 2014 AND THE OTHER LENS IN (B)(6), 2015. THE CATARACTS WERE REMOVED AND IOLS WERE IMPLANTED. THE PATIENT'S CURRENT VISION IN BOTH EYES IS 20/20 AND 20/30. THE PATIENT WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION. THIS REPORT IS FOR THE PATIENT'S LEFT EYE (OS). Manufacturer Narrative: (B)(4). (OTHER): LENS NOT RETURNED. (B)(4). . Manufacturer Narrative: PER MEDICAL REVIEW - MICL DIRECTIONS FOR USE (DFU) STATE THAT THE VISIAN ICL IS INDICATED FOR ADULTS 21-45YEARS. THE DFU FURTHER STATES THAT THE RELATIONSHIP BETWEEN THE MICL AND FUTURE LENS OPACITIES IS YET UNDETERMINED. CATARACTS ARE KNOWN TO OCCUR WITH GREATER FREQUENCY IN DIABETICS, PATIENTS WITH CERTAIN OTHER SYSTEMIC DISORDERS, IN MYOPES, AND WITH INCREASING AGE. THE MICL WAS USED OFF-LABEL FOR THIS PATIENT, AND THERE IS INSUFFICIENT INFORMATION TO DETERMINE IF THE DEVELOPMENT OF CATARACTS WAS DEVICE-RELATED. BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4776261&pc=MTA | 2023826-2015-00563 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY REPORT NO: (B)(4). THE PATIENT REPORTED HAD MICL IMPLANTABLE COLLAMER LENS IMPLANTED IN BOTH EYES IN MID 2012. THE PATIENT WAS HAVING PROBLEMS WITH HER VISION (CLOUDY/BLURRY) AND WAS TOLD CATARACTS WERE DEVELOPING IN BOTH EYES. THE PATIENT EVENTUALLY HAD BOTH LENSES EXPLANTED, ONE LENS IN (B)(6), 2014 AND THE OTHER LENS IN (B)(6) 2015. THE CATARACTS WERE REMOVED AND IOLS WERE IMPLANTED. THE PATIENT'S CURRENT VISION IN BOTH EYES IS 20/20 AND 20/30. THE PATIENT WAS UNABLE TO PROVIDE ANY ADDITIONAL INFORMATION. THIS REPORT IS FOR THE PATIENT'S RIGHT EYE (OD). Manufacturer Narrative: PT WEIGHT: UNK. DATE OF EVENT: UNK. EXPIRATION DATE - UNK. IMPLANT AND EXPLANT DATES: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. MFG DATE: UNK. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: PER MEDICAL REVIEW - MICL DIRECTIONS FOR USE (DFU) STATE THAT THE VISIAN ICL IS INDICATED FOR ADULTS 21-45 YEARS. THE DFU FURTHER STATES THAT THE RELATIONSHIP BETWEEN THE MICL AND FUTURE LENS OPACITIES IS YET UNDETERMINED. CATARACTS ARE KNOWN TO OCCUR WITH GREATER FREQUENCY IN DIABETICS, PATIENTS WITH CERTAIN OTHER SYSTEMIC DISORDERS, IN MYOPES, AND WITH INCREASING AGE. THE MICL WAS USED OFF-LABEL FOR THIS PATIENT, AND THERE IS INSUFFICIENT INFORMATION TO DETERMINE IF THE DEVELOPMENT OF CATARACTS WAS DEVICE-RELATED. BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721163&pc=MTA | 2023826-2017-01196 | 2017-06-07 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/4.5/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT, BLURRED VISION, AND LENS REPOSITIONING. LENS WAS "REMOVED AND CATARACT PROCEDURE HAVE BEEN DONE DUE TO CHANGES ON NATURAL LENS". THE PATIENT'S POST-OP CONDITION IS "WITH APHAKIC LENS AND HIS VISION IS 6/12". Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT, FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE), AND THE LENS MISSING PIECE OF HAPTIC. THE PATIENT'S POST-OP VISION IS 6/12. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6874039&pc=MTA | 2023826-2017-01447 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANT A 13.2 MM VTICMO13.2, -7.0/+4.5/107 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ON (B)(6) 2015, THE LENS WAS EXPLANTED DUE TO POST TRAUMATIC DISLOCATION IN A/C AND TRAUMATIC CATARACT DUE TO FIREWORKS MUCH LATER THAN THE IMPLANTATION. IT HAD NO CONNECTION TO THE SURGERY. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND PRESENCE OF FIBERS ON LENS SURFACE. THE PATIENT ALSO HAD BLURRED VISION. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE:TRAUMA CAUSED BY FIRWORK. NO CONNECTION TO SURGERY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6681259&pc=MTA | 2023826-2017-01089 | 2016-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.0/+1.5/40 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO REFRACTIVE SURPRISE AND THE DEVELOPMENT OF A CATARACT. A REPOSITIONING OF THE LENS, ENLARGEMENT OF THE INCISION (IN ORDER TO PERFORM CATARACT SURGERY), PHACO-EMULSIFICATION, AND IOL IMPLANTATION WAS PERFORMED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND THE PATIENT'S POST OP UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25 WITH STATUS OF THE EYE BEING PSEUDOPHAKIC. Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S. Manufacturer Narrative: "HOSPITALIZATION-INITIAL OR PROLONGED" WAS CHECKED BY MISTAKE. THERE WAS NO LENS REPOSITIONING. (B)(4). Manufacturer Narrative: THE PRODUCT WAS MANUFACTURED IN (B)(4) , NOT IN THE U.S.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978092&pc=MTA | 2023826-2017-01620 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -3.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS OPACITY-ASCO. SURGEON ALSO PERFORMED PHACO-EMULSIFICATION AND IOL IMPLANTATION. SURGEON STATES THAT THE PATIENT "DEVELOPED ASCO AFTER 5 MONTHS SURGERY." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS). BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/25 AT THE LAST EXAM WHICH WAS A DECREASE COMPARED TO BASELINE BCVA WHICH WAS 20/22. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021462&pc=MTA | 2023826-2017-01676 | 2017-09-06 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICM5_12.1, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, LOSS OF BCVA, FROM 20/40 TO 20/100, AND NS AND CORTICAL LENS OPACITY WAS NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. ON (B)(6) 2017, THE PATIENT'S POST-OP UCVA IS 20/30 AND THE STATUS OF THE STATUS IS PERFECT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH ANY CATARACT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5762131&pc=MTA | 2023826-2016-00817 | 2016-02-09 | Injury | STAAR SURGICAL COMPANY | 2016-06-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -11.00/ +6.0/100 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL) BLURRED VISION, THE CATARACT WAS REMOVED, AND A TORIC IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438479&pc=MTA | 2023826-2018-00544 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT THE PROBLEMS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM 12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. THE LENS WAS EXCHANGED FOR A LONGER LENS OF SIMILAR MODEL ON (B)(6) 2018 AND THE "VAULT WAS NORMAL AROUND 400UM APROX.". THE FINAL BCVA VALUE FOR PATIENT IS 0.8. EXPLANTED DATE: IT IS CORRECTED TO (B)(6) 2018. (B)(4). DEVICE MANUFACTURE DATE: IT IS CORRECTED TO 08-MAR-2011. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10190944&pc=MTA | 2023826-2020-01388 | 2019-05-12 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: A CASE REPORT WAS RECEIVED AND INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -22.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND THE PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT UNDERWENT FEMTOSECOND LASER ASSISTED CATARACT SURGERY (FLACS) COMBINED WITH ICL EXTRACTION AND AN IOL WAS IMPLANTED. VISUAL ACUITY WAS REDUCED FOR THE LAST 2 YEARS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513549&pc=MTA | 2023826-2015-00149 | 2014-11-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2014 LOW VAULT AND LENS OPACITY WAS OBSERVED AND THE PATIENT EXPERIENCED BLURRED VISON. THE LENS WAS EXCHANGED FOR A LONGER LENGTH WHICH RESOLVED THE VAULTING ISSUE, HOWEVER, THE CATARACT PROGRESED. THE SECOND LENS WAS REMOVED, THE CATARACT WAS EXTRACTED AND AN IOL WAS IMPLANTED. THE PATIENT'S LAST UCVA WAS 20/60. Manufacturer Narrative: .(B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RECEIVED DRY WITH A TORN HAPTIC. THERE WAS EVIDENCE OF A CLEAR SURGICAL RESIDUE. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). IN CASES OF POSTOPERATIVE LOW VAULTING STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS OR ROTATION ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPUSLAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, MEDICAL REVIEW AND EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. THE MOST PROBABLE CAUSE OF THIS EVENT IS LOW VAULT SECONDARY TO A SHORT LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5252472&pc=MTA | 2023826-2015-01553 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS HAD A LOW VAULT, REFRACTIVE SURPRISE AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE SURGEON IS PLANNING TO EXPLANT AND EXCHANGE FOR A LONGER LENS. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE: (B)(4) PREVIOUS CODE INCORRECT. NO INDUCED CATARACT WAS REPORTED. DEVICE CODE: (B)(4) (IMPLANTED)-NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. ADDITIONAL DATA: PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, EXPLANT AND EXCHANGE. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTION TORN, HAPTIC BROKEN, AND HAIR ON LENS SURFACE. LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS, DATE UNKNOWN. PROBLEM WAS RESOLVED. PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY IS 20/25. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074702&pc=MTA | 2023826-2016-01568 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR, POSTERIOR SUBCAPSULAR, NUCLEAR AND CORTICAL), WHICH WAS NOTED ON (B)(6) 2016. THE PATIENT EXPERIENCED BLURRED VISION. THE PATIENT WAS SCHEDULED FOR YAG CAPSULOTOMY 4 WEEKS AFTER LENS EXPLANT. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20 -3. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE/FIBERS ON THE LENS SURFACE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074711&pc=MTA | 2023826-2016-01569 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR AND POSTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2016. THE PATIENT EXPERIENCED BLURRED VISION. THE PATIENT WAS SCHEDULED FOR YAG CAPSULOTOMY. AT LAST POST-OP VISIT ON (B)(6) 2016, THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20 -2. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WERE FIBERS ON THE LENS. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT AND THE LENS WAS FOUND TO BE IN SPECIFICATION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6742631&pc=MTA | 2023826-2017-01242 | 2016-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE SURGEON ALSO NOTED THAT THE PATIENT'S VISUAL ACUITY WAS 20/60 ON (B)(6) 2016. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) AND EXPLANTED ON (B)(6) 2016. THE CATARACT WAS FIRST OBSERVED ON (B)(6) 2016 AND WAS LOCATED IN THE NUCLEAR PART OF THE EYE. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. (B)(4). DATE OF EVENT - (B)(6) 2016. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7870372&pc=MTA | 2023826-2018-01382 | 2017-06-01 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -15.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC); PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. ACCORDING TO THE SURGEON, THE EXACT CAUSE OF THE CATARACT IS UNKNOWN BUT MAY HAVE OCCURRED DUE TO HIGH MYOPIA. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFGUE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL DATA RECEIVED ON 15-FEB-2019: " THE PATIENT ALREADY HAD IOL IN THE LEFT EYE DUE TO CATARACT FORMATION. PATIENT CAME BACK TO THE CLINIC LAST SUMMER COMPLAINING BLURRY VISION AND DOCOTR FOUND NSC, NOT ASCO. HE ADVISED TO REMOVE ICL AND IMPLANTED IOL, (B)(6) 2018" (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848712&pc=MTA | 2023826-2019-01374 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: ON 03 JULY 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES OPACITIES AFFECTING VISUAL ACUITY IN 10 EYES. INCIPIENT ANTERIOR SUBCAPSULAR OPACITIES IN CONNECTION WITH A DETERIORATION OF BCVA BY AT LEAST 2 ROWS IN 3 MYOPIC EYES. PROGRESSING CATARACT PRONOUNCEDLY REDUCING VISION IN 7 EYES (5 MYOPIC EYES AND 2 HYPERMETROPIC), FOR WHICH EXPLANTATION OF THE ICL AND CATARACT SURGERY WITH IMPLANTATION OF PC IOL WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ON 03 JULY 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES INCIPIENT ANTERIOR SUBCAPSULAR OPACITIES IN CONNECTION WITH A DETERIORATION OF BCVA BY AT LEAST 2 ROWS IN 3 MYOPIC EYES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. ICM V4 USED FOR MYOPIA, TICM V4 USED FOR ASTIGMATISM. CLAIM#: (B)(4). Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT CORRECTED TO OTHER SERIOUS IN INITIAL MDR. COMMON DEVICE NAME PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA & PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. PATIENT CODE 3191 NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE:UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848714&pc=MTA | 2023826-2019-01375 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: ON 03 JULY 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES CATARACT SIGNIFICANTLY AFFECTING VISUAL ACUITY OCCURRED IN 7 EYES (5 MYOPIC EYES AND 2 HYPEROPIC EYES). ICL REMOVAL AND CATARACT SURGERY WITH IMPLANTATION OF POSTERIOR CHAMBER INTRAOCULAR LENS (PC IOL) WAS PERFORMED. ALL PATIENTS IN WHOM ICL EXPLANTATION AND CATARACT SURGERY WITH IMPLANTATION OF PC IOL WAS PERFORMED WERE SATISFIED, BCVA REACHED 1.0. ONLY 2 EYES OF ONE HIGHLY MYOPIC PATIENT WAS THERE DECENTRATION OF PC IOL FOLLOWING EXPLANTATION OF ICL AND IMPLANTATION OF PC IOL. THE PC IOL WAS EXPLANTED, PATIENT REMAINED APHAKIC. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: COMMON DEVICE NAME: PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA & PHAKIC TORIC, INTRAOCULAR LENS, PRODUCT CODE: QCB. DEVICE INFORMATION: ICM V4 USED FOR MYOPIA, ICH V3 USED FOR HYPEROPIA, TICM V4 USED FOR ASTIGMATISM. PATIENT CODE 2137 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. DEVICE CODE 1401 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7438480&pc=MTA | 2023826-2018-00543 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) IN 2011. ON (B)(6) 2018, LOW VAULT, BLURRED VISION, AS WELL AS ASC AND CORTICAL LENS OPACITY WERE NOTICED. SURGERY IS PLANNED TO CORRECT PROBLEMS. Manufacturer Narrative: THE BEST CORRECTED VISUAL ACUITY (BCVA) FOR THE PATIENT IS CURRENTLY 1.0. (B)(4) Manufacturer Narrative: "SURGERY IS PLANNED TO CORRECT PROBLEMS." IN INITIAL MDR SHOULD BE UPDATED TO "THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. NO INFORMATION REGARDING THE DETAILS OF THE SECONDARY SURGERY HAS BEEN PROVIDED DESPITE THE REQUESTS FOR DETAILS." PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: IMPLANTED 2011. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938159&pc=MTA | 2023826-2020-00798 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Cataract; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) IN (B)(6) 2019. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE VAULT IS MINIMAL. LASIK ENHANCEMENT WAS PERFORMED 9 MONTHS LATER. THE PATIENT HAS NOW DEVELOPED AN ANTERIOR CORTICAL CATARACT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER ON (B)(6) 2019. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D6: IMPLANTED DATE - (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611974&pc=MTA | 2023826-2019-00792 | 2019-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -4.50/0.5/150 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOSS OF BCVA, BLURRED VISION, AND REFRACTIVE CHANGE OVERTIME WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PHACOEMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS SCHEDULED FOR (B)(6) 2019 BUT THE PATIENT PASSED AWAY ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -1.75 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA WAS OBSERVED. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS FOLLOWING ICL IMPLANTATION. PHACOEMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION WAS PERFORMED ON THE RIGHT EYE 3 WEEKS AFTER EXPLANTATION. THE PATIENT PASSED AWAY ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9377871&pc=MTA | 2023826-2019-02249 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; No Code Available | N/A | 1 | Event Description: ON (B)(6) 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN INDIAN JOURNAL OF OPHTHALMOLOGY, ' CHALLENGES DURING FEMTOSECOND LASER ASSISTED CATARACT SURGERY WITH POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS.' THE ARTICLE REPORTS ONE CASE OF CATARACT FOLLOWING ICL WITH SUBSEQUENT EXPLANT AND FLACS (FEMTOSECOND LASER-ASSISTED CATARACT SURGERY). PER CASE REPORT A (B)(6) MAN WITH HIGH MYOPIA, WHO HAD UNDERGONE ICL SURGERY PRESENTED WITH PROGRESSIVELY DECREASED VISION IN HIS RIGHT EYE FOR THE PAST 2 YEARS. SLIT LAMP EXAMINATION SHOWED ICL IN SITU WITH A PATENT PERIPHERAL IRIDOTOMY AND NORMAL INTRAOCULAR PRESSURE. ICL WAS NOT IN CONTACT WITH THE LENS, LENS VAULT APPEARED NORMAL. NUCLEAR CATARACT WAS OBSERVED. THE ICL WAS EXPLANTED AND FLACS WAS PERFORMED. A THREE- PIECE HYDROPHOBIC FOLDABLE IOL WAS IMPLANTED. PATIENT REGAINED UNCORRECTED DISTANT VISION 20/20 AT 1-MONTH POST CATARACT SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE 3191: WAS REPORTED WITH CORRECT INFORMATION IN MANUFACTURER NARRATIVE IN INITIAL MEDWATCH REPORT, BUT CODE SHOULD HAVE BEEN SELECTED IN H6 SECTION. H6 - PATIENT CODE: (B)(6) SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. H6 - DEVICE CODE: 1401 SHOULD BE CORRECTED TO 2993. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANTED: UNK. EXPLANTED:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10164930&pc=MTA | 2023826-2020-01278 | 2019-05-12 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2020-06-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; No Code Available | NA | 1 | Event Description: A CASE REPORT WAS RECEIVED AND INDICATED THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENSES IN THE PATIENT OU. SIX YEARS AFTER THE LENSES WERE IMPLANTED, THE PATIENT DEVELOPED ANTERIOR SUBCAPSULAR CATARACTS OU. THE PATIENT UNDERWENT FEMTOSECOND LASER ASSISTED CATARACT SURGERY (FLACS) COMBINED WITH ICL EXTRACTION. IOLS WERE IMPLANTED OU. THE LENS IN THE LEFT EYE (OS) TOUCHED THE ANTERIOR CAPSULE DUE TO "0" VAULT. THE VAULT IN THE RIGHT EYE (OD) WAS SHALLOW. VISUAL ACUITY WAS REDUCED FOR THE LAST 2 YEARS OU. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D3: MANUFACTURER NAME, CITY AND STATE. G2: MANUFACTURING SITE ADDRESS/PHONE FOR DEVICES. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722066&pc=MTA | 2023826-2022-01932 | 2022-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Blurred Vision; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. IT WAS REPORTED THAT THE PATIENT HAS LENS OPACITY (CATARACT ASC). IT WAS ALSO NOTED "THE INVESTIGATOR DISCUSSED THE DEVELOPMENT OF ANTERIOR SUBCAPSULAR OPACITIES WITH THE PATIENT AND THE DECISION WAS MADE TO MONITOR THE VISION FOR NOW". THE PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4614" AND "2199" SHOULD BE REMOVED AND CODE "4627" AND "4625" SHOULD BE ADDED. H6-HEALTH EFFECT- CLINICAL CODE: "2140" AND "2137" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "1401" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S LEFT EYE (OS) ON 03/JUNE/2020 AS A BACKUP LENS. ON 09/MAY/2022 LENS OPACITY CATARACT (ASC) WAS OBSERVED. THE PATIENT ALSO EXPERIENCED BLURRY/HAZY VISION. ON 29/AUG/2023 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED. THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "SUBJ INITIALLY REPORTED BLURRY & HAZY VA OS AT V7 (09MAY22). PI NOTED AN ASC OS. PI & SUBJ DECIDED TO MONITOR. AT V8 (28JUL23), SUBJ REPORTED VA HAD WORSENED, ESP IN BRIGHT LIGHTS. PI & SUBJ DECIDED TO PROCEED W/ CATARACT SX OS.SAE EVENT SUMMARY 2 : SUBJ HAD UNEVENTFUL ICL EXPLANT & SIMULTANEOUS CATARACT EXTRACTION SURGERY ON (B)(6) 2023. UCVA OS IMPROVED TO 20/12.5 AT 1 DAY PO. SITE WILL CONT TO FOLLOW SUBJ TO ENSURE HEALING AND STABILITY". THE CAUSE OF THE EVENT WAS THE DEVICE. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069167&pc=MTA | 2023826-2017-01792 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069168&pc=MTA | 2023826-2017-01791 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017, ANTERIOR LENS OPACITY WAS NOTICED AS WELL AS BLURRY VISION AT DISTANCE WITH SIGNIFICANT GLARE AND HALOES AND LOW VAULT AS THE ICL IS RUBBING AGAINST THE NATURAL LENS. THE LENS REMAINS IMPLANTED. THE LENS IS CURRENTLY BEING PLANNED TO BE EXCHANGED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6247963&pc=MTA | 2023826-2017-00071 | 2011-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -16.00/+1.50/094 DIOPTER, IN THE PATIENT'S LEFT EYE. THE LENS WAS EXPLANTED DUE TO LOW VAULTING, REFRACTIVE SURPRISE, LENS OPACITY, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNTING OR SIGNIFICANT ENDOTHELIAL CELL LOSS, GLARE/HALOES, AND BLURRED VISION. THE PATIENT UNDERWENT PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6877264&pc=MTA | 2023826-2017-01463 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM ICM12.0 IMPLANTABLE COLLAMER LENS, -21.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2016 THE LENS WAS EXPLANTED. THE SURGEON NOTED LOW VAULT, LENS ROTATION, LENS OPACITY ASCO, AND "ARTIPHAKIA" AS THE STATUS OF THE EYE. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION IS ALSO REPORTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. THE PATIENT'S DATE OF BIRTH IS CORRECTED FROM "(B)(6) 1986" TO "(B)(6) 1983". BLURRED VISION AND PRESENCE OF GLARES/ HALOES EXPERIENCED BY THE PATIENT, SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.9MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6129538&pc=MTA | 2023826-2016-01682 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Cataract; Blurred Vision; Visual Impairment; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND SURGEON REPOSITIONED THE NEW LENS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 12.5/+2.00/03 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION (NOT ASSOCIATED WITH LOW VAULTING), LENS OPACITY, LENS DISLOCATION/SUBLUXATION, REFRACTIVE CHANGE OVER TIME AND BLURRED VISION. THE SURGEON EXCHANGED THE LENS FOR A LONGER LENS AND REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS OPACITY, LENS DISLOCATION/SUBLUXATION, AND BLURRED VISION. REPOSITIONING OF THE LENS WAS PERFORMED AND THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. (B)(4). Manufacturer Narrative: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/2.0/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED LENS ROTATION, DISLOCATION, REFRACTIVE CHANGE, AND BLURRED VISION. THE SURGEON REPOSITIONED THE LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE SURGEON REPORTED LENS OPACITY (ASCO) AND LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS ON (B)(6) 2016 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8089155&pc=MTA | 2023826-2018-01717 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Capsular Bag Tear; No Code Available | N/A | 1 | Event Description: THE ARTICAL INDICATED THAT THE SURGEON IMPLANTED A VTICH, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). REPORTING THAT PERIPHERAL ANTERRIOR LENS CAPUSLE DAMAGE AND LARGE PERIPHERAL IRIDECTOMY (PI) WAS OBSERVED WITH INTRAOPTERATIVE PI DURING VTICH IMPLANTAION. AT 2 WEEKS POST-OP, UCVA WAS 20/20. BOTH THE TRAUMATIC CATARACT AND LARGE PI DID NOT AFFECT VISION AND PATIENT IS ASYMPTOMATIC. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: AN ARTICLE INDICATED THAT A SURGEON IMPLANTED A VTICH IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). INTRAOPERATIVE PI WAS PERFORMED. IT WAS REPORTED THAT THE ANTERIOR LENS CAPSULE WAS DAMAGED ALONG WITH PRESENCE OF TRAUMATIC CATARACT AND LOW VAULT. A LARGE PI WAS OBSERVED. BOTH THE TRAUMATIC CATARACT AND LARGE PI DID NOT AFFECT VISION AND THE PATIENT IS ASYMPTOMATIC. SUPPOSEDLY THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5694861&pc=MTA | 2023826-2016-00799 | 2016-04-14 | Injury | STAAR SURGICAL COMPANY | 2016-06-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Loss of Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED ASC (ANTERIOR SUBCAPSULAR CATARACT) ON (B)(6) 2016 AND SUBSEQUENTLY LOST BCVA (BEST-CORRECTED VISUAL ACUITY). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572298&pc=MTA | 2023826-2016-00432 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Blurred Vision; Visual Impairment | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/X010 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THERE WAS LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015, AND THE PATIENT EXPERIENCED LOSS OF BCVA. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/200. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5816274&pc=MTA | 2023826-2016-00974 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, LENS OPACITY (CATARACT), AND LOSS OF VISION. THE PATIENT EXPERIENCED AN ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, BLURRED VISION, IRIS ATROPHY, AND CORNEAL EDEMA. THE ANTERIOR CHAMBER IRRIGATION OF REMAINING VISCO FLUIDS WAS PERFORMED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/200. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND SPOT ON LENS SURFACE. DIMENSIONAL INSPECTION CONLUDED THAT THE LENS IS IN SPECIFICATION. (B)(4). Manufacturer Narrative: NO REPORTED LOSS OF VISION. DEVICE EVALUATION: VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661678&pc=MTA | 2023826-2016-00727 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Cataract, Induced; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. CATARACT SURGERY WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7886291&pc=MTA | 2023826-2018-01403 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE LENS IS SCHEDULED TO BE REMOVED DUE TO CATARACT FORMATION. IT IS NOT KNOWN WHETHER THE LENS HAS BEEN REMOVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED INFORMATION: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ON (B)(6) 2018, THE LENS WAS REMOVED FROM THE PATIENT'S EYE DUE TO CATARACT FORMATION. ON THE SAME DAY, THE PATIENT THEN UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. IN THE SURGEON'S OPINION THE CAUSE OF THE CATARACT FORMATION IS THE ICL. THE PATIENT AND VISION ARE CURRENTLY STABLE. ADDITIONAL REPORT SOURCE ADDED. (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553032&pc=MTA | 2023826-2016-00401 | 2014-08-13 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -19.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND NO VISIBLE DAMAGE; MEASUREMENT FOUND THE LENS WITHIN SPECIFICATIONS. PREVIOUS (B)(4) (NEW INCISION) WAS INCORRECT; ONLY SECONDARY SURGERY FOR EXPLANT WAS REPORTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336746&pc=MTA | 2023826-2018-00379 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. Manufacturer Narrative: CORRECTED DATA: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT EXTRACTION WITH IOP PLACEMENT) CONCLUSION CODE: OFF-LABEL USE OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336750&pc=MTA | 2023826-2018-00381 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>(B)(6) YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4) . CONCLUSION CODE: OFF-LABEL USE OVER (B)(6) YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336751&pc=MTA | 2023826-2018-00384 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT EXTRACTION WITH IOP PLACEMENT). CONCLUSION CODE: OFF-LABEL USE (OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336752&pc=MTA | 2023826-2018-00392 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: INITIAL FIRST SENTENCE STATES THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. CORRECTION-IT SHOULD STATE THAT THE SURGEON IMPLANTED A V4C IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) Manufacturer Narrative: PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT EXTRACTION WITH IOP PLACEMENT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336755&pc=MTA | 2023826-2018-00380 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: INITIAL FIRST SENTENCE STATES THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. CORRECTION-IT SHOULD STATE THAT THE SURGEON IMPLANTED A V4C IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT EXTRACTION WITH IOP PLACEMENT) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336760&pc=MTA | 2023826-2018-00383 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: CORRECTED DATA: ARTICLE STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATENT'S EYE. MANUFACTURER NAME, CITY AND STATE IS UNKNOWN. MANUFACTURING SITE ADDRESS/PHONE FOR DEVICES IS UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE: OFF-LABEL USE (OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336761&pc=MTA | 2023826-2018-00385 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE (OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336762&pc=MTA | 2023826-2018-00386 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>(B)(6) YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE (OVER (B)(6) YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM). (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336763&pc=MTA | 2023826-2018-00382 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) Manufacturer Narrative: (B)(4). CONCLUSION CODE: OFF-LABEL USE OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336766&pc=MTA | 2023826-2018-00390 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: MODEL #: IT IS CORRECTED FROM "ICM120V4" TO UNK. ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND INTRAOCULAR LENS PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE PRIMARILY DUE TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE ACD < 3 MM AND AGE > 45 YEARS. NO FURTHER INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE (OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336769&pc=MTA | 2023826-2018-00388 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: MODEL #: IT IS CORRECTED FROM "ICM120V4" TO UNK. ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND INTRAOCULAR LENS PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE PRIMARILY DUE TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE ACD < 3 MM AND AGE > (B)(6) YEARS. NO FURTHER INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: SERIAL NUMBER, EXPIRATION DATE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). OFF-LABEL USE (OVER 45 YEARS OF AGE AT TIME OF IMPLANT, ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336770&pc=MTA | 2023826-2018-00389 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: MODEL #: IT IS CORRECTED FROM "ICM120V4" TO UNK. ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND INTRAOCULAR LENS PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE PRIMARILY DUE TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE ACD < 3 MM AND AGE > 45 YEARS. NO FURTHER INFORMATION IS AVAILABLE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7336772&pc=MTA | 2023826-2018-00387 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICM120V4 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXTRACTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM AND AGE>45YRS. NO FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339938&pc=MTA | 2023826-2018-00364 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) Manufacturer Narrative: EXPIRATION DATE: UNK. THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339954&pc=MTA | 2023826-2018-00365 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE GREATER THAN (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339957&pc=MTA | 2023826-2018-00366 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339969&pc=MTA | 2023826-2018-00367 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE MANUFACTURER NAME, CITY AND STATE IS UNKNOWN. THE MANUFACTURING SITE ADDRESS/PHONE FOR DEVICES IS UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339971&pc=MTA | 2023826-2018-00368 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4).CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7342611&pc=MTA | 2023826-2018-00369 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: EXPIRATION DATE: UNK, THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343204&pc=MTA | 2023826-2018-00370 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: EXPIRATION DATE: UNK. THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343233&pc=MTA | 2023826-2018-00371 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343271&pc=MTA | 2023826-2018-00372 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343281&pc=MTA | 2023826-2018-00373 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343296&pc=MTA | 2023826-2018-00374 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE>45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4) Manufacturer Narrative: EXPIRATION DATE: UNK. THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334594&pc=MTA | 2023826-2018-00346 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: "MICROSTAAR" INJECTOR IS ADDED TO THE INITIAL MDR. REPORT SOURCE: "(B)(6) " IS ALSO ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM AND IN A PATIENT OVER THE AGE OF (B)(6). (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334595&pc=MTA | 2023826-2018-00344 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF 45 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334596&pc=MTA | 2023826-2018-00347 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM AND IN A PATIENT OVER THE AGE OF 45 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334597&pc=MTA | 2023826-2018-00348 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334598&pc=MTA | 2023826-2018-00349 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: MICROSTAAR INJECTOR IS ADDED TO THE INITIAL MDR. REPORT SOURCE: "(B)(6)" IS ADDED TO THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM AND IN A PATIENT OVER THE AGE OF (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334601&pc=MTA | 2023826-2018-00352 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334602&pc=MTA | 2023826-2018-00353 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334609&pc=MTA | 2023826-2018-00345 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE: (B)(4) SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANT. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM AND IN A PATIENT OVER THE AGE OF 45 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334614&pc=MTA | 2023826-2018-00350 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: MICROSTAAR INJECTOR IS ADDED TO THE INITIAL MDR. REPORT SOURCE: "(B)(6 " IS ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334616&pc=MTA | 2023826-2018-00343 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENTS OVER THE AGE OF 45 YEARS OLD. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7334622&pc=MTA | 2023826-2018-00351 |  | Injury | Unknown Manufacturer | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT HAS NOT BEEN FORTHCOMING. Manufacturer Narrative: (B)(4) Manufacturer Narrative: NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343521&pc=MTA | 2023826-2018-00399 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) < 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343634&pc=MTA | 2023826-2018-00400 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA 510(K):THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) < 3.0MM. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343658&pc=MTA | 2023826-2018-00401 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE: PER DFU THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) < 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343739&pc=MTA | 2023826-2018-00402 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF ACD<3MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) < 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343748&pc=MTA | 2023826-2018-00404 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343749&pc=MTA | 2023826-2018-00403 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: MICROSTAAR INJECTOR IS ADDED TO THE INITIAL MDR. REPORT SOURCE: "(B)(6)" IS ADDED TO THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7343765&pc=MTA | 2023826-2018-00405 |  | Injury | Unknown Manufacturer | 2018-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7345378&pc=MTA | 2023826-2018-00406 |  | Injury | Unknown Manufacturer | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE GREATER THAN (B)(6). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7345478&pc=MTA | 2023826-2018-00407 |  | Injury | Unknown Manufacturer | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7345498&pc=MTA | 2023826-2018-00408 |  | Injury | Unknown Manufacturer | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7345513&pc=MTA | 2023826-2018-00409 |  | Injury | Unknown Manufacturer | 2018-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED WITH CATARACT EXTRACTION AND IOL PLACEMENT DUE TO AN ANTERIOR SUBCAPSULAR CATARACT (ASC). PATIENT REPORTED VISION AS FAIR OR WORSE DUE PRIMARILY TO CATARACT. PATIENT WAS CONSIDERED TO BE "OUTSIDE OF FDA GUIDELINES" BECAUSE OF AGE >45 YRS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: "MICROSTAAR" INJECTOR IS ADDED TO THE INITIAL MDR. REPORT SOURCE: "(B)(6)" IS ALSO ADDED TO THE INITIAL MDR. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140257&pc=MTA | 2023826-2016-01700 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING A CATARACT AND RECEIVING PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND THE LENS SCRATCHED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS OVER THE AGE OF (B)(6). CORRECTED DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING CATARACT. SECONDARY SURGICAL INTERVENTION, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION, WAS PERFORMED ON PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6140262&pc=MTA | 2023826-2016-01699 | 2016-10-18 | Injury | STAAR SURGICAL COMPANY | 2016-12-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM513.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT DEVELOPING A CATARACT AND RECEIVING PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN PATIENTS OVER THE AGE OF 45 YEARS OLD. CORRECTED DATA: THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT DEVELOPING CATARACT. SECONDARY SURGICAL INTERVENTION, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION, WAS PERFORMED ON PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293149&pc=MTA | 2023826-2015-01656 | 2013-08-17 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -15.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2013. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAIR ON SURFACE BUT NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: DIOPTER: -15.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378618&pc=MTA | 2023826-2015-01669 | 2014-11-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/087 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2015) UNCORRECTED VISUAL ACUITY WAS 20/15. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5266255&pc=MTA | 2023826-2015-01603 | 2014-12-27 | Injury | STAAR SURGICAL COMPANY | 2015-12-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM12.1, -11.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. A ANTERIOR SUBCAPSULAR CATARACT OPACITY WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A VICMO LENS, SAME LENGTH AND SIMILAR DIOPTER SIZE ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. THE CATARACT IS BEING MONITORED. POST-OP, PATIENT'S UCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5293173&pc=MTA | 2023826-2015-01654 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-12-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM 115V4 -13.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2007. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Event Description: CATARACT SURGERY AND IOL IMPLANTATION WERE PERFORMED ON (B)(6) 2015. Manufacturer Narrative: (B)(4).CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). Manufacturer Narrative: DATE OF EVENT: PREVIOUS DATE INCORRECT. DESCRIBE EVENT OR PROBLEM: PREVIOUS DATE INCORRECT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304862&pc=MTA | 2023826-2015-01615 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED AND OBSERVED NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (TISSUE). A DIMENSIONAL INSPECTION WAS PERFORMED AND FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305142&pc=MTA | 2023826-2015-01614 | 2015-01-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5378594&pc=MTA | 2023826-2015-01668 | 2015-04-03 | Injury | STAAR SURGICAL COMPANY AG | 2016-01-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. CORRECTIONAL DATA: PATIENT CODE - (B)(4) PREVIOUS CODE NOT APPLICABLE, UNKNOWN CAUSE FOR LENS OPACITY. PATIENT CODE - (B)(4) (NEW INCISION) - CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. PATIENT CODE - (B)(4) (EXPLANTED) - SHOULD BE INCLUDED. SHOULD NOT BE LISTED AS DEVICE CODE. DEVICE CODE - (B)(4) (EXPLANTED) - SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5292426&pc=MTA | 2023826-2015-01655 | 2015-08-20 | Injury | STAAR SURGICAL COMPANY | 2015-12-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 -19.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2005. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED, CATARACT SURGERY AND A IOL WAS IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015, BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: DIOPTER: -19.00; DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549348&pc=MTA | 2023826-2016-00387 | 2015-09-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 -22.00 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2013. AN ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED AND ON (B)(6) 2015, THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -22.00. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919743&pc=MTA | 2023826-2016-01203 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE REPORTER INDICATED THERE WAS ADHESION AND DEGENERATION OF THE IRIS. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/400. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (NOT POSSIBLE TO SPECIFY). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5671142&pc=MTA | 2023826-2016-00741 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+0.5./X160 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (CORTICAL), NOTED ON (B)(6) 2015. THE PATIENT HAD CATARACT SURGERY AND AN INTRAOCULAR LENS WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. LENS RETURNED BENT WAS ALSO NOTED. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5931126&pc=MTA | 2023826-2016-01226 | 2016-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM ICM120V4 12.0 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT RECEIVED PHACO-EMULSIFICATION (CATARACT SURGERY), AND AN IOL WAS IMPLANTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5934314&pc=MTA | 2023826-2016-01225 | 2016-03-04 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0 MM ICM120V4 12.0 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT RECEIVED PHACO-EMULSIFICATION (CATARACT SURGERY), AND AN IOL WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMGE AND DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923611&pc=MTA | 2023826-2016-01208 | 2016-05-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (SPOTS). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) SECONDARY SURGERY. EXPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5918280&pc=MTA | 2023826-2016-01181 | 2016-06-14 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (HAIRS) AND NO VISIBLE DAMAGE TO LENS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PMA 510(K): - THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5972078&pc=MTA | 2023826-2016-01311 | 2016-07-11 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). Manufacturer Narrative: SERIAL # (B)(4), EXPIRATION DATE - 10/31/2018. DEVICE MANUFACTURE DATE - 11/07/2015. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. CONCLUSION CODE, AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OLD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553040&pc=MTA | 2023826-2016-00403 | 2015-06-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919728&pc=MTA | 2023826-2016-01201 | 2015-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -03.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS DEVICE IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PATIENT CODE: (B)(4) - PREVIOUS CODE NOT APPLICABLE; NO INDUCED CATARACT REPORTED. DEVICE CODE: (B)(4) (EXPLANTED) - CODE NOT APPLICABLE; SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5919733&pc=MTA | 2023826-2016-01202 | 2015-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -03.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED.. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE THAT IS NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553044&pc=MTA | 2023826-2016-00405 | 2015-11-20 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955097&pc=MTA | 2023826-2016-01269 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM12.1 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON 03/14/2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/21. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT, LENS HAPTIC TORN/BROKEN/BENT/DEFORMED, AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5917886&pc=MTA | 2023826-2016-01178 | 2016-02-11 | Injury | STAAR SURGICAL COMPANY AG | 2016-08-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 . THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS OPACITY. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED ON THE PATIENT, AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4): ONLY SECONDARY INTERVENTION, LENS EXPLANT WAS REPORTED. PATIENT ALSO HAD A PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION. DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE THAT IS NOT POSSIBLE TO SPECIFY. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: EVALUATION FOUND LENS ARE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579726&pc=MTA | 2023826-2016-00480 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), OBSERVED ON (B)(6) 2016. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE CORRECT EVENT DATE REPORTED WAS (B)(6) 2016'. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID AND IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND LENS IS IN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5814335&pc=MTA | 2023826-2016-00916 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-07-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR)/PERIPHERAL TOUCH. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND SPOTS ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5971116&pc=MTA | 2023826-2016-01228 | 2016-08-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS OPACITY. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. CORRECTED INFORMATION: (B)(4). THE REPORTER DID NOT STATE THE CATARACT WAS INDUCED. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5931212&pc=MTA | 2023826-2016-01224 | 2015-06-13 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; Cataract, Induced; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5 MM ICM115V4 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULT AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT IS NOT MANUFACTURED IN THE U.S. AND NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8002463&pc=MTA | 2023826-2018-01537 | 2018-05-17 | Injury | STAAR SURGICAL COMPANY | 2018-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO A NUCLEAR SCLEROTIC CATARACT AND A CAPSULAR TEAR. THE REPORTER INDICATED DURING THE SURGERY, THE PATIENT'S BILATERAL SUSPENSORY LIGAMENT WAS LOOSE AND THE LEFT ANTERIOR SUSPENSORY LIGAMENT WAS BROKEN. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENT WAS NERVOUS). Manufacturer Narrative: CORRECTED DATA: (B))(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908388&pc=MTA | 2023826-2019-01481 | 2019-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: H6- DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908395&pc=MTA | 2023826-2019-01482 | 2019-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Cataract, Induced; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED EARLY ONSET CATARACT AND THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO DEVELOPMENT OF A CATARACT (ASC) AND REDUCED VISION. CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO DEVICE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, CATARACT SURGERY (PHACOEMULSIFICATION), IOL IMPLANTATION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: PATIENT CODE 1766 SHOULD BE CORRECTED TO PATIENT CODE 1767 IN THE INITIAL MDR. ADDITIONAL INFORMATION: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464314&pc=MTA | 2023826-2023-00587 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Decompensation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). POST-OP THE EYE HAS LOW ECC COUNTING, ANTERIOR AND SUBCAPSULAR CATARACT. THE EYE IS PENDING CATARACT SURGERY AND ENDOTHELIAL TRANSPLANTATION/DUE TO CORNEAL DECOMPENSATION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK WEIGHT: UNK ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK EXPIRATION DATE UNK IMPLANT DATE: UNK EXPLANT DATE: UNK HEALTH IMPACT- CLINICAL CODE: 4581 - LOW ECC COUNTING CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464332&pc=MTA | 2023826-2023-00586 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Decompensation; Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED DUE TO POOR VISION, CORNEAL EDEMA, IRIS-CORNEA SYNECHIA, LOW ECC COUNTING, ANTERIOR AND POSTERIOR SUBCAPSULAR CATARACT AND DEFORMED PUPIL. THE LENS WAS REPLACED WITH A SHORTER LENS (12.1MM). THE RIGHT EYE IS PENDING CATARACT SURGERY AND ENDOTHELIAL TRANSPLANTATION/ DUE TO CORNEAL DECOMPENSATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - SYNECHIA IRIS-CORNEA, DEFORMED PUPIL, POOR VISION, LOW ECC COUNTING. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15009430&pc=MTA | 2023826-2022-02285 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022; AS A REPLACEMENT LENS FOR A PREVIOUS REFRACTIVE SURPRISE. THE PATIENT DEVELOPED LENS OPACITY (ASC) WHICH WAS ASSESSED ON (B)(6) 2022; CORNEAL/STROMAL OEDEMA AND DESCEMET FOLDS AND PHOTOPHOBIA WAS ALSO REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. PHACO-EMULSIFICATION WAS PERFORMED AND AN IOL FROM A DIFFERENT MANUFACTURER WAS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. POSTOPERATIVE COMMENTS NOTED "THE UCVA AND BCVA HAS NOT INCREASED YET BECAUSE THERE ARE STILL CORNEAL & STROMAL OEDEMA AND DESCEMET FOLDS. THE PATIENT IS UNDER CLOSE MONITORING". THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR AND NOTED "DURING THE EXCHANGE THE CRYSTALLINE LENS WAS TOUCHED WITH THE ICL MANIPULATOR BY THE SURGEON". Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464587&pc=MTA | 2023826-2023-00588 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) IN 2022. THE LENS WAS REMOVED DUE TO POOR VISION, SINEQUIAS OF THE IRIS TO THE CORNEA AND CORNEAL EDEMA WITH VERY LOW ECC COUNTING. THE LENS WAS REPLACED WITH A SHORTER LENS. POST-OP, THERE IS AN ANTERIOR SUBCAPSULAR CATARACT AND OVAL PUPIL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - SUTURE WAS OBSERVED. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM ICL IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) IN 2022. THE LENS WAS REMOVED DUE TO ANTERIOR SUBCAPSULAR CATARACT, PUPIL OVALIZATION, POOR VISION, SINEQUIAS OF THE IRIS TO THE CORNEA AND CORNEAL EDEMA WITH VERY LOW ECC COUNTING. THE LENS WAS REPLACED WITH A SHORTER LENS (12.1MM, -07.00 DIOPTER, IMPLANT VERTICAL), THEN ONE WEEK AFTER EXCHANGE NORMAL. A SUTURE WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - POOR VISION, SINEQUIAS OF IRIS TO CORNEA, LOW ECC COUNTING, OVAL PUPIL. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987375&pc=MTA | 2023826-2018-01533 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Corneal Edema; Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/060 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), THE LAST WEEK OF (B)(6) 2018. THE REPORTER INDICATED THE PATIENT HAD A FIXED DILATED PUPIL POST-OP, SEVERE CORNEAL EDEMA, WHICH HAS BEEN PERSISTENT UNTIL NOW AND AN ANTERIOR SUBCAPSULAR CATARACT. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, BLURRED VISION AND PUPIL/IRIS ATROPHY. THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED ONE MONTH AFTER THE SURGERY AND AN IOL WAS IMPLANTED. THE PUPIL REMAINED DILATED AND THE CORNEA WAS EDEMATOUS AND WHITE. THE PATIENT HAS HAD A CORNEAL TRANSPLANT. THE REPORTER INDICATED THE INTRAOCULAR PRESSURE HAS BEEN STABLE AND NORMAL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENTS VISION IS STILL HAND MOTION POST OPERATIVELY AND INTRAOCULAR PRESSURE IS 12MMHG. THE FIRST CHECKUP AFTER KERATOPLASTY INDICATED ALL IS GOOD. CORRECTED DATA: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495099&pc=MTA | 2023826-2023-03441 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Corneal Edema; Hyphema; Intraocular Infection; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED IMPLANTABLE COLLAMER LENS IMPLANTATION (ICL) VERSUS SMALL INCISION LENTICULE EXTRACTION (SMILE) IN LOW TO MODERATE MYOPIA: STUDY PROTOCOL FOR A RANDOMIZED, NON-INFERIORITY TRIAL - THIS STUDY AIMED TO EVALUATED WHETHER ICL/TICL IS COMPARABLE TO SMILE FOR LOW TO MODERATE MYOPIA IN TERMS OF REFRACTIVE OUTCOMES AT 3 AND 18 MONTHS POSTOPERATIVELY. IN THIS STUDY IT WAS INDICATED ADVERSE EVENTS SUCH AS LENS IMPAIRMENT, IRIS PROLAPSE, HYPHEMA, OCULAR HYPERTENSION, TRANSIENT CORNEA EDEMA, CORNEAL ENDODERMIS DAMAGE, VAULT ABNORMALITY, SURGERY RELATED CATARACT AND INTRAOCULAR INFECTION, HAD OCCURRED FOR PATIENT WITH IMPLANTABLE COLLAMER LENSES. Manufacturer Narrative: CORRECTED DATA: B5- ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. THE ARTICLE DESCRIBES A PROSPECTIVE STUDY THAT HAS NOT YET OCCURRED. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - LENS IMPAIRMENT, IRIS PROLAPSE, OCULAR HYPERTENSION, CORNEAL ENDODERMIS DAMAGE, CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10356766&pc=MTA | 2023826-2020-01609 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Corneal Edema; Retinal Detachment; Blurred Vision; Capsular Bag Tear; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY OF 87 EYES OF 55 MYOPIC PATIENTS, WHO UNDERWENT BILENSECTOMY (PHAKIC INTRAOCULAR LENS PIOL - EXPLANTATION WITH CATARACT SURGERY AND IOL IMPLANTATION), FOR POSTERIOR CHAMBER PIOL, WITH A FOLLOW UP TIME OF 12 MONTHS. THE MAIN REASON FOR BILENSECTOMY WAS CATARACT DEVELOPMENT, FOLLOWED BY MISCALCULATION OF LENS SIZE AND CORNEAL EDEMA. INTRAOPERATIVE COMPLICATIONS WERE ONE POSTERIOR CAPSULE RUPTURE AND 3 EYES REQUIRED USE OF PUPIL EXPANDERS FOR ADEQUATE PUPIL DILATION. POSTOPERATIVELY, ONE EYE DEVELOPED RETINAL DETACHMENT. GOOD SAFETY AND VISUAL OUTCOMES WERE OBSERVED 1 YEAR AFTER BILENSECTOMY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5 - CORRECTED TO "AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY OF 87 EYES OF 55 MYOPIC PATIENTS, WHO UNDERWENT BILENSECTOMY (PHAKIC INTRAOCULAR LENS PIOL - EXPLANATION WITH CATARACT SURGERY AND IOL IMPLANTATION), FOR POSTERIOR CHAMBER PIOL, WITH A FOLLOW UP TIME OF 12 MONTHS. THE MAIN CAUSE OF BILENSECTOMY WAS DUE TO CATARACT DEVELOPMENT, CORNEAL EDEMA, IOL DISLOCATION, REFRACTIVE SURPRISE. INTRAOPERATIVE COMPLICATIONS WERE ONE POSTERIOR CAPSULE RUPTURE AND 3 EYES REQUIRED USE OF PUPIL EXPANDERS FOR ADEQUATE PUPIL DILATION. POSTOPERATIVELY, THERE WAS ONE CASE OF HYDRO REACTIVE MYDRIATIC PUPIL REQUIRING PUPILLOPLASTY , RETINAL DETACHMENT, AND POSTERIOR CAPSULE OPACIFICATION, AND TWO CASES OF PIGMENT OVER THE TRABECULAR MESHWORK AND HIGH IOP REQUIRING A COMBINED PROCEDURE WITH TRABECULECTOMY. GOOD SAFETY AND VISUAL OUTCOMES WERE OBSERVED 1 YEAR AFTER BILENSECTOMY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED." IN INITIAL MDR. H6 - PATIENT CODE: 2137 - REFRACTIVE SURPRISE SHOULD BE ADDED IN INITIAL MDR. H6 - DEVICE CODE: 2923 - LENS DISLOCATION SHOULD BE ADDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MFR DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11944308&pc=MTA | 2023826-2021-01751 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Dry Eye(s); Pupillary Block; Retinal Detachment | P030016 | 1 | Event Description: ON (B)(6) 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE ENTITLED, ' SAFETY OF PHAKIC INTRAOCULAR COLLAMER LENS (VISIAN ICL) IMPLANTATION IN 95 HIGHLY MYOPIC SPECIAL- NEEDS CHILDREN.' THE ARTICLE REPORTS THE FOLLOWING ADVERSE EVENTS FOLLOWING ICL IMPLANTATION IN CHILDREN WITH SPECIAL NEEDS: ADVERSE EVENTS FOR THE VISIAN ICL COHORT OF 95 CHILDREN AND 160 IMPLANTED EYES. MINOR EVENTS: STEROID-RESPONSE OHT 2% (3) EYES, CILIARY MUSCLE SPASM 1.2% (2) EYES, ENDOTHELIAL CELL LOSS 4.0% PER YEAR IN 31 EYES MEASURED. MAJOR EVENTS: PUPILLARY BLOCK 4% (7) EYES, WOUND LEAK 0.6% (1) EYES, CATARACT 0.6% (1) EYES, RETINAL DETACHMENT 0.6% (1) EYES. ADDITIONAL SURGICAL INTERVENTION (REVISION OF THE PERIPHERAL IRIDOTOMY, WOUND LEAK REPAIR, VICRYL SUTURES, LENS EXPLANT AND PEDIATRIC CATARACT SUGERY, PARS PLANA VITRECTOMY, INTRAOCULAR LASER RETINOPEXY, GAS/FLUID EXCHANGE) AND MEDICAL INTERVENTION (DORZOLAMIDE-TIMOLOL, CYCLOPENTILATE 1% DROPS, INTRAVENOUS ACETAZOLAMIDE ) WERE PERFORMED/PRESCRIBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: UNK, ESUBMITTER DOES NOT ALLOW YOU TO PUT UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148280&pc=MTA | 2023826-2021-02469 | 2021-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Endophthalmitis; Blurred Vision | N/A | 1 | Event Description: A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE SURGEON REPORTS PERMANENT LOSS OF BCVA AND ENDOPHTHALMITIS. DIAGNOSTICS WERE PERFORMED; NO ORGANISM WAS FOUND. A VITRECTOMY, INTRAVITREAL ANTIBIOTICS, AND A VITREOUS TAP WERE PERFORMED (B)(6) 2021. ON (B)(6) 2021 A VITRECTOMY WAS PERFORMED AND THE LENS WAS EXPLANTED, PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED AS THE PATIENT HAD DEVELOPED A CATARACT. PROBLEM IS RESOLVED. Manufacturer Narrative: INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8519015&pc=MTA | 2023826-2019-00607 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Endophthalmitis; Vitrectomy; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -16.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ONE DAY POST-OP THE PATIENT HAD EXAGGERATED INFLAMMATION AND EXUDATES IN THE EYE. THE PATIENT WAS DIAGNOSED WITH ACUTE ENDOPHTHALMITIS AND IMMEDIATELY WAS INJECTED WITH INTRAOCULAR ANTIBIOTICS. VITREOUS BIOPSY SAMPLE REVEALED PSEUDOMONAS INFECTION. AUTOCLAVE INDICATORS PASSED. OT AND AIR CULTURES REPORTED NEGATIVE. INSTRUMENTS AND CONSUMABLES REPORTED NEGATIVE FOR ANY MICROBIOLOGICAL GROWTH. VITREO-RETINAL SURGERY, CATARACT EXTRACTION, AND ICL EXPLANT PERFORMED. BCVA=6/18. INFECTION IS CURRENTLY INACTIVE. Manufacturer Narrative: REPORTER INDICATED THAT VITRECTOMY SURGERY WAS PERFORMED AND THE PATIENT'S STATUS IS STABLE, WITH VISION AT 6/24P. CLAIM#: (B)(4). Manufacturer Narrative: METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936885&pc=MTA | 2023826-2020-02829 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Cataract; Eye Injury; Flashers; Glaucoma; Visual Impairment | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE PATIENT REPORTS THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE LEFT (OS) EYE IN (B)(6) 2015. REPORTEDLY, THE PATIENT EXPERIENCED AN INCREASE IN ASTIGMATISM, PERIPHERAL VISION LOSS, POOR NIGHT VISION, AND FLASHES. THE PATIENT STATES THE SURGEON HAS NOW DIAGNOSED GLAUCOMA, CATARACT, AND OPTIC NERVE DAMAGE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; (B)(4); CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11448857&pc=MTA | 2023826-2021-00612 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Eye Injury; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN BMC OPHTHALMOLOGY WAS RECEIVED ON (B)(6) 2021, ENTITLED 'REASONS FOR EXPLANTATION OF PHAKIC INTRAOCULAR LENSES AND ASSOCIATED PERIOPERATIVE COMPLICATIONS: CROSS-SECTIONAL EXPLANT REGISTRY ANALYSIS.' THE ARTICLE STATES THAT 4 EYES WERE EXPLANTED DUE TO DEVELOPMENT OF A CATARACT: 1 ANTERIOR SUBCAPSULAR; 2 NUCLEAR; 1 POSTERIOR SUBCAPSULAR; AND 1 OF THE 4 EYES ALSO HAD CORNEAL ENDOTHELIAL DAMAGE. AN ADDITIONAL 1 EYE WAS EXPLANTED DUE TO OVERSIZING LEADING TO INTERMITTED IOP INCREASE AND PIGMENT DISPERSION. 2 ADDITIONAL EYES WERE EXPLANTED DUE TO PHOTOPIC PHENOMENA. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11435324&pc=MTA | MW5099834 | 2019-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-03-05 | MTA | STAAR SURGICAL ICL IMPLANT | Insufficient Information | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain |  | 1 | Event Description: I HAD ICL SURGERY ON (B)(6) 2009. IN 2011, I STARTED TO HAVE ISSUES WITH MY VISION. I HAD TO START WEARING EYE GLASSES TO DRIVE AND THEN ENDED UP WITH -1.25 VISION TO THIS DATE. IN (B)(6) 2019 I STARTED TO HAVE AN ITCHING/POKING FEELING IN MY RIGHT EYE. MY EYE CONTINUED TO WATER AND ITCH AS IF SOMETHING WAS POKING ME. I VISITED SEVERAL EYE DOCTORS, INCLUDING MY SURGEON WHO HAD DONE MY ICL SURGERY. IT FEELS/FELT AS IF SOMETHING WAS WRONG WITH MY ICL. THE PAIN IS BEHIND MY EYE. THE SURGEON STATED THAT I HAVE A SMALL CATARACT NOW ON MY EYE AND THAT IS CAUSING MY BLURRY VISION AND THE ISSUE. I AM HAVING BLURRY VISION, BUT I AM STILL WITNESSING THE WATERY/ITCHING SENSATION. THE SURGEON DR. LIU ADVISED I NEED TO GET MY ICL IMPLANT TAKEN OUT AND INPUT A CATARACT LENSES. I STATED THAT I DO NOT FEEL THIS IS RIGHT. SOMETHING IS NOT RIGHT WITH THE CURRENT ILC LENSES INSIDE MY EYE. I HAVE SINCE DEALT WITH THIS CONTINUED EFFECT FOR THE PAST TWO YEARS. EYE DOCTORS CANNOT FIGURE IT OUT. IT IS NOT MY CATARACT CAUSING THIS SENSATION. IT'S ALMOST LIKE MY ICL IS TRYING TO POKE OUT OF MY EYE OR COME OUT. I AM FINALLY GETTING A SECOND OPINION THIS NEXT WEEK WITH ANOTHER EYE SURGEON TO CHECK MY ICL. TO SEE IF THAT IS IN FACT THE ISSUE, AS MY RIGHT EYE IS STILL ITCHING, WATERY. I CAN'T PUT A CONTACT IN MY RIGHT EYE AS IT IS TOO BOTHERSOME. I WANT TO REPORT THIS AS I FEEL AFTER THIS MANY YEARS PERHAPS THE LENSES HAS GONE BAD OR IT NEEDS REPLACED AFTER A CERTAIN TIME. WHEN I DID ICL BACK IN 2009 IT WAS NEW AND JUST BECOMING A THING. PERHAPS THE ICL'S BACK THEN ARE NOT AS UP TO DATE AS WHAT THEY HAVE CREATED NOW. STAAR SURGICAL ICL IMPLANTS. IMPLANT ICL: MICL12.1 RIGHT EYE. DIOPTER. -10.0/ OPTIC DIAMETER 4.9-5.8MM OVERALL DIAMETER: 12.1M. SN: (B)(4) LEFT EYE: DIOPTER.-10.5/ OPTIC DIAMETER 4.9- 5.8MM OVERALL DIAMETER: 12.2M SN: (B)(4). FDA SAFETY REPORT ID # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620596&pc=MTA | 2023826-2021-00861 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Cataract; Foreign Body Sensation in Eye; Itching Sensation; Blurred Vision; Excessive Tear Production; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2009. REPORTEDLY, THE PATIENT DEVELOPED LENS OPACITY CORTICAL, REFRACTIVE CHANGE OVER TIME, THE PATIENT HAD TO WEAR GLASSES TO DRIVE, A WATERY/ITCHING/POKING FEELING AND PAIN IN THE EYE. BLURRY VISION ALSO REPORTED. THE SURGEON STATES THE PATIENT'S ACD CHANGED FROM 3.93 IN 2009 TO 3.39 IN 2020. THE PATIENT'S AXIAL LENGTH INCREASED. "NO ONE HAS CONTROL OVER THAT!" Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17421037&pc=MTA | 2023826-2023-03198 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Glaucoma; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO A PATIENT'S EYE. THE SURGEON INDICATED THERE MAY BE PREMATURE CATARACT, LOSS OF ECC, CHRONIC GLAUCOMA WITH A HYPEROPIC IOL PATIENT WITH A DEEP ANTERIOR CHAMBER WITH AN ICL IMPLANTED UNEVENTFULLY WITH NO PROBLEMS FOR AN INITIAL COUPLE OF YEARS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - LOSS OFF ENDOTHELIAL CELL COUNT LOSS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16138193&pc=MTA | MW5114267 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | MTA | ICL LENS VISIAN ICL | Adverse Event Without Identified Device or Use Problem | Cataract; Glaucoma; Headache; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: I DID ICL SURGERY FOR RIGHT EYE ON (B)(6) 2022 AND LEFT EYE ON (B)(6) 2022. AS I WAKE UP FROM SURGERY FOR LEFT EYE I HAD INSUPPORTABLE HEADACHES EYEACHE TURNS OUT THAT I HAD OCULAR HYPERTENSION. THE SURGEON PUT ME ON DIAMOX / MANNITOL TILL THE (B)(6) 2022 THEN HE DECIDED TO REMOVE THE ICL LENS THIS CAUSED ME CATARACT LOSS OF ENDOTHELIAL CELLS IN LEFT EYE AND GLAUCOMA.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7318875&pc=MTA | 2023826-2018-00305 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-03-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; Glaucoma; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: CASE REPORT STATED THE PATIENT HAD A MICL IMPLANTABLE COLLAMER LENS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) "8 YEARS EARLIER". THE PATIENT WAS FOUND TO HAVE A DECLINE IN VISUAL ACUITY FROM AN ANTERIOR SUBCAPSULAR CATARACT AND ELEVATED IOP. THERE WERE SIGNS OF PIGMENT DISPERSION. DIFFUSE OPTIC NERVE THINNING WAS CONSISTENT WITH ADVANCED GLAUCOMATOUS VISUAL FIELD DEFECTS. THE PATIENT WAS TREATED WITH TOPICAL GLAUCOMA MEDICATIONS AND THE PHAKIC ICL WAS EXPLANTED CONCURRENTLY WITH CATARACT SURGERY TO PREVENT FURTHER VISUAL FIELD LOSS. THE PATIENT'S IOP REMAINED CONTROLLED ON TOPICAL GLAUCOMA MEDICATIONS POST-OPERATIVELY AND HIS BEST-CORRECTED VISUAL ACUITY IN 20/20. HE CONTINUED TO NOTE SUBJECTIVELY DIMINISHED VISION AND ADVANCED VISUAL FIELD DEFECTS THAT PERSISTED AFTER CATARACT SURGERY. Manufacturer Narrative: THE PATIENT'S IOP REMAINED CONTROLLED ON TOPICAL GLAUCOMA MEDICATIONS POST-OPERATIVELY. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION:- NOT CONSIDERED OFF-LABEL DUE TO AGE OF THE PATIENT BEING UNDER 45 YEARS OLD AT THE TIME OF THE SURGERY. (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT OVER THE AGE OF 45 YEARS OLD. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6220084&pc=MTA | 2023826-2017-00009 | 2016-03-17 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Glaucoma; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -14.50/+2.5/087 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENS AND PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DESCRIPTION: PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/13 AS OF 04/21/2016. DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672769&pc=MTA | 2023826-2023-03648 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Glaucoma; Pupillary Block; Uveitis; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS RECEIVED FROM AAO, 2009, FOR "PHAKIC INTRAOCULAR LENS IMPLANTATION FOR THE CORRECTION OF MYOPIA". THE ARTICLE REPORTED THE FOLLOWING COMPLICATIONS - ENDOTHELIAL CELL LOSS, CATARACT FORMATION, SECONDARY GLAUCOMA (PUPILLARY BLOCK, PIGMENT DISPERSION) IRIS ATROPHY (PUPIL OVALIZATION), AND TRAUMATIC DISLOCATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ENDOTHELIAL CELL LOSS; PIGMENT DISPERSION; IRIS ATROPHY; LENS DISLOCATION. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: AN ARTICLE WAS RECEIVED FROM AAO 2009 FOR "PHAKIC INTRAOCULAR LENS IMPLANTATION FOR THE CORRECTION OF MYOPIA". THE ARTICLE REPORTED THE FOLLOWING COMPLICATIONS - ELEVATED IOP, LENS OPACITY (CATARACT), GLARE/HALOES, PIGMENT DISPERSION, GLAUCOMA, PUPIL BLOCK, LENS DISLOCATION (TRAUMATIC DISLOCATION), LOSS OF BCVA, TRANSIENT UVEITIS, ENDOTHELIAL CELL LOSS, IRIS ATROPHY/PUPIL OVALIZATION, UNSPECIFIED REOPERATIONS (SSI). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6: HEALTH EFFECT CLINICAL CODE: 4581 - LOSS OF BCVA. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833461&pc=MTA | 2023826-2021-04351 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Glaucoma; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A JOURNAL ARTICLE WAS RECEIVED FROM THE JOURNAL OF CATARACT AND REFRACTIVE SURGERY, FOR "UNITED STATES MILITARY IMPLANTABLE COLLAMER LENS SURGICAL OUTCOMES: AN 11 YEAR RETROSPECTIVE REVIEW". THE ARTICLE SITED A STUDY OF 3105 EYES OF MYOPIC PATIENTS, IMPLANTED WITH ICL LENSES. IN THE STUDY, THE FOLLOWING ADVERSE EVENTS OCCURRED - 14 EYES HAD ANTERIOR SUBCAPSULAR CATARACTS, REQUIRING ICL REMOVAL AND CATARACT EXTRACTION; 2 EYES HAD GLAUCOMA, REQUIRING LONG TERM TREATMENT; 3 EYES HAD RETINAL DETACHMENTS; 8 EYES HAD ACUTE ANGLE CLOSURE GLAUCOMA; 3 EYES HAD TRAUMATIC INCISION OPENING; 6 EYES HAD KERATOCONUS AND 135 EYES HAD ICL EXPLANTATION. THE PERCENTAGE OF ECL FROM BASELINE IS 7.1% AT 3 YEARS POSTOPERATIVELY REACHES STATISTICAL SIGNIFICANCE AT 5 YEARS (21.8% DECLINE, P<0.0001). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK SEX: UNK WEIGHT: UNK ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK, EXPIRATION DATE UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK, DEVICE MANUFACTURE DATE: UNK, (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11651554&pc=MTA | 2023826-2021-00885 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. GLARE/HALOES WERE OBSERVED. THE LENS REMAINS IMPLANTED. SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -10.50 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2021 DUE TO LOW VAULT, LENS OPACITY (ASC), BLURRED VISION AND GLARE/HALOES. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.SEE MFR # 2021-00884 FOR CLAIM AGAINST THE INITIAL LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DERIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. AFTER LENS EXPLANT, PHACO-EMULSIFICATION (CATARACT SURGERY) & IOL IMPLANTATION WAS PERFORMED AND THE PROBLEM WAS RESOLVED. HEALTH IMPACT CODE 4625. ADDITIONAL SURGERY (PHACOEMULSIFICATION & IOL IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4937371&pc=MTA | 2023826-2015-00893 | 2015-03-23 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING AND LENS OPACITY, WHICH WAS FIRST NOTED ON (B)(6) 2015. THE REPORTER INDICATED IT WAS AN ANTERIOR SUBCAPSULAR OPACITY. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION, GLARE/HALOS AND BLURRED VISION. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF REMAINING VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863567&pc=MTA | 2023826-2017-01432 |  | Injury | STAAR SURGICAL COMPANY | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Halo; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED HAD A VISIAN ICL IMPLANTED SEVERAL YEARS AGO (2009). FOR THE PAST FEW YEARS HE HAS HAD INCREASED HALOS AND STARBURSTS IN THE RIGHT EYE (OD). HIS OPTOMETRIST PRESCRIBED A LOW LEVEL REFRACTION. THE ISSUE IS GETTING WORSE AND HE HAS BEEN EXPERIENCING HAZINESS IN VISION AND FLARING OF OBJECTS IN DIM LIGHTING AND BRIGHT DAYLIGHT. THE PATIENT HAS THE BEGINNING OF A CATARACT IN ANTERIOR BORDER OF CRYSTALLINE LENS, WHICH IS MILD. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814249&pc=MTA | 2023826-2020-00475 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Cataract; Headache; Blurred Vision; Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, BLURRED VISION, LOSS OF BEST CORRECTED VISUAL ACUITY AND LENS OPACITY (POSTERIOR SUBCAPSULAR) WAS NOTED. THE PATIENT IS SENSITIVE TO BRIGHT LIGHT AND HEADACHE WHEN READING. THE LENS REMAINED IMPLANTED AND PATIENT WILL RETURN IN 3 MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED WITH AN INTRAOCULAR LENS ON (B)(6) 2020. CATARACT SURGERY WAS ALSO PERFORMED ON (B)(6) 2020, DURING LENS REPLACEMENT. THE REPORTER INDICATED THAT THE PATIENT WAS SEEN ON (B)(6) 2020, AND THAT SHE IS "STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." H6 - HEALTH EFFECT IMPACT CODE: 4625 - CATARACT SURGERY, CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090066&pc=MTA | 2023826-2018-01697 | 2018-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Hemorrhage/Bleeding; Hyphema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -9.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASC). REPORTEDLY, DURING THE INITIAL PROCEDURE, HEMORRHAGE OF AN IRIS VESSEL OCCURED. TO MANAGE THE IRIS BLEED THE SURGEON INTRODUCED BUBBLES INTO THE EYE. REPORTEDLY, THE BUBLLES WERE STILL PRESENT 3 DAYS LATER AND THE ICL WAS IN CONTACT WITH THE ANTERIOR CAPSULE. THE SURGEON STATES THE CAUSE OF THE ICL TOUCHING THE ANTERIOR CAPSULE WAS DUE TO THE BUBBLES, NOT THE ICL DEVICE. THE CAUSE OF THE IRIS HEMORRHAGE AND LENS OPACITY (ASC) IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED WITH THE LENS EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATED BY MFR - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CORRECTED DATA: EVENT - "HEMORRHAGE OF AN IRIS VESSEL OCCURRED" SHOULD BE CORRECTED TO "HEMORRHAGE OF AN IRIS VESSEL OCCURED LEADING TO HYPHEMA" CLAIM #(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6305835&pc=MTA | 2023826-2017-00223 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND WAS PRESCRIBED ORAL MEDICATIONS OR EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO INSIDE OF EYE WAS INFLAMED. INFORMATION RECEIVED FROM THE FACILITY CONFIRMED THE PATIENT REPORT AND THE DATE OF THE CATARACT DIAGNOSIS WAS (B)(6) 2016. THEY INDICATED PATIENT HAS EPISCLERITIS EXTERNALLY WHICH REQUIRED PRED FORTE TO CLEAR. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107373&pc=MTA | 2023826-2019-01752 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Inflammation; Pain; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. A COUPLE OF HOURS AFTER THE INITIAL SURGERY, THE PATIENT EXPERIENCED PAIN AND A NEGATIVE EFFECT ON HER VISION. A CATARACT HAD DEVELOPED AND THE LENS WAS EXPLANTED. THE CATARACT WAS REMOVED AND A MONOFOCAL LENS WAS IMPLANTED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 1932: INFLAMMATION. DEVICE CODE 1494:OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6828273&pc=MTA | 2023826-2017-01379 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Decentration; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Appropriate Term/Code Not Available | Cataract; Inflammation; Swelling; Blurred Vision; No Code Available | NA | 1 | Event Description: ON 30-JUL-2017 STAAR WAS MADE AWARE OF A JOURNAL ARTICLE REGARDING A CASE OF HYPERMATURE CATARACT FORMATION FOLLOWING IMPLANTATION OF AN IMPLANTABLE COLLAMER LENS WITH AN AQUAPORT. THE ARTICLE STATES THE PATIENT WAS A (B)(6) MAN AND A 12.6 MM LENS WAS INSERTED INTO HIS LEFT EYE (OS) AT ANOTHER (UNKNOWN) CLINIC 4 MONTHS PREVIOUSLY. THE PATIENT COMPLAINED OF BLURRY VISION. MEDICAL RECORDS SHOWED NO INTRAOPERATIVE SURGICAL COMPLICATIONS. UPON EXAMINATION, THE FOLLOWING WAS NOTED: ANTERIOR SUBCAPSULAR OPACITY, SWELLING OF THE LENS, VAULTING OF THE LENS WAS 1/4 OF THE CORNEAL THICKNESS, THE ICL WAS DECENTERED INFERONASALLY AND IN A VERTICAL POSITION. IT IS STATED THAT, AFTER ONE MONTH, THE SUBCAPSULAR OPACITY PROGRESSED RAPIDLY TO A HYPERMATURE CATARACT AND THE ICL WAS ROTATED TO 60 DEGREES FROM THE HORIZONAL MERIDIAN. THERE WAS A FLARE OF 1+ IN THE ANTERIOR CHAMBER. CATARACT SURGERY WAS PERFORMED AND THE ICL WAS EXPLANTED. UPON EXPLANTATION, THE LENS SHOWED WHITISH INFILTRATES ON ITS SURFACE. THE REASON FOR THE RAPID CATARACT PROGRESSION IS NOT CLEAR. UPON CONTACT WITH 2 OF THE 3 SURGEONS WHO WROTE THE ARTICLE, THE FOLLOWING INFORMATION WAS OBTAINED: THIS CASE OCCURRED 3-4 YEARS PRIOR TO ARTICLE PUBLICATION. THE SURGEON HAD A "STRONG SUSPICION" THAT THERE WAS A SURGICAL COMPLICATION ON THIS CASE AND THAT THE LENS WAS TOUCHED DURING THE OPERATION HOWEVER, THIS COULD NOT BE PROVEN NOR CONFIRMED AS THE ORIGINAL OPERATION SURGEON DENIED ANY SURGICAL COMPLICATION. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10714478&pc=MTA | 2023826-2020-02494 | 2020-01-15 05:00 | Injury | STAAR SUGICAL COMPANY | 2020-10-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; Insufficient Information | NA | 1 | Event Description: IT WAS REPORTED THAT A LENS IS BEING RETURNED. THE REASON FOR RETURN IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS OPACITY (PSC, OBSERVED: (B)(6) 2020) AND CATARACT SURGERY WAS PERFORMED WITHIN THE SAME SURGERY AS REMOVAL. THE REPORTER STATED "PATIENT WAS SEEN ON (B)(6) 2020, STABLE FOLLOWING SURGERY, HAS BEEN INFORMED ABOUT THE MILD POSTERIOR CAPSULE OPACIFICATION AND THE NEED FOR A YAG IF THIS WORSENS. SHE DEVELOPED A CATARACT FROM LENS IMPLANTATION. THE CATARACT HAS BEEN REMOVED AND AN INTRAOCULAR LENS IMPLANTED. PATIENT WAS SEEN ON (B)(6) 2020, STABLE FOLLOWING SURGERY, VISION IS BETTER, PRESENTLY THIS LOOKS CLEARER THAN RIGHT AFTER HER PROCEDURE." FOR CAUSE OF EVENT THE REPORTER STATED "ICL RELATED - INITIAL LENS 13.2 WAS OVERSIZED AND THE EDGES MAY HAVE INCISED THE CAPSULE. THE 12.6 LENS WAS USED TO REPLACE THIS AT THE SAME SITTING BUT SHE DEVELOPED A CATARACT WITHIN A FEW DAYS. THE 12.6 WAS ALSO TOO SHALLOW. IDEALLY SHE WOULD HAVE HAD A 12.9." H6- WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED FROM THE REPORTER THAT PREVIOUS INFORMATION PROVIDED WAS INTENDED FOR CLAIM# (B)(4), MEDWATCH REPORT# 2023826-2020-00475, THERE IS NO COMPLAINT FOR LENS SN# (B)(6). REPORTED INDICATED "LEFT EYE HAS ICL IN AND VISION WAS GOOD STRAIGHT AWAY POST OP AND IS STILL NOW, NO ISSUES". CLAIM#: (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969934&pc=MTA | 2023826-2020-02885 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | P030016 | 1 | Event Description: RECEIVED A COMPLAINT FROM (B)(6), THE PATIENT REPORTED CATARACT AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16448570&pc=MTA | 2023826-2023-00733 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "VAULT-CORRELATED EFFICACY AND SAFETY OF IMPLANTABLE COLLAMER LENS V4C IMPLANTATION FOR MYOPIA IN PATIENTS WITH SHALLOW ANTERIOR CHAMBER DEPH"; INDICATED THAT A STUDY WAS CONDUCTED OF 163 EYES OF 94 PATIENTS FOR AT LEAST 24 MONTHS TO EVALUATE EFFICACY AND SAFETY OUTCOMES AFTER IMPLANTATION OF THE VISIAN IMPLANTABLE COLLAMER LENS (V4C) IN MYOPIA PATIENTS WITH SHALLOW ANTERIOR CHAMBER DEPTH. THE ARTICLE INDICATES THAT HIGH IOP OF 9 EYES AND ANTERIOR CAPSULAR OPACITY OF 5 EYES WERE OBSERVED. THE 9 EYES WITH ELEVATED INTRAOCULAR PRESSURE WERE LESS THEN 30 MMHG; ALL OF THEM COULD BE CONTROLLED BY IOP LOWERING EYEDROPS. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15694587&pc=MTA | 2023826-2022-03875 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULTING; ANTERIOR SUBCAPSULAR CATARACT; IOP OF 38 AND ADDITIONALLY NOTED IOP HAS "BEEN ALL OVER THE BOARD THE PAST FEW CHECKUPS". IT WAS ALSO INDICATED THAT IT COULD BE A POSSIBLE REACTION TO THE INTRACAMERAL SOLUTION OF NOREPINEPHRINE AND LIDOCAINE. MULTIPLE MEDICATION WAS USED. THERE IS A PLAN TO REMOVE THE CATARACT. PATIENT IS CURRENTLY DOING WELL WITH 20/20 IN THE DISTANCE IN HIS RIGHT DOMINATE EYE AND -1.00 WITH 20/40 IN THE LEFT EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15695845&pc=MTA | 2023826-2022-03876 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LOW VAULTING; ANTERIOR SUBCAPSULAR CATARACT; IOP OF 38 AND ADDITIONALLY NOTED IOP HAS "BEEN ALL OVER THE BOARD THE PAST FEW CHECKUPS". IT WAS ALSO INDICATED THAT IT COULD BE A POSSIBLE REACTION TO THE INTRACAMERAL SOLUTION OF NOREPINEPHRINE AND LIDOCAINE. MULTIPLE MEDICATION WAS USED. THERE IS A PLAN TO REMOVE THE CATARACT. PATIENT IS CURRENTLY DOING WELL WITH 20/20 IN THE DISTANCE IN HIS RIGHT DOMINATE EYE AND -1.00 WITH 20/40 IN THE LEFT EYE. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11899203&pc=MTA | 2023826-2021-01685 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; Intraocular Pressure Increased | NA | 1 | Event Description: ON (B)(6) 2021 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN SAGE, EUROPEAN JOURNAL OF OPHTHALMOLOGY, ENTITLED, 'POSTERIOR CHAMBER COLLAMER PHAKIC INTRAOCULAR LENS IMPLANTATION: COMPARISON OF EFFICACY AND SAFETY FOR LOW AND MODERATE-TO-HIGH MYOPIA'. THE STUDY ARTICLE REPORTS THE FOLLOWING ADVERSE EVENTS IN 8 EYES THAT UNDERWENT ICL IMPLANTATION: 3 CASES OF EARLY POST-OP IOP (MANAGED WITH TOPICAL HYPOTENSIVE MEDICATION), 4 ICL EXCHANGE (3 DUE TO LOW VAULT ASSOCIATED WITH ROTATION), 1 CASE OF ASYMPTOMATIC ANTERIOR SUBCAPSULAR CATARACT (WITHOUT CDVA IMPAIRMENT). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516827&pc=MTA | 2023826-2021-03729 | 2018-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND CATARACT. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: 1494: OFF-LABEL USE: ACD<3.0MM CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18342236&pc=MTA | 2023826-2023-05535 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased | P030016 | 1 | Event Description: A COPY OF AN ARTICLE FROM INT OPHTHALMOL, (B)(6) 2023 WAS RECEIVED REGARDING "FIVE-YEAR FOLLOW-UP OF A POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS WITH A CENTRAL HOLE FOR CORRECTION OF MYOPIA". THE ARTICLE INVOLVED 241 EYES THAT UNDERWENT ICL IMPLANTATION. POST-OP, THERE WERE ISOLATED CASES OF ELEVATED IOP, WHICH WITH TOPICAL TREATMENT, NONE PERSISTED LONGER THAN 1 WEEK. IT WAS REPORTED TWO EYES OF ONE PATIENT AT 21 MONTHS FROM INITIAL SURGERY, SHOWED CATARACT FORMATION. THE ICLS WERE REMOVED AND IOLS WERE IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B:UNK. H4: UNK. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - IOL IMPLANTED. H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL TREATMENT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17755064&pc=MTA | 2023826-2023-03982 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A COPY OF A CLINICAL CASE REPORT WAS FOUND, REGARDING "ICL EXPLANT DUE TO OCULAR HYPERTENSION, REDUCED ENDOTHELIAL CELL COUNT AND ANTERIOR SUB CAPSULAR CATARACT DUE TO PATIENT SELECTION ERROR". THE ARTICLE INVOLVED A 24-YEAR-OLD FEMALE WITH HIGH MYOPIA, FOLLOWING POST ICL PLACEMENT IN LEFT EYE (OS). AFTER POST-OP PERIOD, THE PATIENT PRESENTED WITH OCULAR HYPERTENSION THAT WAS CONTROLLED WITH ANTIHYPERTENSIVES, SHE ACHIEVED IOP CONTROL A MONTH AFTER THE OPERATION, BUT AFTER A FEW MONTHS POST-OP, SHE PRESENTED WITH A LOW ENDOTHELIAL CELL COUNT AND AN ANTERIOR SUBCAPSULAR CATARACT, FOR WHICH IT WAS DECIDED TO EXPLANT THE ICL . FINALLY, CATARACT SURGERY WAS PERFORMED AND A MONOFOCAL IOL WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - REDUCED ENDOTHELIAL CELL COUNT. H6: HEALTH EFFECT IMPACT CODE: 4644 - ANTIHYPERTENSIVES. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - CATARACT SURGERY, IOL IMPLANTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840575&pc=MTA | 2023826-2021-04392 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: AN ARTICLE PUBLISHED IN THE JOURNAL OF CATARACT AND REFRACTIVE SURGERY WAS RECEIVED ON 29-OCT-2021, ENTITLED 'LONG TERM EFFICACY AND SAFETY PROFILES FOLLOWING POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION IN EYES WITH >/= 10-YEAR FOLLOW-UP.' THE ARTICLE DESCRIBES 2 EYES OF 1 PATIENT REQUIRED ICL EXCHANGE DUE TO MYOPIC PROGRESSION, 8.7 YEARS AFTER THE FIRST SURGERY. THREE EYES OF TWO PATIENTS UNDERWENT CATARACT SURGERY AFTER ICL IMPLANTATION (MEAN 10.1 YEARS). ALL CATARACTS WERE CLASSIFIED AS NUCLEAR CATARACTS ACCORDING TO THE LENS OPACITIES CLASSIFICATION SYSTEM III, WHICH SUGGEST THAT THEY WERE AGE RELATED RATHER THAN RELATED TO THE ICL LENS IMPLANT. TWO EYES LOST 2 OR MORE LINES OF CDVA, ONE REVERSIBLY DUE TO DEVELOPMENT OF CATARACT (FINAL CDVA AFTER ICL EXPLANTATION PLUS CATARACT SURGERY WAS 20/25), AND ANOTHER DUE TO DEVELOPMENT OF A MYOPIC MACULAR HOLE. 2 EYES HAD A MILD TRANSIENT INCREASE IN IOP BUT NO TREATMENT WAS REQUIRED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: A1-(B)(6). A2- (B)(6) 1986; (B)(6) 1968; (B)(6) 1965. B3-(B)(6) 2017; (B)(6) 2017; (B)(6) 2017; (B)(6) 2018. B5-ADDITIONAL INFORMATION: REGARDING THE PATIENT REQUIRING ICL EXCHANGE DUE TO MYOPIC PROGRESSION: A LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE (B)(6) 2008 AND RIGHT (OD) EYE (B)(6) 2008. ON (B)(6) 2017 AND (B)(6) 2017 RESPECTIVELY THE LENS WAS REMOVED AND REPLACED WITH AN ALTERNATE LENS AND PROBLEM RESOLVED. THE CAUSE OF BOTH EVENTS IS A PATIENT-RELATED FACTOR: "PROGRESSIVE REFRACTIVE DEFECIT." REGARDING TWO EYES OF ONE PATIENT WHO UNDERWENT CATARACT SURGERY: A LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2008 AND RIGHT (OD) EYE ON (B)(6) 2008. ON (B)(6) 2018 AND (B)(6) 2014 RESPECTIVELY THE LENSES WERE EXPLANTED AND CATARACT SURGERY WAS PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. PROBLEM RESOLVED. REGARDING ONE EYE OF ONE PATIENT WHO UNDERWENT CATARACT SURGERY: A LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2008. ON (B)(6) 2018 THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. BOTH PSC AND SN CATARACTS WERE REPORTED. THE CAUSE IS REPORTED AS OTHER. THE PROBLEM IS RESOLVED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11098622&pc=MTA | 2023826-2020-03254 | 2020-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND PUNTIFORM OPACITY IN THE POSTERIOR CAPSULE. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR: NON RECOMMENDED OVD AND EXCESSIVE USE OF OVD DUE TO A NON TOTAL DILATION DURING SURGERY. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. REPORTEDLY, THE IOP WAS MANAGED AFTER THE EVACUATION OF OVD. PUNTIFORM OPACITY HAS DECREASED AFTER 48 HOURS OF SURGERY. Manufacturer Narrative: B5-UPDATED INFORMATION: REPORTEDLY, THE PROBLEM IS NOW RESOLVED. (B)(4). Manufacturer Narrative: B5. ELEVATED IOP IS AN EXPECTED OR FORESEEABLE SIDE EFFECT. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301799&pc=MTA | 2023826-2023-00386 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Improper or Incorrect Procedure or Method | Cataract; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER (PATIENT CURRENT OBSERVING DOCTOR) STATES THAT A SURGEON AT A DIFFERENT CLINIC IMPLANTED A IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE AND POWER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE CURRENT OBSERVING DOCTOR SAW PATIENT FOR THE FIRST TIME ON (B)(6) 2022, WHERE HE UNFORTUNATELY BECAME AWARE OF SEVERAL MEDICAL ERRORS. PER EMAIL THE FINDINGS IN THE EYE AFTER THE GLAUCOMA ATTACK WERE AS FOLLOWS-"TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT."TEPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT. OCULAR BACKGROUND FINDINGS NORMAL." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE IMPLANTING SURGEON HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2023-00326 DOR FELLOW EYE (OD). Manufacturer Narrative: CLINICAL CODE 4581: TEMPORAL CORNEAL INCISION, ANTERIOR CHAMBER SHALLOWER, IRIS TEMPORALLY DEPIGMENTED AND "RAGGED", PIGMENT DISPERSION IN THE ANTERIOR CHAMBER, WIDE AND PLEGIC PUPIL, ICL SPECKLED WITH PIGMENT, ANTERIOR SUBCAPSULAR LENS OPACITY- CATARACT CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5530607&pc=MTA | 2023826-2016-00341 | 2016-02-23 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE SURGEON IMPLANTED A 12.5MM ICH125V3 IMPLANTABLE COLLAMER, +06.50 DIOPTER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2008. OCULAR HYPERTENSION, BLURRED VISION WAS NOTED AND A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND A IOL WAS IMPLANTED ON (B)(6) 2016. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/40. IOL IS IN GOOD POSITION. HEALING NICELY AND CONTINUE WITH MEDICATION AS PRESCRIBED. SEE MFR. #2023826-2016-00342 FOR LEFT EYE. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT AND WAS FOUND TO BE IN SPECIFICATION. CONCLUSION CODE(S) 67: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DIOPTER: +06.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5530608&pc=MTA | 2023826-2016-00342 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE SURGEON IMPLANTED A 12.5MM ICH125V3 IMPLANTABLE COLLAMER, +03.50 DIOPTER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2008. OCULAR HYPERTENSION, BLURRED VISION WAS NOTED AND A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND A IOL WAS IMPLANTED ON (B)(6) 2016. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/30+2. IOL IS IN GOOD POSITION. HEALING NICELY AND CONTINUE WITH MEDICATION AS PRESCRIBED. SEE MFR. #2023826-2016-00341 FOR RIGHT EYE. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT AND WAS FOUND TO BE IN SPECIFICATION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: DIOPTER: +03.50; (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669737&pc=MTA | 2023826-2020-02329 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE. THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - LOW VAULT, TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10669743&pc=MTA | 2023826-2020-02328 | 2020-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Cataract; Intraocular Pressure Increased; Iritis; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO INTRAOCULAR PRESSURE SPIKE (NOT CONTROLLABLE WITH DROPS) AND LENS WAS TOO LARGE, WHICH WAS CAUSING THE LENS TO BOW TOWARDS THE IRIS, INCREASING THE PRESSURE . THERE IS NO LENS IN THE EYE CURRENTLY, SURGEON WILL PROCEED WITH CATARACT EXTRACTION AS ANTERIOR CHANGES TO THE NATURAL LENS OCCURRED AFTER EXPLANTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A CONTACT CASE. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: B5 - "LENS WAS TOO LARGE" AMMENDED TO "LENS WAS TOO SMALL" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE 4581: NO CODE AVAILABLE (BOWING OF IRIS) SHOULD BE INCLUDED IN INITIAL MDR. H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (PHACOEMULSIFICATION, IOL IMPLANTATION) SHOULD BE INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: B5 - TRANSIENT LOSS OF BCVA, GLARE/HALOES, IRITIS, ICL-LENS CONTACT, AND ICL-IRIS CONTACT WERE ALSO REPORTED. THE PROBLEM IS REPORTED AS RESOLVED AFTER PHACOEMULSIFICATION AND IOL IMPLANTATION. HOWEVER, MILD PUPIL CORECTOPIA IS REPORTED POST EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE SIZE MAY HAVE BEEN TOO SMALL." H6 - HEALTH EFFECT IMPACT CODE 4581: NO CODE AVAILABLE (ICL-LENS TOUCH, ICL-IRIS CONTACT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5646258&pc=MTA | MW5062182 | 2016-03-15 | Injury | Unknown Manufacturer | 2016-05-06 | MTA | ICL LENS | Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; Nausea; Pain; Vomiting |  | 1 | Event Description: I WAS ADMITTED TO HAVE AN ICL SURGERY IN THE ASSIGNED DATE. DURING THE SURGERY IN WHICH I HAD LOCAL ANESTHESIA, I FELT PRESSURE IN THE RIGHT EYE, YET THE DOCTOR IGNORED MY COMPLAINT. THE LEFT EYE WENT OUT PERFECTLY. AFTER THE SURGERY WAS DONE, I WAS ADMITTED HOME AND NOT HOSPITALIZED FOR AT LEAST 4 HOURS UNDER THE DOCTOR'S PERMISSION. AGAIN I FELT NAUSEOUS AND VOMITED TWICE ON THE WAY HOME. MY PARENTS CALLED THE DOCTOR TO TELL HIM THAT I WAS SUFFERING PAIN IN MY RIGHT EYE, BUT HE SAID IT WAS NATURAL. I STAYED HOME IN BED 3 DAYS WITHOUT REVISION, DURING THIS PERIOD, I HAD A VERY HIGH IOP. I WAS BARELY ABLE TO MOVE. ATER 3 DAYS, I WENT TO THE DOCTOR FOR REVISION UPON HIS REQUEST. UNFORTUNATELY, IT WAS TOO LATE CAUSE I ALREADY HAD TRAUMATIC MYDRIASIS. AND NOW I AM DEVELOPING CATARACT AND PAINFUL RIGHT EYE. I AM NOW BEING FOLLOWED WITH ANOTHER OPHTHALMOLOGIST. HE SAID THAT I NEED CATARACT SURGERY AND A CLEAR LENS EXTRACTION.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8853045&pc=MTA | 2023826-2019-01326 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - DECENTRATION OF THE ICL WITH OCCURRENCE OF SUBCAPSULAR OPACITY IN THE PERIPHERY OF THE LENS IN 1 EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO: 'AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED ONE EYE (OD) WITH AN IOP GREATER THAN 50MMHG, LENS OPACITY, AND ADDITIONAL ADVERSE EVENTS WAS OBSERVED SOON AFTER SURGERY. IN THE PATIENT'S RIGHT EYE (OD), THE 'ELEVATION OF IOP WAS CONDITIONED BY THE DECENTRATION OF THE ICL UPWARD, WITH SHALLOWING OF THE INTERIOR CHAMBER IN ITS LOWER HALF'. THE PATIENT WAS TREATED MEDICALLY AND WITH SECONDARY LASER IRIDOTOMY AND REPOSITIONING OF THE LENS. REPORTEDLY, 'SPONTANEOUS DECENTRATION OF THE ICL IN [OD] REOCCURRED WITH ONSET OF SUBCAPSULAR OPACITY OF THE LENS IN THE PERIPHERY'. THE ICL WAS EXPLANTED AND THE PATIENT WENT ON TO HAVE LASER REFRACTIVE SURGERY. AT LAST FOLLOW UP SEVEN YEARS AFTER SURGERY THE IOP REMAINED STABLE AT 13MMHG. THE AUTHOR STATED THAT THE 'POSTOPERATIVE COURSE WAS CONDITIONED BY THE IMPLANTATION OF THE INCORRECT SIZE OF ICL' AND THAT 'SOON AFTER THE PROCEDURE, EXCESSIVE VENTRAL BUCKLING OCCURRED DUE TO THE INFLUENCE OF THE INAPPROPRIATE LENGTH OF THE ICL, WITH THE DEVELOPMENT OF ACUTE PUPILLARY BLOCK'. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.' CORRECTED TO PHAKIC INTRAOCULAR LENS & PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: MTA & QCB (COHORT INCLUDED SEVERAL LENS MODELS). CORRECTED TO ICM V4, VICMO, TICM V4, VTICM, VTICMO (COHORT INCLUDED SEVERAL LENS MODELS). SHOULD HAVE INCLUDED LITERATURE IN INITIAL MDR. SHOULD HAVE INCLUDED PATIENT CODE 3191 (EXCESSIVE VENTRAL BUCKLING, ACUTE PUPILLARY BLOCK; SECONDARY INTERVENTION, TREATED WITH MEDICATION, SECONDARY LASER IRIDOTOMY, REPOSITIONING, EXPLANT, LASER REFRACTIVE SURGERY) IN INTIAL MDR. METHOD CODE 4117 IS NOT APPLICABLE IN INITIAL MDR. THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7155887&pc=MTA | 2023826-2017-01982 | 2017-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4, IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, ON (B)(6) 2015 IN THE PATIENT'S RIGHT EYE (OD). LENS OPACITY (ASCO) WAS OBSERVED ON (B)(6) 2017,ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE CHANGE OVERTIME DUE TO CATARACT WAS ALSO NOTICED. THE LENS WAS EXPLANTED ON (B)(6) 2017, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS REPORTED. Manufacturer Narrative: FINAL PATIENT STATUS IS REPORTED AS "GOOD RESULT, PATIENT IS OK" WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: REPORTER STATED THAT "THE LENS IS NOT AVAILABLE (DISCARDED)". DATE OF THIS REPORT: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: IT IS CORRECTED FROM "30-NOV-2017" TO "29-NOV-2017" IN THE INITIAL MDR. DEVICE MANUFACTURE DATE: IT IS CORRECTED FROM "19-JUN-2017" TO "19-JUN-2014". (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772247&pc=MTA | 2023826-2019-01176 | 2019-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Cataract; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "DUE TO FLOPPY IRIS AND HIGH PRESSURE RESULTING IN ANGLE CLOSURE", THE LENS WAS REMOVED ON (B)(6) 2019. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 1426, SHOULD BE REMOVED FROM SUPPLEMENTAL #2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD HAVE BEEN ADDED IN SUPPLEMENTAL MEDWATCH REPORT. WORK ORDER SEARCH: "NO SIMILAR COMPLAINTS" SHOULD BE CORRECTED TO "THREE SIMILAR COMPLAINTS." RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION :THE REPORTER INDICATED THAT EXCESSIVE VAULT (INITIALLY 1.5, THEN CHANGED TO 5) WAS OBSERVED. PUPIL BLOCK WITH ELEVATED IOP (55 MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (55 MMHG) WERE OBSERVED ON (B)(6) 2019, LENS OPACITY (ASC) OBSERVED ON (B)(6) 2019. REPORTEDLY ENLARGEMENT OF PI AND ADDITION OF NEW PI WERE SURGICALLY PERFORMED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND "TOPOMAX RELATED SHIFT? SURGICALLY ENLARGED PI FAILED TO RESOLVE EDEMATUOUS CB WORSENED BY TOPOMAX?" PER REPORTER ON (B)(6) 2019, "THERE IS CATARACT FORMATION AND PATIENT IS SCHEDULED FOR CE WITH IOL IMPLANTATION." (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052146&pc=MTA | 2023826-2023-02080 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A COPY OF A REVIEW ARTICLE OF THE PUBLISHED LITERATURE IN PUBMED WAS RECEIVED. THE ARTICLE CONSIDERED THE OUTCOMES PUBLISHED BY OTHER AUTHORS IN THE LITERATURE. THE ARTICLE WAS "IMPLANTABLE COLLAMER POSTERIOR CHAMBER INTRAOCULAR LENSES: A REVIEW OF POTENTIAL COMPLICATIONS". THE ARTICLE INDICATED CATARACT WAS THE MAJOR POSTOPERATIVE COMPLICATION REPORTED (136 IN 2592 EYES). 21 CATARACTS WERE REPORTED AS SURGICALLY INDUCED, 46 EYES HAD POOR VAULT AND CATARACT SURGERY WAS CARRIED OUT IN 38 OF EYES. EARLY ACUTE INTRAOCULAR PRESSURE INCREASE WAS ALSO REPORTED TO BE RELATIVELY FREQUENT, WHEREAS ACUTE PUPILLARY BLOCK WAS LESS FREQUENT AND MOSTLY RESOLVED WITH ADDITIONAL IRIDOTOMIES. REPORTED ENDOTHELIAL CELL LOSS VARIED FROM 9.9% AT 2 YEARS TO 3.7% 4 YEARS POSTOPERATIVELY. Manufacturer Narrative: HEALTH IMPACT- CLINICAL CODE: 4581 - ENDOTHELIAL CELL LOSS. HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ADDITIONAL PIS, CATARACT SURGERY. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6859975&pc=MTA | 2023826-2017-01422 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Intraocular Pressure Increased; Pupillary Block; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LENS OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. THE LENS WAS EXCHANGED. A POSSIBLE IRIDOPLASTY IS PLANNED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. THE REPORTER ALSO STATED THAT THE SOFTWARE DOES NOT ALWAYS PREDICT CORRECT SIZE IMPLANT. (B)(4). CORRECTION DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, (PUPILLARY BLOCK) ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD, ANTERIOR LEN OPACITY, WATER VACUOLES, AND FIXED DILATED PUPIL. ON (B)(6) 2017, AN IRIDOTOMY ENLARGEMENT WAS DONE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND CLEAR RESIDUE ON LENS AND A PIECE OF THE HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007063&pc=MTA | 2023826-2023-02047 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Cataract; Intraocular Pressure Increased; Retinal Detachment; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE TITLED "CAUSES OF PHAKIC IMPLANTABLE COLLAMER LENS EXPLANTATION/EXCHANGE AT (B)(6) HOSPITAL", INDICATED THAT OR 787 ICL IMPLANTS, 30 IMPLANTS WERE EXPLANTED. THE CAUSES OF THE EXPLANT INCLUDED INCORRECT LENS SIZE, CATARACT, HIGH RESIDUAL ASTIGMATISM, RHEGMATOGENOUS RETINAL DETACHMENT, GLARE, CORRECTIVE MEASURES MAINLY INCLUDED EXCHANGE WITH AN APPROPRIATELY SIZED LENS, PHACOEMULSIFICATION AND POSTERIOR CHAMBER INTRAOCULAR LENS IMPLANTATION, OR REPLACEMENT OF ICL WITH CORRECT POWER. INCORRECT ICL SIZE WAS THE MOST COMMON CAUSE OF EXPLANTATION. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "IT ADDITIONALLY NOTES THERE WERE CASES THAT INCLUDED ELEVATED INTRAOCULAR PRESSURE, LENS ROTATION, SHALLOW AC, PAS IN THIS STUDY." -SHOULD HAVE BEEN INCLUDED IN THE INITIAL MEDWATCH REPORT. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SHALLOW AC, PAS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11758278&pc=MTA | 2023826-2021-01312 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED FROM (B)(6), DATED (B)(6) 2021 THAT CITED A CLINICAL STUDY OF "STUDY OF COMPLICATIONS OF PHAKIC INTRAOCULAR LENSES IN CORRECTION OF MYOPIA". WITHIN THE STUDY OF 52 EYES USING THE IMPLANTABLE COLLAMER LENS (ICL), IT WAS FOUND ONE EYE SUFFERED A CATARACT, WHICH WAS VISUALLY INSIGNIFICANT, WITHOUT THE NEED OF CATARACT SURGERY. THREE EYES SAW AN INCREASE OF INTRAOCULAR PRESSURE (IOP) , WHICH WAS CONTROLLED WITH TOPICAL MEDICATIONS. SIX EYES OF ICL PATIENTS COMPLAINED OF GLARE AND HALO. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE : UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233814&pc=MTA | 2023826-2019-00072 | 2018-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE INITIAL LENS WAS NOT EXCHANGED FOR ANOTHER LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8233815&pc=MTA | 2023826-2019-00071 | 2018-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Intraocular Pressure Increased; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PROBLEMS WITH PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE, LENS DISLOCATION/SUBLUXATION, GLARE/HALOS AND ANTERIOR SUBCAPSULAR CHANGES (LENS OPACITY). THE REPORTER INDICATED THE LENS HAPTICS WERE POSTERIOR TO SULCUS AND OPTIC WAS AGAINST THE NATURAL LENS. A SECONDARY YAG WAS PERFORMED FOR ENLARGEMENT OF THE PERIPHERAL IRIDECTOMIES. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR (THE PATIENTS ANATOMY SEEMED TO BE THE CAUSE OF THE ICL FAILURE). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. SEE MFR REPORT # 2023826-2019-00073 FOR INITIAL LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926670&pc=MTA | 2023826-2016-00374 | 2015-12-17 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED AT FOLLOW-UP VISITS, THE LENS HAPTIC WAS FOLDED AND THE PUPIL WAS OVALISED. THE LENS WAS TILTED ON ONE SIDE AND WAS RE-POSITIONED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: FOLLOW-UP #4 FOR MDR # 2023826-2017-00374 WAS SUBMITTED IN ERROR. THE REPORTED INFORMATION WAS MEANT TO BE FOLLOW-UP #4 FOR MDR # 2023826-2016-00374. AS A RESULT, ONLY THE INITIAL MDR # 2023826-2017-00374 PERTAINS TO CLAIM# 710106. FOLLOW-UP #4 FOR MDR # 2023826-2016-00374 WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT HAPTIC WAS BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6024717&pc=MTA | 2023826-2016-01448 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Intraocular Pressure, Delayed, Uncontrolled; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO PRESSURE AND PAIN FOR PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRY VISION. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. POST OP THE PATIENT HAD TRACE CORTICAL OPACITY AND BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE VISUAL ACUITY WAS MORE STABLE BUT THE PUPIL WAS DILATED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291699&pc=MTA | 2023826-2017-00197 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6291665&pc=MTA | 2023826-2017-00196 |  | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Missing Value Reason | Cataract; Loss of Vision | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -4.0 DIOPTER, IN THE RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358993&pc=MTA | 2023826-2017-00252 | 2015-09-16 | Injury | STAAR SURGICAL COMPANY | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED AND REPLACED. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THEY CONFIRMED THE PATIENTS REPORT. THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2015 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE LENS WAS EXPLANTED ON (B)(6) 2015, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP VA WAS 20/25. THE FACILITY REPORTED THE EVENT WAS RELATED TO THE DEVICE. Manufacturer Narrative: (B)(4). METHOD: LENS WORK ORDER SEARCH RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6408305&pc=MTA | 2023826-2017-00414 | 2016-02-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2016 AND IT WAS AN ANTERIOR SUBCAPSULAR CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE DATE OF EXPLANT WAS (B)(6) 2016, THE PATIENTS POST-OP VA WAS 20/20 AND PROGNOSIS WAS EXCELLENT. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4601777&pc=MTA | 2023826-2015-00318 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED HAD LOST VISION DUE TO THE DEVELOPMENT OF A CATARACT AND THE LENS WAS EXPLANTED. THE LENS WAS IMPLANTED ON (B)(6) 2013. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED AND THE CATARACT REMOVED ON (B)(6) 2015. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - EVEN THOUGH THE POSITION OF ICL IMPLANTATION COULD CONTRIBUTE TO CATARACT, OTHER FACTORS COULD ALSO CAUSE THE DEVELOPMENT OF CATARACT SUCH AS AGING, TRAUMA (EYE RUBBING, EXTERNAL FORCES), DIABETES, SMOKING AND DRINKING, DRUG USE (STEROIDS) AND RADIATION EXPOSURE. ACCORDING TO THE LITERATURE, MYOPIA PATIENTS ARE AT EARLIER AGE OF DEVELOPING CATARACT THAN PEOPLE OF NORMAL REFRACTION. HIGH DEGREE MYOPIA PATIENT COULD BE 10 YEARS YOUNGER THAN NORMAL PEOPLE TO GET CATARACT. GIVEN THE AGE AND HIGH DEGREE OF MYOPIA (-15D) OF THIS PATIENT IT IS LIKELY THAT THE MEDICAL DEVICE IS NOT THE ROOT CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12215415&pc=MTA | 2023826-2021-02797 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATES A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -16.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. REPORTEDLY ON (B)(6) 2021 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE INITIAL LENS WAS IMPLANTED (B)(6) 2016. ON (B)(6) 2021 LENS OPACITY PSC/NS WAS OBSERVED. THE LENS WAS REMOVED, PHACO-EMULSIFICATION PERFORMED AND AN IOL WAS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS "POSTERIOR SUBCAPSULAR POLAR AGE-RELATED CATARACT." H6-HEALTH EFFICATED IMPACT CODE 4650 (PHACO-EMULSIFICATION, IOL IMPLANTATION). CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17722429&pc=MTA | 2023826-2023-03743 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM512.6, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS OPACITY(CATARACT) ASC OBSERVED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PATIENT IS HAPPY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5 - THE REPORTER ALSO INDICATED THAT PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS ALSO PERFORMED THAT RESOLVED THE PROBLEM. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416011&pc=MTA | 2023826-2020-01713 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S EYE. A CATARACT MAY HAVE DEVELOPED AND THE SURGEON IS PLANNING TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THIS WAS A REQUEST FOR INFORMATION AND WAS NOT A REAL CASE. THERE WAS NO PRODUCT MALFUNCTION OR PATIENT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505351&pc=MTA | 2023826-2020-02057 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE ( LENS TOO SMALL) WAS REPORTED TO BE THE CAUSE OF THIS EVENT. THE REPORTER STATED " LENS IS TOO CLOSE TO NATURAL LENS NOT SITTING CORRECTLY." LENS EXCHANGE IS PLANNED BUT NOT YET TAKEN PLACE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED AT A LATER DATE THAT THE PATIENT DEVELOPED CATARACTS. LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10505372&pc=MTA | 2023826-2020-02056 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. DEVICE ( LENS TOO SMALL) WAS REPORTED TO BE THE CAUSE OF THIS EVENT. THE REPORTER STATED " LENS IS TOO CLOSE TO NATURAL LENS NOT SITTING CORRECTLY." LENS EXCHANGE IS PLANNED BUT NOT YET TAKEN PLACE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PATIENT DEVELOPED CATARACTS. THE LENS STILL REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062446&pc=MTA | 2023826-2023-04853 | 2023-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY ASC THAT WAS OBSERVED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16869280&pc=MTA | 2023826-2023-01634 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE IN SPAIN. A SURGEON FROM IRELAND REPORTEDLY STATED, "RECENTLY EXPLANTED AN ICL WHICH HAD BEEN IMPLANTED IN MADRID 10 YEARS AGO PRESENTING WITH MINIMAL VAULT AND CATARACT." Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: "1766" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222347&pc=MTA | 2023826-2021-02864 | 2021-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY (ASC, OBSERVED: (B)(6) 2021). "ANTERIOR SUBCAPSULAR CATARACT FORMATION TO EDGE OF VISUAL AXIS" WAS REPORTED. IT WAS REPORTED THAT THE DOCTOR NOTICED THE LENS OPACITY WAS WORSENING AND WANTED TO EXCHANGE ICL FOR A DIFFERENT SIZE, SO IT WOULDN'T RUB ON THE LENS. THE REPORTER STATED, "LENS OPACITY IS STILL PRESENT, WE HOPE IT STAYS STABLE AND WE DON'T NOTICE ANY IMMEDIATE CHANGES." CAUSE OF THIS EVENT WAS REPORTED TO BE THE DEVICE. CAUSE DETAILS "SMALL ICL." BCVA WAS REPORTED TO BE (B)(6) 2021. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222348&pc=MTA | 2023826-2021-02865 | 2021-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM,MICL12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. IT WAS REPORTED THAT AN ICL EXCHANGED IS SCHEDULED FOR THE PATIENT, CAUSE IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS OPACITY (ASC, OBSERVED: (B)(6) 2021). "ANTERIOR SUBCAPSULAR CATARACT FORMATION TO EDGE OF VISUAL AXIS" WAS REPORTED. IT WAS REPORTED THAT THE DOCTOR NOTICED THE LENS OPACITY WAS WORSENING AND WANTED TO EXCHANGE ICL FOR A DIFFERENT SIZE, SO IT WOULDN'T RUB ON THE LENS. THE REPORTER STATED, "LENS OPACITY IS STILL PRESENT, WE HOPE IT STAYS STABLE AND WE DON'T NOTICE ANY IMMEDIATE CHANGES." CAUSE OF THIS EVENT WAS REPORTED TO BE THE DEVICE. CAUSE DETAILS "SMALL ICL." BCVA WAS REPORTED TO BE 20/70 ON (B)(6) 2021. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFO:UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11012227&pc=MTA | 2023826-2020-03022 | 2020-10-15 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2020-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Material Opacification; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2003. LENS OPACITY (ASC WAS OBSERVED ON 15/OCT/2020. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.0MM, ICM120V4, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN 2003. LENS OPACITY (ASC) WAS OBSERVED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A LENS OF A DIFFERENT MODEL WITH A LONGER LENGTH ON (B)(6) 2020 WHICH RESOLVED THE PROBLEM. REPORTEDLY, THE STATUS OF THE EYE IS "ALL FINE" AND THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271733&pc=MTA | 2023826-2023-05506 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. A PLANNED LENS REMOVAL AND REPLACMENT WAS REPORTED, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354764&pc=MTA | 2023826-2019-00272 | 2013-07-10 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL 12.1, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 DUE TO LENS OPACITY, CORTICAL, BEING OBSERVED (B)(6) 2018. IT WAS REPORTED THE LENS WAS REMOVED TO PREVENT THE PROGRESSION THE CORTICAL CATARACT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8354778&pc=MTA | 2023826-2019-00273 | 2013-07-25 | Injury | STAAR SURGICAL COMPANY AG | 2019-02-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2013. IT WAS REPORTED THAT PATIENT IS DEVELOPING EARLY CORTICAL CATARACTS AND THE ICL WILL BE REMOVED AND EXCHANGED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM MICL12.1, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. IT WAS REPORTED THAT PATIENT STARTED DEVELOPING EARLY CORTICAL CATARACTS. THE ICL WAS REMOVED ON (B)(6) 2019. THE ICL WAS EXCHANGED FOR A SAME MODEL LARGER LENGTH LENS WITH SAME DIOPTER. THE REPORTER WILL NOT KNOW IF THE ISSUE WAS RESOLVED UNTIL FURTHER FOLLOW UP. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153415&pc=MTA | 2023826-2016-01727 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. AS THE SURGEON WAS IMPLANTING THE LENS, THE PATIENT TURNED AND THE TIP OF THE SFC-45 FP CARTRIDGE TORE THE ANTERIOR CAPSULE. A CATARACT DEVELOPED AND THE LENS WAS EXPLANTED ON (B)(6) 2016. CATARACT SURGERY WAS PERFORMED. Manufacturer Narrative: (B)(4). RESULTS CODES: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6560729&pc=MTA | 2023826-2017-00806 | 2017-04-06 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) AND THE SURGEON NOTED A CATARACT. THE LENS WAS CUT TO REMOVE WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. A SUTURE WAS REQUIRED. THE PATIENT'S POST-OP UCVA WAS 20/30. THE REPORTER STATED THERE WAS NO PROBLEM WITH THE LENS VAULT AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7207877&pc=MTA | 2023826-2018-00107 | 2017-12-19 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.1, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2009. ON (B)(6) 2017, THE LENS WAS REMOVED BECAUSE THE PATIENT DEVELOPED A CATARACT. CATARACT REMOVAL WAS PERFORMED AT THE SAME TIME OF ICL REMOVAL AND A MX60 ( BAUSCH + LOMB) INTRAOCULAR LENS WAS IMPLANTED. IT IS REPORTED THAT INCIDENT/INJURY WAS NO PRODUCT RELATED, THERE WAS NO PROBLEM WITH THE LENS. PATIENT IS DOING GOOD. Manufacturer Narrative: LENS WAS RETURNED IN A SPECIMEN CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER: IN FOLLOW-UP #1, CORRECTED TO 10-JAN-2018. (B)(4) Manufacturer Narrative: CONCLUSION: (B)(4) "THE LENS WAS REMOVED BECAUSE THE PATIENT HAD A CATARACT. NO PROBLEMS WITH LENS." (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7288989&pc=MTA | 2023826-2018-00264 | 2018-01-16 | Injury | STAAR SURGICAL COMPANY | 2018-02-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON EXPLANTED A MICL13.2, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS FROM THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED DUE TO A CATARACT. CATARACT SURGERY WAS PERFORMED AT THE SAME TIME OF ICL REMOVAL AND A SYMFONY (AMO) INTRAOCULAR LENS WAS IMPLANTED. IT IS REPORTED THAT THE CAUSE OF EVENT WAS NOT PRODUCT RELATED, THERE WAS NO PROBLEM WITH THE LENS. PATIENT STATUS WAS REPORTED AS "GOOD". Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL 13.2; -11.5 DIOPTER; IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED FROM THE EYE ON (B)(6) 2018 DUE TO CATARACT. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8274513&pc=MTA | 2023826-2019-00158 | 2018-05-10 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR/NUCLEAR SCLEROTIC/CORTICAL). PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8210581&pc=MTA | 2023826-2019-00012 | 2018-11-01 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 LENS OPACITY (ASC) WAS OBSERVED AND THE LENS WAS REMOVED ON 15 NOV 2018. IT WAS REPORTED THAT BECAUSE THE PATIENT IS RELATIVELY OLD AND OPACITY OF THE LENS IS MILD, THE PATIENT IS BEING CORRECTED WITH GLASSES. Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED ON 15-FEB-2019: "THIS WAS THE SURGEON'S FIRST CASE OF ICL IMPLANTATION. DR. (B)(6) DECIDED TO EXPLANT BOTH LENSES DUE TO TWO REASONS: PATIENT'S OVER REFRACTION (OR) WAS +1.50, OU AND OBVIOUSLY THIS PRESBYOPIC PATIENT WAS NOT HAPPY AT ALL; DR. (B)(6) FOUND ASCO IN THE PERIPHERY IN BOTH LENSES. AFTER EXPLANTATIONS, PATIENT'S VISION IS CORRECTED BY SPECTACLE AND HER BCVA IS 0.8, OU. NO FURTHER SURGICAL OPTION IS PLANNED. DR. ADVISED FUTURE IOL IMPLANTATION IN CASE SHE DEVELOPS CATARACT." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00011 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00012 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8210647&pc=MTA | 2023826-2019-00011 | 2018-11-01 | Injury | STAAR SURGICAL COMPANY | 2019-01-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -4.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 LENS OPACITY (ASC) WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2018. IT WAS REPORTED THAT BECAUSE THE PATIENT IS RELATIVELY OLD AND OPACITY OF THE LENS IS MILD, THE PATIENT IS BEING CORRECTED WITH GLASSES. Manufacturer Narrative: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. SEE MDR# 2023826-2019-00012 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. SEE MDR# 2023826-2019-00011 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: ADDITIONAL INFORMATION RECEIVED ON 15-FEB-2019: "THIS WAS THE SURGEON'S FIRST CASE OF ICL IMPLANTATION. DR. (B)(6) DECIDED TO EXPLANT BOTH LENSES DUE TO TWO REASONS: PATIENT'S OVER REFRACTION (OR) WAS +1.50, OU AND OBVIOUSLY THIS PRESBYOPIC PATIENT WAS NOT HAPPY AT ALL; DR. (B)(6) FOUND ASCO IN THE PERIPHERY IN BOTH LENSES. AFTER EXPLANTATIONS, PATIENT'S VISION IS CORRECTED BY SPECTACLE AND HER BCVA IS 0.8, OU. NO FURTHER SURGICAL OPTION IS PLANNED. DR. ADVISED FUTURE IOL IMPLANTATION IN CASE SHE DEVELOPS CATARACT." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10396107&pc=MTA | MW5096019 | 2018-12-01 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 | MTA | VISIAN ICL | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available |  | 1 | Event Description: ICL SURGERY (B)(6) 2016. GLARES SINCE DAY ONE. CATARACTS FORMED WITHIN 2 YEARS. I AM ONLY IN MY (B)(6). FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795448&pc=MTA | 2023826-2019-01237 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE LEFT (OS) EYE ON (B)(6) 2019 DUE TO AN ANTERIOR SUBCAPSULAR CATARACT. THE SURGERY IS REPORTED TO HAVE BEEN SUCCESSFUL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN CONTACT LENS CONTAINER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: H6 - RESULT CODE 114 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#2. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795471&pc=MTA | 2023826-2019-01236 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE RIGHT (OD) EYE ON (B)(6) 2019 DUE TO AN ANTERIOR SUBCAPSULAR CATARACT. THE SURGERY IS REPORTED TO HAVE BEEN SUCCESSFUL. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN CONTACT LENS CONTAINER WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RETURNED WITH RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 09/12/2019. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANTERIOR SUBCAPSULAR CATARACT IN THE LEFT EYE (OS). NO CATARACT WAS PRESENT IN THE RIGHT EYE, BUT THE PATIENT STILL UNDERWENT CATARACT SURGERY AND EXPLANT IN THE RIGHT EYE (OD). REPORTEDLY THE PATIENT HAS 'HEALED' AND 'HE IS FINE'. PATIENT CODE 1766 NO LONGER APPLICABLE. CATARACT NOT PRESENT IN RIGHT EYE (OD), PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, CATARACT SURGERY), LENS WORKS ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6 - RESULT CODE 114 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#2. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8528238&pc=MTA | 2023826-2019-00631 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON THE SAME DAY IN A DIFFERENT SURGERY DUE TO LENS OPACITY (ASC). THE REPORTER STATED " THE DOCTOR NOTED THE OPACITY AFTER IMPLANTATION. THE OPACITY IS MORE RESOLVED THAN BEFORE, IT WAS A LITTLE IMPROVED AFTER BEING TREATED WITH MEDICATION. IMMEDIATELY AFTER THE REMOVAL OF THE LENS BCVA WAS DECREASED BUT RECOVERED. AT THE LAST VISIT BCVA WAS 20/20 ((B)(6) 2019)". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION: LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736590&pc=MTA | 2023826-2019-01119 | 2019-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED AND THE SURGEON PERFORMED PHACO-EMULSIFICATION AND AN IOL WAS IMPLANTED. THE SURGEON REPORTS LENS OPACITY-CORTICAL. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K #: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276016&pc=MTA | 2023826-2019-02092 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: LENS WAS EXPLANTED DUE TO CATARACT SURGERY. PATIENT CODE 3191: NO CODE AVAILABLE (CATARACT SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: IT WAS REPORTED THAT THEY FIRST NOTED THE CATARACT ON (B)(6) 2019. WHEN THE CATARACT WAS REMOVED IT WAS REPLACED WITH AN IOL FROM A DIFFERENT MANUFACTURER. POST-OP VISIT WAS REPORTED AS THE PATIENT WAS DOING VERY WELL. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9276024&pc=MTA | 2023826-2019-02093 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-11-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2007. THE LENS WAS EXPLANTED DUE TO DEVELOPMENT OF A CATARACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT RECEIVED CATARACT SURGERY. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (CATARACT SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THE CATARACT WAS NOTICED ON 19-SEPT-2019. WHEN THE CATARACT WAS REMOVED IT WAS REPLACED WITH AN IOL FROM ANOTHER MANUFACTURER. POST-OP REPORT STATED THE "PATIENT WAS DOING VERY WELL". CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6153412&pc=MTA | 2023826-2016-01728 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE AND NOTED THE CAPSULE WAS OPEN. THE SURGEON NOTED ONE OF THE PERIPHERAL IRIDECTOMIES MAY HAVE TORN AND THE CAPSULE WAS OPENING UP. THE SURGERY WAS ABORTED AND THE PATIENT HAD CATARACT SURGERY. THE REPORTER WAS UNABLE TO REPORT ANY ADDITIONAL INFORMATION. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8396926&pc=MTA | 2023826-2019-00356 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN IMPLANTABLE COLLAMER LENS, DUE TO THE DEVELOPMENT OF A CATARACT. THE LENS WAS IN THE PATIENTS EYE FOR 15 YEARS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS EXPLANTED FROM THE RIGHT EYE (OD) ON 11/13/2018. THE REPORTER INDICATED PRIOR TO EXPLANTING THE LENS, THE VAULT WAS GOOD AND THERE WAS NO LENS ROTATION. THE CATARACT WAS A CORTICAL (ANTERIOR), NUCLEAR AND POSTERIOR SUBCAPSULAR. THE CATARACT WAS AGE-RELATED. THERE WAS NO PATIENT INJURY WHEN REMOVING THE LENS FROM THE EYE. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED AFTER THE CATARACT WAS REMOVED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414865&pc=MTA | 2023826-2019-00396 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN IMPLANTABLE COLLAMER LENS FROM THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018, DUE TO THE DEVELOPMENT OF A CATARACT AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE REPORTER INDICATED IT WAS A CORTICAL (ANTERIOR), NUCLEAR AND POSTERIOR SUBCAPSULAR CATARACT. THE LENS WAS IN THE PATIENTS EYE FOR 15 YEARS. THE REPORTER INDICATED THE CATARACT WAS AGE-RELATED. THE LENS WAS DISCARDED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031566&pc=MTA | 2023826-2019-01695 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT & REFRACTIVE SURGERY IN JUNE 2019 TITLED, '10-YEAR CLINICAL OUTCOMES AFTER IMPLANTATION OF A POSTERIOR CHAMBER PHAKIC LENS FOR MYOPIA.' THE ARTICLE STATES THAT LENS OPACITY WAS DEVELOPED IN 21 EYES, INCLUDING 10 EYES WITH ASC AND 11WITHOUT ASC. PHACOEMULSIFICATION COMBINED WITH PREVIOUS ICL REMOVAL WAS PERFORMED IN 6 EYES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: PHACOEMULSIFICATION, LENS EXPLANT SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441175&pc=MTA | 2023826-2019-02347 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2013. REPORTEDLY, THE PATIENT DEVELOPED A CATARACT. THE LENS WAS EXPLANTED AND CATARACT SURGERY WAS PERFORMED. AN IOL FROM ANOTHER MANUFACTURER WAS IMPLANTED ON (B)(6) 2019. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957640&pc=MTA | 2023826-2020-00850 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | UNK | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND THEN REMOVED "BECAUSE PATIENT DEVELOPED A CATARACT." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140605&pc=MTA | 2023826-2020-01204 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS OPACITY NS, DATE ASSESSED (B)(6) 2019. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "THE PROBLEM WAS RESOLVED." SHOULD BE ADDED. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703060&pc=MTA | 2023826-2020-00268 |  | Injury | STAAR SURGICAL COMPANY AG | 2020-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 14.5 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CATARACT DEVELOPED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE INCISION WAS NOT ENLARGED TO REMOVE THE LENS AND SUTURES WERE NOT REQUIRED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, WRAPPED IN GAUZE. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON LENS. BUT NO VISIBLE DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7278307&pc=MTA | 2023826-2018-00231 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V3 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED 9 YEARS AFTER BEING IMPLANTED IN THE PATIENTS EYE, DUE TO NUCLEAR SCLEROSIS. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: EXPIRATION DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794667&pc=MTA | 2023826-2018-01196 | 2018-05-16 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE. ANOTHER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), DURING THE SAME SURGERY. THE LENS WAS REMOVED DUE TO LENS TEAR/BREAK DURING INJECTION INTO THE EYE AND THE SURGEON DAMAGED THE CAPSULE. THE SURGEON PERFORMED PHACOEMULSIFICATION, WITH IOL IMPLANTATION. AN INTRAOCULAR LENS (IOL) WAS IMPLANTED IN A SECOND SURGERY AND THE PROBLEM WAS RESOLVED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: PREVIOUS DATE NOT APPLICABLE, LENS WAS NOT RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592014&pc=MTA | 2023826-2015-00279 | 2012-01-12 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011 AND GLARES/HALOES WERE OBSERVED ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(4) 2014 AND WAS NOT REPLACED. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/20. SEE MFR REPORT # 2023826-2015-00278 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: PT : UNK. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH, DEVICE HISTORY AND MEDICAL REVIEWS. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE AND A HAPTIC WAS TORN. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH WERE REMEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. A DEVICE HISTORY REVIEW WAS PERFORMED AND CONCLUDED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTER EVENT. MEDICAL REVIEW - GLARES/HALOS ARE SUBJECTIVE PERCEPTION OF STARBURST OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING THE DAY. POTENTIAL CAUSES ARE: RESIDUE REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PREOPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ECT. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF A WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS HE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY AND MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885902&pc=MTA | 2023826-2015-00759 | 2013-04-10 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012 AND LENS OPACITY WAS NOTED ON (B)(6)2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND NOT REPLACED. THE PATIENT'S LAST BCVA WAS 20/22. SEE MFR# 2023826-2015-00758 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885877&pc=MTA | 2023826-2015-00758 | 2014-05-02 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012 AND LENS OPACITY WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 AND NOT REPLACED. THE PATIENT'S LAST BCVA WAS 20/22. SEE MFR# 2023826-2015-00759 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839567&pc=MTA | 2023826-2015-00628 | 2014-08-09 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXCHANGED FOR A SAME SIZE BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH, MEDICAL REVIEW RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUBCAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN THE 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED OR THE ICL ITSELF. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600860&pc=MTA | 2023826-2015-00309 | 2014-10-01 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (CATARACT). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION SURGERY AND AN IOL WAS IMPLANTED. POST-OP, THE PATIENT WAS PSEUPHAKIC AND THE UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4)..
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839504&pc=MTA | 2023826-2015-00587 | 2014-10-25 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). THE REPORTER INDICATED THERE WAS NO FOLLOW-UP VISIT AFTER LENS REMOVAL. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUBCAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN THE 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED OR THE ICL ITSELF. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592739&pc=MTA | 2023826-2015-00233 | 2015-01-03 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LENS OPACITY (NUCLEAR SCLEROSIS). THE CATARACT WAS REMOVED BY PHACO-EMULSIFICATION SURGERY AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ICL RELATED OPACITIES ARE USUALLY ANTERIOR LENS LOCATED (ASCO = ANTERIOR SUB-CAPSULAR OPACITY). NUCLEAR SCLEROSIS IS KNOWN TO HAVE A STRONG ASSOCIATION WITH AXIAL MYOPIA. SEVERAL FACTORS MAY PLAY A ROLE IN CATARACT DEVELOPMENT (AGING, DEGREE OF MYOPIA, SYSTEMIC DISEASES, MEDICATIONS, SURGICAL MANIPULATION, INFLAMMATION, TRAUMA, ETC). A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING PROCESS OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4888053&pc=MTA | 2023826-2015-00763 | 2015-03-04 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. LENS OPACITY AND LOW VAULT WERE NOTED ON (B)(6) 2015 AND THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NOT REPLACED. PATIENT'S LAST UCVA WAS 20/25. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4639048&pc=MTA | 2023826-2015-00387 | 2015-03-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON EXPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 DUE TO THE DEVELOPMENT OF A TRAUMATIC CATARACT . THE LENS WAS IMPLANTED ON (B)(6) 2011. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. A SUTURE WAS REQUIRED. THE PATIENT IS DOING WELL. THE LENS WAS DISCARDED BY THE FACILITY. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND TWO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS DISCARDED BY FACILITY. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - EXTERNAL FACTORS SUCH AS TRAUMA COULD CAUSE CATARACT. IT IS LIKELY THAT THE MEDICAL DEVICE IS NOT THE CAUSE OF THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4618715&pc=MTA | 2023826-2015-00338 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS WAS REPOSITIONED APPROXIMATELY 2 WEEK AGO AND THE LENS POSITION LOOKED GOOD. THE PATIENT'S VISION CONTINUES TO IMPROVE WITH UNCORRECTED VISUAL ACUITY OF 20/30. THERE ARE SOME CORTICAL CHANGES, WHICH MAY BE THE BEGINNING OF A CATARACT. THE PUPIL IS SLIGHTLY ATONIC, WHICH IS CAUSING PROBLEMS WITH THE PATIENT'S NIGHT VISION (GLARE). THE LENS REMAINS IMPLANTED, WITH THE VAULT AT 30% AND APPEARS SYMMETRIC. Manufacturer Narrative: (B)(4) - (NEW INCISION); (SECONDARY SURGERY); CATARACT, INDUCED; (GLARE); (REPOSITIONED). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - LENS EDGE EFFECTS ARE A SOURCE OF OPTICAL ABERRATIONS-SUCH AS GLARE AND HALOS. NIGHT GLARE RESULTS WHEN PUPILS DILATE LARGER THAN THE ICLS, AND THIS PATIENT HAS A MILDLY ATONIC PUPIL OS. THE THREE-YEAR FDA STUDY REPORTED MARKED/SEVERE NIGHT DRIVING DIFFICULTIES IN 16.7% OF PATIENTS WITH LOW LENS OPTIC DIAMETER. THIS COMPLAINT IS LIKELY RELATED TO THE ICLS. BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4822542&pc=MTA | 2023826-2015-00645 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-06-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON EXPLANTED AN MICL IMPLANTABLE COLLAMER LENS DUE TO THE PATIENT DEVELOPED A CATARACT. THE PATIENT HAD THE MICL IMPLANTED IN 2008. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2015. THE CATARACT WAS REMOVED AND AN INTRAOCULAR LENS WAS IMPLANTED. THE REPORTER INDICATED THE INCIDENT/INJURY WAS NOT PRODUCT-RELATED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, ICL WAS EXPLANTED 7 YEARS POSTOPERATIVELY TO ADDRESS CATARACT DEVELOPMENT. ACCORDING TO CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY (CCQS), RECEIVED ON (B)(6) 2016, ICL AND CATARACT WERE REMOVED AT THE SAME SURGERY DATE AND IOL WAS SUCCESSFULLY IMPLANTED. THE VA WAS 20/20 POSTOPERATIVELY. THE SAME FORM STATED THAT CATARACT DEVELOPMENT WAS NOT PRODUCT RELATED. IT SHOULD BE NOTED THAT CATARACT HAD DEVELOPED BILATERALLY. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4) NEW INCISION. SECONDARY SURGERY. EXPLANTED. DEVICE EVALUATED BY MANUFACTURER?: LENS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY FACTORS SUCH AS AGING, TRAUMA, DIABETES, RADIATION, SMOKING, ALCOHOL USE AND EVEN SOME MEDICATIONS (EX. STEROID) COULD CAUSE CATARACT. GIVEN THE LATE ONSET OF THE EVENT, THESE FACTORS COULD NOT BE RULED OUT. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275441&pc=MTA | 2023826-2018-00230 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V3 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENTS EYE 10 YEARS AGO, DUE TO POST-TRAUMATIC ICL DISLOCATION, WITH ANTERIOR SUBCAPSULAR CATARACT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. Manufacturer Narrative: CORRECTED INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A V3 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENT'S EYE FOR 10 YEARS DUE TO POST-TRAUMATIC ICL DISLOCATION WITH ANTERIOR SUBCAPSULAR CATARACT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8848713&pc=MTA | 2023826-2019-01376 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; No Code Available | N/A | 1 | Event Description: ON 03 JULY 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES ICL DISLOCATION IN 2 EYES OF A HIGHLY MYOPIC PATIENT. REPOSITIONING OF ICL WAS PERFORMED, WITH SUBSEQUENT EXPLANTATION DUE TO REPEAT DISLOCATION AND INCIPIENT CATARACT. DURING OBSERVATION AN INCREASE IN THE AXIAL LENGTH OF 2MM WAS RECORDED. SUBSEQUENTLY CATARACT SURGERY WAS PERFORMED ON THIS PATIENT WITH IMPLANTATION OF PC IOL. AT THE LAST FOLLOW-UP EXAMINATION UCVA WAS 1.0 AND 0.9. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ON 03 JULY 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN CZECH AND SLOVAK JOURNAL OF OPHTHALMOLOGY ENTITLED, 'INCIDENCE OF CATARACT FOLLOWING IMPLANTATION OF POSTERIOR- CHAMBER PHAKIC LENS ICL (IMPLANTABLE COLLAMER LENS) LONG TERM RESULTS.' THE ARTICLE ANALYZES THE RESULTS OF ICL IMPLANTATION IN 34 PATIENTS FROM 1998 TO 2013. THE ARTICLE REFERENCES ICL DISLOCATION IN 2 EYES OF A HIGHLY MYOPIC PATIENT. REPOSITIONING OF ICL WAS PERFORMED, WITH SUBSEQUENT EXPLANTATION DUE TO REPEAT DISLOCATION AND INCIPIENT CATARACT. DURING OBSERVATION AN INCREASE IN THE AXIAL LENGTH OF 2MM WAS RECORDED. SUBSEQUENTLY CATARACT SURGERY WAS PERFORMED ON THIS PATIENT WITH IMPLANTATION OF PC IOL. THE PC IOL ALSO DECENTERED AND REQUIRED EXPLANT. AT THE LAST FOLLOW-UP EXAMINATION BCVA WAS 1.0 AND 0.9. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. ICM V4 USED FOR MYOPIA, TICM V4 USED FOR ASTIGMATISM. PATIENT CODE 3191: CATARACT SURGERY, IOL IMPLANT, IOL EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358894&pc=MTA | 2023826-2017-00260 | 2013-04-16 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.50 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AND THE PATIENT REPORTS NO RESTRICTIONS WITH ASCO. Manufacturer Narrative: CONCLUSION CODE: PER DFU, PATIENT'S AGE IS OVER 45 YEARS OF AGE WHICH IS OFF LABEL USE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6358899&pc=MTA | 2023826-2017-00261 | 2013-11-19 | Injury | STAAR SURGICAL COMPANY AG | 2017-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICMO13.2, -8.00 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AS THE PATIENT IS STILL HAPPY AND FEELS NO LIMITATIONS FROM THE ASCO. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED LENS OPACITY (ASCO). THE LENS REMAINS IMPLANTED IN THE EYE AS THE PATIENT IS STILL HAPPY AND FEELS NO LIMITATIONS FROM THE ASCO. THE SURGEON HAS NO PLANS TO EXPLANT OR EXCHANGE THE ICL. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY ON (B)(6) 2016 WAS 20/25. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS OVER 45 YEARS OF AGE AND WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300934&pc=MTA | 2023826-2018-00292 | 2015-10-28 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -11.0/+1.0/150 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS LENS OPACITY NS. REPORTEDLY, PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS PERFORMED. THE SURGEON STATES "CATARACT SURGERY PERFORMED. ICL REMOVED DUE TO CATARACT. NOW PERFECT VISION." Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6303775&pc=MTA | 2023826-2017-00143 | 2016-06-04 | Injury | STAAR SURGICAL COMPANY | 2017-02-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LENS OPACITY ASCO (ANTERIOR SUBCAPSULAR OPACITY) WHICH WAS OBSERVED ON (B)(6) 2016. ON PATIENT'S LAST VISIT ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/30 WHICH IS THE SAME AS PREOPERATIVE DATA AS OF (B)(6) 2015. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE'S CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FIBERS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6721184&pc=MTA | 2023826-2017-01190 | 2016-09-12 | Injury | STAAR SURGICAL COMPANY | 2017-07-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+4/87 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS OPACITY (POSTERIOR SUBCAPSULAR CATARACT). ON (B)(6) 2017 THE LENS WAS EXPLANTED FOR PHACO-EMULSIFICATION AND IOL IMPLANTATION. THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. THE PRODUCT HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL AND HAS SOME MOISTURE. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THIS PRODUCT IS MANUFACTURED IN THE U.S, BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7557824&pc=MTA | 2023826-2018-00812 | 2018-03-16 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON NOTING LENS OPACITY ASC. REPORTEDLY, ON (B)(6) 2018 CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED WHICH RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4)-NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY & IOL IMPLANTATION). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389922&pc=MTA | 2023826-2017-00343 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ON (B)(6) 2017 THE REPORTER INDICATED AWARENESS "OF A CASE OF V4C LENS + LENS OPACITY." THE LENS WAS REMOVED AND PHACO IOL IMPLANTATION WAS PERFORMED. AT THE DATE OF MDR SUBMISSION, ATTEMPS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, DIOPTER -11.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LENS OPACITY. PHACO IOL IMPLANTATION WAS PERFORMED. Manufacturer Narrative: PATIENT INFORMATION: UNK. EVENT DATE: UNK. SUSPECT MEDICAL DEVICE: UNK. INITIAL REPORTER ADDRESS: UNK. PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. MFG DATE: UNK. (B)(4). Manufacturer Narrative: (B)(4): LENS EXCHANGE; SHOULD STATE: LENS EXPLANT. CONCLUSION: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))]. DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740169&pc=MTA | 2023826-2017-01236 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTED STATED THAT THE LENS IS SCHEDULED TO BE REMOVED ON (B)(6) 2017 DUE TO THE PATIENT FORMING CATARACTS. THE REPORTER INDICTED THAT CATARACT SURGERY IS BEING PERFORMED AND A TECNIS IOL FROM A DIFFERENT MANUFACTURER IS BEING IMPLANTED. REPORTED STATED THAT IT IS UNKNOWN IF THE ADVERSE EVENT WAS RELATED TO THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL AND THERE WERE NO ISSUES WITH THE REMOVAL OF THE ICL LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740468&pc=MTA | 2023826-2017-01237 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2015. THE REPORTED STATED THAT THE LENS IS SCHEDULED TO BE REMOVED ON (B)(6) 2017 DUE TO THE PATIENT FORMING CATARACTS. THE REPORTER INDICTED THAT CATARACT SURGERY IS BEING PERFORMED AND A TECNIS IOL FROM ANOTHER MANUFACTURER IS BEING IMPLANTED. REPORTED STATED THAT IT IS UNKNOWN IF THE ADVERSE EVENT WAS RELATED TO THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL AND THERE WERE NO ISSUES WITH THE REMOVAL OF THE ICL LENS. (B)(4). Manufacturer Narrative: SCHEDULED TO BE EXPLANTED ON (B)(6) 2017. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484256&pc=MTA | 2023826-2018-00646 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, ICL EXPLANT, CATARACT EXTRACTION WITH IOP IMPLANTATION) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484260&pc=MTA | 2023826-2018-00647 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION: UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484266&pc=MTA | 2023826-2018-00644 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: PATIENT INFORMATION: UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7484287&pc=MTA | 2023826-2018-00649 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT INFORMATION: UNK. DATE OF EVENT-UNK. SERIAL NUMBER, EXPIRATION DATE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7485001&pc=MTA | 2023826-2018-00648 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481813&pc=MTA | 2023826-2018-00635 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED A ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDTIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED. PATIENT CODE: (B)(4), SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481815&pc=MTA | 2023826-2018-00633 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED A ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDTIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED PATIENT CODE: (B)(4) -SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481816&pc=MTA | 2023826-2018-00630 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED A ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDTIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED PATIENT CODE: (B)(4) -SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481817&pc=MTA | 2023826-2018-00631 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED A ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDTIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SERIAL NUMBER REPORTED PATIENT CODE: (B)(4)-SECONDARY SURGERY, LENS EXPLANT, CATARACT SURGERY, IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481821&pc=MTA | 2023826-2018-00634 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED A ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7481824&pc=MTA | 2023826-2018-00632 |  | Injury | Unknown Manufacturer | 2018-05-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | N/A | 1 | Event Description: ARTICLE STATED THAT THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS IN A PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO CATARACT DEVELOPMENT. PATIENT UNDERWENT CATARACT SURGERY AND IOL IMPLANTATION. ATTEMPTS AT OBTAINING ADDITIONAL INFORMATION HAVE NOT BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488506&pc=MTA | 2023826-2018-00664 |  | Injury | Unknown Manufacturer | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: NO SERIAL NUMBER REPORTED. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (MANUFACTURING SITE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488565&pc=MTA | 2023826-2018-00665 |  | Injury | Unknown Manufacturer | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488613&pc=MTA | 2023826-2018-00666 |  | Injury | Unknown Manufacturer | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488638&pc=MTA | 2023826-2018-00667 |  | Injury | Unknown Manufacturer | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7488657&pc=MTA | 2023826-2018-00668 |  | Injury | Unknown Manufacturer | 2018-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6879893&pc=MTA | 2023826-2017-01480 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2017-09-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -14.0 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON THE SAME DATE, INTRAOPERATIVELY, THE LENS WAS EXPLANTED DUE TO LENS OPACITY (ASCO). REPORTEDLY, THE TOP OF THE LENS WAS NICKED BY THE SURGEON IN ERROR, WHICH LED TO "CAPSULE MEMBRANE DAMAGE." THE ICL WAS THEN REMOVED AND REPLACED WITH AN IOL. THE PATIENT HAS NO PRE-EXISTING CONDITION OF LENS OPACITY. ACCORDING TO THE SURGEON, THE LENS CAPSULE MEMBRANE DAMAGE IS THE MAIN CAUSE OF THE SITUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE ON LENS). (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6805946&pc=MTA | 2023826-2017-01283 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY DUE TO LENS OPACITY (ASCO). PHACO- EMULSIFACTION WAS PERFORMED AND AN IOL WAS IMPLANTED. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS FOUND TO BE 20/20. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: IN THE PROCESS OF THE IOL IMPLANT, DUE TO THE LACK OF PATIENT'S COOPERATION (TURN THE HEAD) LEAD TO THE OPERATION CAUSED BY NATURAL CRYSTALS BECOME CLOUDY. THE IOL IMPLANTATION SOLVED THE PROBLEM AND THE PATIENT'S CONDITION IS GOOD. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577503&pc=MTA | 2023826-2016-00438 | 2016-01-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 -08.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. A LENS OPACITY (CORTICAL) WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016, CATARACT SURGERY WAS PERFORMED AND A IOL WAS IMPLANTED. PATIENT'S LAST VISIT ON 02/22/2016, BCVA WAS 20/20. SEE MFR. #2023826-2016-00447 FOR PREVIOUS SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS OVER THE AGE OF 45 YEARS OLD AND WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -08.50. DEVICE NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289852&pc=MTA | 2023826-2017-00169 | 2016-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-01-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER STATED CAPSULAR TEAR AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY) WAS OBSERVED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON THE SAME DAY. PHACO-EMULSIFICATION (CATARACT SURGERY) WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) THERE WAS A CAPSULAR TEAR FOR THE ICL LENS THAT WAS EXPLANTED. THE SECONDARY SURGERY WHICH WAS A PHACO-EMULSIFICATION (CATERACT SURGERY) AND IOL IMPLANTATION HAPPENED DUE TO THE PATIENT HAVING CATARACT. (B)(4) LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6080846&pc=MTA | 2023826-2016-01598 | 2014-11-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), NOTED ON (B)(6) 2014. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/23. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS UNDER THE AGE OF 21 YEARS. A 16.0 DIOPTER IS INCORRECT, SHOULD BE 16.5 DIOPTER. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077381&pc=MTA | 2023826-2016-01531 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PMA/510(K) : NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074345&pc=MTA | 2023826-2016-00912 | 2016-04-13 | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM TICM115V4 IMPLANTABLE COLLAMER LENS, -22.50/+3.0/095 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT UNDER THE AGE OF 21 YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404792&pc=MTA | 2023826-2017-00401 | 2016-07-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -17.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED DRY IN A SMALL VIAL. THERE WAS NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2016, THE PATIENT EXPERIENCED LENS OPACITY (ASCO). (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6289966&pc=MTA | 2023826-2017-00149 | 2016-08-06 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM (B)(4) -12.5 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT REPORTING LOW VAULT AND LATE LENS OPACITY (ASCO). ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXCHANGED AGAIN DUE TO LOW VAULT, REFERENCE MDR# 2023826-2017-00194. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/15. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048826&pc=MTA | 2023826-2016-01515 | 2016-08-15 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS - -00.82/100.41/57 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS OPACITY (ANTERIOR SUBCAPSULAR OPACITY). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6566884&pc=MTA | 2023826-2017-00810 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/82 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE SURGEON NOTICED LOW VAULTING AND LENS OPACITY (ASCO) EARLY POSTOPERATIVELY. THE LENS WAS EXCHANGED FOR A LONGER LENS OF A DIFFERENT MODEL AND DIOPTER AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/14 (BETTER THAN PREOPERATIVE). Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE THE LENS WAS RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: (B)(4). AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS OVER (B)(6)YEARS OF AGE AND WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6590465&pc=MTA | 2023826-2017-00825 | 2016-08-16 | Injury | STAAR SURGICAL COMPANY | 2017-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/88 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED LOW VAULT AND LENS OPACITY (ASCO). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND WAS EXCHANGED WITH A LONGER LENS WITH A DIFFERENT MODEL AND DIOPTER, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED SCOTCHED ON A PAPER. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FOREIGN MATERIAL ON LENS SURFACE (FIBERS AND REST OF SCOTCH). IT WAS NECESSARY TO REHYDRATE BECAUSE OF RETURNED SCOTCHED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE- AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) UNDER 3.0MM AND FOR PATIENTS OVER THE AGE OF (B)(6) AT THE TIME OF IMPLANTATION. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6117014&pc=MTA | 2023826-2016-01622 | 2016-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICH12.1 IMPLANTABLE COLLAMER LENS, 0.00/3.5/114 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS OPACITY, AND ANTERIOR SUBCAPSULAR CATARACT. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON IT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGICAL INTERVENTION, AND LENS EXPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214422&pc=MTA | 2023826-2016-01843 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS OPACITY, AND ANTERIOR SUBCAPSULAR CATARACT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS MEASUREMENTS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6214458&pc=MTA | 2023826-2016-01844 | 2016-11-24 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS OPACITY, AND ANTERIOR SUBCAPSULAR CATARACT. Manufacturer Narrative: ADDITIONAL EVENT DESCRIPTION - PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25 AS OF (B)(6) 2016. DEVICE EVALUATION: LENS WAS RETURNED BENT AND DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE LENS OPTIC WAS TORN AND THE HAPTIC WAS BROKEN. FOREIGN MATERIAL (FIBERS) WAS FOUND ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6365366&pc=MTA | 2023826-2017-00265 | 2017-01-12 | Injury | STAAR SURGICAL COMPANY | 2017-02-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, (DIOPTER UNKNOWN) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY (ASCO). AS OF THE DATE OF MDR REPORTING THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7495937&pc=MTA | 2023826-2018-00713 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2018-05-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS OPACITY-ASC. THE SURGEON THEN PERFORMED CATARACT SURGERY WITH A TRIFOCAL LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND THE LENS IS MISSING A PIECE OF HAPTIC. A-9.5 DIOPTER. REPORT SOURCE - USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7177027&pc=MTA | 2023826-2018-00008 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2018-01-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR), OBSERVED ON (B)(6) 2017. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. CATARACT SURGERY WAS PERFORMED 4 DAYS LATER. THE PATIENT'S CURRENT BCVA IS 20/40. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN OFF AND MISSING. THE LENS WIDTH WAS MEASURED AND WAS VERIFIED. WAS UNABLE TO VERIFY DIAMETER/LENGTH DUE TO MISSING HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7339941&pc=MTA | 2023826-2018-00412 | 2017-12-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE LENS OPACITY (CATARACT). CATARACT SURGERY (PHACOEMULSIFICATION) WAS PERFORMED AND AN INTRAOCULAR LENS (IOL) WAS IMPLANTED. THE REPORTER INDICATED THE PATIENT IS ALLERGIC TO SULPHUR MEDICATIONS HENCE DIAMOX WAS NOT PRESCRIBED. THE CATARACT COULD HAVE POSSIBLY FORMED DUE TO FLUCTUATIONS IN INTRAOCULAR PRESSURE (IOP) AND AC DEPTH POST-OP. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7278488&pc=MTA | 2023826-2018-00232 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V3 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED 3 YEARS AFTER BEING IMPLANTED IN THE PATIENTS EYE, DUE TO SHALLOW VAULT WITH ANTERIOR SUBCAPSULAR CATARACT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTATED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, PHACOEMULSIFICATION SURGERY WITH IOL IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10311086&pc=MTA | 2023826-2020-01496 |  | Injury | STAAR SURGICAL COMPANY | 2020-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATE THAT A 13.2MM, MICL13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULT AND ASC CATARACT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND PHACO-EMULSIFICATION AND IOL IMPLANTATION WAS PERFORMED, PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4887820&pc=MTA | 2023826-2015-00762 | 2010-12-10 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. LENS OPACITY AND LOW VAULT WERE NOTED ON (B)(6) 2010. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885665&pc=MTA | 2023826-2015-00757 | 2011-05-14 | Injury | STAAR SURGICAL COMPANY | 2015-07-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS OPACITY (ANTERIOR SUBCAPSULAR). THE LENS WAS EXPLANTED AND AN IOL WAS IMPLANTED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - 02 (EVALUATION ANTICIPATED BUT NOT YET DONE) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592589&pc=MTA | 2023826-2015-00235 | 2014-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND LENS OPACITY (CORTICAL). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592661&pc=MTA | 2023826-2015-00277 | 2014-07-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A CLEAR SURGICAL RESIDUE AND THERE WAS NO VISIBLE DAMAGE WAS OBSERVED. BOTH THE WIDTH AND LENGTH WERE REMEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASE WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL INCREASE THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5655201&pc=MTA | 2023826-2016-00685 | 2014-09-18 | Injury | STAAR SURGICAL COMPANY AG | 2016-05-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4, -20 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT AND AN ANTERIOR SUBCAPSULAR CATARACT WERE OBSERVED ON (B)(6) 2014. THE LENS WAS EXPLANTED, CATARACT SURGERY WAS PERFORMED, AND AN IOL WAS IMPLANTED ON (B)(6) 2015. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592687&pc=MTA | 2023826-2015-00281 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LASTEST UCVA WAS 20/30. SEE MFR REPORT #2023826-2015-00282 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592699&pc=MTA | 2023826-2015-00282 | 2014-11-18 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING AND ANTERIOR SUBCAPSULAR OPACITY. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LASTEST UCVA WAS 20/20. SEE MFR REPORT #2023826-2015-00281 FOR THE PATIENT'S OTHER EYE. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH A TORN HAPTIC AND A CLEAR SURGICAL RESIDUE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE WIDTH AND LENGTH OF THE LENS WAS RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT ANTERIOR SUB-CAPSULAR CATARACT OCCURS IN 1.3% OF PATIENTS WHERE AN ICL HAS BEEN IMPLANTED, FOLLOWING THE RECOMMENDATIONS FOR USE. THIS USUALLY OCCURS WITHIN 1ST 6 MONTHS POST-OP. THE MOST PROBABLE CAUSE HAS BEEN DETERMINED TO BE CRYSTALLINE LENS TOUCH DURING SURGERY WHICH COULD HAVE BEEN EITHER FROM THE INSTRUMENTS USED ON THE ICL ITSELF. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. STAAR RECOMMENDS THE ICL TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR IN CASES WHERE THERE IS A TRACE ANTERIOR SUB-CAPSULAR CATARACT WITH NO PROGRESSION. STAAR BELIEVES THAT THE ACTION TO REMOVE AND/OR REPLACE THE ICL CREASES THE RISK FOR ANTERIOR SUB-CAPSULAR CATARACT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154946&pc=MTA | 2023826-2015-01267 | 2015-02-21 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. LENS OPACITY (ASCO) WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP VISIT ON (B)(6) 2015, UCVA WAS 20/26. Manufacturer Narrative: (B)(6). (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -16.5. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4727980&pc=MTA | 2023826-2015-00472 | 2015-03-12 | Injury | STAAR SURGICAL COMPANY | 2015-04-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. LENS OPACITY (ANTERIOR) WAS NOTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER INDICATED THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO SIZING. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT LOW VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF LOW VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, PRODUCT EVALUATION AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. THE REPORTER INDICATED THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO SIZING. AS REPORTED ON THE MEDICAL REVIEW, IT IS LIKELY THE INSUFFICIENT VAULTING WAS A CONSEQUENCE OF WRONG LENS USE. HOWEVER, NO DEFINITE ROOT CAUSE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161950&pc=MTA | 2023826-2015-01278 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/24. SEE MFR #2023826-2015-01277 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -12.00. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSION: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5161958&pc=MTA | 2023826-2015-01277 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. PATIENT VISIT ON (B)(6) 2015 - UCVA WAS 20/24. SEE MFR #2023826-2015-01278 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. WORK ORDER SEARCH: TWO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4) (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5095987&pc=MTA | 2023826-2015-01180 | 2015-08-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2009. LOW VAULT, REFRACTIVE CHANGE OVER TIME AND LENS OPACITY WAS OBSERVED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SURGERY IS PENDING TO EXPLANT AND EXCHANGE FOR LONGER LENS. UNCORRECTED VISUAL ACUITY WAS 20/10 ON (B)(6) 2015. SEE MFR #20238236-2015-01179 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). LENS EXPLANTED, BUT NOT RETURNED. Manufacturer Narrative: DIOPTER: -16.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS REMAINS IMPLANTED. (B)(4). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5707624&pc=MTA | 2023826-2016-00696 | 2016-04-01 | Injury | STAAR SURGICAL COMPANY | 2016-06-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT HAD A MILD SUP. ANTERIOR CORTICAL OPACITY. Manufacturer Narrative: (B)(4). ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059292&pc=MTA | 2023826-2016-01530 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS - -18.5/+1.5/78 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT, LENS ROTATION AND CATARACT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/13. Manufacturer Narrative: THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/13 AND UNCORRECTED VISUAL ACUITY WAS 20/13. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199510&pc=MTA | 2023826-2015-01455 | 2015-02-28 | Injury | STAAR SURGICAL COMPANY | 2015-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Cataract; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER -12.50 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. LOW VAULT AND A ANTERIOR SUBCAPSULAR CATARACT WAS OBSERVED ON (B)(6) 2015. THE ICL WAS EXPLANTED, CATARACT SURGERY AND IOL IMPLANTATION WAS PERFORMED ON (B)(6) 2015. PATIENT'S LAST VISIT ON (B)(6) 2015 - BCVA WAS 20/20 AND UCVA WAS 20/32. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. INITIAL MDR DATE 10/13/2015. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.50. (B)(4). CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4592752&pc=MTA | 2023826-2015-00227 | 2013-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-03-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO LOW VAULTING, LENS ROTATION AND LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS NOTED ON (B)(6) 2013. THE PATIENT HAD CATARACT SURGERY AND AN IOL WAS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PT WEIGHT: UNK. (B)(4) DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - CLINICAL STUDIES HAVE SHOWN THAT TORIC ICL ROTATION OCCURS IN 0.5% OF PATIENTS WHERE A TICL HAS BEEN IMPLANTED. SEVERAL FACTORS MAY CONTRIBUTE TO THIS PHENOMENON (I.E. PATIENTS ANATOMICAL STRUCTURE, A SHORT LENS, HAPTIC POSITION, ETC.). ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR SULCUS OR CILIARY SULCUS CYST, ETC). ALTHOUGH SEVERAL FACTORS MAY PLAY A ROLE IN CATARACT DEVELOPMENT (AGE, DEGREE OF MYOPIA, SYSTEMIC DISEASES, MEDICATIONS, SURGICAL MANIPULATION, ETC.) WE CANNOT RULE OUT LOW VAULTING AS A CONTRIBUTING FACTOR. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850306&pc=MTA | 2023826-2017-01413 | 2017-05-23 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.5/107 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR), WHICH WAS OBSERVED ON (B)(6) 2017. POST CATARACT OPERATION PATIENT IS HAPPY WITH THE VISION. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF LABEL USE (ACD<3.0MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7562910&pc=MTA | 2023826-2018-00836 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS OPACITY (ANTERIOR SUBCAPSULAR). CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, CATARACT SURGERY AND IOL IMPLANT CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6042655&pc=MTA | 2023826-2016-01499 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, IN HER LEFT EYE (OS). THE PATIENT REPORTED A CATARACT HAD DEVELOPED AND THE ICL WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6042656&pc=MTA | 2023826-2016-01498 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Cataract; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS, IN HER RIGHT EYE (OD). THE PATIENT REPORTED A CATARACT HAD DEVELOPED AND THE ICL WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817542&pc=MTA | 2023826-2017-01371 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED AN MICL IMPLANTABLE COLLAMER LENS, WAS EXPLANTED FROM THE PATIENT'S RIGHT EYE ON (B)(6) 2017 DUE TO THE DEVELOPMENT OF A CATARACT. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL. THE LENS WAS IMPLANTED APPROXIMATELY FIVE YEARS AGO BY ANOTHER DOCTOR/FACILITY. THE REPORTER WAS UNABLE TO PROVIDE IMPLANT DATE OR LENS/SERIAL NUMBER INFORMATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED APPROXIMATELY 5 YEARS AGO (2012). (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6817585&pc=MTA | 2023826-2017-01370 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED AN MICL IMPLANTABLE COLLAMER LENS, WAS EXPLANTED FROM THE PATIENT'S LEFT EYE ON (B)(6) 2017 DUE TO THE DEVELOPMENT OF A CATARACT. THE CATARACT WAS REMOVED AND AN IOL WAS IMPLANTED. THE PATIENT IS DOING WELL. THE LENS WAS IMPLANTED APPROXIMATELY FIVE YEARS AGO BY ANOTHER DOCTOR/FACILITY. THE REPORTER WAS UNABLE TO PROVIDE IMPLANT DATE OR LENS/SERIAL NUMBER INFORMATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED APPROXIMATELY 5 YEARS AGO (2012). (B)(4). Manufacturer Narrative: (B)(4). CONCLUSION CODE: OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT) (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469578&pc=MTA | 2023826-2018-00575 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469582&pc=MTA | 2023826-2018-00576 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469599&pc=MTA | 2023826-2018-00577 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469623&pc=MTA | 2023826-2018-00578 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED, PHACOEMULSIFICATION AND IOL PLACEMENT (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7469630&pc=MTA | 2023826-2018-00579 |  | Injury | STAAR SURGICAL COMPANY AG | 2018-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENT IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) UNDERWENT ICL EXPLANTATION DUE TO CATARACT DEVELOPMENT. PATIENT ALSO UNDERWENT CATARACT SURGERY BY PHACOEMULSIFICATION AND IOL PLACEMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962469&pc=MTA | 2023826-2019-01587 | 2018-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Cataract; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE DEVELOPMENT OF A CATARACT. THE CATARACT WAS REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. B5 - THE LENS WAS EXPLANTED ON (B)(6) 2019. THE CAUSE OF THE EVENT WAS REPORTED AS THE ICL. PATIENT HAD A PRE-EXISTING CONDITION OF PROBABLY AMBLYOPIA. POSTOPERATIVE REPORT STATED PATIENT IS "DOING WELL" ON (B)(6) 2019 LAST VISIT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889749&pc=MTA | 2023826-2019-01444 | 2019-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -06.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND LENS REMAINS IMPLANTED. REPORTER INDICATED "THE VAULT WAS 271 MICROMETERS ONE WEEK AFTER SURGERY ((B)(6) 2019), BUT AFTER THAT THE VAULT DECREASED." SURGEON CONTINUES TO MONITOR AND IS "CONSIDERING EXPLANTATION OR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION THAT SHOULD HAVE BEEN REPORTED IN INITIAL MDR: THE SURGEON ALSO FOUND PUNCTATE CATARACT DURING A POST-OP VISIT 6 MONTHS AFTER IMPLANTATION ((B)(6) 2019). ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT. PATIENT CODE 1766 ADDED, PATIENT CODE 2692 NOT APPLICABLE. RESULT CODE 213 ADDED, RESULT CODE 114 NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889750&pc=MTA | 2023826-2019-01445 | 2019-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -06.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LOW VAULT WAS OBSERVED AND LENS REMAINS IMPLANTED. REPORTER INDICATED "THE VAULT WAS 241 MICROMETERS ONE WEEK AFTER SURGERY ((B)(6) 2019), BUT AFTER THAT THE VAULT DECREASED." SURGEON CONTINUES TO MONITOR AND IS "CONSIDERING EXPLANTATION OR EXCHANGE." CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION THAT SHOULD HAVE BEEN REPORTED IN INITIAL MDR: THE SURGEON ALSO FOUND PUNCTATE CATARACT DURING A POST-OP VISIT 6 MONTHS AFTER IMPLANTATION ((B)(6) 2019). ADVERSE EVENT OR PRODUCT PROBLEM CORRECTED TO ADVERSE EVENT. PATIENT CODE 1766 ADDED, PATIENT CODE 2692 NOT APPLICABLE. RESULT CODE 213 ADDED, RESULT CODE 114 NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8505691&pc=MTA | 2023826-2019-00589 |  | Injury | Unknown Manufacturer | 2019-04-11 04:00 | MTA | UNK | Adverse Event Without Identified Device or Use Problem | Cataract; Patient Problem/Medical Problem | UNK | 1 | Event Description: PATIENT HAS HAD IMPLANT FOR 5 YEARS AND DEVELOPED CATARACTS. WILL BE HAVING CATARACT SURGERY IN TWO WEEKS. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17979759&pc=MTA | 2023826-2023-04656 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | Cataract; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: A STUDY ARTICLE WAS PUBLISHED IN PREPRINTS.ORG TITLED: "REFRACTIVE SURGERY TO CORRECT VISUAL IMPAIRMENTS IN 267 CHILDREN WITH AUTISM SPECTRUM AND RELATED NEURO DEVELOPMENTAL DISORDERS: IMPROVEMENTS IN VISION AND BEHAVIOR" ON 31-JUL-2023. THE ARTICLE REPORTS THAT PIOL'S (UNKNOWN BRAND/MODEL/SIZE) WERE IMPLANTED INTO THE EYES OF PATIENTS UNDER 18 YEARS OF AGE DIAGNOSED WITH "ASD OR NDD WITH ASD-LIKE ASSOCIATED BEHAVIORS: DOWN SYNDROME, ANGELMAN SYNDROME, CORNELIA DE LANGE SYNDROME, FETAL ALCOHOL SYNDROME, RHETT SYNDROME, OR SELF-INJURIOUS BEHAVIOR." TWO EYES DEVELOPED CATARACT AND REQUIRED EXPLANTATION. TWO EYES REQUIRED ADDITIONAL PI DUE TO PUPIL BLOCK AND OCULAR HYPERTENSION. THE ARTICLE CONCLUDES THAT REFRACTIVE ERROR AND VISUAL ACUITY IMPROVED SUBSTANTIALLY. Manufacturer Narrative: H6-CLINICAL CODE: (B)(4) -OCULAR HYPERTENSION. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350594&pc=MTA | 2023826-2022-00298 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Cataract; Pupillary Block; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "NOVEL ICL: LOW ADVERSE EVENTS IN HIGHLY MYOPIC SPECIAL-NEEDS CHILDREN" INDICATED THAT A STUDY WAS CONDUCTED TO EVALUATE THE SAFETY OF THE VISIAN INTRAOCULAR COLLAMER LENS (ICL) IN HIGHLY MYOPIC SPECIAL-NEEDS PEDIATRIC CASES. IT NOTES THE STUDY INCLUDES 95 CHILDREN (160 EYES) WITH IMPLANTATION OF AN ICL OF OVER 5 YEARS. THE STUDY RESULT INDICATES A MINOR ADVERSE EVENT WAS OCULAR HYPERTENSION IN RESPONSE TO STEROIDS THAT OCCURRED IN 3 EYES WHEN THE STEROIDS WERE STOPPED. THE MOST FREQUENT ADVERSE EVENT WAS POSTOPERATIVE PUPILLARY BLOCK THAT REQUIRED PI REVISION IN 7 EYES. IN 1 EYE A WOUND LEAK; A CATARACT IN 1 EYE; TRAUMATIC RETINAL DETACHMENT IN 1 CHILD. THE ARTICLE CONCLUDES THE RATE OF ADVERSE EVENTS WAS LOW IN THIS HIGHER RISK; DIFFICULT TO MANAGE POPULATION OF SPECIAL NEEDS CHILDREN. Manufacturer Narrative: THIS MDR IS A DUPLICATE OF MFR # 2023826-2021-01751. PLEASE DISREGARD THE CURRENT REPORT. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275445&pc=MTA | 2023826-2018-00228 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure | Cataract; Retinal Detachment; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V4 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENTS EYE 3 YEARS AGO, DUE TO INVERSE ICL, A RHEGMATOGENOUS RETINAL DETACHMENT AND ANTERIOR SUBCAPSULAR CATARACT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. THE REPORTER ATTRIBUTED THE RETINAL DETACHMENT TO HIGH MYOPIA. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V4 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENT'S EYE FOR 3 YEARS DUE TO INVERSE ICL; A RHEGMATOGENOUS RETINAL DETACHMENT AND ANTERIOR SUBCAPSULAR CATARACT. PHACOEMULSIFICATION SURGERY WAS PERFORMED AND AN IOL (INTRAOCULAR LENS) WAS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704662&pc=MTA | 2023826-2015-00424 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED DISCOMFORT, LIGHT SENSITIVITY AND ANTERIOR UVEITIS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704665&pc=MTA | 2023826-2015-00425 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Cataract; Uveitis; Halo; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. DURING EXPLANTATION, ANTERIOR SUBCAPSULAR OPACITIES SUDDENLY DEVELOPED. THE PATIENT EXPERIENCED PIGMENT DISPERSION, GLARE/HALOS, INFLAMMATION, ANTERIOR UVEITIS, DISCOMFORT, LIGHT SENSITIVITY, PAIN AND TEARING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT IS BEING MONITORED AND IS DOING BETTER IN TERMS OF EYE DISCOMFORT BUT HAS COMPLAINED OF BLURRED VISION. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577133&pc=MTA | 2023826-2017-00836 | 2017-04-18 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO DEVELOPMENT OF A CATARACT OR COULD BE A FILM BETWEEN NATURAL LENS AND ICL OR THE START OF CATARACTS. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE DATE OF CATARACT DIAGNOSIS WAS (B)(6) 2017 AND THE TYPE OF CATARACT WAS ANTERIOR SUBCAPSULAR. THE REPORTER INDICATED THE EVENT WAS RELATED TO THE DEVICE. NO MEDICATIONS OR TREATMENTS WERE REQUIRED. IMPLANTED DATE: (B)(6) 2011. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6426147&pc=MTA | 2023826-2017-00471 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED EYE IS WORSE FOR READING. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY REPORTED THEY CONFIRMED THE PATIENT REPORT. DATE OF LAST VISIT WAS (B)(6) 2016 AND REPORTED PATIENT HAD PRESBYOPIA, WITH A MILD CATARACT. NO MEDICATIONS OR TREATMENTS WERE REQUIRED AND THE EVENT WAS NOT RELATED TO THE DEVICE. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546330&pc=MTA | 2023826-2017-00772 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED REDUCED VISUAL ACUITY AND WAS WEARING GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011. THE FACILITY REPORTED THE PATIENT HAD DECREASED VISUAL ACUITY DUE TO A CATARACT HAD DEVELOPED. SURGERY IS PLANNED AT A FUTURE DATE FOR REMOVAL OF CATARACT. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE FACILITY REPORTED THE PATIENT HAD LASIK ON (B)(6) 2011 AND WAS PRESCRIBED MEDICATION FOR NIGHT DRIVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323983&pc=MTA | 2023826-2019-00238 |  | Injury | STAAR SURGICAL COMPANY AG | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2009. THE REPORTER INDICATED A CATARACT HAS DEVELOPED AND VISION IS SLOWLY CLOUDING. THE PLAN IS TO EXPLANT THE LENS BUT CURRENTLY THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: REPORT # MW5082245. AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. OCCUPATION: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17872920&pc=MTA | 2023826-2023-04387 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS REMOVED AND REPLACED FOR A SMALLER LENGTH LENS DUE TO LENS OPACITY. THE PROBLEM WAS RESOLVED. REPORTEDLY, "QUIET ICL VAULT GOOD." Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2140" SHOULD BE REMOVED AND "1766" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3001" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS OPACITY (CATARACT). ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "QUIET. ICL VAULT GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873747&pc=MTA | 2023826-2018-01363 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01364 FOR FELLOW EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7873772&pc=MTA | 2023826-2018-01364 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A SHALLOW VAULT. THE PATIENT RETURNED FOR A POST-OP VISIT A COUPLE OF WEEKS AGO WITH WORSENING VISION AND A LOT OF HAZE. THE PATIENT HAD ANTERIOR SUBCAPSULAR CATARACTS OU, OD IS WORSE THAN OS. OD IS CORRECTABLE TO 20/25 AND OS IS ONLY SUPERIOR SO NOT SENSITIVE TO CATARACT QUITE YET. THE PATIENT IS NOT HAPPY WITH HER VISION AND IS BEING MONITORED. THE LENS REMAINS IMPLANTED. SEE MFR. REPORT # 2023826-2018-01363 FOR FELLOW EYE. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8678980&pc=MTA | 2023826-2019-00867 | 2019-04-23 | Injury | STAAR SURGICAL COMPANY | 2019-06-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances | UNK | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM THE FDA - # MW5086195. THE PATIENT REPORTED THAT THEY HAD VISIAN ICL PROCEDURE EARLY 2017. SINCE THEN PATIENT HAS DEVELOPED SIGNIFICANT CATARACTS WHICH REQUIRES AN IOP PROCEDURE TO CORRECT. CURRENTLY THE PATIENT IS EXPERIENCING LOSS OF CONTRAST IN BRIGHT LIGHT, IMPAIRED DEPTH PERCEPTION, REDUCED FIELD OF VISION AND REDUCED NIGHT VISION. REPORTEDLY THE SURGEON WANTS TO PERFORM ADDITIONAL IOL PROCEDURE TO CORRECT THE OUTCOME OF THE ICL PROCEDURE. THE REPORTER INDICATED THAT VYVANSE, FINASTERIDE, TRAZODONE, AND VITAMIN D, B12, MAGNESIUM, ZYRTEC ARE CURRENTLY BEING TAKEN. NO FURTHER INFORMATION PROVIDED TO FOLLOW UP ON. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082685&pc=MTA | 2023826-2021-04921 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cataract; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED FOR "URRETS-ZAVALIA SYNDROME AFTER IMPLANTATION OF A PHAKIC INTRAOCULAR LENS". THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS OVERSIZED AND WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS LENGTH. A FEW WEEKS LATER, THE PATIENT HAD AN ATONIC, MID-DILATED PUPIL, WITH NO DIRECT OR CONSENSUAL RESPONSE TO LIGHT AND A PRONOUNCED ANTERIOR AND POSTERIOR SUBCAPSULAR CATARACT. THE PATIENT COMPLAINED OF VISUAL DISCOMFORT, WITH SIGNIFICANT GLARE SENSITIVITY THE LENS WAS EXPLANTED AND PHACOEMULSIFICATION WAS PERFORMED , WITH THE IMPLANT OF 2 OFFSET ANIRIDIA RINGS. AN IOL LENS WAS IMPLANTED. MEDICATION WAS ADMINISTERED. FOUR WEEKS POST-OP, THERE WAS A SIGNIFICANT REDUCTION OF GLARE AND IOP WAS CONTROLLED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6863580&pc=MTA | 2023826-2017-01430 | 2017-08-08 | Injury | STAAR SURGICAL COMPANY AG | 2017-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Cataract; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -15.0 DIOPTER, ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO A CATARACT, AND REPLACED WITH AN INTRAOCULAR LENS (IOL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN SPECIMEN CUP WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE OF HAPTIC AND CLEAR RESIDUE ON LENS. LENS HAS AN "80" STAMPED /EMBEDDED. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362600&pc=MTA | 2023826-2019-00300 |  | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IS STILL UNHAPPY WITH HIS VISION. THE DOCTOR FOUND THE START OF CATARACTS AT HIS LAST EXAM AND RECOMMENDED REMOVING THE LENS AT THE SAME TIME. PATIENT HAS SOUGHT A 2ND OPINION. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT - CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. (DEVICE MANUFACTURER DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8362754&pc=MTA | 2023826-2019-00301 |  | Injury | STAAR SURGICAL COMPANY | 2019-02-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Cataract; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAS A PATIENT WHO HAS NOT BEEN HAPPY WITH HIS VISIAN ICL. PER REPORTER, PATIENT COMPLAINS OF A WHITE BAR OR STREAK THAT COMES ACROSS HIS VISION WHEN HIS PUPIL IS DILATED. IT SEEMS TO BE TRIGGERED WHEN LOOKING AT COMPUTER SCREENS OR OUTDOOR LIGHTS. HAVE TRIED PUPIL REDUCING DROPS WITHOUT SUCCESS. PATIENT IS BOTHERED SO MUCH BY THIS HE WANTS THESE REMOVED. HE HAD HIS IMPLANTS BACK IN (B)(6) OF 2018. THE IMPLANTS ARE BOTH MICL 13.2 -11.50. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT- CLOSURE OF THE IRIDOTOMY HAS BEEN PERFORMED AND THE PATIENT IS NOW WAITING HEALING. PATIENT CODE 3191: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: UNK, NO SERIAL NUMBER REPORTED. (IMPLANTED DATE): (B)(6) OF 2018. UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8555318&pc=MTA | MW5086195 | 2019-04-23 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 | MTA | VISIAN ICL | Adverse Event Without Identified Device or Use Problem | Cataract; Visual Impairment |  | 1 | Event Description: HAD VISIAN ICL PROCEDURE PERFORMED BY (B)(6) IN EARLY 2017. HAVE SINCE DEVELOPED SIGNIFICANT CATARACTS WHICH REQUIRE AN IOP PROCEDURE TO CORRECT. CURRENTLY EXPERIENCING LOSS OF CONTRAST IN BRIGHT LIGHT, IMPAIRED DEPTH PERCEPTION, REDUCED FIELD OF VISION, AND REDUCED NIGHT VISION. AFTER FULL EXAM, SURGEON IS QUOTING ADDITIONAL (B)(6) OUT OF POCKET TO PERFORM IOL PROCEDURE TO CORRECT OUTCOME OF ICL PROCEDURE. FDA SAFETY REPORT ID # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791357&pc=MTA | 2023826-2018-01238 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Cataract; Vitrectomy; No Code Available | UNK | 1 | Event Description: ARTICLE STATED THAT THE PATIENT'S PAST OCULAR HISTORY WAS SIGNIFICANT FOR PLACEMENT OF VISIAN ICL SEVERAL YEARS EARLIER FOR THE TREATMENT OF HIGH MYOPIA. SHORTLY AFTER IMPLANTATION, THE RIGHT ICL DISLOCATED, REQUIRING REMOVAL, AT WHICH TIME THE CATARACT WAS ALSO EXTRACTED AND A POSTERIOR CHAMBER IOL WAS PLACED. SEVERAL MONTHS LATER, THE IOL ALSO DISLOCATED, REQUIRING REMOVAL AND A PARS PLANA VITRECTOMY, LEAVING THE RIGHT EYE (OD) APHAKIC. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: NO SERIAL NUMBER REPORTED (B)(4). SECONDARY SURGICAL INTERVENTION, EXPLANT AND CATARACT SURGERY WITH IOL IMPLANTATION. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9233429&pc=MTA | 2023826-2019-02023 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Material Opacification; Adverse Event Without Identified Device or Use Problem | Cataract; Vitreous Floaters; Retinal Detachment; Blurred Vision; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: RECEIVED A VOLUNTARY EVENT REPORT FROM FDA: # MW5089673. THE PATIENT REPORTED ON (B)(6)2019, A "SERIOUS INJURY" REPORT WITH OUTCOME OF "DISABILITY/PERMANENT DAMAGE". THE PATIENT STATED "GOT VISIAN ICL PLASTIC LENS IMPLANTS IN BOTH EYES. IN LEFT EYE OVER A YEAR LATER I HAD SYMPTOMS OF FOGGY VISION, NEW FLOATERS AND TODAY MY VISION IS ALMOST GONE. THERE'S A HUGE BLOB OF WATER BLOCKING MY SIGHT. WHEN I MOVE MY EYE I CAN SEE THE WATER MOVE." PER THE SAME REPORT CONCOMITANT MEDICAL PRODUCTS LISTED AS: "ALPHAGAN AND PRENATAL VITAMINS", WITH PATIENT PROBLEMS OF: "IMPAIRED VISION, VITREOUS FLOATERS" AND DEVICE PROBLEMS: "ADVERSE EVENTS WITHOUT IDENTIFIED DEVICE OR USE PROBLEM" WAS REPORTED. PER PATIENT EMAIL OF 30 SEP 2019 "VISIAN ICL WAS REMOVED ON THE LEFT EYE (B)(6) TO PREPARE FOR VITRECTOMY FOR RETINAL DETACHMENT REPAIR". INFORMATION RECEIVED FROM THE IMPLANTING SURGEON INDICATED THAT A 13.2MM, MICL13.2, -16.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017 AND REMOVED ON (B)(6) 2019. LOSS OF BCVA, LENS OPACITY (NS), RETINAL DETACHMENT, BLURRED VISION, AND FLOATERS WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION/TREATMENT (PRK, ICL REMOVAL + CE/IOL, RETINAL REPAIR) WERE PERFORMED. ON PATIENTS LAST VISIT ON (B)(6) 2019, THE PATIENTS BCVA WAS 20/80-, UCVA 20/200. CLEAR CORNEA, SUPERIOR PI, PC/IOL WITH MILD TO MODERATE PCO, S/P VITRECTOMY, 45 % GAS, AND RETINA ATTACHED WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). PER RETINAL DETACHMENT REPAIR SURGEON AND ICL REMOVAL AND CATARACT SURGERY SURGEON THIS INCIDENT DID NOT INVOLVE AN ADVERSE EVENT RELATED TO THE VISIAN ICL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRK TREATMENT WAS PERFORMED ON (B)(6) 2018. ICL REMOVAL AND CE/IOL PROCEDURES WERE COMPLETED DURING ONE SURGICAL EVENT BY A 2ND SURGEON. RD REPAIR WAS PERFORMED BY A 3RD SURGEON ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE: 1426 SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT # 2 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906066&pc=MTA | 2023826-2018-01390 |  | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Clouding, Central Corneal; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM513.2, -14.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2018 PATIENT WAS OBSERVED WITH SOME CENTRAL CORNEAL EDEMA WITH NO INFLAMMATION OR NO INFECTION. PATIENT IS RECOVERING WELL, AND WILL BE REVIEWED IN NEXT APPOINTMENT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN SPECIMEN CUP WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. ON (B)(6) 2018 THE PATIENT WAS OBSERVED WITH SOME CENTRAL CORNEAL EDEMA WITH NO INFLAMMATION OR NO INFECTION. PATIENT IS RECOVERING WELL, AND THE CENTRAL CORNEAL EDEMA HAS BEEN RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298841&pc=MTA | 2023826-2017-00208 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Conjunctivitis; Glaucoma; Inflammation; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HIS RIGHT EYE (OD). THE PATIENT REPORTED WAS PRESCRIBED ORAL MEDICATIONS/EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO THE INSIDE OF HIS EYE WAS INFLAMED. ALSO WAS PRESCRIBED EYE DROPS FOR 3 CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THEY WERE UNABLE TO CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2015. THE PATIENT HAD DEVELOPED VIRAL CONJUNCTIVITIS AT 4 MONTHS POST-OP AND THIS WAS NOT RELATED TO THE ICL. THE FACILITY REPORTED THE PATIENT HAD NO HISTORY OF ELEVATED IOP OR GLAUCOMA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: PHYSICIAN'S REPORT STATED ALL PATIENT REPORTED ADVERSE EVENTS DID NOT OCCUR. CONJUNCTIVITIS DID OCCUR BUT NOT DEVICE RELATED. CORRECTED DATA: MFR DATE: THE CORRECTED DATE IS 02/28/2013. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6298876&pc=MTA | 2023826-2017-00209 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Conjunctivitis; Glaucoma; Inflammation; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HIS LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED ORAL MEDICATIONS/EYE DROPS FOR AT LEAST 3 CONSECUTIVE MONTHS DUE TO THE INSIDE OF HIS EYE WAS INFLAMED. ALSO WAS PRESCRIBED EYE DROPS FOR 3 CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED THEY WERE UNABLE TO CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2015. THE PATIENT HAD DEVELOPED VIRAL CONJUNCTIVITIS AT 4 MONTHS POST-OP AND THIS WAS NOT RELATED TO THE ICL. THE FACILITY REPORTED THE PATIENT HAD NO HISTORY OF ELEVATED IOP OR GLAUCOMA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CONCLUSION: PHYSICIAN'S REPORT STATED ALL PATIENT REPORTED ADVERSE EVENTS DID NOT OCCUR. CONJUNCTIVITIS DID OCCUR BUT NOT DEVICE RELATED. MFG DATE: THE CORRECTED DATE IS 02/28/2013. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11934181&pc=MTA | 2023826-2021-01747 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Abrasion | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO13.7, -01.75 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT REFRACTIVE CHANGE OVERTIME. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "THERE IS STILL AN EROSION ON CENTRAL POSITION OF THE CORNEA".ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT THE EROSION IS GONE AND THAT THE CORNEA IS FINE. LAST EXAM WAS (B)(6) 2021 - RX: +0,25/-0,25/112 UCVA 20/20. IT WAS REPORTED THAT THE ICL WAS NOT THE CAUSE OF THE EROSION. CLAIM # (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894693&pc=MTA | 2023826-2023-01550 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED TITLED "DESCEMET'S MEMBRAN EDOTHELIAL KERATOPLASTY FOR CORNEAL EDOTHELIAL FAILURE SECONDARY TO THREE TYPES OF PHAKI INTERAOCULAR LENS - RETROSPECTIVE STUDY", NOTED TWO EYES (12.5%) HAD POSTERIOR-CHAMBER COLLAMER PIOLS (PC-PIOL). THE PURPOSE OF THE STUDY WAS TO REPORT CORNEAL ENDOTHELIAL FAILURE SECONDARY TO PHAKIC INTRAOCULAR LENS IMPLANTATION (PIOL). THE ARTICLE CONCLUDES CLINICALLY SIGNIFICANT VISUAL IMPROVEMENT, AND WITH A HIGH PROPORTION OF EYES REACHING HIGH LEVELS OF BCVA AND VERY GOOD REFRACTIVE OUTCOMES. REPORTEDLY, BOTH EYES PRESENTED WITH CORNEAL DECOMPENSATION. "ONE PRESENTED TO US ALREADY PSEUDOPHAKIC AND WITH CORNEAL DECOMPENSATION AND WE HAVE NO INFORMATION OF THE LENS THAT HAD BEEN IMPLANTED (THE FELLOW EYE HAD AN ICL WITH A CENTRAL PORT), AND THE OTHER PATIENT HAD AN ICL OF -16.00D." Manufacturer Narrative: H6- HEALTH EFFECT IMPACT CODE 4581: ENDOTHELIAL CELL FAILURE FOLLOWING PIOL WITH SUBSEQUENT KERATOPLASTY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11944799&pc=MTA | 8040449-2021-00003 | 2021-04-14 04:00 | Injury | OPHTEC BV | 2021-06-07 04:00 | MTA | ARTISAN MYOPIA 6/8.5 | Output Problem | Corneal Decompensation | P030028 | 1 | Event Description: AN ARTISAN MYOPIA PIOL WAS IMPLANTED ON (B)(6) 2013. DURING A CONSULT ON (B)(6) 2020, IT WAS OBSERVED THAT THE CORNEA WAS DECOMPENSATED IN THE LEFT EYE REQUIRING A CATARACT SURGERY WITH IOL EXPLANTATION, COMBINED WITH A CORNEA TRANSPLANTATION. PRIOR TO THE EXPLANTATION, THE ENDOTHELIAL CELL COUNT WAS 1000 CELLS/MM2. A CORNEAL TRANSPLANT WAS PERFORMED (DSAEK). THE LENS WAS REPLACED WITH A REGULAR CATARACT IOL.Event Description: AN ARTISAN MYOPIA PIOL WAS IMPLANTED ON (B)(6) 2013. DURING A CONSULT ON (B)(6) 2020 IT WAS OBSERVED THAT THE CORNEA WAS DECOMPENSATED IN THE LEFT EYE REQUIRING A CATARACT SURGERY WITH IOL EXPLANTATION COMBINED WITH A CORNEA TRANSPLANTATION. PRIOR TO THE EXPLANTATION, THE ENDOTHELIAL CELL COUNT WAS 1000 CELLS/MM2. Manufacturer Narrative: TRACEABILITY REVIEW SHOWS THAT THE PRODUCT WAS MANUFACTURED IN ACCORDANCE WITH SPECIFICATIONS. NO OTHER EVENTS WERE REPORTED FOR PRODUCTS FROM THIS LOST. NO ADDITIONAL CLINICAL INFORMATION IS AVAILABLE. Manufacturer Narrative: THE RETURNED PRODUCT WAS INVESTIGATED. PRODUCT MEETS SPECIFICATIONS. AS STATED IN THE INITIAL REPORT, REVIEW OF PRODUCT DATA (TRACEABILITY REVIEW) SHOWS THAT PRODUCT WAS MANUFACTURED IN ACCORDANCE WITH SPECIFICATIONS. NO OTHER EVENTS WERE REPORTED REGARDING PRODUCTS FROM THIS LOTS. ADDITIONAL INFORMATION, INCLUDING CLINICAL INFORMATION, WAS REQUESTED IN ORDER TO ANALYZE THE CASE IN DETAIL. NO FURTHER INFORMATION WAS PROVIDED. NO RELEVANT INFORMATION, INCLUDING OTHER PREVIOUS ECC COUNTS, BIOMETRY DATA OF THE EYE, COULD BE PROVIDED. THEREFORE THE ROOT CAUSE COULD NOT BE DETERMINED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791689&pc=MTA | 2023826-2018-01181 | 2018-07-10 | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Use of Device Problem; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Corneal Decompensation; Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/2.0/129 (SPHERE/CYLINDER/AXIS) IMPLANTBLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. IT WAS REPORTED THAT THE SURGEON REMOVED THE LENS AND IMPLANTED IT AGAIN. ONCE THE LENS WAS INSIDE THE EYE, THE SURGEON REALIZED THAT THE LENS WAS DAMAGED AND REMOVED THE LENS AGAIN. CORNEAL EDEMA AND CORENAL DECOMPENSATION WERE OBSERVED. THE REPORTER STATED THAT THE CAUSE OF THE ADVERSE EVENTS WAS THE SURGEON/SURGICAL PROCEDURE. THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE EVENT. PATIENT IS AWAITING ICL REPALCEMENT AND THE EYE IS OKAY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INPSECTION FOUND THE HAPTIC TORN AND DEFORMED; AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: NEW INFORMATION: "SURGEON IMPLANTED THE NEW LENS (SAME PARAMETERS AS THE FIRST ONE SINCE THE PROBLEM WAS THAT HE BROKE THE LENS) VTICMO 13.2 -18/+2.0/108 . PATIENT'S AV IS SATISFYING, VAULT IS NORMAL". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8359111&pc=MTA | 2023826-2019-00298 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2, 3.50/4.00/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP WITH ANGLE CLOSURE, CORNEAL DECOMPENSATION AND UNREACTIVE PUPIL. IT WAS REPORTED THAT AT 1 DAY POST-OP, "PAINFUL ACUTE GLAUCOMA DEVELOPED WITH AN INTRAOCULAR PRESSURE BEYOND MEASUREABLE VALUE". ON DAY 3, THE LENS WAS EXPLANTED. 2 DAY POST-EXPLANT BCVA WAS 20/100 WITH IOP OF 22MMHG (BASELINE WAS 21MMHG) AND " UNREACTIVE DILATED PUPIL SEVERE CORNEAL EDEMA". ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2019: "THE PATIENTS PUPIL LOOKS LIKE FIXED AND HE MAY NEED A SECOND OPERATION FOR IT. YES, UNREACTIVE DILATED PUPIL AND SEVERE CORNEAL EDEMA HAVE RESOLVED. NO , THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. YES, PRE-OP PERIPHERAL IRIDOTOMIES WERE PERFORMED ONE DAY BEFORE SURGERY. 2 IRIDOTOMIES HAS BEEN DONE AT 11:00 O'CLOCK AND 13:30 O'CLOCK". Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159605&pc=MTA | 2023826-2018-01833 | 2018-11-07 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT; CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, INFLAMMATION, AND SHALLOWING OF ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM. (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159607&pc=MTA | 2023826-2018-01834 | 2018-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Decompensation; Corneal Edema; Inflammation; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, LEADING TO CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WAS REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION ( AC WASHOUT, ICL "REPOSTIONING", REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM, AND MILD MYDRIASIS, LIKELY CONTRIBUTING TO GLARE COMPLAINT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS OF -7.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. EXCESSIVE VAULT, CORNEAL DECOMPENSATION, UNCONTROLLED IOP, NARROWING OF THE ANGLES, CORNEAL EDEMA, AND SHALLOWING OF THE ACD WAS OBSERVED. PAIN AND BLURRY VISION WERE REPORTED AS ADDITIONAL SYMPTOMS. ADDITIONAL SURGICAL INTERVENTION (AC WASHOUT, ICL REPOSITIONING, REPEAT YAG PI) WAS PERFORMED. IT WAS REPORTED THAT IOP LOWERING DROPS AND PILLS WERE PRESCRIBED. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE EXPLANT RESOLVED THE HIGH VAULT AND IOP. PATIENT IS DOING BETTER BUT DOES HAVE PIGMENT ON ENDOTHELIUM AND MILD MYDRIASIS LIKELY CONTRIBUTING TO THE GLARE COMPLAINT. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, NARROWING OF ANGLES, SHALLOWING OF ACD, AC WASHOUT, ICL REPOSITIONING, YAG PI, IOP LOWERING DROPS AND PILLS, EXPLANT, MYDRIASIS (B)(4). Manufacturer Narrative: (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11231056&pc=MTA | 2023826-2021-00199 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. POST-OP, THERE WAS SEMI-MYDRIASIS (DILATION OF PUPIL) AND CORNEAL EDEMA. A POSSIBLE CAUSE OF THE EVENT WAS OVERSIZE LENS (ICL). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: AT 2 MONTHS POST-OP, THERE IS NO MORE CORNEAL EDEMA OR HIGH IOP, ONLY SEMI-MYDRIASIS. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMEMSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878560&pc=MTA | 2023826-2022-00910 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICM5_13.2; -10.50 DIOPTER; INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED NON-CONFORMITY - UNRESPONSIVE HYPETONUS DESPITE SURGICAL PROCEDURE AND COMPLICATIONS. THE PATIENT EXPERIENCED LOSS OF BCVA; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; CORNEAL DECOMPENSATION; BLURRED VISION. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THE POSTOPERATIVE NOTES INDICATED MILD TRANSPARENT CRYSTALLINE; CORNEAL EDEMA "ABOUT PRESENT PUPIL REACTIVE WAITING TO PERFORM VISUAL FIELD". IT WAS ALSO NOTED THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7791893&pc=MTA | 2023826-2018-01244 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-08-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/+3.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO HIGH IOP, MID DILATED PUPIL, CORNEAL EDEMA, CLOSE ANGEL GLAUCOMA AND ANGLE BLOCK. THE SURGEON PERFORMED A PARS PLANA VITRECTOMY, PHACO-EMULSIFICATION (CATARACT SURGERY) WITH IOL IMPLANTATION AND IRIS SUTURE. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP VISIT VISUAL ACUITY WAS 20/40 WITHOUT CORRECTION, IOP WAS NORMAL WITH A REGULAR PUPIL. Manufacturer Narrative: CORRECTED DATA: ADDED 'OTHER SERIOUS' TO INITIAL REPORT. REPORTER ALSO INDICATES CORNEAL DECOMPENSATION. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6229803&pc=MTA | 2023826-2017-00020 | 2016-12-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/0.96 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND ELEVATED IOP. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/096 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, PIGMENT DISPERSION, AND CORNEAL EDEMA. ON (B)(6) 2016 THE LENS WAS EXPLANTED. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATION PENDING.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7109978&pc=MTA | 2023826-2017-01877 | 2017-10-30 | Injury | STAAR SURGICAL COMPANY | 2017-12-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.0/+3.5/171 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON REPORTS EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, AND CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: REPORTER STATED PATIENT IS OKAY. THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: 24-OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0 MM (2.92 MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13529754&pc=MTA | 2023826-2022-00498 | 2022-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Glaucoma; Intraocular Pressure Increased; Pain; Pupillary Block; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -2 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; PUPIL BLOCK; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; CORNEAL DECOMPENSATION; MEDICATION USED; PAIN; GLAUCOMA AND IT WAS NOTED THAT YAG SURGERY WAS PERFORMED AS WELL AS PARACENTESIS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. POST OP RESULTS REPORTED THE PROBLEM WAS RESOLVED AND WITH THE LENS REMOVED THE PATIENT NOW HAS CORNEAL EDEMA; MID DILATED PUPIL; IRIS TRAUMA. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "ICL SEEMS TO BIG FOR THIS EYE DESPITE CORRECT MEASUREMENT AND CALCULATION PREOPERATIVELY". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID, WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED CORRECTED VA IS 6/6 WITH SOME DEGREE WITH IRIS ATROPHY. H6 - HEALTH EFFECT CLINICAL CODE: 4581- IRIS ATROPHY. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7899789&pc=MTA | 2023826-2018-01401 | 2017-06-24 | Injury | STAAR SURGICAL COMPANY | 2018-09-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Inflammation; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. AT ONE DAY POST-OP, INFLAMMATION (TASS) PRESENTED IN THE EYE. CORTISONE WAS ADMINISTERED AND THE INFLAMMATION CLEARED UP. THE PATIENT ALSO EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND CORNEAL DECOMPENSATION. THE LENS WAS EXPLANTED ON (B)(6) 2018. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE POST-OP CONDITION OF THE PATIENT'S EYE IS AS FOLLOWS - SYNECHIAE OF CHAMBER ANGLE (PARTIAL), SOME SMALL SUBCAPSULAR VACUOLES IN PERIPHERY OF CRYSTALLINE LENS AND SLIGHT MYDRIASIS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON THE LENS. (B)(4). DEVICE HISTORY RECORD REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6540579&pc=MTA | 2023826-2017-00740 | 2017-01-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE DR. SAID THE PROBLEM IS SLOWLY GETTING BETTER. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS EXCHANGED RESOLVED THE PROBLEM FOR THE PATIENT. LENS RETURNED IN LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, DEBRIS AND CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, IRIS DISTORTION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, CORNEAL DECOMPENSATION, UNREACTIVE PUPIL, ANGLE CLOSURE, GLARE/HALOES, BLURRED VISION, AND LOSS OF BCVA. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8529658&pc=MTA | 2023826-2019-00632 | 2019-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, PUPIL BLOCK WITH IOP (INTRAOCULAR PRESSURE), ANGLE CLOSURE, CORNEAL DECOMPENSATION, LENS DISLOCATION/SUBLUXATION, AND UNREACTIVE (FIXED) PUPIL. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. SURGEON IS NOT SURE YET IF THE EXCHANGE LENS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID , IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148265&pc=MTA | 2023826-2021-02424 | 2021-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Corneal Decompensation; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER STATES AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS HIGH VAULT, ELEVATED IOP, AND PRESSURE-LOWERING MEDICATION ADMINISTERED. Manufacturer Narrative: A 13.2MM MICL13.2 ICL WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM WAS NOT RESOLVED. SEE MFR REP# 2023826-2021-03007, FOR EXCHANGED LENS. THE SURGEON REPORTS ADDITIONAL SYMPTOMS: ANGLE CLOSURE, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), PUPIL BLOCK, CORNEAL DECOMPENSATION, BLURRED VISION, ENLARGEMENT OF PI, ADDITION OF PI-YAG. THE CAUSE IS RELATED AS PATIENT-RELATED: "LENS TOO LONG FOR PATIENT'S WTW." H6-HEALTH EFFECT CLINICAL CODE 4581 (ANGLE CLOSURE, SIGNIFICANT REDUCTION OF ICA). DEVICE CODE 1494 (ACD<3.0MM). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN SPECIMEN CUP, DRY, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DEBRIS THROUGHOUT THE LENS AND RESIDUE ON THE HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT - UNK. DEVICE MANUFACTURE DATE - UNK CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125559&pc=MTA | 2023826-2021-02334 | 2020-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Intraocular Pressure Increased; Retinal Detachment; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE (IOP), CORNEAL DECOMPENSATION AND RETINAL DETACHMENT. PHACO-EMULSIFICATION (CATARACT SURGERY), IOL IMPLANTATION AND RETINAL REPAIR WAS PERFORMED. THE PROBLEM WAS RESOLVED AND THE RETINA IS IN PLACE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS -18.00 DIOPTER IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO RETINAL DETACHMENT. PHACO-EMULSIFICATION (CATARACT SURGERY); IOL IMPLANTATION AND RETINAL REPAIR WAS PERFORMED. THE PROBLEM IS RESOLVED. H6 - HEALTH EFFECT CLINICAL CODE 4581,1937,1790 IS NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7300916&pc=MTA | 2023826-2018-00281 | 2017-12-13 | Injury | STAAR SURGICAL COMPANY | 2018-02-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Decompensation; Low Blood Pressure/ Hypotension; Inflammation; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), CORNEAL DECOMPENSATION, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, BLURRED VISION, INFLAMMATION, AND "CHOROIDAL DETACHMENT, SEVERE HYPOTENSION." REPORTER STATES THAT HYPOTENSION IS "BEING RESOLVED" AND "IN FACT THE PATIENT IS VERY GOOD." Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. MEASUREMENT OF THE HYDRATION FOUND MOISTURE CONTENT ON A HIGH LEVEL; BUT NOT EXCEEDING LIMITS. THE OVERALL LENS DIAMETER WAS THEN FOUND OUT OF TOLERANCE. THE LENS SN: (B)(4) WAS A 12.6MM IMPLANTABLE LENS THAT IS 0.22MM LARGER THAN THE THEORETICAL; THE TOLERANCE IS +/-0.20MM. SINCE THE WORK ORDER SEARCH DID NOT FIND ADDITIONAL SIMILAR ERRORS TO THIS COMPLAINT EVENT, IT CAN BE CONCLUDED THAT THE OOS CONDITION IS NOT LIKELY DUE TO MANUFACTURING OR PACKAGING. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: MEASUREMENT OF THE HYDRATION FOUND THE LENS TO BE OVERHYDRATED. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6659008&pc=MTA | 2023826-2017-01001 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Decompensation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, DIOPTER -3.5/+2/10 IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT BEGAN TO EXPERIENCE CORNEAL DECOMPENSATION, IRIS ATROPHY, IRIS HERNIA, LENS ROTATION, AND IRIS PROLAPSE. SECONDARY INTERVENTIONS INCLUDE ANTERIOR SUBINCISIONAL SINECHIOTOMY, ENLARGEMENT OF INCISION, ANTERIOR CHAMBER IRRIGATION, IRIS SUTURE, AND ADDITIONAL SUTURE. THE LENS REMAINS IMPLANTED AND IS UNLIKELY TO BE REMOVED, ACCORDING TO CLAIM REPORTER. THE INCISION WAS ENLARGED TO PERFORM THE IRIS SUTURE AND THE ANTERIOR CHAMBER IRRIGATION. ONCE PERFORMED, THE INCISION WAS SUTURED AND AN ASTIGMATISM DEVELOPED BECAUSE OF THE STITCHES' TENSION. THE SURGEON HAS REQUESTED MEDICAL CONSULTATION TO DETERMINE IF FURTHER ICL ROTATION (ICL HAD ALREADY ROTATED FROM ORIGINAL POSITION) WOULD COMPENSATE FOR THE ASTIGMATISM. Manufacturer Narrative: PREVIOUS: REPOSITIONING PENDING MEDICAL CONSULTATION. UPDATE: LENS WAS REPOSITIONED (B)(6) 2017. THE PROBLEM HAS BEEN RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072306&pc=MTA | 2023826-2021-02037 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Decompensation; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: A (B)(6) WOMAN PRESENTED WITH FOGGY VISION AND INTOLERANCE TO BRIGHT LIGHT SINCE 10 MONTHS WITH HER BEST-CORRECTED VISUAL ACUITY (BCVA) IN THE RIGHT EYE 20/30 . SHE UNDERWENT IMPLANTABLE COLLAMER LENS INSERTION 10 MONTHS AGO IN THE RIGHT EYE AND REPORTED BLURRY VISION FROM FIRST DAY POSTOPERATIVE WHICH DID NOT IMPROVE. FOGGY VISION, PHOTOPHOBIA, WATERING AND PAIN IN THE RIGHT EYE WORSENED OVER A PERIOD OF 7 MONTHS DURING WHICH SHE TOOK A BREAK FROM WORK AND WAS UNDER CONSERVATIVE MANAGEMENT WITH TOPICAL PRED-NISOLONE 1% IN TAPERING DOSE, MOXIFLOXACIN 0.5%AND HYPERTONIC SALINE NACL 5% DROPS. DMEK WAS PERFORMED AND AT 8 MONTHS SHE HAD A BCVA OF 20/20 AND ECD 1504 CELLS/MM2 WITH A REDUCTION OF 38.18% CELLS AND IOP OF 14 MM HG. THE CORNEA WAS CLEAR AND THE ICL WAS IN A NORMAL POSITION WITH ADEQUATE VAULT AND A CLEAR CRYSTALLINE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT HAD CORNEAL DECOMPENSATION. SHOULD HAVE ALSO BEEN INCLUDED. H6 - HEALTH EFFECT CLINICAL CODE 4581: PHOTOPHOBIA, ENDOTHELIAL CELL LOSS. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910957&pc=MTA | 2023826-2020-00694 | 2020-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT'S CONDITION AFTER IMPLANT IS MARKED AS "MILD CORNEAL EDEMA." REFERENCE MFR FILE# (B)(4) FOR INITIAL IMPLANT. Manufacturer Narrative: CORRECTED DATA: B1 - ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2 - OTHER SERIOUS (IMPORTANT MEDICAL EVENTS) SHOULD BE REMOVED AS IT IS N/A. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT- UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910961&pc=MTA | 2023826-2020-00693 | 2020-02-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT'S CONDITION AFTER IMPLANT IS MARKED AS "MODERATE CORNEAL EDEMA." REFERENCE MFR FILE# (B)(4) FOR INITIAL IMPLANT. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED. (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616805&pc=MTA | 2023826-2021-00844 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE REPORTER INDICATED THERE WAS PATIENT INJURY, THE PATIENT EXPERIENCED TRANSIENT VISUAL OBSCURATION DUE TO CORNEAL EDEMA. THE CORNEAL EDEMA DEVELOPED BECAUSE THE CORNEA WAS STRESSED WHEN THE INITIAL DAMAGED LENS WAS REMOVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND WRAPPED IN GAUZE WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN, AND FOREIGN DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558785&pc=MTA | 2023826-2021-00702 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11558800&pc=MTA | 2023826-2021-00701 |  | Injury | STAAR SURGICAL COMPANY | 2021-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS CORNEAL EDEMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15510501&pc=MTA | 2023826-2022-03374 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 AS REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. IT WAS REPORTED THAT THE PROBLEM RESOLVED BUT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) CORNEAL EDEMA WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-PER UPDATED INFORMATION THE CORNEAL EDEMA WAS NOTICED THE DAY AFTER SURGERY AND WAS TEMPORARY, RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7213982&pc=MTA | 2023826-2018-00140 | 2017-12-20 | Injury | STAAR SURGICAL COMPANY | 2018-01-23 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES CORNEAL EDEMA THAT WAS "BEYOND STANDARD AFTER OCULAR SURGERY." REPORTEDLY, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION STATES THAT ON THE FOLLOW UP VISIT (B)(6) 2017, THE CORNEA IS "RECOVERED TO TRANSPARENT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF-LABEL USE [LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8339700&pc=MTA | 2023826-2019-00259 | 2019-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse | Corneal Edema | P030016 | 1 | Event Description: THE SURGEON REPORTED THAT HE ATTEMPTED TO IMPLANT A 13.2MM, MICL13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE BUT IT WOULD NOT UNFOLD. " I NEVER COMPLETED THE INSERTION INTO THE EYE BECAUSE THE LENS WOULD NOT UNFOLD ABOUT 65% OUT OF THE INJECTOR AND I FELT IT NOT PRUDENT TO PROCEED GIVEN THE 3 ATTEMPTS TO ALLOW IT TO UNFOLD. THE LENS WAS NEVER FULLY IMPLANTED, THE LENS WAS ALLOWED TO COME OUT OF THE INJECTOR ABOUT 60-70% AND ALLOWED TO UNFOLD BUT IT NEVER DID. I USED MANEUVERS SUCH AS "JIGGLING" THE INJECTOR. I REMOVED THE INJECTOR EACH TIME, RELOADED, AND INSERTED 2 ADDITIONAL TIMES. NO LASTING INJURY TO THE PATIENT. JUST MORE CORNEAL EDEMA THAN USUAL, BUT IT RESOLVED AS EXPECTED. CAUSE OF CORNEAL EDEMA IS UNCLEAR, LIKELY SURGERY BUT GIVEN THE FACT THAT THE PRODUCT MADE THE SURGERY MORE COMPLICATED THAN NECESSARY, COULD BE PRODUCT AS WELL, BUT NOT DIRECTLY. NO, CORNEAL EDEMA DID NOT REQUIRE TREATMENT BEYOND WHAT IS STANDARD AFTER OCULAR SURGERY." Manufacturer Narrative: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575681&pc=MTA | 2023826-2016-00416 | 2016-03-07 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse; Torn Material | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS WOULD NOT UNFOLD. THE SURGEON ATTEMPTED TO UNFURL THE LENS BUT THE CHAMBER SHALLOWED. IT WAS DECIDED TO REMOVE THE LENS AND THE LENS WAS TORN. THE SURGEON INSERTED THE BACKUP LENS AND THE LENS WOULD NOT UNFOLD AND THE CHAMBER SHALLOWED. HE WAS ABLE TO UNFOLD THE LENS BUT NOTED THE LENS WAS UPSIDE DOWN. SEE MFR. REPORT # 2023826-2016-00417. THE LENS WAS REMOVED AND A 3RD LENS WAS IMPLANTED WITH NO PROBLEM. THERE WAS NO PATIENT INJURY WITH THE REMOVAL OF THE LENS BUT POST-OP THE PATIENT HAD CORNEAL EDEMA, WHICH IS IMPROVING. THE SURGEON INDICATED THE EVENT WAS DUE TO THE PATIENT'S ANATOMY AND WAS NOT PRODUCT RELATED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULT: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575683&pc=MTA | 2023826-2016-00417 | 2016-03-07 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Fold, Unfold or Collapse; Unintended Movement | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS WOULD NOT UNFOLD. THE SURGEON ATTEMPTED TO UNFURL THE LENS BUT THE CHAMBER SHALLOWED. THE SURGEON WAS ABLE TO UNFOLD THE LENS BUT NOTED THE LENS WAS UPSIDE DOWN. IT WAS DECIDED TO REMOVE THE LENS AND A 3RD LENS WAS IMPLANTED WITH NO PROBLEM. THIS WAS THE 2ND OF 2 LENSES THAT WAS INSERTED INTO THE PATIENT'S EYE AND WOULD NOT UNFOLD - SEE MFR. REPORT # 2023826-2016-00416. THERE WAS NO PATIENT INJURY WITH THE REMOVAL OF THE LENS BUT POST-OP THE PATIENT HAD CORNEAL EDEMA, WHICH IS IMPROVING. THE SURGEON INDICATED THE EVENT WAS DUE TO THE PATIENT'S ANATOMY AND WAS NOT PRODUCT RELATED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11934936&pc=MTA | 2023826-2021-01752 | 2021-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -5.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTER OPINON THE CAUSE OF THIS EVENT WAS THE DEVICE " ICL WAS TOO LONG DESPITE STAAR SIZING RECOMMENDATION." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) " HEALING WELL, GOOD POSITION, MILD CORNEAL WOUND EDEMA CONTRIBUTING TO BLUR DAY 1 PST OPT" WAS REPORTED. IT WAS REPORTED THAT THE CORNEAL EDEMA DID NOT REQUIRE TREATMENT BEYOND WHAT IS STANDARD AFTER OCULAR SURGERY AND THAT AT PATIENTS 1 WEEK POST-OP APPOINTMENT THE EDEMA HAD RESOLVED. THE REPORTER STATED " CAUSE OF EDEMA WAS THE INCISION SITE WHICH IS VERY NORMAL." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION - LENS WAS RETURNED DRY AND IN MICROCENTRIFUGE VIAL. VISIBLE INSPECTION FOUND OPTIC SPLIT AND RESIDUE/DEBRIS ON PRODUCT. UNABLE TO PERFORM DIMENSIONAL INSPECTION AS THE RETURNED LENS IS SPLIT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13873873&pc=MTA | 2023826-2022-00852 | 2022-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AS AN EXCHANGE LENS ON (B)(6) 2022. TRANSIENT LOSS OF BCVA WAS OBSERVED. "PATIENT 3 DAY POST-OP CORNEA EDEMA, HE WILL IMPROVE AS EDEMA LOWERS." CAUSE IS REPORTED AS UNKNOWN. SEE MFR # 2023826-2022-00853 FOR INITIAL LENS. Manufacturer Narrative: H1: CORRECTED. CLAIM #(B)(4). Manufacturer Narrative: RACE-ETHNICITY: UNK. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15335115&pc=MTA | 2023826-2022-02953 | 2022-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.1MM, VICH12.1, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT (OD) EYE. THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY. PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. IT WAS REPORTED THAT ON (B)(6)2022, 3 DAYS AFTER SURGERY CORNEAL EDEMA WAS NOTED. HOWEVER, IT DID NOT REQUIRE TREATMENT BEYOND WHAT IS STANDARD AFTER OCULAR SURGERY. THE EDEMA WAS RELATED TO THE SURGERY ITSELF (REMOVAL AND REPLACEMENT) AND NOT THE LENS. Manufacturer Narrative: EXPIRATION DATE: 03/31/2025. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637149&pc=MTA | 2023826-2021-03959 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ACCIDENTALLY IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE UPSIDE DOWN. REPORTEDLY, THE LENS WAS IMMEDIATELY REMOVED; DUE TO CORNEAL EDEMA THE DOCTOR WANTS TO WAIT TO IMPLANT A REPLACEMENT LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT- UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE, EXPLANT DATE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE- UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870355&pc=MTA | 2023826-2016-01056 | 2016-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Torn Material; Appropriate Term/Code Not Available | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS TORE DURING INJECTION INTO THE EYES. THE LENS WAS REPOSITIONED AND REMOVED. THE PATIENT EXPERIENCED CORNEAL EDEMA. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PRODUCT PROBLEM TO ADVERSE EVENT: INITIAL SELECTION WAS INCORRECT; EVENT IS AN ADVERSE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513711&pc=MTA | 2023826-2015-00120 | 2015-01-20 | Malfunction | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS BUT THE LENS FLIPPED AND WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. AT THE PATIENT'S SIX DAY POST-OP, THE CORNEAL EDEMA WAS RESOLVED AND THE VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4) - CORNEAL EDEMA; UNINTENDED MOVEMENT. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. EVALUATION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - THIS COMPLAINT FILE STATES THAT THE LENS "FLIPPED UPSIDE DOWN DURING INSERTION". THERE IS NO INFORMATION TO DETERMINE IF THERE WAS ANY MALFUNCTION OF THE INSERTION SYSTEM, BUT A WORK-ORDER SEARCH SHOWED NO SIMILAR COMPLAINTS. THERE IS NO INFORMATION TO DETERMINE IF THE FLIP WAS DUE TO IMPROPER LENS LOADING, OR SURGEON HANDLING EITHER. THE PATIENT DEVELOPED CORNEAL EDEMA POSTOPERATIVELY, BUT IT CLEARED BY THE SIX-DAY VISIT. THERE IS NO REPORT OF INCREASED IOP OR REDUCED ECD. THE HAPTIC FOOT WAS VERY LIKELY SEVERED DURING EXPLANTATION. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES THAT COULD BE IDENTIFIED AS THE CAUSE OF THIS COMPLAINT. THE CAUSE OF THE LENS TO NOT UNFOLD CORRECTLY, IS POSSIBLY DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A POSSIBLE CAUSE OF THE EVENT MAY BE DUE TO IMPROPER LOADING OR HANDLING PRACTICES IN THE OPERATING ROOM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8632085&pc=MTA | 2023826-2019-00842 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema | N/A | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 12.1MM, VICM5 12.1, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS, "WAS FOLDED IN HALF AND DID NOT OPEN WELL." REPORTEDLY THE LENS UNFOLDED UPSIDE DOWN IN THE EYE, THE LENS TORE DURING REMOVAL AND THE LENS WAS EXCHANGED ON (B)(6) 2019, PROBLEM RESOLVED. NO PATIENT INJURY. ON 20/APR/2019 PATIENT HAS CORNEAL EDEMA. PER THE REPORTER ON (B)(6) 2019, PATIENT'S "CORNEAL EDEMA IS RESOLVED. PATIENT IS HAPPY NOW." CAUSE OF THE EVENT IS REPORTED AS "THE LENS IS TOO STICKY AND LESS SPACE IN ACD TO MANEUVER." Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10915417&pc=MTA | 2023826-2020-02806 | 2020-10-29 04:00 | Injury | STAAR SUGICAL COMPANY | 2020-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -12.0 DIOPTER, IMPLANTABLE COLLAMER LENS FLIPPED AND DOCTOR TOOK A BITE OUT OF IT WHILE EXPLANTING. A BACK UP LENS WAS IMPLANTED. CORNEAL EDEMA WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: G3: CORRECTED TO ON (B)(6) 2020 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007847&pc=MTA | 2023826-2023-04805 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -4.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED WITH PATIENT' INJURY - AS CORNEAL EDEMA OCCURRED AFTER THE LENS WAS REMOVED AND THE PATIENT DECIDED THEY DID NOT WANT TO WEAR ICLS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12686764&pc=MTA | 2023826-2021-04113 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | Corneal Edema | NA | 1 | Event Description: IT WAS REPORTED THAT A DOCTOR ACCIDENTALLY IMPLANTED AN ICL UPSIDE DOWN IN THE EYE AND THIS LENS WAS IMMEDIATELY REMOVED. BECAUSE THE PATIENT HAS CORNEAL EDEMA, THE DOCTOR WANTS TO WAIT BEFORE IMPLANTING THE LENS AGAIN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627866&pc=MTA | 2023826-2023-03723 | 2023-08-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THE PATIENT'S VISION IS FOGGY BUT BETTER THAN IT WAS 2 HOURS AGO, 1+ INJECTION, 1+ CORNEAL EDEMA. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - FOGGY VISION, +1 INJECTION. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180362&pc=MTA | 2023826-2022-02628 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. POSTOPERATIVE NOTES INDICATED THE PATIENT HAD VERY SLIGHT CORNEAL EDEMA; BUT IT WAS CLARIFIED THIS WAS A NORMAL POSTOP CONDITION AND THE PATIENT IS OKAY. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED, HAPTIC DEFORMED, AND HAPTIC STRETCHED OUT ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942421&pc=MTA | 2023826-2021-04638 | 2021-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS TRANSIENT LOSS OF BCVA AND THE PATIENT WAS UNHAPPY WITH THE RESULT (VA). REPORTEDLY THIS WAS DUE TO POSTOP CORNEAL EDEMA. THE PROBLEM HAS BEEN RESOLVED. POSTOP UCVA IS AS GOOD AS PREOP BCVA AND CORNEAL EDEMA IS GONE FROM THE EYE. "THE PATIENT HAS A SPECIAL ATTITUDE AND VERY HIGH EXPECTATIONS. SO SHE IS STILL NOT 100% HAPPY WITH THE COURSE OF THE SURGERY AND THE OUTCOME." Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042387&pc=MTA | 2023826-2018-01606 | 2018-09-26 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. INTRA-OPERATIVELY, THE LENS WAS EXPLANTED. THE SURGEON REPORTS BLURRED VISION AND CORNEAL EDEMA. ADDITIONALLY STATED, THE "LENS WAS INJECTED UPSIDE DOWN WHEN THE DOCTOR IMPLANTED. THE DOCTOR FOUND CORNEAL EDEMA WHEN WAS REMOVED THE LENS." THE REPORTER INDICATED THAT THE LENS WAS INJECTED INTO THE PATIENT'S EYE IN SUCH A WAY THAT THE IRIS-FACING LENS SURFACE WAS INSERTED FACING THE CAPSULE. ALSO INDICATED IS, "THREE MINUTES HAS ELAPSED FROM THE LOADING OF THE LENS TO THE INJECTION. DURING THIS TIME THE CARTRIDGE WAS KEEPING UNDERWATER. THE INCISION SIZE WAS NARROW." THE SURGEON INDICATES THAT THE CAUSE OF THE EVENT IS NOT KNOWN. THE CORNEAL EDEMA IS RESOLVING. Manufacturer Narrative: "THE DOCTOR HAS IMPLANTED A NEW LENS ON (B)(6) 2018 BECAUSE THE PATIENT'S CORNEAL CONDITION (BLURRED VISION) HAS BEEN IMPROVED". THE PROBLEM WAS RESOLVED AFTER THE NEW LENS WAS IMPLANTED. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS FOREIGN MATERIAL (FIBER) ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17393103&pc=MTA | 2023826-2023-03145 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Corneal Edema; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER TORE DURING INJECTION, INTRAOPERATIVELY THE LENS WAS REMOVED WITH PATIENT INJURY OF POSTOP CORNEAL EDEMA. SURGEON ALSO NOTED PATIENT HAD BLURRED VISION. A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED ON THIS SAME DATE. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND THERE IS DEVICE CAUSALITY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10749903&pc=MTA | 2023826-2020-02525 | 2020-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-10-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Corneal Edema; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -6.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS REFRACTIVE SURPRIS, IRIS PROLAPSE, EDEMA, AND IRIS HERNIA. REPORTEDLY EXCESSIVE MANIPULATION OCCURRED DURING IMPLANT SURGERY WHICH PROBABLY LED TO THE SYMPTOMS. 1-WEEK POSTOP THE UCVA IS 20/20 AND THE SYMPTOMS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378746&pc=MTA | 2023826-2018-00466 | 2018-02-26 | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -5.5/+3.5/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REFRACTIVE SURPRISE WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOPLEGIC REFRACTION FOR THE LENS CALCULATION. LENS REMAINS IMPLANTED. LENS EXCHANGE IS PLANNED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, IMPLANTABLE COLLAMER LENS OF -5.5/+3.5/066 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. REFRACTIVE SURPRISE ALONG WITH BLURRED VISION WAS NOTICED IN THE PATIENT DUE TO USER ERROR. THE SURGEON USED THE MANIFEST REFRACTION INSTEAD OF THE CYCLOGENIC REFRACTION FOR LENS CALCULATION. LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS "LITTLE PERIPHERAL CORNEAL OEDEMA". EXPLANTED DATE: (B)(6) 2018. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267530&pc=MTA | 2023826-2018-00214 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/+2.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE AND CORNEAL EDEMA; INCIDENT DATE STATED AS (B)(6) 2018. REPORTEDLY, THE CORNEAL EDEMA DISAPPEARED AT 1-WEEK POST-OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. CLAIMANT STATES THE LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MMVTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/+2.0/082 SPHERE/ CYLINDER/ AXIS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON NOTES REFRACTIVE SURPRISE; DOUBLE VISION AND CORNEAL EDEMA ON (B)(6) 2018. REPORTEDLY THE CORNEAL EDEMA DISAPPEARED AT 1 WEEK POST OP AND NO TREATMENT BEYOND WHAT IS STANDARD FOR OCULAR SURGERY WAS PERFORMED. THE LENS WAS EXCHANGED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. IMPLANTED DATE: IT IS CORRECTED FROM "(B)(6) 2017" TO "(B)(6) 2017". EXPLANTED DATE: (B)(6) 2018. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: IT IS CORRECTED TO "THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED. ALTHOUGH LENS SN: (B)(4) WAS FOUND OUT OF TOLERANCE FOR SPHERIC REFRACTIVE DIOPTER, THERE IS NO REASONABLE PROBABLE CAUSE IN MANUFACTURING THAT WOULD HAVE RESULTED IN A REFRACTIVE CHANGE. THE SPHERIC REFRACTION OF THIS LENS WAS CALCULATED TO BE -13.96D AT THE TIME OF RELEASE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8445704&pc=MTA | 2023826-2019-00471 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Corneal Clouding/Hazing; Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, - 08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO SEVERE ANTERIOR SEGMENT REACTION. THE REPORTER INDICATED THE INFLAMMATION WAS TASS. AT THE POST-OP VISIT ON (B)(6) 2019, THE TOTAL CORNEA WAS HAZE AND EDEMA WAS IMPROVING WITH INTENSE TOPICAL AND SYSTEMIC STEROID MEDICATIONS. THE PATIENTS CURRENT STATUS IS IMPROVEMENT IN VA 6/6P, OCULAR STATUS IS NORMAL AND CORNEA IS GOOD AND CLEAR. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7156006&pc=MTA | 2023826-2017-01974 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-12-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Corneal Touch; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: ADDITIONAL INFORMATION FROM SURGEON: REPORTEDLY, PRIOR TO ICL REMOVAL, THE LENS WAS PUSHING THE IRIS AND TOUCHING THE CORNEA. THE CORNEA WAS "STARTING TO GET OEDEMATOUS." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552498&pc=MTA | 2023826-2019-00660 | 2018-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Corneal Ulcer; Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON THE SAME DAY, CORNEAL EDEMA AND INFLAMMATORY CELLS DEVELOPED IN THE ANTERIOR CHAMBER. DIAGNOSTIC TESTS WERE PERFORMED AND THE SURGEON REPORTS ENDOPHTHALMITIS. NO AQUEOUS SAMPLES WERE TAKEN. STEROID EYE DROPS, ORAL MEDICINE, AND EYE DROPS WITH HYPOTENSIVE EFFECT WERE ADMINISTERED. CURRENTLY, THE PATIENT IS RECOVERED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT CODE: 7196 IS NOT APPLICABLE AND SHOULD BE CORRECTED TO PATIENT CODE: 3191: NO CODE AVAILABLE (INFLAMMATORY CELLS) IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: PATIENT CODE 1791: CORNEAL EDEMA .PATIENT CODE 1796: INFLAMMATORY CELLS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652094&pc=MTA | 2023826-2016-00549 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Corneal Edema; Dry Eye(s); Loss of Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAD LOST VISION DUE TO PROBLEM WITH THE CORNEA, SUCH AS CORNEAL EDEMA (SWELLING). THE PATIENT HAD EPITHELIAL BASEMENT CELL DYSTROPHY, WHERE SCRAPING OF THE CORNEA WAS NEEDED. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY CONFIRMED THE PATIENT REPORT. THE DATE OF THE LAST VISIT WAS (B)(6) 2015. THE FACILITY INDICATED THE PATIENTS CONDITION WAS PRE-EXISTING AND THE PATIENTS VISION WILL FLUCTUATE/ONGOING. PATIENTS BCVA WAS 20/20. THE PATIENT ALSO HAS DRY EYE, WHICH ALSO IS PRE-EXISTING. THE PATIENTS PROGNOSIS IS GOOD. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6760836&pc=MTA | 2023826-2017-01145 | 2017-07-03 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Corneal Edema; Edema; Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -4.5/1.5/86 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE NEXT DAY, SIGNS OF TOXIC ANTERIOR SEGMENT SYNDROME, INCLUDE TEARING, PAIN, BLURRY VISION, EYELID EDEMA, HYPEREMIA, CORNEAL EDEMA, FIXED PUPIL AND FIBRIN CLOT IN THE ANTERIOR CHAMBER WAS NOTICED. THERAPY WAS STARTED WITH ORAL PREDNISOLONE (60 MG/DAY) AND ACETAZOLAMIDE 250 MM (1+1+1), TROPICAL PREDNISOLONE 1 DROP PER HOUR, TRAPICAMIDE, OFLOXACINE, AND BROMOFENAC. THE FOLLOWING DAY, THE PATIENT EXPERIENCE LESS PAIN, INTENSE HYPEREMIA, LESS EYELID EDEMA BUT A VERY CLOUDY WHITISH CORNEA. CORNEAL PERIPHERY WAS PARTIALLY PARED AND A BLOOD CLOT WAS SEEN IN THE INFERIOR A.C. AN ANTERIOR CHAMBER WASHING WAS PERFORMED AS WELL AS THE TEMPORAL INCISION WITCH BEING SUTURED WITH 2 10/00 NYLON STICHES AND INTRAVITREOUS INJECTION OF VANCOMICINE 1MG/0.1 ML AND CEFTAZIDIME 1.0 MG/0.1 ML. ACCORDING TO THE SURGEON, THE CORNEA IS CLEARING UP AND WOULD NEED TO MAKE ANOTHER SURGERY TO REMOVE THE MEMBRANE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5639098&pc=MTA | 2023826-2016-00630 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Corneal Edema; Endophthalmitis; Hypopyon; Intraocular Pressure Increased; Pupillary Block; Vitrectomy; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE PATIENT PRESENTED WITH POST-OP ENDOPHTHALMITIS. THE PATIENT UNDERWENT A VITRECTOMY WITH ANTIBIOTIC INJECTION. THE PATIENT EXPERIENCED HYPOPYON, PUPILLARY BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND CORNEAL EDEMA. THE LENS REMAINS IMPLANTED AND THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS NOT DEVICE RELATED, MOST LIKELY DUE TO INTRODUCTION OF BACTERIA DURING OR AFTER SURGERY.Event Description: ADDITIONAL INFORMATION RECEIVED - THE CUSTOMER REPORTED THE LENS WAS EXPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: LENS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593077&pc=MTA | 2023826-2019-00757 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination; Misfocusing; Off-Label Use; Appropriate Term/Code Not Available | Corneal Edema; Inflammation; Blurred Vision; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE CLINICAL STUDY THAT FOUND 3 PATIENTS IMPLANTED WITH THE POSTERIOR PHAKIC IMPLANTABLE COLLAMER LENS (ICL) AND PRESENTED WITH TOXIC ANTERIOR SEGMENT SYNDROME (TASS). ALL 3 PATIENTS WERE TREATED WITH MEDICATION AND THE INFLAMMATION DISAPPEARED. A POSSIBLE CAUSE WAS THE HANDLING OF THE ICLS OR INSTRUMENT TIPS WITH POWDERED-GLOVES AS REPORTED BY THE SURGEON. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: PATIENT #1 EXPERIENCED DECREASED VISION, CONJUNCTIVAL HYPERAEMIA AND MILD CORNEAL EDEMA. MULTIPLE WHITE DEPOSITS WERE LOCATED ON THE SURFACE OF THE ICL. PATIENT HAD KERATOCONUS AND PREVIOUS INTRACORNEAL SURGERY WHICH IS OFF-LABEL USE WITH THE ICL LENS. PATIENT #2 EXPERIENCED CONJUNCTIVAL HYPERAEMIA, CORNEAL EDEMA AND DECREASE IN VISION. PATIENT #3 PRESENTED WITH 20/20 VISION BUT WITH WHITE DEPOSITS ON THE LENS OPTIC. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6531883&pc=MTA | 2023826-2017-00727 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Corneal Edema; Inflammation; Intraocular Pressure Increased; Pupillary Block; Uveitis; Blurred Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/101 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, LOSS OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND AN UNREACTIVE PUPIL (FIXED). THE PATIENT EXPERIENCED GLARE/HALOS, BLURRED VISION AND INFLAMMATION. ALSO, MILD UVEITIS, PIGMENT DISPERSION AND HAD DEVELOPED URETTS ZAVALIA SYNDROME (PERSISTENT PUPIL DILATION). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP SHOWED CORNEAL EDEMA, SHALLOW ANTERIOR CHAMBER AND NORMAL INTRAOCULAR PRESSURE. Manufacturer Narrative: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED AND THE HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389853&pc=MTA | 2023826-2017-00347 | 2017-02-02 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Edema; Inflammation; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS, VAULTED FORWARD. IN ADDITION, THE PATIENT EXPERIENCED IRIS ATROPHY AND A DESCEMET MEMBRANE TEAR. ON (B)(6) 2017, THE LENS WAS EXCHANGED AND REPOSITIONED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. THE LENS WILL BE RETURNED. BASED ON SURGEON'S COMMENTS DOCTOR IS MONITORING PATIENT'S PROGRESS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16/2/111 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. THE PATIENT EXPERIENCED UNREACTIVE PUPIL, GLARE, INFLAMMATION, AND SWOLLEN CORNEA. THE INFERIOR EDGE OF ICL CAUGHT ABOVE IRIS WHICH CAUSED THE LENS TO VAULT FORWARD SO LENS REPOSITIONING WAS PERFORMED. IN ADDITION, PATIENT EXPERIENCED IRIS ATROPHY AND DESCEMET MEMBRANE TEAR. ON (B)(6)2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING. AS OF (B)(6)2017 THE PATIENT WAS STILL EXPERIENCING POSTOPERATIVE SEQUELAE BUT CORNEA WAS CLEAR. THE LENS WILL NOT BE RETURNED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). RESULT: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WAS REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850587&pc=MTA | 2023826-2023-01613 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICMO 13.7, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IT WAS ALSO REPORTED THAT CORNEAL EDEMA WAS NOTICED ON (B)(6) 2023, HOWEVER IT DID NOT REQUIRE TREATMENT BEYOND WAS IS STANDARD AFTER OCULAR SURGERY AND THE CORNEAL EDEMA RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677856&pc=MTA | 2023826-2021-04046 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | Corneal Edema; Intraocular Pressure Increased; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED DURING THE SURGERY AND THE PATIENT COMPLAINED OF BLURRED VISION. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS SOME CORNEAL EDEMA DUE TO THE INCREASED IOP AND BLURRED EYES. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711487&pc=MTA | 2023826-2017-01177 | 2017-06-16 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, DIOPTER -18.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. FOLLOW UP ON (B)(6) 2017, THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT, BLURRED AND DIMNESS OF VISION, ELEVATED IOP, IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, CORNEA OEDEMATUS. NEXT DAY, THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS REPOSITIONING, AND IRIDECTOMY WAS PERFORMED. FOLLOW UP ON (B)(6) 2017, THE PATIENT'S RIGHT EYE CORNEA GOT CLEARER WITH FIXED DILATED ATONIC PUPIL, PIGMENTATION IN ICL LENS, AND ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE IOP WAS CONTROLLED, AND ANGLES OPEN, PUPIL SEMI DILATED 4MM TO 5MM. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISISION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). IRIS PIGMENTATION, FIXED SEMI DILATED PUPIL, ANGLE CLOSURE, DIMNESS OF VISION, SECONDARY SURGERY, YAG, LENS REPOSITIONING, IRIDECTOMY, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS, LENS EXPLANT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594294&pc=MTA | 2023826-2018-00898 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/156 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "SHALLOW AC," DECREASE IN UA, AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER "TOW TIME AC WASH" AND "AC REFORMATION." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN UPDATED INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION: LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.76MM)]. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8903520&pc=MTA | 2023826-2019-01478 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTERD INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION WAS OBSERVED. ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. THE REPORTER STATED THAT THE PATIENT IS HAPPY WITH THE RESULTS AND DONT WANT TO CHANGE THE LENS . IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: PHOTOPHOBIA, PATIENT CODE 1791: CORNEAL OEDEMA, METHOD CODE 4117: LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783776&pc=MTA | 2023826-2021-04214 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Eye Pain | NA | 1 | Event Description: A JOURNAL ARTICLE WAS RECEIVED FROM THE JOURNAL OF OPHTHALMOLOGY, DATED 06-OCT-2021, FOR "PURE ICL IMPLANTATION: A NOVEL OPHTHALMIC VISCOSURGICAL DEVICE-FREE METHOD". THE ARTICLE SITED A STUDY OF 40 PATIENTS (80 EYES) OF MYOPIC PATIENTS, IMPLANTED WITH ICL LENSES. IN THE STUDY FOR BOTH EYES, ONE EYE UNDERWENT TRADITIONAL ICL IMPLANTATION AND THE OTHER EYE UNDERWENT OVD-FREE (PURE) ICL IMPLANTATION. A RISE IN INTRAOCULAR PRESSURE (IOP) WAS NOTED 2 HOURS AFTER SURGERY IN 8 EYES IN THE TRADITIONAL GROUP BUT NONE IN THE PURE GROUP. ONE PATIENT IN THE TRADITIONAL GROUP SUFFERED MILD EYE PAIN AND CORNEAL EDEMA. THE INTRAOCULAR PRESSURE WAS LOWERED BY APPROPRIATE MANAGEMENT. WITHIN 24 HOURS, THE IOP HAD DECREASED IN THE 8 EYES. THE LENSES REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THEY BELIEVED THE CAUSE WAS DUE TO VISCOELASTIC AGENT IN THE POSTERIOR CHAMBER WAS NOT COMPLETELY FLUSHED. NONE OF THE PATIENTS HAD THEIR ICLS REMOVED OR REPLACED. INTRAOCULAR PRESSURE FLUCTUATIONS WERE NOT OBSERVED IN ANY OF THE PROCEDURES WITHOUT VISCOELASTIC AGENTS. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE : UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5688628&pc=MTA | 2023826-2016-00670 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-31 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED ELEVATED IOP, HAZY VISION, PIGMENT DISPERSION, AND MILD EDEMA. LASER PERIPHERAL IRIDOTOMY (LASER PI) WAS PERFORMED TO LOWER INTRAOCULAR PRESSURE. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5745383&pc=MTA | 2023826-2016-00851 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Difficult to Remove; Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A 12.6MM ICL. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER STATED THE PATIENT HAD CORNEAL EDEMA AS A RESULT OF DIFFICULT EXPLANT SURGERY. THE PATIENT EXPERIENCED ELEVATED IOP AND NARROWING OF THE ANGLE. THE REPORTER INDICATED THAT THE EVENT WAS PRODUCT-RELATED. Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM) - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND SMALL TEARS ON THE LENS HAPTIC. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850665&pc=MTA | 2023826-2017-01401 | 2017-08-03 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/X092 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2017. ON (B)(6)2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANGLE CLOSURE WITH ELEVATED IOP. THE SURGEON REPORTED ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. SURGEON ALSO REPORTS CORNEAL EDEMA/MIDRIASIS. THE SURGEON STATES THAT THE PATIENT STATUS IS OKAY. Manufacturer Narrative: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7070542&pc=MTA | 2023826-2017-01801 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AT THE POST-OP VISIT ON (B)(6) 2017, CORNEAL EDEMA WAS OBSERVED, BUT THE PATIENT HAS FULLY RECOVERED AND IS NOW FINE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/28. Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF ONE LENS HAPTIC BROKEN. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594288&pc=MTA | 2023826-2018-00897 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/014 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THREE DAYS POSTOP IT IS NOTED THAT THE SITUATION HAS IMPROVED SLIGHTLY, "TOW TIME AC WASH AND AC REFORMATION WITH PI (SURGICAL) AND AIR BUBBLE PLACEMENT." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A "STERIPACK". SPELLING OF "STERIPACK" WAS INCORRECT PRIOR. (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER STATED "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." DEVICE EVALUATION - LENS WAS RETURNED DRY IN A STERIPECLE WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594291&pc=MTA | 2023826-2018-00896 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS THE FOLLOWING: ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ADDITION OF NEW PI (SURGICAL), ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC," AND CORNEAL EDEMA. STATES THE PROBLEM STARTED TO RESOLVE AFTER TWO AC WASHES AND SURGICAL PI WITH "AIR BUBBLE IN AC." ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED FROM REPORTER INDICATED: "THE LAST STATUS WAS AT 3 DAYS POST ICL REMOVAL: CORNEAL EDEMA + ISCHEMIC IRIS; DIFFERENT GRADES. SURELY NO EXCHANGE LENSES HAVE BEEN IMPLANTED." LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT, IN A STERIPACK. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE [(SECONDARY SURGERY, LENS EXPLANT, ADDITION OF NEW PI (SURGICAL), SIGNIFICANT REDUCTION OF ICA, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, "FLAT AC"] CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.93MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602196&pc=MTA | 2023826-2016-00514 | 2016-01-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/50. THE PATIENT HAD CORNEAL EDEMA POST-OP. Manufacturer Narrative: DEVICE MFR DATE: 09/29/2015 TO 10/08/2015, PREVIOUS DATE IS INCORRECT. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602203&pc=MTA | 2023826-2016-00515 | 2016-01-02 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/67. THE PATIENT HAD CORNEAL EDEMA POST-OP. Manufacturer Narrative: CORRECTED DATA: MFR DATE: 09/14/2015 TO 09/21/2015, PREVIOUS DATE IS INCORRECT. (B)(4). ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652092&pc=MTA | 2023826-2016-00637 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD CORNEAL EDEMA AND THE POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652096&pc=MTA | 2023826-2016-00638 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE REPORTER INDICATED THERE WAS ENDOTHELIAL TOUCH. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD CORNEAL EDEMA AND THE POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8903395&pc=MTA | 2023826-2019-01477 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTERD INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. THE REPORTER STATED THAT THE PATIENT IS SATISFIED WITH THE RESULTS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 OF -12.0 DIOPTER; IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT; REFRACTIVE SURPRISE; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; CORNEAL OEDEMA; BLURRED VISION; HERNIATED IRIS; IRIS ATROPHY; IRIS REPAIR; AND ADDITION OF SUTURE; ADDITION OF NEW PI (YAG) WAS REPORTED AS SECONDARY INTERVENTION. REPORTER STATED THAT PATIENT IS SATISFIED WITH THE RESULTS. PATIENT CODE 3191: HERNIATED IRIS; IRIS ATROPHY; IRIS REPAIR; ADDITION OF SUTURE; PATIENT CODE 1791:CORNEAL OEDEMA, DEVICE CODE 1401:REFRACTIVE SURPRISE, METHOD CODE 4117:LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957614&pc=MTA | 2023826-2020-00848 | 2019-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (30 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT PRESENTED WITH HIGH INTRAOCULAR PRESSURE 4 MONTHS AFTER THE OPERATION, WHICH CONTINUED UNTIL NOW. THE LENS WAS REMOVED. BUT AFTER THE REMOVAL OF THE LENS, THE INTRAOCULAR PRESSURE WAS STILL HIGH" AND CORNEAL EDEMA WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER STATED "CORNEAL EDEMA HAS NOT BEEN RESOLVED AT THIS MOMENT. THERE IS NO INFECTION. CORNEAL EDEMA ONLY BECAUSE OF HYPERTENSION. NO OTHER INTERVENTION PERFORMED BUT INTENDED TO HAVE SURGICAL INTERVENTION ON HYPERTENSION." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- (B)(6) 2021" SHOULD BE CORRECTED TO "07-APR-2020" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957617&pc=MTA | 2023826-2020-00849 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (26MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT PRESENTED WITH HIGH INTRAOCULAR PRESSURE 4 MONTHS AFTER THE OPERATION, CANNOT BE CONTROLLED BY MEDICINE TREATMENT. THE LENS WAS REMOVED AT 8TH MONTH POST OPERATION. BUT, THE INTRAOCULAR PRESSURE WAS STILL HIGH AFTER REMOVAL OF LENS" AND CORNEAL EDEMA WAS OBSERVED. PATIENT NEED FURTHER SURGICAL INTERVENTION TO TREAT HIGH IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5- THE REPORTER STATED "CORNEAL EDEMA HAS NOT BEEN RESOLVED AT THIS MOMENT. THERE ARE NO INFECTION. CORNEAL EDEMA ONLY BECAUSE OF HYPERTENSION. NO OTHER INTERVENTION PERFORMED BUT INTENDED TO HAVE SURGICAL INTERVENTION ON HYPERTENSION." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- (B)(6) 2021 SHOULD BE CORRECTED TO "07-APR-2021" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10- "(B)(6) 2021" SHOULD BE CORRECTED TO "(B)(6) 2020" IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11043753&pc=MTA | 2023826-2020-03159 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Optical Nerve Damage; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED HIGH VAULT AND HIGH IOP (INTRAOCULAR PRESSURE). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5:THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -5.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (47MMHG), PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PIS) WAS PERFORMED. THE REPORTER STATED "THE PATIENT WAS CYCLOPLEGED WITH ATROPINE FOR AN EXTENDED PERIOD OF TIME UNTIL THE IOP COULD BE STABILIZED AND THE CORNEAL EDEMA IMPROVE WITH STEROIDS, ORAL AND TOPICAL IOP LOWERING MEDICATIONS. THIS TOOK SEVERAL WEEKS TO RESOLVE". IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/20--, VAULT:1.5CT, IOP:12MMHG). FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "HE SUBSEQUENTLY DEVELOPED UNEXPLAINED SYMPTOMATIC RIGHT SIDED OPTIC NERVE EDEMA AND WAS SENT TO THE ER FOR FULL NEUROLOGICAL AND OPHTHALMOLOGICAL EVALUATION. MRI WAS NEGATIVE EXCEPT FOR ENHANCEMENT OF THE RIGHT OPTIC NERVE POTENTIALLY CONSISTENT WITH OPTIC NEURITIS. THE HEAD NEUROLOGIST DID NOT FEEL THAT IT WAS CLASSIC NEURITIS BUT HE WAS TREATED FOR IT REGARDLESS". ADDITIONAL INFORMATION REPORTED "HE WAS GIVEN IV STEROIDS IN THE HOSPITAL AND SENT HOME WITH A FEW DAYS OF HIGH DOES ORAL STEROIDS." NO OTHER SYMPTOMS HAVE DEVELOPED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL CAUSING ANGLE CLOSURE AND PUPILLARY BLOCK. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: CORRECTED DATA: B5 - SHOULD HAVE BEEN SPECIFIED THAT A YAG PI WAS PERFORMED. H6 - 4581: ANGLE CLOSURE, PIGMENT DISPERSION, ENLARGMENT OF PI (YAG), SHOULD HAVE BEEN SPECIFIED. CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DOB: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. COMMON DEVICE NAME: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # 431266
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497527&pc=MTA | 2023826-2019-00578 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-09 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Pain | NA | 1 | Event Description: BMC OPTHALMOLOGY ARTICLE PUBLISHED IN 2019 WHICH CITED A STUDY TITLED, 'COMPARISON OF VISUAL QUALITY AFTER EVO-ICL IMPLANTATION AND SMILE TO SELECT THE APPROPRIATE SURGICAL METHOD FOR HIGH MYOPIA.' THE ARTICLE STATES THAT AFTER IMPLANTATION, TWO EYES EXPERIENCED HIGH INTRAOCULAR PRESSURE (IOP), OCULAR PAIN, AND CORNEAL EDEMA. PATIENT WAS GIVEN INTRAVENOUS INFUSION OF MANNITOL AND SYMPTOMS DISAPPEARED AFTER 24 HOURS. ANOTHER PATIENT EXPERIENCED IOP OF 28-31MMHG TWO WEEKS AFTER SURGERY. CARTEOLOL HYDROCHLORIDE EYE DROPS WERE ADMINISTERED 2X DAILY AND THE IOP RECOVERED 5 DAYS LATER. ONE MONTH AFTER SURGERY, IOP WAS STABLE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14605731&pc=MTA | 2023826-2022-01798 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: AN ARTICLE PUBLISHED "INTERFACE FLUID SYNDROME AFTER BIOPTICS SURGERY" INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AFTER LASIK SURGERY. 1 MONTH POSTOPERATIVELY THE PATIENT COMPLAINED OF DECREASED VISION; PAIN; AND PHOTOPHOBIA IN THE EYE; HYPERTENSION (ELEVATED IOP). PATIENT WAS PRESCRIBED MEDICATION. AT A LATER DATE SLIT LAMP EXAMINATION INDICATED SEVERE CORNEAL EDEMA AS WELL. IT WAS LATER REPORTED THAT AFTER MEDICATION USE THE PATIENT'S EYE WAS CLEAR; IOP WAS STABLE; MEDICATION WAS DISCONTINUED AND THE PATIENT WAS DISCHARGED. THE ARTICLE CONCLUDES NOTING "WE BELIEVE THAT THE ICL WAS NOT THE DIRECT CAUSE OF IFS; BUT RATHER IT WAS THE PROLONGED AGGRESSIVE TOPICAL STEROID USE AFTER ICL IN A STEROID RESPONDER WHO HAD UNDERGONE LASIK. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7097942&pc=MTA | 2023826-2017-01842 | 2017-10-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, ANGLE CLOSURE, NARROWING OF THE ANGLE, SHALLOWING OF THE ACD, BLURRY VISION, PAIN AND CORNEAL EDEMA. PARACENTESIS WAS DONE TO LOWER THE PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2017, THE PATIENT'S POST-OP BCVA WAS 20/25. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS TYPES EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366387&pc=MTA | 2023826-2018-00456 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. THE LENS WAS FOUND WITH NO VISIBLE DAMAGE. DIMENSIONAL EVALUATION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7366474&pc=MTA | 2023826-2018-00455 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, NARROW ANGLES, ELEVATED INTRAOCULAR PRESSURE (IOP) AND CORNEAL EDEMA. THE PATIENT EXPERIENCED INTERMITTENT DECREASE OF VISUAL ACUITY, BLURRY VISION AND EYE PAIN. THE PATIENT RECEIVED MEDICAL TREATMENT FOR THE HIGH IOP. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE STATUS OF THE EYE POST-OP IS ASYMPTOMATIC. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114926&pc=MTA | 2023826-2020-01139 |  | Injury | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Adverse Event Without Identified Device or Use Problem | Corneal Edema; Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF GLAUCOMA IN MAY 2020 TITLED, 'ACUTE ANGLE CLOSURE FOLLOWING IMPLANTABLE COLLAMER LENS FOR MYOPIA.' THE ARTICLE STATES THAT NINE HOURS AFTER LENS IMPLANTATION, PATIENT REPORTED INTENSE PAIN AND BLURRED VISION IN RIGHT EYE (OD). EXAMINATION REVEALED AN ACUITY OF HAND MOTIONS AND A MID-DILATED PUPIL, NON-REACTIVE TO LIGHT. CORNEAL EDEMA, CONJUNCTIVAL INJECTION AND A SHALLOW ANTERIOR CHAMBER. SECONDARY ANGLE CLOSURE WAS MADE WITH POSSIBLE CAUSE OF PUPIL BLOCK, AQUEOUS MISDIRECTION, OR NON-PUPIL BLOCK. A DROP OF TRAVOPROST (0.004%) AND BRIMONIDINE TARTRATE (0.2%) TIMOLOL MALCATE. IOP WAS UNCHANGED, ORAL GLYCERINE WAS PRESCRIBED AS WELL. PARS PLANA VITRECTOMY AND SURGICAL IRIDECTOMY WAS PERFORMED. THE ANTERIOR CHAMBER DEEPENED WITHOUT LENS REMOVAL AND PATIENT'S SYMPTOMS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). PATIENT CODE 3191: NO CODE AVAILABLE (MILD DILATED PUPIL, CONJUNCTIVAL INJECTION, SHALLOW ANTERIOR CHAMBER, ACUITY OF HAND MOTIONS). CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9426953&pc=MTA | 2023826-2019-02295 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Corneal Edema; Intraocular Pressure Increased; Pain; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE MEDICINE IN SEPTEMBER 2019 TITLED,' EVALUATION OF VISUAL QUALITY AFTER EVO-ICL IMPLANTATION FOR HYPERMYOPIA.' THE ARTICLE STATES THAT OCULAR PAIN AND CORNEAL EDEMA WERE OBSERVED IN 1 EYE WITH HIGH INTRAOCULAR PRESSURE AFTER EVO-ICL IMPLANTATION. THE SYMPTOMS DISAPPEARED 24 HOURS AFTER INTRAVENOUS INFUSION OF MANNITOL. INTRAOCULAR PRESSURE IN ANOTHER PATIENT WAS ELEVATED TWO WEEKS AFTER SURGERY. CARTEOLOL HYDROCHLORIDE EYE DROPS WERE ADMINISTERED TWICE PER DAY, AND INTRAOCULAR PRESSURE RECOVERED 5 DAYS LATER. ON MONTH AFTER SURGERY, INTRAOCULAR PRESSURE WAS STABLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B1- "ADVERSE EVENT" SHOULD BE MARKED IN THE INITIAL MDR. B2- "REQUIRED INTERVENTION" SHOULD BE MARKED IN THE INITIAL MDR. H1- "SERIOUS INJURY" SHOULD BE MARKED IN THE INITIAL MDR. H6- PATIENT CODE 3191- MEDICAL INTERVENTION SHOULD BE IN THE INITIAL MDR. DEVICE CODE 3189 IS NOT APPLICABLE AND SHOULD BE CODE 2993 IN THE INITIAL MDR. RESULT CODE 213 IS NOT APPLICABLE AND SHOULD BE CODE 114 IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439246&pc=MTA | 2023826-2017-00486 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT EXPERIENCED IRIS HERNIATION, AND REDUCTION SURGERY. ON THE NEXT DAY, THE PATIENT EXPERIENCED OCULAR PAIN AND BLURRY VISION, WITH SHALLOW ANTERIOR CHAMBER, MODERATE DIFFUSE CORNEAL EDEMA AND OCULAR HYPERTENSION. THE CENTRAL PORT OF THE LENS WAS BLOCKED BY IRIS PIGMENT. IRIDECTOMY SURGERY WAS PERFORMED, AND THE PROBLEM WAS RESOLVED. ENLARGEMENT OF PI OR ADDITION OF NEW PI SURGERY WAS ALSO PERFORMED. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED, WILL NOT BE RETURNED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872434&pc=MTA | 2023826-2017-01443 | 2017-08-20 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, ELEVATED IOP, NARROWING OF THE ANGLE, BLURRY VISION, PAIN, PUPILLARY BLOCK, AND CORNEAL EDEMA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND VIGAMOX AND PREDNISOLONE WERE USED AS STANDARD POST OPERATIVE PROTOCOL. THE PROBLEMS WERE RESOLVED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT TYPE EVENT WAS FOUND. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221040&pc=MTA | 2023826-2015-01452 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2015-11-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; Pain; Pupillary Block; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01451 FOR RIGHT EYE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -14.00/0.5/103 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE PUPIL, ANGLE CLOSURE, AND POST-OP PAIN CAUSED BY ELEVATED IOP AND CORNEAL EDEMA. THE ICL WAS EXPLANTED. NO OTHER LENS WAS IMPLANTED. (B)(4). CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6658366&pc=MTA | 2023826-2017-00957 | 2017-05-02 | Injury | STAAR SURGICAL COMPANY | 2017-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PUPILLARY BLOCK, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND CORNEAL EDEMA. THE REPORTER INDICATED MULTIPLE PARACENTISIS PERFORMED AND THE PERIPHERAL IRIDECTOMIES WERE REDONE. THE PATIENT WAS PRESCRIBED MEDICATION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENTS POST-OP CONDITION IS - TRACE K EDEMA, 1+ CELLS WITH SHALLOW ANTERIOR CHAMBER, DILATED PUPIL WITH CORNEAL IRIS TOUCH PERIPHERALLY, AND BEST-CORRECTED VISUAL ACUITY 20/200. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE PROBLEM WAS NOT RESOLVED AFTER EXPLANTATION OF THE LENS. THE PATIENT CONTINUED TO HAVE INTRAOCULAR PRESSURE (IOP) INCREASE POST EXPLANT OF LENS. AT ONE WEEK POST-OP, THE PATIENT UNDERWENT AHMED VALVE SHUNT INSERTION FOR IOP, WITH GLAUCOMA SPECIALIST. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11234650&pc=MTA | 2023826-2021-00225 | 2020-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (43MMHG), GLARE/HALOES AND BLURRED VISION WAS OBSERVED. ADDITIONAL SURGICAL INTERVENTION (PI) WAS PERFORMED. IT WAS REPORTED THAT THE PROBLEM RESOLVED, LAST VISIT DATE (B)(6) 2020 (BCVA:20/20, UCVA:20/25-, VAULT:1.0 CT, IOP:14MMHG). THE REPORTER STATED "HE DEVELOPED A STEROID RESPONSE WITH CORNEAL EDEMA (B)(6) 2020 AND WAS SWITCHED TO LOTEMAX FROM PREDNISOLONE ACETATE 1%. HAS SINCE BEEN TAPERED OFF ALL STEROIDS AND IOP LOWERING MEDICATIONS SUCCESSFULLY." IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT MAY HAVE BEEN RETAINED OVD BEHIND THE IOL. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE, SHOULD HAVE BEEN SPECIFIED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6501830&pc=MTA | 2023826-2017-00562 | 2017-03-30 | Injury | STAAR SURGICAL COMPANY | 2017-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | P990013 | 1 | Event Description: THE REPORTER STATED THAT A MICL13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PUPILLARY BLOCK CAUSING ELEVATED IOP. THE REPORTER INDICATED THE PATIENT HAD EXCESSIVE VAULT, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ACD. THE PATIENT'S REPORTED CONDITION AT THE TIME OF THE LAST VISIT ON (B)(6) 2017 WAS 20/70 BCVA, NORMAL IOP AND CORNEAL EDEMA. Manufacturer Narrative: AT THE TIME OF SURGERY PATIENT WAS (B)(6) AND PER DFU INDICATIONS VISIAN IMPLANTABLE COLLAMER LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENT IN ADULTS (B)(6) OF AGE. THERE IS NO SIGNIFICANT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4). Manufacturer Narrative: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN GAUZE WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND MISSING PIECE OF THE HAPTIC. (B)(4). WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221055&pc=MTA | 2023826-2015-01451 | 2015-09-30 | Injury | STAAR SURGICAL COMPANY | 2015-11-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL WAS NOTED. A YAG, ENLARGEMENT OF PI AND ANTERIOR CHAMBER IRRIGATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. THE ICL WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01452 FOR LEFT EYE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORNEAL EDEMA WAS REPORTED. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395494&pc=MTA | 2023826-2023-05778 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Corneal Edema; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE FLIPPED INTO THE PATIENTS EYE. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7749218&pc=MTA | 2023826-2018-01164 | 2018-07-05 | Injury | STAAR SURGICAL COMPANY | 2018-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Edema; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THA 12.6MM MICL12.6, -8.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED WITH ADDITIONAL SUTURE DUE TO IRIS ATROPHY AND PUPIL DILATION. ADDITIONAL INFORMATION REPORTED: CHAMBER WELL FRAMED. IRIS ATROPHY AT INCISION SITE. PUPIL SIZE NORMAL. SUPERIOR CORNEAL EDEMA GRADE 2. OS VISUAL FIELD PASS. SLIGHT RESTRICTION INFERIOR NOT IN VISUAL AXIS. WILL FOLLOW UP WITH VISUAL FIELD IN 1 MONTH. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN A CONTACT LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: CORNEAL EDEMA RESOLVED. PATIENT WAS TREATED WITH DUREZOL 0.05% AND ZYMAXID 0.5% AS PART OF NORMAL POST-OP REGIMEN. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ADDED: IN THE SURGEON'S OPINION THE CAUSE OF THE EVENT IS A PATIENT-RELATED FACTOR. ADDITIONAL REPORT SOURCES ADDED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PATIENT CODE (B)(4): SECONDARY SURGICAL INTERVENTION, EXPLANT AND ADDITIONAL SUTURE (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711466&pc=MTA | 2023826-2017-01180 | 2017-05-09 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Corneal Edema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2, DIOPTER +10.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOSS OF ANTERIOR CHAMBER DEPTH, IRIS DEPIGMENTATION, PIGMENT DISPERSION, AND CONEAL OEDEMA. THE LENS WAS EXPLANTED IN THE SAME SURGERY. THE REPORTER ALSO MENTIONED THAT THE IMPLANTATION PROCEDURE COULD NOT GO A HEAD DUE TO CHAMBER FLATTENING, IRIS HERNIA. FOLLOW UP ON (B)(6) 2017, CORNEAL OEDEMA IS SLOWLY IMPROVING, AND VISUAL AXIS IS NOT AFFECTED AT THIS STAGE. AS OF THE DATE OF MDR SUBMISISON THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS RETURNED IN TWO PIECES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281448&pc=MTA | 2023826-2018-00257 | 2017-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.00/135 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2017. ALSO REPORTED IS SIGNIFICANT ENDOTHELIAL CELL LOSS. CLAIM REPORTER STATES "BIG CORNEAL EDEMA AFTER SURGERY. THICKER PAQUIMETRY...REFRACTION MODIFIED." REPORTEDLY THE EYE HEALED PROPERLY AFTER SURGERY. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS VIAL. VISUAL INSPECTION FOUND A TEAR IN THE HAPTIC AND OPTIC. THE LENS WAS ALSO MISSING A PIECE OF ITS HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152526&pc=MTA | 2023826-2017-01958 | 2017-10-03 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS "PUPIL EXPANDED AND NOT CONSTRICTED (ABNORMAL PUPILLARY MYDRIASIS)." SURGEON ALSO REPORTS INTRAOCULAR INJECTIONS OF "SANPILO (PILOCARPINE HYDROCHLORIDE)" AND STATES THAT CORNEAL EDEMA IS "OBSERVED AT PUPILLARY MYDRIASIS TREATMENT." THE SURGEON REPORTS THAT DURING A FOLLOW-UP VISIT, THE ICL WAS REMOVED AS HE THOUGHT THE ICL POSSIBLY WAS NOT ALLOWING THE PUPIL TO CONSTRICT. HOWEVER, UPON REMOVAL THERE WERE NO IRREGULARITIES FOUND ON THE ICL SO IT WAS "IMPLANTED AGAIN." ON (B)(6) 2017 THE LENS WAS FINALLY EXPLANTED AND REPORTEDLY THE PUPIL DID NOT CONSTRICT AFTER EXPLANT. NO ADDITIONAL INFORMATION IS AVAILABLE AS OF THE TIME OF MDR SUBMISSION. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE ON LENS) AND THE LENS MISSING A PIECE OF HAPTIC. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064998&pc=MTA | 2023826-2016-01561 | 2016-10-17 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -3.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED AT THE TIME OF SURGERY, THE PUPIL WAS AT A MINIMAL DILATION OF 8MM AND THE ICL WAS EXTREMELY VAULTED IN THE ANTERIOR CHAMBER, SO THE SURGEON DECIDED TO REMOVE THE LENS DURING THE SAME SURGERY. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THERE WAS CORNEAL EDEMA AT ONE DAY POST-OP. VISION WAS 20/20. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6893656&pc=MTA | 2023826-2017-01485 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CONDITION WAS: UCVA - 20/30 +2 (WILL IMPROVE), CORNEAL EDEMA WILL RESOLVE, AND INTRAOCULAR PRESSURE (IOP) WAS NORMAL . Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MINI CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS CURVED. THERE WAS EVIDENCE OF RED AND WHITE RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978082&pc=MTA | 2023826-2017-01612 | 2017-09-26 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.50 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON THE SAME DATE THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, FIXED PUPIL, AND PIGMENT DISPERSION. IT IS UNKNOWN IF THIS OCCURRED INTRAOPERATIVELY OR POST-OPERATIVELY. NO ADDITIONAL INFORMATION IS KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. REPORTEDLY, "CORNEAL EDEMA RESOLVED, IOP NORMAL." (B)(4). CONCLUSION CODE UPDATED: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6) YRS)] (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7161175&pc=MTA | 2023826-2018-00003 | 2017-12-07 | Injury | STAAR SURGICAL COMPANY | 2018-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND CORNEAL EDEMA. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7594282&pc=MTA | 2023826-2018-00895 | 2018-04-16 | Injury | STAAR SURGICAL COMPANY | 2018-06-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/013 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE SURGEON NOTES ADDITION OF NEW PI (SURGICAL). ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A STERIPACK AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTIMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.0/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE SURGEON NOTES ADDITION OF NEW PI (SURGICAL). THE LAST PATIENT STATUS SHOWED CORNEAL EDEMA AND ISCHEMIC IRIS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371773&pc=MTA | 2023826-2016-00066 | 2015-12-15 | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT HAD EDEMA ONE-DAY POST-OP. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906883&pc=MTA | 2023826-2018-01445 | 2018-08-02 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Corneal Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.6MM VICH 11.6 IMPLANTABLE COLLAMER LENS, +7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND LENS ROTATION. LENS REPOSITIONING AND MILD CORNEAL EDEMA IS ALSO REPORTED. THE CAUSE IS LISTED AS DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) : NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, LENS REPOSITIONING) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824672&pc=MTA | 2023826-2020-00424 | 2020-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Corneal Edema; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND CORNEAL EDEMA WERE OBSERVED. THE LENS WAS INTRAOPERATIVELY REMOVED WITH ENLARGEMENT OF PI AND YAG. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. ON PATIENT'S (B)(6) 2020 "MILD PERSISTENT CORNEAL EDEMA OBSERVED AND PATIENT CONSIDERING PLACING SMALL ICL." PER THE SURGEON ON (B)(6) 2020, "PATIENT PREFERS TO USE GLASSES AND CONTACT, NOT TO ATTEMPT ANOTHER ICL SURGERY." Manufacturer Narrative: CORRECTED DATA: B5 -"POSTOPERATIVE REPORT INDICATED PATIENT HAD A LOSS OF BCVA, FROM 20/20 TO 20/50." SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1. B5 - ON (B)(6) 2020 THE REPORTER INDICATED PATIENT HAD ENLARGED PI (YAG) SURGERY PERFORMED 1-2 HOURS AFTER IMPLANT. SURGEON THEN EXPLANTED THE LENS ON THE SAME DATE AND INDICATED "I ENLARGED THE PI'S IN BOTH EYES TO INSURE THEY WERE PATENT BUT THIS DID NOT LEAD TO DEEPENING OF THE ANTERIOR CHAMBER OVER THE NEXT HOUR. I FELT THAT SHE WAS AT GREAT RISK FOR PUPILLARY BLOCK AND FELT REMOVAL WAS THE SAFEST COURSE". H6 - H6 - PATIENT CODE 3191: NO CODE AVAILABLE ( LOSS OF BCVA,) SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - PATIENT ALSO EXPERIENCED NARROW ANGLES EVEN THOUGH THEY HAD NORMAL IOP. POSTOPERATIVE REPORT INDICATED PATIENT HAD A LOSS OF BCVA, FROM 20/20 TO 20/50. H6 - PATIENT CODE 3191: NO CODE AVAILABLE ( LENS EXPLANT, YAG, NARROW ANGLES, LOSS OF BCVA) CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 18-FEB-2020 SHOULD BE CORRECTED TO 17-FEB-2020 IN INTIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A POUCH WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. ADDITIONAL INFORMATION: D10 - RETURNED TO MANUFACTURER: (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6462705&pc=MTA | 2023826-2017-00551 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Prolapse; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON INSERTED AND REMOVED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, DURING THE SAME SURGERY DUE TO THE LENS STARTED TO VAULT AND WAS TOO LONG. THE CHAMBER WAS SHALLOW AND THE IRIS HAD BEGUN TO PROLAPSE THROUGH THE INCISION. THE LENS WAS REMOVED WITH NO PATIENT COMPLICATIONS. THERE WAS IRIS ATROPHY AND SOME CORNEAL EDEMA. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. AFTER A WEEK, THE PATIENT HAD IMPROVED AND VA WAS 20/25. THE REPORTER STATED THE EVENT WAS DUE TO A SIZING ISSUE. Manufacturer Narrative: DEVICE EVALUATION - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A TEAR/CUT FROM ONE HAPTIC THROUGH THE LENS OPTIC. THE LENS WAS RETURNED IN LIQUID. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6675039&pc=MTA | 2023826-2017-01045 | 2017-05-16 | Injury | STAAR SURGICAL COMPANY | 2017-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, DIOPTER -6.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE, SIMILAR DIOPTER LENS DUE TO THE PATIENT EXPERIENCING LOW VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOS, AND SWOLLEN CORNEA. SECONDARY SURGICAL INTERVENTIONS INCLUDE LENS REPOSITIONING AND AN ADDITIONAL SUTURE. THE SURGEON STATED THAT, DUE TO CORNEAL SWELLING, IT WAS TOO SOON TO DETERMINE WHETHER THE PROBLEM WAS RESOLVED. THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: IT IS TOO SOON TO DETERMINE WHETHER THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. LENS HAS NOT BEEN RETURNED FOR EVALUATION. THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. THE LENS HAS BEEN RETURNED FOR EVALUATION. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842746&pc=MTA | 2023826-2020-00518 | 2020-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Corneal Edema; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND GLARE/HALOS. THE PROBLEM WAS RESOLVED HOWEVER, PATIENT'S CONDITION AT FINAL VISIT IS MARKED AS "MILD CORNEAL EDEMA." THE CAUSE OF THE EVENT IS REPORTED AS "ANATOMICAL." REFERENCE MFR FILE# (B)(4) FOR 2ND IMPLANT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1791 SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1, AS IT IS REPORTED ON MFR FILE# (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310632&pc=MTA | 2023826-2020-01552 | 2016-03-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Touch; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED HAPTICS CONTACTED THE PERIPHERAL CORNEAL ENDOTHELIUM. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE REPORTER INDICATED THEY DIDN'T HAVE EXPERIENCE IN THE EARLY TIME WHEN ICL FLOATED IN THE ANTERIOR CHAMBER WITH THE "OVD FREE" TECHNIQUE. THIS PATIENT WAS PART OF A STUDY COMPARING THE SAFETY OF IMPLANTABLE COLLAMER LENS (ICL) WITH AND WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE (OVD). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS IMPLANTED ON (B)(6) 2016. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H1: IN INITIAL MDR AND SUPPLEMANTAL REPORT SHOULD BE CORRECTED FROM SERIOUS TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310711&pc=MTA | 2023826-2020-01511 | 2016-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Corneal Touch; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER INDICATED HAPTICS CONTACTED THE PERIPHERAL CORNEAL ENDOTHELIUM. THE LENS REMAINED IMPLANTED. THE REPORTER INDICATED THEY DIDN'T HAVE EXPERIENCE IN THE EARLY TIME WHEN ICL FLOATED IN THE ANTERIOR CHAMBER WITH THE "OVD FREE" TECHNIQUE. THIS PATIENT WAS PART OF A STUDY COMPARING THE SAFETY OF IMPLANTABLE COLLAMER LENS (ICL) WITH AND WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE (OVD). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5291931&pc=MTA | 2023826-2015-01634 | 2015-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-12-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Cyclitis; Vitrectomy; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/X087 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT EXPERIENCED CYCLITIS. AFTER THE LENS WAS EXPLANTED, THE PATIENT WAS TRANSFERRED TO TERTIARY HOSPITAL FOR PAS PLANA VITRECTOMY (PPV). Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS NOT RETURNED. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14654834&pc=MTA | 2023826-2022-01796 | 2021-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Cyst(s) | P030016 | 1 | Event Description: AN ARTICLE FROM ACTA OPHTHALMOLOGICA 2022 WAS RECEIVED IN REGARD TO "BLOCK EXCISION AND TECTONIC CORNEOSCLERAL GRAFTING FOR EPITHELIAL IMPLANTATION CYST AFTER INTRAOCULAR CONTACT LENS IMPLANTATION". THE ARTICLE INVOLVED A 32-YEAR-OLD WOMAN WITH CORRECTION OF MYOPIA FOLLOWING POST BILATERAL ICL PLACEMENT. THREE YEARS LATER IN JANUARY 2021, THE WOMAN PRESENTED WITH A PROGRESSIVELY GROWING TRANSPARENT IRIS CYST IN HER LEFT EYE. BLOCK EXCISION INCLUDING COMPLETE CYST WITH ADJACENT IRIS, AND CILIARY BODY AND TECTONIC CORNEOSCLERAL GRAFT WERE SUCCESSFULLY PERFORMED WITHOUT UNCONTROLLED RUPTURE OF THE CYST OR FURTHER INTERVENTION ON EITHER LENS OR ICL. ONE DAY POST-OP, IOP WAS 5 MMHG AND UCVA WAS 0.7 LOGMAR. EIGHT MONTHS AFTER CYST REMOVAL SURGERY, CRYSTALLINE LENS REMAINED CLEAR AND ICL REMAINED WELL CENTERED. SINCE ICL IMPLANTATION WAS ANECDOTALLY UNEVENTFUL, MINOR INGROWTH OF CONJUCTIVAL EPITHELIUM THROUGH AN IMPERFECT TUNNEL IS SUSPECTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- "EIGHT MONTHS AFTER CYST REMOVAL SURGERY CRYSTALLINE LENS REMAINED CLEAR AND ICL REMAINED WELL CENTERED. PUPILLOPLASTY WAS ALSO PERFORMED." CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152455&pc=MTA | 2023826-2017-01962 | 2017-10-09 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Device Dislodged or Dislocated | Cyst(s); Blurred Vision; Halo; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -1.5/+3.5/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, LENS ROTATION, REFRACTIVE SURPRISE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) (ANGLE CLOSURE), HALOES, VISUAL DISCOMFORT, AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS. ALSO REPORTED IS SURGICAL ENLARGEMENT/ADDITION OF NEW PI, LOW BCVA AND UCVA, AND LENS REPOSITIONING. THE SURGEON STATES THAT "PROBABLY THE LENS IS TOO BIG FOR THE EYE" BUT BELIEVES THAT THE VISUAL PROBLEMS OF THE PATIENT ARE NOT RELATED TO THE LENS BUT TO THE CORNEA. HE BELIEVES THAT CHANGING THE ICL WILL NOT SOLVE THE VISUAL DISCOMFORT AND THAT THE LENS IS NOT DOING ANY HARM TO THE PATIENT. THE LENS REMAINS IMPLANTED. THE PATIENT IS "NOT HAPPY WITH HIS VISUAL ACUITY." Manufacturer Narrative: THE SURGEON NOTED THAT "THE IRIS CYST WAS NOT SUSPECTED BEFORE THE SURGERY. THE SUSPICION WAS RISEN AFTER THE FIRST BY THE FACT THAT THE VAULT WAS UNEVEN AND THAT WITH DILATION ALWAYS GETS MORE EVEN AND WITH A NORMAL HEIGHT. AND THE CYST SUSPICION WAS EVEN STRONGER AFTER THE SECOND SURGERY SINCE RE-ACCOMMODATION OF THE ICL DID NOT CHANGE THE BEHAVIOR OF THE LENS." WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: PREVIOUS CODE IS NOT APPLICABLE, KERATOCONUS IS NO LONGER CONSIDERED A CONTRAINDICATION. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5375453&pc=MTA | 2023826-2016-00050 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Cyst(s); Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS HAD AN EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND LENS ROTATION. THE PATIENT EXPERIENCED GLARES/HALOS. THE REPORTER INDICATED THE LENS ROTATION MAY HAVE BEEN CAUSED BY A CILIARY CYST. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS IMPLANTED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338922&pc=MTA | 2023826-2019-02167 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); Intraocular Pressure Increased; Pain; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY THAT FOUND PATIENT IMPLANTED WITH A 12.1MM IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER IN THE PATIENTS RIGHT EYE (OD). AT ONE DAY POST-OPERATIVE, THE PATIENT SUFFERED PAINFUL EYE. THE PATIENT HAD SEVERE CORNEAL EDEMA. SLIT-LAMP EXAMINATION SHOWED IRIDO-CORNEAL CONTACT AT TEMPORAL MID-PERIPHERY, SHALLOW ANTERIOR CHAMBER WITH CENTRAL VAULTING OF ICL, PARTIAL OBSTRUCTION OF PERIPHERAL IRIDOTOMY, INTRAOCULAR PRESSURE WAS NOT MEASURED BUT THE EYE WAS VERY FIRM. ORAL AND SYSTEMATIC ANTIGLAUCOMA MEDICATIONS INCLUDING INTRAVENOUS (IV) MANNITOL 20% WERE ADMINISTERED. ADDITIONAL LASER PROCEDURE TO REOPEN PERIPHERAL IRIDOTOMY WAS PERFORMED. AN ULTRASOUND BIOMICROSCOPY (UBM) SHOWED A TEMPORAL POSTERIOR IRIS CYST AT 8 O'CLOCK, WHICH WAS PUSHING THE PIOL-IRIS COMPLEX FORWARD CAUSING IRIDO-CORNEAL ANGLE CLOSURE. PATIENT WAS READMITTED TO SURGERY FOR REPOSITIONING OF THE ICL. THE NEXT DAY, PATIENT DID NOT COMPLAIN OF SEVERE OCULAR PAIN. CORNEAL EDEMA IMPROVED MARKEDLY AND IOP WAS 20 MMHG. AT 1 MONTH FOLLOW-UP VISIT, UCVA WAS 0.9 BILATERALLY. CORNEAL EDEMA DISAPPEARED AND IOP WAS 20 MMHG. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANTED: UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753514&pc=MTA | 2023826-2020-00376 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | Cyst(s); Iritis; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER INDICATED THE LENS HAD A HIGH VAULT AND BOWED FORWARD IRIS. THERE WAS SOME PERSISTENT IRITIS AND MAYBE VERTICAL TILT. THE PATIENT IS DOING OK ON PREDNISOLONE, WHICH WILL BE TAPERED OFF OVER NEXT FEW WEEKS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED AFTER REVIEWING AN ARCSCAN, IT SHOWED A POSSIBLE CYST CAUSING THE TILT, SO THIS IS NOT LIKELY DUE TO THE DEVICE. NO EXPLANTATION IS PLANNED FOR NOW AND THE PATIENT IS HAPPY. CLAIM # : (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7047145&pc=MTA | 2023826-2017-01747 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND MASSIVE CYST TEMPORAL. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7047225&pc=MTA | 2023826-2017-01750 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Cyst(s); No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+4.0/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND MASSIVE CYST TEMPORAL. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE CHANGED: DEVICE FAILURE RELATED TO PATIENT CONDITION (MASSIVE CYST TEMPORAL). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10026173&pc=MTA | 2023826-2020-01021 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Cyst(s); No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD A CILIARY CYST AND THE LENS WAS IN AN ECCENTRIC POSITION. THE PATIENT REQUESTED THE LENS BE EXPLANTED. THE CYST WAS PRE-EXISTING. THE LENS WAS OUT OF POSITION, THE CENTRAL HOLE OF THE LENS WAS LOCATED IN THE LOWER HALF OF THE PUPILLARY REGION. THE SURGEON DOES NOT PLAN TO IMPLANT OTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10026180&pc=MTA | 2023826-2020-01020 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | Cyst(s); No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT HAD A CILIARY CYST AND THE LENS WAS IN AN ECCENTRIC POSITION. THE PATIENT REQUESTED THE LENS BE EXPLANTED. THE CYST WAS PRE-EXISTING. THE LENS WAS OUT OF POSITION, THE CENTRAL HOLE OF THE LENS WAS LOCATED IN THE LOWER HALF OF THE PUPILLARY REGION. THE SURGEON DOES NOT PLAN TO IMPLANT OTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATIN: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938957&pc=MTA | MW5064619 | 2016-08-10 | Injury | STAAR SURGICAL | 2016-09-01 | MTA | VISIAN ICL | Insufficient Information | Dehydration; Pain; Visual Impairment; Vomiting |  | 1 | Event Description: I WENT IN FOR A "BUSY" AND ICL ON BOTH EYES ON (B)(6) 2016. ON (B)(6) 2016, I AWOKE IN EXCRUCIATING PAIN AND WAS UNABLE TO SEE OUT OF ONE EYE. I WAS DIRECTED TO TWO DIFFERENT DOCTORS OFFICES AND THEN FINALLY AN EMERGENCY ROOM IN (B)(6) AND THEN ADMITTED INTO (B)(6) HOSPITAL. THEY PERFORMED A SURGICAL IRIDOTOMY IN MY LEFT EYE TO RELEASE THE PRESSURE THAT WENT UP TO 62 AND MY LEFT EYE. I WAS LASERED DOZENS OF TIME TRYING TO RELEASE THE PRESSURE. THE PAIN WAS EXCRUCIATING. I COULDN'T KEEP FOOD DOWN. I WAS VOMITING EVERY 15 MINUTES. I COULD NOT BE IN LATE. I WAS DEHYDRATED. AFTER THE SURGERY ON (B)(6) I WAS RELEASED FROM THE HOSPITAL ON (B)(6). I WAS SCHEDULED TO RETURN BACK TO (B)(6) HOSPITAL IN (B)(6) FOR PROCEDURE BUT IT WAS TOO PAINFUL TO BE DONE. I RETURNED BACK TO MY OPHTHALMOLOGIST AND THEN A GLAUCOMA SPECIALIST AND THEY WERE TREATING ME TWICE A WEEK FOR A COUPLE OF WEEKS. I AM CURRENTLY SEEING ABOUT 20/60 IN MY LEFT EYE AND I'M NOT IN AS MUCH PAIN. THE DOCTORS ARE STILL WAITING TO SEE IF THE LENS WILL HAVE TO BE REMOVED, OR IF MY VISION WILL JUST BE IMPAIRED. MY VISION IN MY RIGHT EYE DOES NOT ALLOW ME TO SEE THINGS CLOSE UP. THE DOCTORS ARE SAYING THAT I NEED TO WEAR READING GLASSES WHICH DEFEATS THE WHOLE PURPOSE OF GETTING THE BUSY AND ICL SO THAT YOU DO NOT NEED GLASSES. ON PRIOR TO THE SURGERY, I WAS SEEING FINE WITH MY CONTACTS NEVER NEEDED READING GLASSES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9759552&pc=MTA | 2023826-2020-00316 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Deposits; Inflammation; Intraocular Pressure Increased; Vomiting; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD), IN (B)(6) 2008. AT POST-OP 3 HOURS, THE PATIENT HAD SEVERE NAUSEA AND VOMITING, BUT NO EYE PAIN. THE INTRAOCULAR PRESSURE WAS ELEVATED. POST-OP THERE WAS ANTERIOR CHAMBER INFLAMMATION AND APPEARANCE OF DIFFUSE WHITISH PRECIPITATES ON THE ANTERIOR SURFACE OF THE CRYSTALLINE LENS, HYPOTONY AND MID-DILATED MILDLY REACTIVE PUPIL. WITH ADMINISTRATION OF TOPICAL AND SYSTEMIC STEROIDS, THE ANTERIOR CHAMBER REACTION SUBSIDED, AND THE ANTERIOR CAPSULAR DEPOSITS GRADUALLY RESOLVED PERIPHERALLY WITH SOME REMAINING CENTRALLY OVER THE COURSE OF SEVERAL WEEKS. THE PATIENT'S VISUAL ACUITY AT 6 MONTHS WAS 20/20. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO RETAINED VISCOELASTIC IN THE EYE, USE OF MYOSTAT AND SOME IRIS MANIPULATION IN A HYPEROPIC PATIENT. THE EVENT WAS NOT DUE TO THE ICL LENS. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. EXPLANTED: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MANUFACTURER DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788234&pc=MTA | 2023826-2021-01341 | 2018-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Discomfort | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER TORE DURING INJECTION INTO THE PATIENT'S LEFT (OS) EYE. THIS OCCURRED ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE LENS, THE PATIENT COMPLAINING OF DISCOMFORT AND THE LENS WAS SLIGHTLY OFF-CENTER. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "ACCORDING TO THE SURGEON THE LENS WAS STUCK AND IT HAD DIFFICULTY OPENING AND ONE 'EAR' WAS SLIGHTLY BROKEN DURING INJECTION...HOWEVER THE SURGEON LEFT IT IMPLANTED. THE PATIENT COMPLAINING OF DISCOMFORT AND THE LENS BEING OFF-CENTER, THE SURGEON DECIDED TO CHANGE THIS BROKEN LEN WITH A NEW ONE." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS ON LENS. CLAIM #(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633327&pc=MTA | 2023826-2023-03538 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -2.25 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH SAME LENGTH, SLIGHTLY STRONGER LENS ON (B)(6) 2023 DUE TO PATIENT DISSATISFACTION OF VISION ACUITY. PROBLEM RESOLVED. REPORTEDLY, UCVA IMPROVED AFTER EXCHANGE. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271651&pc=MTA | 2023826-2021-03030 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THERE WAS CONTACT WITH THE IRIS, CAUSING PIGMENT DISPERSION. THE PATIENT HAD DISCOMFORT AND DESIRE TO TOUCH EYE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS UNDER OBSERVATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271699&pc=MTA | 2023826-2021-03029 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND THERE WAS CONTACT WITH THE IRIS, CAUSING PIGMENT DISPERSION. THE PATIENT HAD DISCOMFORT AND DESIRE TO TOUCH EYE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. POST-OPERATIVE OBSERVATION SHOWED THE DEVELOPMENT OF CATARACT. THE REFRACTIVE RESULT IS STABLE. SEE MFR. REPORT # 2023826-2021-04028 FOR REPLACEMENT LENS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 721074.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090694&pc=MTA | 2023826-2021-04995 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING FOREIGN BODY SENSATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS UNKNOWN BUT NOTED THE DEVICE DID NOT FAIL TO PERFORM AND "PATIENT IS UNCOMFORTABLE WITH THE LENS IN PLACE". Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15906913&pc=MTA | 2023826-2022-04370 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 -9.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO DIZZINESS AND EYE DISCOMFORT. THE EXPLANT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). (B)(4). CLAIM # (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12381179&pc=MTA | 2023826-2021-03389 | 2021-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient Device Interaction Problem | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT FELT DEPRESSED AND UNCOMFORTABLE AFTER THE OPERATION. THE PATIENT ASKED FOR THE LENS TO BE REMOVED AND DID NOT WANT ANOTHER LENS IMPLANTED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12381180&pc=MTA | 2023826-2021-03388 | 2021-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient Device Interaction Problem | Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT FELT DEPRESSED AND UNCOMFORTABLE AFTER THE OPERATION. THE PATIENT ASKED FOR THE LENS TO BE REMOVED AND DID NOT WANT ANOTHER LENS IMPLANTED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749254&pc=MTA | 2023826-2021-04225 | 2021-07-04 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Discomfort; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICMO_12.1 -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCING DISCOMFORT AND PAIN, THE REPORTER STATED " PATIENTS IGNORANT STARRY EYES ARE UNCOMFORTABLE, PAIN". THIS EXPLANT RESOLVED THE PROBLEM. CAUSE WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107540&pc=MTA | 2023826-2022-01096 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Discomfort; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 FOR RELIEF OF EYE PAIN AND DISCOMFORT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8441987&pc=MTA | 2023826-2019-00440 | 2018-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE REPORTER INDICATED THE PATIENT HAD EYE ACHES. PATIENT CODE: 3191 - EYE ACHES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8129216&pc=MTA | 2023826-2018-01760 | 2018-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.00/1.5/107 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. IT WAS REPORTED THAT ON (B)(6) 2018, PATIENT COMPLAINED OF DISCOMFORT. ON (B)(6) 2018 SURGEON TRIED TO REPOSITIONING THE LENS AND NOTICED THE LENS DAMAGED, THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY AND THE PATIENTS CONDITION IS GOOD. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172148&pc=MTA | 2023826-2018-01844 | 2018-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409308&pc=MTA | 2023826-2019-00385 | 2019-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS), IT TORE. THE LENS WAS IMPLANTED ON (B)(6) 2019. REPORTEDLY, THE PATIENT COMPLAINED OF DISCOMFORT AND WHEN THE SURGEON EXCHANGED THE LENS ON (B)(6) 2019 WITH AN ALTERNATE LENS, IT WAS DISCOVERED THAT THE LENS WAS TORN. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595431&pc=MTA | 2023826-2019-00771 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "DISCOMFORT AND ACHE." REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595441&pc=MTA | 2023826-2019-00770 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO "DISCOMFORT AND ACHE." REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, "ACHE"). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881045&pc=MTA | 2023826-2020-00669 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT DISCOMFORT (DUE TO PRESBYOPIA). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10148278&pc=MTA | 2023826-2020-01248 | 2020-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED DISCOMFORT, ON (B)(6) 2020 THE LENS WAS REMOVED AND REPLACED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. IT WAS CLARIFIED ON A FOLLOW UP VISIT, THAT THE SURGEON HAD FOUND A SCRATCH ON THE INTIAL LENS IMPLANT. THE PATIENT ALSO EXPERIENCED A FOREIGN BODY SENSATION. INJECTOR MODEL-MSI-PF; LOT#-UNK. PATIENT CODE 3191: NO CODE AVAILABLE (FOREIGN BODY SENSATION). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5305163&pc=MTA | 2023826-2015-01610 | 2014-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SUBJECTIVE DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). THE PATIENT IS UNDER THE AGE OF (B)(6) YEARS OLD. (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536608&pc=MTA | 2023826-2016-00356 | 2015-12-07 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT COMPLAINED OF SUBJECTIVE DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.86MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5533631&pc=MTA | 2023826-2016-00355 | 2015-12-07 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT COMPLAINED OF SUBJECTIVE DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4) AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS MODEL IS CONTRAINDICATED IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH(ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM. THE PATIENT HAD AN ACD OF 2.87MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6629980&pc=MTA | 2023826-2017-00958 | 2016-12-19 | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/X100 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS DISLOCATION/SUBLUXATION, AND UNCOMFORTABLE. THE SURGEON MENTIONED THAT THE LENS WAS ROTATED 100 DEGREE, AND WAS REPOSITIONED. AFTER REPOSITIONING, THE LENS ROTATED A SECOND TIME AND DISPLACED 100 DEGREE. AS OF THE DAY OF MDR SUBMISSION THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT AND A PIECE OF PAPER STUCK TO THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN, MISSING PIECE OF HAPTIC, AND LENS TORN IN TWO PIECES WITH DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8133688&pc=MTA | 2023826-2018-01718 | 2018-08-25 | Injury | STAAR SURGICAL COMPANY | 2018-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -9.5/+2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO PATIENT'S RIGHT EYE ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED WITH LOW VAULT WAS OBSERVED THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BECAUSE THE PATIENT EXPERIENCED DISCOMFORT. CURRENTLY, THE PATIENT'S CONDITION IS GOOD. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661066&pc=MTA | 2023826-2017-01026 | 2016-11-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ACCORDING TO THE REPORTER, THE DR. FOUND A LIKELY SMALL CRACK ON THE LENS BY THE OBSERVATION, AND THE PATIENT EXPERIENCED UNCOMFORTABLE VISION. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME SIZE AND DIOPTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A SMALL VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RED AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR THE UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516767&pc=MTA | 2023826-2018-00739 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT. REPORTEDLY, DESPITE A UCVA OF 1.0, "THE PATIENT JUST FELT UNCOMFORTABLE." THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXPLANT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7516768&pc=MTA | 2023826-2018-00737 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT. REPORTEDLY, DESPITE A UCVA OF 1.0, "THE PATIENT JUST FELT UNCOMFORTABLE." THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND THE LENS MISSING THE TORN PIECE. THERE WAS PRESENCE OF DEBRIS AND RESIDUE ON THE LENS SURFACE. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7335037&pc=MTA | 2023826-2018-00341 | 2018-01-04 | Injury | STAAR SURGICAL COMPANY | 2018-03-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 VTICM5_12.1, -4.50/+1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO PATIENT BEING DISSATISFIED WITH THE VISUAL ACUITY WHEN "LOOKING AT SOMETHING AT NEAR DISTANCE." THE PATIENT ALSO EXPERIENCED SIGNIFICANT DISCOMFORT. THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS PRODUCT. VISUAL OBSERVATION FOUND THE LENS WITH A BENT HAPTIC AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539111&pc=MTA | 2023826-2018-00763 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO DISCOMFORT. AFTER THE REMOVAL, THERE WAS NO DISCOMFORT. THERE IS NO PLANS YET TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7539112&pc=MTA | 2023826-2018-00764 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXPLANTED DUE TO DISCOMFORT. AFTER THE REMOVAL, THERE WAS NO DISCOMFORT. THERE IS NO PLANS YET TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE A CURVED SHAPE AND PRESENCE OF RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258308&pc=MTA | 2023826-2017-00084 | 2016-10-11 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/154 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6)2016, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, DISCOMFORT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS EXCHANGE RESOLVED THE PROBLEM. AS OF THE DATE OF THIS SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/40 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/40. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254837&pc=MTA | 2023826-2019-00091 | 2018-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Discomfort; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE SURGEON EXPLANTED THE LENS BECAUSE IT IS "TOO BIG" AND THE PATIENT IS UNCOMFORTABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2019, A SHORTER LENS WAS IMPLANTED. THE PROBLEM IS RESOLVED. THE LENS WAS RETURNED IN A PLASTIC BOX, TAPED TO GAUZE. THE LENS WAS STUCK ON THE GAUZE, METHOD CODE ADDED. RESULT CODE UPDATED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304868&pc=MTA | 2023826-2015-01609 | 2014-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information; Appropriate Term/Code Not Available | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SUBJECTIVE DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSION OF INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CONCLUSION - PATIENT IS UNDER THE AGE OF 21 YEARS OLD. DIMENSIONAL INSPECTION INDICATED LENS IS IN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6143453&pc=MTA | 2023826-2016-01715 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY | 2016-12-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Misfocusing | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT HAD AN ICL WITH AN UNKNOWN SERIAL NUMBER AND DIOPTER IMPLANTED IN BANGALORE 28 MONTHS PRIOR TO (B)(6) 2016, AND FOR THE LAST TWO MONTH THE PATIENT HAS EXPERIENCED SOME DISCOMFORT SEEING OBJECTS IN THE RIGHT EYE, WHICH NEED MORE FOCUS. PATIENT VISITED SAME SURGEON WHO IMPLANTED ICL. EXAM INDICATES THAT EYE POWER AT -0.50 AND -0.25 "IN BOTH EYES", SPECIFIC EYE/POWER WAS NOT CLEAR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION : (UNABLE TO CONFIRM COMPLAINT): ATTEMPTED TO CONTACT PATIENT FOR ADDITIONAL INFORMATION, NONE FORTH COMING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6790278&pc=MTA | 2023826-2017-01272 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17/+6/86 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF A PRE-EXISTING CONDITION,KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -16.5/5.5/078 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF PRE-EXISTING CONDITION, KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). DEVICE WAS IMPLANTED ON A PATIENT WITH A PRE-EXISTING CONDITION (KERATOCONUS EYE) WHICH IS CONTRAINDICATED FOR USE IN THE DFU. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6790320&pc=MTA | 2023826-2017-01269 | 2016-12-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/5.5/078 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS BCVA DUE TO PRESENCE OF A PRE-EXISTING CONDITION, KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -17/+6/86 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED DISCOMFORT AND LOSS OF BCVA DUE TO PRESENCE OF PRE-EXISTING CONDITION KERATOCONUS EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). DEVICE WAS IMPLANTED ON A PATIENT WITH A PRE-EXISTING CONDITION (KERATOCONUS EYE) WHICH IS CONTRAINDICATED FOR USE IN THE DFU. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042009&pc=MTA | 2023826-2018-01592 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.5/+2.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO DISCOMFORT. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT IN THE SURGEONS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8042208&pc=MTA | 2023826-2018-01591 | 2018-09-20 | Injury | STAAR SURGICAL COMPANY | 2018-11-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/+1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO DISCOMFORT. THIS EXPLANT RESOLVED THE PROBLEM. IT WAS REPORTED THAT IN THE SURGEONS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172145&pc=MTA | 2023826-2018-01847 | 2018-11-16 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO DISCOMFORT. REPORTEDLY, THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6031493&pc=MTA | 2023826-2016-01444 | 2016-06-13 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Dizziness | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCING DIZZINESS. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN/BROKEN/BENT DEFORMED. FOREIGN MATERIAL (UNIDENTIFIED SPOTS) WAS PRESENT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS MEASUREMENTS WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.93MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182707&pc=MTA | 2023826-2018-01762 | 2018-07-12 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dizziness; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -9.00/+2.0/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT EXPERIENCED DIZZINESS AND FEELING SEASICKNESS AFTER INITIAL SURGERY. IT WAS REPORTED THAT IT WAS NOT REFRACTIVE SURPRISE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8182688&pc=MTA | 2023826-2018-01763 | 2018-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dizziness; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -9.00/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT EXPERIENCED DIZZINESS AND FEELING SEASICK AFTER INITIAL SURGERY. IT WAS REPORTED THAT IT WAS NOT REFRACTIVE SURPRISE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10909725&pc=MTA | 2023826-2020-02742 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s) | P030016 | 1 | Event Description: RECEIVED COMMENT FROM (B)(6) FROM PATIENT WHO REPORTED RIGHT EYE (OD) WAS DRY. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H1: "SERIOUS INJURY" SHOULD BE "MALFUNCTION" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: EXPIRATION DATE: UNK. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10909768&pc=MTA | 2023826-2020-02743 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-11-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s) | P030016 | 1 | Event Description: RECEIVED COMMENT FROM INSTAGRAM FROM PATIENT WHO REPORTED LEFT EYE (OS) WAS DRY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H1: "SERIOUS INJURY" SHOULD BE "MALFUNCTION" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200457&pc=MTA | 2023826-2021-00188 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s) | P030016 | 1 | Event Description: ON 23 DEC 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE ENTITLED, ' VISUAL OUTCOMES AFTER SEQUENTIAL POSTERIOR CHAMBER PHAKIC IOL WITH CORNEAL REFRACTIVE SURGERY (BIOPTICS) FOR THE TREATMENT OF MYOPIC ASTIGMATISM.' THE ARTICLE REFERENCES 3 EYES (2 PATIENTS) REQUIRING PUNCTUAL PLUGS DUE TO DRY EYE WITH THE VISIAN ICL .ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: ON 23-12-2020, A NOTIFICATION OF AN ARTICLE TITLED "VISUAL OUTCOMES AFTER SEQUENTIAL POSTERIOR CHAMBER PHAKIC IOL WITH CORNEAL REFRACTIVE SURGERY (BIOPTICS) FOR THE TREATMENT OF MYOPIC ASTIGMATISM". THE ARTICLE NOTES A 100 EYES STUDY AND PATIENTS EXPERIENCED DRY EYES AFTER PC-PIOL IMPLANTATION; WHICH 3 EYES OF TWO PATIENTS REQUIRED PUNCTAL PLUGS AND ALSO 6 EYES OF THREE PATIENTS AFTER LASER ABLATION SURGERY ALSO EXPERIENCED DRY EYES. IT ALSO MENTIONS PRK AND LASIK WAS PERFORMED ON PATIENT'S IN THE STUDY. THE ARTICLE CONCLUSION INDICATES BIOPTICS FOR HIGH MYOPIC ASTIGMATISM WAS SAFE AND EFFECTIVE. REVERSE BIOPTICS ALTHOUGH NOT TRADITION COULD PROVIDE SIMILAR RESULTS. HEALTH EFFECT IMPACT CODE: 4625 - LASIK/PRK. PUNCTAL PLUGS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15583914&pc=MTA | 2023826-2022-03589 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Dry Eye(s) | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VICMO12.6, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) DRYNESS WAS REPORTED. FOR PATIENTS CURRENT CONDITION IT WAS REPORTED THAT THE EYE IS S/P ICL AND DOING WELL, NO INTERVENTION WAS NECESSARY AND IT HAS IMPROVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5171627&pc=MTA | 2023826-2015-01391 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Dry Eye(s); Blurred Vision; Halo; No Code Available | P030016 | 1 | Event Description: A PATIENT REPORTED THROUGH (B)(6) HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED AND IS EXPERIENCING DRY, ITCHY EYE. VISION IS BLURRY AND NOT AS SHARP. ALSO GLARE AND HALOS AT NIGHT. THE PATIENT HAS REPORTED VISION HAS IMPROVED OVER TIME. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.2 MICL13.2 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT HAS DRY EYES AND HAS BEEN PRESCRIBED DROPS. AT THE POST-OP VISIT ON (B)(6) 2016, THE PATIENT STILL REPORTED HALOS/GHOSTING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. PER MEDICAL REVIEW - THE PATIENT REPORTED THAT HE WAS EXPERIENCING SUBJECTIVE VISUAL DISTURBANCES ("HALOS/GLARE''), ESPECIALLY DURING NIGHT TIME, FOLLOWING ICL IMPLANTATION ONE WEEK BEFORE. NO REPORT OF SECONDARY SURGICALLY OR MEDICALLY INTERVENTION PERFORMED. LENS REMAINS IMPLANTED. ACCORDING TO REPORT, FROM THE FACILITY, DATED ON (B)(6) 2016 ICL WAS IMPLANTED BILATERALLY AT THE SAME SURGERY DATE AND THE EXACT CAUSE OF THE EVENT WAS NOT PROVIDED. THE PATIENT HAD ASTIGMATISM AS WELL AS DRY EYE SYMPTOMS SO HE WAS ADVISED TO CONSIDER PRK FOR ASTIGMATISM AND TO USE EYE DROPS MORE THAN TWICE A DAY. VISUAL DISTURBANCES HAVE BEEN IDENTIFIED AS POTENTIAL COMPLICATIONS AFTER ICL IMPLANTATION. THE OPTICAL DIAMETER OF THE ICL RANGES FROM 4.9 MM TO 5.8 MM AND THE PRECAUTIONS SECTION OF THE DFU INSTRUCTS PHYSICIANS THAT "PRIOR TO SURGERY, THE SURGEON MUST PROVIDE PROSPECTIVE PATIENTS WITH A COPY OF THE PATIENT INFORMATION BOOKLET FOR THIS PRODUCT AND INFORM THESE PATIENTS OF THE POSSIBLE BENEFITS AND COMPLICATIONS ASSOCIATED WITH THE USE OF THIS DEVICE". IT FURTHER PRECAUTIONS THAT "THE EFFECT OF PUPIL SIZE ON VISUAL SYMPTOMS IS NOT KNOWN". THE DFU INDICATES THE "SMALLER THE OPTIC DIAMETER, THE GREATER THE INCIDENCE OF SUBJECTIVE PATIENT SYMPTOMS". REASSESSMENT WILL BE PERFORMED WHEN (IF) ADDITIONAL INFORMATION IS RECEIVED FROM THE IMPLANTING SURGEON. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980211&pc=MTA | 2023826-2020-02954 | 2017-10-01 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Dry Eye(s); Fatigue; Headache; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT REPORTS "HEADACHES DUE TO DRY EYE, MGD." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EYELID ECZEMA SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED BY THE PATIENT THAT THEY ALSO EXPERIENCED "BAD INTERACTION WITH HEAD", FATIGUE (TIREDNESS), DISCOMFORT AND "PULLING OF THE EYE AND FACE". PATIENT IS SEEKING TO HAVE THE LENS REMOVED. LENS REMAINS IMPLANTED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - "BAD INTERACTION WITH HEAD", "PULLING OF EYE AND FACE". CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; SUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10980212&pc=MTA | 2023826-2020-02955 | 2017-10-01 04:00 | Malfunction | STAAR SUGICAL COMPANY | 2020-12-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Dry Eye(s); Fatigue; Headache; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS LEFT EYE (OS). THE PATIENT REPORTS "HEADACHES DUE TO DRY EYE, MGD." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - EYELID ECZEMA, SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED BY THE PATIENT THAT THEY ALSO EXPERIENCED "BAD INTERACTION WITH THE HEAD", FATIGUE (TIREDNESS), DISCOMFORT AND "PULLING OF THE EYE AND FACE". PATIENT IS SEEKING TO HAVE THE LENS REMOVED. LENS REMAINS IMPLANTED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - " BAD INTERACTION WITH THE HEAD", "PULLING OF THE EYE AND FACE". CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15375297&pc=MTA | 2023826-2022-03042 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Dry Eye(s); Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE PATIENT INDICATED THAT THEY HAD AN IMPLANTABLE COLLAMER LENS IMPLANTED INTO THEIR EYE. THE PATIENT REPORTED THAT THEY HAD EYE FLOATERS; HALOES; AND EYE DRYNESS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - EYE FLOATERS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9580722&pc=MTA | MW5092155 | 2017-09-23 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 | MTA | ICL LENSE (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Headache; Pain |  | 1 | Event Description: THE DAY AFTER SURGERY ONE OF MY PUPILS STOOD DILATED FOR ABOUT 6 MONTHS I COULD NOT DRIVE NOR HAVE BRIGHT LIGHTS ON, I REQUIRED PILOCARPINE DROPS IN CONSTRICT CV THE PUPIL. I ALSO HAVE HORRIBLE DRY EYES, I HAD TO GET PLUGS AND ALSO DEVELOPED HORRIBLE MIGRAINES WHICH I TAKE NUMEROUS MEDICINE FOR. I HAVE CONSTANT EYE PAIN, DRYNESS, AND HORRIBLE MIGRAINE HEADACHES; IF YOU HAD THE MONEY I WOULD TAKE THEM OUT. I HAVE NO SOCIAL LIFE, I'M IN CONSTANT PAIN AND THE SIDE EFFECTS FROM THE MEDICINE IS ALSO BAD. I HAVE HAD NO RELIEF AND WISH I NEVER HAD THEM. REASON FOR USE: CORRECT VISION. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753270&pc=MTA | 2023826-2020-00365 | 2017-09-24 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Headache; Pain; No Code Available | P030016 | 1 | Event Description: PATIENT COMPLAINT SUBMITTED TO FDA-THE PATIENT STATES AN ICL WAS IMPLANTED ON (B)(6) 2017. REPORTEDLY, THE DAY AFTER SURGERY THE PUPIL STAYED DILATED FOR ABOUT 6 MONTHS. PATIENT REPORTS CONSTANT EYE PAIN, DRYNESS, AND MIGRAINE HEADACHES. ALSO REPORTED: UNABLE TO DRIVE NOR HAVE BRIGHT LIGHTS ON, PILOCARPINE DROPS REQUIRED ALONG WITH VARIOUS OTHER MEDICATION PRESCRIBED FOR MIGRAINES, DRY EYES THAT HAD TO GET "PLUGS" FOR, SIDE EFFECTS FROM MEDS ARE "BAD." THE LENS REMAINS IMPLANTED. UNABLE TO OBTAIN ADDITIONAL INFORMATION. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753277&pc=MTA | 2023826-2020-00366 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-26 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Headache; Pain; No Code Available | P030016 | 1 | Event Description: PATIENT COMPLAINT SUBMITTED TO FDA-THE PATIENT STATES AN ICL WAS IMPLANTED ON (B)(6) 2017. PATIENT REPORTS CONSTANT EYE PAIN, DRYNESS, AND MIGRAINE HEADACHES. ALSO REPORTED: UNABLE TO DRIVE NOR HAVE BRIGHT LIGHTS ON, VARIOUS MEDICATIONS PRESCRIBED FOR MIGRAINES, DRY EYES THAT HAD TO GET "PLUGS" FOR, SIDE EFFECTS FROM MEDS ARE "BAD." THE LENS REMAINS IMPLANTED. UNABLE TO OBTAIN ADDITIONAL INFORMATION. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774155&pc=MTA | 2023826-2019-01205 | 2015-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Misfocusing | Dry Eye(s); Headache; Pain; Red Eye(s); Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S LEFTT EYE (OS) IN (B)(6) OF 2015. HYPERMETROPIA OF +1, ASTIGMATISM OF "AROUND" 0.5, PAIN, AND RED EYE IS REPORTED. LASIK SURGERY WAS PERFORMED (B)(6) 2017 WHICH RESOLVED THE PROBLEM. HOWEVER, RED EYE STILL OCCURS APPROXIMATELY EVERY 2 MONTHS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: "HEADACHE, DRY EYE AND NAUSEA IS REPORTED." SHOULD BE ADDED TO THE INITIAL MDR. PATIENT CODE 3191: NAUSEA. (B)(4). Manufacturer Narrative: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8574636&pc=MTA | 2023826-2019-00729 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A 13.2MM, MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IMPLANTED INTO HIS RIGHT EYE (OD) ON (B)(6) 2018. THE PATIENT STATES THAT HE IS EXPERIENCING CHRONIC EYE SYNDROME SINCE THE IMPLANT WAS PLACED. HE INDICATED THAT HE DID NOT HAVE THIS PRIOR TO THE ACTUAL SURGERY. HE STATES THAT WHEN HE TAPES HIS EYE SHUT THE CONDITION IS NOT AS SEVERE AS WHEN HE SLEEPS WITHOUT THE TAPE, WHEN HE SLEEPS HIS EYE IS SLIGHTLY SLIT. Manufacturer Narrative: "CHRONIC EYE SYNDROME" SHOULD BE CHANGED TO "CHRONIC DRY EYE SYNDROME." CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL SYMPTOMS REPORTED: PAINFUL DRY EYE AND ELEVATED IOP. H6-INVESTIGATION FINDINGS CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "TREATMENT INCLUDES SERUM TEAR THERAPY; MOISTURE CHAMBER GOGGLES; FISH OIL; AND TAPING OF EYELIDS." SHOULD BE ADDED. H6 - 4650 - SERUM TEAR THERAPY. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575042&pc=MTA | 2023826-2019-00730 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: THE PATIENT REPORTED THAT HE HAD A 13.2MM, MICL13.2, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IMPLANTED INTO HIS LEFT EYE (OS) ON (B)(6) 2018. THE PATIENT STATES THAT HE IS EXPERIENCING CHRONIC EYE SYNDROME SINCE THE IMPLANT WAS PLACED. HE INDICATED THAT HE DID NOT HAVE THIS PRIOR TO THE ACTUAL SURGERY. HE STATES THAT WHEN HE TAPES HIS EYE SHUT THE CONDITION IS NOT AS SEVERE AS WHEN HE SLEEPS WITHOUT THE TAPE, WHEN HE SLEEPS HIS EYE IS SLIGHTLY SLIT. Manufacturer Narrative: EVENT DESCRIPTION: "CHRONIC EYE SYNDROME" SHOULD BE CHANGED TO "CHRONIC DRY EYE SYNDROME." CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL SYMPTOMS REPORTED: PAINFUL DRY EYE AND ELEVATED IOP. H6-INVESTIGATION FINDINGS CODE CORRECTED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5. "TREATMENT INCLUDES SERUM TEAR THERAPY; MOISTURE CHAMBER GOGGLES; FISH OIL; AND TAPING OF EYELIDS." SHOULD BE ADDED. H6. 4650 SERUM TEAR THERAPY. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936812&pc=MTA | 2023826-2020-02828 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE PATIENT REPORTS ELEVATED IOP AND PAINFUL DRY EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ON (B)(6) 2020, IT WAS DETERMINED THAT THE CURRENT CASE IS A DUPLICATE OF MFR REPORT# 2023826-2019-00730. PLEASE DISREGARD THE CURRENT REPORT. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936875&pc=MTA | 2023826-2020-02827 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Eye Pain | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE PATIENT REPORTS ELEVATED IOP, AND PAINFUL DRY EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ON (B)(6) 2020, IT WAS DETERMINED THAT THIS CASE IS A DUPLICATE OF MFR REPORT# 2023826-2019-00729. PLEASE DISREGARD THE CURRENT REPORT. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9720640&pc=MTA | MW5093020 | 2019-12-05 | Injury | STARR SURGICAL COMPANY | 2020-02-14 | MTA | IMPLANTABLE CONTACT LENSES | Patient-Device Incompatibility | Dry Eye(s); Intraocular Pressure Increased; Pain |  | 1 | Event Description: I HAD STARR VISION ICL'S IMPLANTED IN MY EYES. ONE WAS DONE (B)(6) 2019 AND THE OTHER (B)(6) 2019. I HAVE HAD EYE PAIN EVER SINCE. THE SURGEON REMOVED THEM FEBRUARY (B)(6) AND (B)(6). I AM STILL HAVING PAIN. I CANNOT BE OUTSIDE IN THE SUN OR WIND. WHEN I AM INSIDE I HAVE TO KEEP THE LIGHTS DIMMED. I'VE BEEN TO MULTIPLE DRS AND SEEMS TO BE DISMISSED. NO ONE CAN FIGURE OUT WHAT IS WRONG. I WAS TREATED FOR DRY EYE FOR OVER MONTH AND I DO NOT BELIEVE THAT IS THE ROOT CAUSE OF THE PAIN. AFTER THE SURGERY MY EYE PRESSURE INCREASED AND I WAS ON COMBIGAN TO HELP. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9772182&pc=MTA | 2023826-2020-00374 | 2019-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. AT ONE-DAY POST-OP, THE PATIENT REPORTED EYE PAIN, WHICH WAS NOT RESOLVED WITH ARTIFICIAL TEARS OR WARM COMPRESSES. THE PATIENT GOT A SECOND OPINION BUT WAS UNABLE TO SOLVE THE SOURCE OF THE PAIN. THE PATIENT REQUESTED THE LENS BE REMOVED AND THE LENS WAS EXPLANTED ON (B)(6) 2020. THERE ARE NO PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT REPORTED POST-OP, HER EYE PRESSURE HAD INCREASED AND WAS TREATED WITH COMBIGAN. WAS ALSO TREATED FOR DRY EYE FOR OVER A MONTH. THE PATIENT'S CURRENT CONDITION STILL HAS OCCASIONAL PAIN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAG, IN A VIAL WITH LIQUID AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: H6 - PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: COMBIGAN) SHOULD HAVE BEEN INCLUDED IN INTIAL MEDWATCH REPORT. ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 1937 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. B5 - IT WAS INDICATED FROM THE REPORTER THAT PRIOR TO SURGERY, IN (B)(6) 2019, THE PATIENT HAD A "HAIR CLIP INJURY" WHICH RESULTED IN A BRUISED EYE LID. DOCTOR FOUND NO EVIDENCE OF AN INJURY TO THE EYE ITSELF. PATIENT RECEIVED SEVERAL OTHER DOCTORS OPINIONS AND AT SOME POINT WAS PRESCRIBED CAMBIGAN TO TREAT ELEVATED IOP. COMBIGAN WAS NOTED TO HAVE BEEN FIRST USED ON (B)(6) 2020 AND WAS DISCONTINUED ON (B)(6) 2020 AND THEN AGAIN RESUMED ON (B)(6) 2020 WHICH REPORTED THE PATIENT HAD 13MMHG EYE PRESSURE. THE DAY BEFORE THEY FOUND HER EYE PRESSURE TO BE 23MMHG. PATIENT CONTINUED TO USE COMBIGAN AS OF (B)(6) 2020 AND HAS BEEN TO A CORNEAL SPECIALIST. WHO IS INVESTIGATING THE POSSIBILITY OF OPHTHALMIC NEURALGIA. CLAIM#: (B)(4). Manufacturer Narrative: REPORT NUMBER - MW5093020. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9772179&pc=MTA | 2023826-2020-00375 | 2020-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, - 11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE PATIENT REPORTED EYE PAIN, WHICH WAS NOT RESOLVED WITH ARTIFICIAL TEARS OR WARM COMPRESSES. THE PATIENT GOT A SECOND OPINION BUT WAS UNABLE TO SOLVE THE SOURCE OF THE PAIN. THE PATIENT REQUESTED THE LENS BE REMOVED AND THE LENS WAS EXPLANTED ON (B)(6) 2020. THERE ARE NO PLANS TO IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT REPORTED POST-OP, HER EYE PRESSURE HAD INCREASED AND WAS TREATED WITH COMBIGAN. WAS ALSO TREATED FOR DRY EYE FOR OVER A MONTH. THE PATIENT'S CURRENT CONDITION IS STILL HAS OCCASIONAL PAIN, ESPECIALLY AT NIGHT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL IN A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6- DEVICE HISTORY REVIEW: THE DEVICE HAS BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- PATIENT CODE 1937 SHOULD BE REMOVED FROM INITIAL MDR AS IT IS NOT APPLICABLE. B5- IT IS ALSO NOTED THAT PRIOR TO SURGERY, IN (B)(6) OF 2019 SHE HAD A "HAIR CLIP INJURY" IN THE RIGHT EYE THAT RESULTED IN A BRUISED EYELID. THE DOCTOR FOUND NO EVIDENCE OF AN INJURY TO THE EYE ITSELF. PATIENT WENT TO SEVERAL OTHER DOCTORS FOR OPINIONS AND WAS AT SOME POINT PLACED ON COMBIGAN TO TREAT ELEVATED PRESSURE RELATED TO STEROID DROP USE. COMBIGAN WAS NOTED TO HAVE BEEN FIRST USED ON (B)(6) 2020 IN THE RIGHT EYE ONLY. COMBIGAN WAS USED IN BOTH EYES UNTIL (B)(6) 2020, AND WAS DISCONTINUED ON (B)(6) 2020 AND THEN AGAIN RESUMED ON (B)(6) 2020 WHICH REPORTED THE PATIENT HAD 34MMHG EYE PRESSURE IN THE LEFT EYE. THE DAY BEFORE THEY FOUND HER EYE PRESSURE TO BE 23MMHG. PATIENT CONTINUED USING COMBIGAN AS OF (B)(6) 2020 AND HAS BEEN TO A CORNEAL SPECIALIST WHO IS INVESTIGATING THE POSSIBILITY OF OPHTHALMIC NEURALGIA. CLAIM# (B)(4). Manufacturer Narrative: REPORT NUMBER - MW5093020. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6375130&pc=MTA | 2023826-2017-00253 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED BLOCKED DUCTS AND THE EYE TEARS UP NON-STOP. THE LENS WAS IMPLANTED ON (B)(6) 2013. THE FACILITY REPORTED PATIENT WAS LAST SEEN (B)(6) 2015, WHERE THE OPHTHALMOLOGIST PLACED SIL PLUGS TO TREAT DRY EYE. THE FACILITY REPORTED THE EVENT WAS NOT DEVICE RELATED. THE PATIENT'S VA POST-OP WAS 20/20. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553057&pc=MTA | 2023826-2016-00420 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN HER RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT REPORTED DRY EYE AND MEIBOMIAN GLAND ATROPHY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE - (B)(6) 2012. ADDITIONAL INFORMATION RECEIVED - THE FACILITY CONFIRMED THE PATIENTS REPORT AND INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT HAD TREATMENT FOR THE DRY EYE AND THERE WERE NO OTHER COMPLICATIONS/HEALTH EFFECTS RELATED TO THE EVENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553220&pc=MTA | 2023826-2016-00421 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.1MM MICL12.1 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT REPORTED DRY EYE AND MEIBOMIAN GLAND ATROPHY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE: (B)(6) 2012. ADDITIONAL INFORMATION RECEIVED. THE FACILITY CONFIRMED THE PATIENTS REPORT AND INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT HAD TREATMENT FOR THE DRY EYE AND THERE WERE NO OTHER COMPLICATIONS/HEALTH EFFECTS RELATED TO THE EVENT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495100&pc=MTA | 2023826-2023-03433 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Uveitis | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED LONG-TERM OUTCOMES AND SAFETY OF THE PHAKIC VISIAN TORIC IMPLANTABLE COLLAMER LENS IN EYES WITH NON-PROGRESSIVE KERATOCONUS WHICH COVERED THE OUTCOME INDICATED A TOTAL OF 52 EYES OF 35 PATIENT'S WERE INCLUDED IN THIS STUDY. OUT OF THESE EYES/PATIENTS 4 EYES OF TWO PATIENTS HAD MILD DRYNESS AND MILD ANTERIOR UVEITIS WHICH PERSISTED FOR 2 MONTHS. IT WAS NOTED THERE WAS NO SERIOUS COMPLICATIONS AND NONE OF THE LENSES WERE EXPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (KERATOCONUS). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284950&pc=MTA | 2023826-2017-00182 | 2012-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED GLARE AROUND LIGHT AT NIGHT. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY REPORTED DATE OF THE LAST VISIT WAS (B)(6) 2013. THE FACILITY INDICATED THE PATIENT REPORTED DRY EYE ON (B)(6) 2012. THE CONDITION WAS RESOLVED BY (B)(6) 2012 WITH ARTIFICIAL TEARS. NO OTHER CONDITIONS WERE NOTED AND PATIENTS VISUAL ACUITY POST-OP WAS 20/25. THE FACILITY INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. Manufacturer Narrative: PATIENT AGE: (B)(6) 1982. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6284953&pc=MTA | 2023826-2017-00181 | 2012-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-01-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED NIGHTTIME GLARE AROUND LIGHTS. THE LENS WAS IMPLANTED ON (B)(6) 2012 AND REMAINS IMPLANTED. THE FACILITY REPORTED DATE OF THE LAST VISIT WAS (B)(6) 2013. THE FACILITY INDICATED THE PATIENT REPORTED DRY EYE ON (B)(6) 2012. THE CONDITION WAS RESOLVED BY (B)(6) 2012 WITH ARTIFICIAL TEARS. NO OTHER CONDITIONS WERE NOTED AND PATIENTS VISUAL ACUITY POST-OP WAS 20/25. THE FACILITY INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. Manufacturer Narrative: (B)(6). WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853774&pc=MTA | 2023826-2017-01410 |  | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HER EYESIGHT WAS NOT AS GOOD. THE PATIENT WAS USING EYE DROPS DUE TO EYES WERE DRY. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE: OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853794&pc=MTA | 2023826-2017-01411 |  | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Dry Eye(s); Visual Disturbances | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED HER EYES HAD CHANGED. THE PATIENT WAS USING EYE DROPS DUE TO EYES WERE DRY. THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CONCLUSION CODE: - OFF LABEL USE (PATIENT ABOVE 45 YEARS OF AGE AT TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652097&pc=MTA | 2023826-2016-00547 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); Visual Impairment | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED NOT BEING ABLE TO SEE TO DRIVE AT NIGHT AND HAS HAD TO GET GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE FACILITY REPORTED THEY COULD NOT CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(4) 2011 AND VISUAL ACUITY WAS 20/20-2. THE REPORTER INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THERE WERE NO OTHER OCULAR COMPLICATIONS AND NO MEDICATIONS OR TREATMENTS WERE REQUIRED. THE PATIENT HAD DRY EYES AND THE VISION WOULD FLUX DAILY DUE TO DRYNESS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5652100&pc=MTA | 2023826-2016-00548 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Dry Eye(s); Visual Impairment | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED NOT BEING ABLE TO SEE TO DRIVE AT NIGHT AND HAS HAD TO GET GLASSES. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THEY COULD NOT CONFIRM THE PATIENT REPORT. THE DATE OF LAST VISIT WAS (B)(6) 2011 AND VISUAL ACUITY WAS 20/20-2. THE REPORTER INDICATED THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. THERE WERE NO OTHER OCULAR COMPLICATIONS AND NO MEDICATIONS OR TREATMENTS WERE REQUIRED. THE PATIENT HAD DRY EYES AND THE VISION WOULD FLUX DAILY DUE TO DRYNESS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO. LENS REMAINS IMPLANTED. (B)(4). LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7555304&pc=MTA | 2023826-2018-00832 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT RELATED FACTOR OF VERTIGO. THE VERTIGO WAS PRE-EXISTING TO ICL IMPLANTATION. AT ONE DAY POST-OP, THE CRYSTALLINE LENS WAS CLEAR AND THERE WAS 1+ INCISIONAL EDEMA. THE VERTIGO HAS BEEN REDUCED AND PATIENT WILL RETURN TO C/L WEAR. THE PATIENT ELECTED TO HAVE THE ICL EXPLANTED . Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S.BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7555311&pc=MTA | 2023826-2018-00831 | 2018-04-04 | Injury | STAAR SURGICAL COMPANY | 2018-05-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT RELATED FACTOR OF VERTIGO. THE VERTIGO WAS PRE-EXISTING TO ICL IMPLANTATION. AT ONE DAY POST-OP, THE CRYSTALLINE LENS WAS CLEAR AND THERE WAS 1+ INCISIONAL EDEMA. THE VERTIGO HAS BEEN REDUCED AND PATIENT WILL RETURN TO C/L WEAR. THE PATIENT ELECTED TO HAVE THE ICL EXPLANTED . Manufacturer Narrative: CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S.BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6278117&pc=MTA | 2023826-2017-00073 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-01-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.00/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. THE PATIENT HAD SOME MILD WOUND EDEMA POST-OP. Manufacturer Narrative: IT SHOULD BE NOTED THAT AT THE TIME OF THE ORIGINAL IMPLANTATION THE PATIENT WAS (B)(6) YEARS OLD AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR STICKY RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6278120&pc=MTA | 2023826-2017-00074 | 2016-11-16 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Edema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -13.5/+1.00/078 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. THE PATIENT HAD SOME MILD WOUND EDEMA POST-OP. Manufacturer Narrative: IT SHOULD BE NOTED THAT AT THE TIME OF THE ORIGINAL IMPLANTATION THE PATIENT WAS (B)(6) AND PER DFU INDICATIONS: "VISIAN® IMPLANTABLE COLLAMER® LENSES (VICL) ARE INDICATED FOR USE IN PHAKIC EYE TREATMENTS IN ADULTS 21-45 YEARS OF AGE", AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN PATIENTS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DEBRIS, CLEAR AND GREY COLORED RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXCHANGED FOR A LONGER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942490&pc=MTA | 2023826-2023-01860 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON 28-MAR-2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE DEVELOPMENT OF ENDOPHTHALMITIS. THE PATIENT WAS TRANSFERRED TO GENERAL HOSPITAL AND THE FOLLOWING TREATMENT WAS PERFORMED - BROMFENAC SODIUM HYDRATE, FLUOROMETHOLONE AND VIGAMOX 0.5% (5ML). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - BROMFENAC SODIUM HYDRATE, FLUOROMETHOLONE, VIGAMOX 0.5% (5ML) H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17950145&pc=MTA | 2023826-2023-04566 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO ENDOPHTHALMITIS. NO DIAGNOSTICS WERE PERFORMED. TREATMENT WAS PERFORMED. REPORTEDLY, "WE TOOK OUT THE LENS AND STRENGTHENED ANTI-INFLAMMATORY TREATMENT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135965&pc=MTA | 2023826-2021-00038 |  | Injury | STAAR SURGICAL COMPANY | 2021-01-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS AGGRESSIVE INFLAMATION AND BELIEVES IT TO BE ENDOPHTHALMITIS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18419441&pc=MTA | 2023826-2023-05833 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis | P030016 | 1 | Event Description: ENDOPHTHALMITIS POST ICL IMPLANTATION WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355967&pc=MTA | 2023826-2016-00031 | 2015-12-09 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Endophthalmitis | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -5.0/+1.0/X088 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. SUSPECTED ENDOPHTALMITIS WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. POSSIBLY PLAN TO EXCHANGE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -5.00/+1.0/X088. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16644574&pc=MTA | 2023826-2023-01124 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Endophthalmitis | UNK | 1 | Event Description: REPORTEDLY DURING A AECOS MEETING IN COLORADO, "A US SURGEON FROM THE AUDIENCE COMMENTED DURING AN ICL PATIENT SELECTION DISCUSSION THAT HE HAD PREVIOUSLY TREATED A PATIENT WITH ICL WHO DEVELOPED A STAPH INFECTION LEADING TO ENDOPHTHALMITIS AND ENUCLEATION. HE DID NOT COMMENT FURTHER REGARDING TIMING OF WHEN THIS OCCURRED, BUT IT APPEARED TO BE LONG AGO WITH VISIAN ICL". Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233918&pc=MTA | 2023826-2021-02880 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Endophthalmitis | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020 . ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO ENDOPHTHALMITIS BEING OBSERVED. NO DIAGNOSTICS WAS PERFORMED. ANTI-INFLAMMATORY WAS PERSCRIBED. PROBLEM RESOLVED. REPORTEDLY, "THERE WAS A LATE SUSPECT OF ENDOPHTHALMITIS/UVEITIES AFTER LENS IMPLANTED." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633180&pc=MTA | 2023826-2019-00880 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. INCIDENT DATE THE SAME DAY, THE SURGEON REPORTS ENDOPHTHALMITIS. EYE DROPS AND "INTERNAL REMEDY" PERFORMED. PATIENT GIVE STEROID INSTALLATION EVERY HOUR, 6 TABLETS OF PREDNISONE FOR 5 DAYS, AND DECADRON SUBCONJUNCTIVAL INJECTION. THE INFLAMMATION WAS THEN RELIEVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL SYMPTOM OF BLURRED VISION IS REPORTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633248&pc=MTA | 2023826-2019-00879 | 2019-02-22 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -12.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. INCIDENT DATE THE SAME DAY, THE SURGEON REPORTS ENDOPHTHALMITIS AND BLURRED VISION. EYE DROPS AND "INTERNAL REMEDY" PERFORMED. PATIENT GIVE STEROID INSTALLATION EVERY HOUR, 6 TABLETS OF PREDNISONE FOR 5 DAYS. THE INFLAMMATION WAS THEN RELIEVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180427&pc=MTA | 2023826-2019-01900 | 2019-03-18 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. IT WAS REPORTED THE PATIENT HAD BLURRED VISION ONE WEEK AFTER SURGERY AND THE SURGEON SUSPECTED POSTOPERATIVE ENDOPHTHALMITIS. THE SURGEON PERFORMED AN ANTERIOR CHAMBER WASHOUT AND INFUSED THE ANTERIOR CHAMBER WITH BSS CONTAINING VEGAMOX DILUTED 20-FOLD. THE INFLAMMATION SUBSIDED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5167176&pc=MTA | 2023826-2015-01272 | 2014-12-17 | Injury | STAAR SURGICAL COMPANY | 2015-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Endophthalmitis; Blurred Vision; No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER -10.00 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT EXPERIENCED AN INFLAMED EYE AND BLURRED VISION. SUSPECTED ENDOPHTALMITIS WAS DIAGNOSIS AND THE LENS WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.00, (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WAS ONE SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948891&pc=MTA | 2023826-2023-04438 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Endophthalmitis; Blurred Vision; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5 12.1. -13.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON 23/SEP/2023. THE LENS WAS EXPLANTED DUE TO EDOPHTHALMITIS, BLURRED VISION, AND TOXIC ANTERIOR SEGMENT SYNDROM(TASS) ON 28/SEP/2023. PATIENT WAS GIVEN INTRAVITREAL ANTIBIOTICS AND VITREOUS TAP WAS STERILE. CAUSE OF EVENT REPORTED AS UNKNOWN. REPORTEDLY, INFLAMMATION REDUCED SIGNIFICANTLY POST EXPLANT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18302000&pc=MTA | 2023826-2023-05513 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: REPORTER INDICATED THAT PATIENT ETURNED TO CLINIC DUE TO OCULAR DISCOMFORT. NOTED FIBROTIC MATERIAL AROUND THE LENS AND HM VISION WAS OBSERVED. THE PATIENT WAS 20/25 ON DAY 1. THE VITREOUS EXCESSED CLEANED BY B SCAN. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM (D4). H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4922992&pc=MTA | 2023826-2015-00830 |  | Injury | STAAR SURGICAL COMPANY AG | 2015-07-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Endophthalmitis; Headache; Intraocular Pressure Increased; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED IN 2013 DUE TO ENDOPHTHALMITIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO (OTHER): LENS NOT RETURNED. (B)(4). OTHER TEXT : LENS NOT RETURNED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE PATIENT EXPERIENCED ACUTE ELEVATED IOP WITH HEADACHE, OCULAR PAIN, VOMITING, PHOTOPHOBIA, AND IRIS PROLAPSE. THIS LED TO WOUND OR CORNEAL ABSCESS WHICH CONSEQUENTLY LED TO ENDOPHTHALMITIS. THE EYE DEVELOPED PHTISIS BULBI AND HAS UNDERGONE OCULAR PROSTHESIS. THE LENS WAS EXPLANTED ON (B)(6) 2013. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(6) - PHOTOPHOBIA. PATIENT CODE: (B)(6) - CORNEAL ABSCESS. PATIENT CODE: (B)(6)- IRIS PROLAPSE. PATIENT CODE: (B)(6)- PHTISIS BULBI. CORRECTED DATA: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9483580&pc=MTA | 2023826-2019-02411 | 2019-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Headache; Pain; Red Eye(s); No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. REPORTEDLY, PATIENT "HAD RED LEFT EYE WITH PAIN AND HEADACHES. SLIT LAMP EXAMINATION SHOWED EXUDATION OF ANTERIOR CHAMBER, THE SURGEON WASHED THE ANTERIOR CHAMBER AND REMOVED THE LENS ON (B)(6) 2019. AN ANTI-INFLAMMATORY THERAPY WAS TAKEN. BACTERIAL CULTURE SHOWED POSITIVE (ANTERIOR CHAMBER AQUEOUS HUMOR) . NOW THE PATIENT IS IN GOOD CONDITION AND HIS VISUAL ACUITY HAS RECOVERED TO PRE-OPERATIVE UCVA AND BCVA IN THE LAST VISIT." CAUSE OF EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - 'ANTIBIOTICS (CERFUROXIME) WERE PRESCRIBED AS TREATMENT.' SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. H6 - PATIENT CODE 3191: MEDICAL INTERVENTION - ANTIBIOTIC TREATMENT (CEFUROXIME), ANTERIOR CHAMBER WASHOUT SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INPECTION FOUND HAPTIC TORN. METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465141&pc=MTA | 2023826-2022-03169 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hyphema; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ENDOPHTHALMITIS. DIAGNOSTICS WERE PERFORMED. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND A VITRECTOMY WERE PERFORMED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: POST-OP, THE PATIENT REPORTED THAT THERE WAS A BLACK OBJECT IN HIS RIGHT EYE. THE EYE FELT FOGGY AND HYPHEMA WAS FOUND. THE LENS WAS REMOVED AND HORMONE THERAPY WAS GIVEN. THE CONDITION DID NOT IMPROVE THE NEXT DAY AND A VITRECTOMY WAS PERFORMED. THE PATIENT IS UNDER OBSERVATION AND TREATMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - FOGGY. H6: HEALTH EFFECT IMPACT CODE: 4644 - HORMONE THERAPY. H6: DEVICE HISTORY RECORD (DHR) REVIEW: 3331 - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: B5: POST-OP, THE PATIENT WAS IMPLANTED WITH SILICONE OIL IN THE GLASSECTOMY SURGERY, THE BEST CORRECTED VISUAL ACUITY WAS CORRECTED TO 20/33 IN THE REVIEW. THE LATEST REVIEW FEEDBACK SHOWED THAT THE PATIENT'S OWN POSTERIOR CAPSULAR CRYSTALS WERE OPACITIES, AND CATARACT SURGERY WOULD BE PERFORMED IN THE FUTURE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. HEALTH IMPACT - ADDITIONAL SURGERY: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7853704&pc=MTA | 2023826-2018-01352 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-09-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Hypopyon; Inflammation; Intraocular Pressure Increased; Pain; Red Eye(s); Blurred Vision; Vitritis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2018 AND OBSERVED TO HAVE INTRAOCULAR INFLAMMATION AND VITRITIS 1 DAY POST-OP. UPON EXAMINATION, THE SURGEON " FOUND REACTION AND MEMBRANES IN AC" AND "WAS OBLIGATED TO KEEP PATIENT INSIDE THE HOSPITAL". THE PATIENT "UNDERGOES INTRA-VITREAL ANTIBIOTIC INJECTION" AND " THE DOCTOR IS WORRIED ABOUT ENDOPHTHALMITIS". PER (B)(6) 2018 EMAIL FROM REPORTER, THE SURGEON HAS EXPLANTED THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). "NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED ON 11 DEC 2018. THE CULTURE IS (B)(6). PATIENT IS STABLE NOW AND THE PROBLEM IS RESOLVED AFTER INTRAVITREAL INJECTION. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION EVENT- ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6)2018: "FEMALE PATIENT SEEKING REFRACTIVE SURGERY RT UNEVENTFUL ICL IMPLANTATION WAS DONE ROUTINE POST OPERATIVE TTT VIGAMOX ED 1×5 ECONOPRED ED 1×5 2ND DAY POST OPERATIVE THE PATIENT PRESENTED WITH DIMINUTION OF VISION, PAIN AND RED EYE. BY EXAM CILIARY INJECTION, HYPOPYON IN AC, HIGH IOP 32, PUPILLARY MEMBRANE. FUNDUS NOT SEEN. ADMISSION OF THE PATIENT U/S WAS FREE. TREAT AS TASS FREQUENT TOPICAL STEROIDS, HOURLY TOPICAL ANTIBIOTICS, CYCLOPLIC, TOPICAL ANTIGLAUCOMA. BY THE END OF THE DAY NO IMPROVEMENT EXCEPT THE IOP DECREASED U/S SHOWED VITRITIS INTRAVITREAL VANCO AND CEFTAZIDIME 1MG WAS GIVEN 2ND DAY THE HYPOPYON INCREASED AND NO IMPROVMENT AC WASH WAS DONE BUT THE MEMBRANES WERE BEHIND ICL SO I DECIDED TO EXPLANT THE ICL AND PEELING THE MEMBRANES IN FRONT OF THE CRYSTALLINE LENS BUT THE YELLOW REFLEX WAS STILL PRESENT A VITRORETINAL COLLEAGUE STARTED TO PUT PPV SET EXAMINED THE POSTERIOR SEGMENT. WE FOUND DENSE INTRAVITREAL MEMBRANES; VITREOUS SAMPLE WITH THE AQUEOUS SAMPLE WAS TAKEN AND SENT TO MICROBIOLOGISTS. PPV WAS DONE AND INTRAVITREAL VANCO AND CEFTAZIDIME WAS GIVEN AT THE END OF THE SURGERY. THE PATIENT IMPROVED 2ND POST OPERATIVE AFTER 1 WEEK SHE REGAINED HER PREOP BCVA." EVENT- ADDITOINAL INFORMATION RECEIVED ON 30 NOV 2018 FROM REPORTER: A 13.2MM VICMO13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6)2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SUSPECTED ENDOPHTALMITIS AND INFLAMMATION. MEDICAL INTERVENTION (INTRAOCULAR INJECTION) WAS REPORTED. PATIENTS CONDITION AT LAST VISIT DATE: (B)(6)2018, RECOVERED POST OPERATIVE BCVA. THE FOLLOWING WAS REPORTED AS ADDITIONAL INFORMATION IN COMPLAINT FORM: " IT IS INDICATED THAT ITS A CASE OF TASS WITH HYPOPYON IN AC AND PUPILLARY MEMBRANES, INTRAVITREAL MEMBRANES WAS ALSO FOUND AND CASE IMPROVED AFTER PPV WITH INTRAVITREAL VANCOMYCIN AND CEFTAZIDIME INJECTION. PATIENT CODE:(B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8724361&pc=MTA | 2023826-2019-01109 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Endophthalmitis; Inflammation; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATES THAT THE SURGEON IMPLANTED A 13.2MM, VICM5-13.2, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE SURGEON FOUND HIGH VAULT AND INFLAMMATION IN THE EYE, THE LENS WAS EXPLANTED, AND IRRIGATION OF THE ANTERIOR CHAMBER WAS PERFORMED. THE EXPLANTED LENS WAS SENT TO AN EXTERNAL TESTING LABORATORY FOR INSPECTION BECAUSE OF THE INFLAMMATION THAT THE EYE DEVELOPED, THE DOCTOR CONFIRMED THAT THERE WAS NO PROBLEM WITH THE LENS. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. PATIENTS CURRENT CONDITION IS GOOD. Manufacturer Narrative: LENS WAS RETURNED IN A STUART TRANSPORT SWAB WITH A GEL LIKE SUBSTANCE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5805752&pc=MTA | 2023826-2016-00903 | 2016-05-19 | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Endophthalmitis; Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED PIGMENT DISPERSION AND SUSPECTED ENDOPHTHALMITIS. THE REPORTER STATED THE IMPLANTED LENS WAS NO LONGER SUITABLE FOR THE PATIENT DUE TO POST-OP INFLAMMATION IN THE ANTERIOR CHAMBER. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND A SPOT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTALISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5806329&pc=MTA | 2023826-2016-00902 | 2016-05-19 | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Endophthalmitis; Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED PIGMENT DISPERSION AND SUSPECTED ENDOPHTHALMITIS. THE REPORTER STATED THE IMPLANTED LENS WAS NO LONGER SUITABLE FOR THE PATIENT DUE TO POST-OP INFLAMMATION IN THE ANTERIOR CHAMBER. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT, THE HAPTIC TORN/BROKEN/BENT/DEFORMED, AND A SPOT ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTALISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14364864&pc=MTA | 2023826-2022-01476 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Inflammation; Pain; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A CASE REPORT FROM (B)(6) JOURNAL OF OPHTHALMOLOGY WAS RECEIVED IN REGARD TO "CHRONIC POSTOPERATIVE CUTIBACTERIUM ACNES ENDOPHTHALMITIS WITH IMPLANTABLE COLLAMER LENS". THE CASE REPORT WAS ABOUT A (B)(6)-YEAR-OLD MALE WHO PRESENTED 3 MONTHS AFTER ICL IMPLANTATION OF THE RIGHT EYE (OD) WITH BLURRY VISION, REDNESS AND OCULAR PAIN IN THE SETTING OF PROLONGED POST-OPERATIVE ANTERIOR CHAMBER (AC) CELL. REDUCED VISUAL ACUITY (VA) AT 20/30-1, KERATIC PRECIPITATES, 1+ AC CELL, AND WHITE ICL PRECIPITATES WERE CONCERNING FOR CHRONIC POST-OPERATIVE ENDOPHTHALMITIS. ANAEROBIC CULTURES FROM A VITREOUS TAP GREW C. ACNES. MULTIPLE INTRA-VITREAL AND INTRACAMERAL INJECTIONS WITH TOPICAL STEROIDS WERE REQUIRED TO MAINTAIN A STABLE VA AT 20/30; HOWEVER, INFLAMMATION PERSISTED AND REMOVAL OF THE ICL AND HIS NATIVE LENS WAS ULTIMATELY REQUIRED. AFTER SEVERAL MONTHS, WITHOUT RECURRENCE OF INFLAMMATION, A SECONDARY SULCUS IOL WAS PLACED WITH RESULTING UNCORRECTED VISUAL ACUITY OF 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17980259&pc=MTA | 2023826-2023-04701 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Inflammation; Red Eye(s); Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS THE PATIENT HAS ANTERIOR CHAMBER INFLAMMATION WITH SYMPTOMS OF MILD BLURRED VISION AND EYE REDNESS 1 MONTH AGO THAT WAS RESOLVED WITH SHORT COURSE OF PREDNISOLONE FOR 1 WEEK. IT WAS LATER REPORTED THE PATIENT HAD SIGNIFICANT INFLAMMATION IN THE RIGHT EYE BUT HER VISION IS STILL VERY GOOD 20/20 AND IOP IS 12MMHG. SURGEON NOTES THEY STARTED PREDNISOLONE THEN QID. LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356138&pc=MTA | 2023826-2018-00439 | 2018-02-21 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; Intraocular Infection; Intraocular Pressure, Delayed, Uncontrolled; Iritis; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON NOTES ACUTE ENDOPHTHALMITIS WITH ACUTE IRITIS. REPORTEDLY, THE PATIENT HAS BEEN TREATED WITH AN INTRAOCULAR INJECTION AND TOPICAL DROPS INCLUDING ANTIBIOTICS. THE CLAIM REPORTER LISTS THE CAUSE OF THE ADVERSE EVENT AS HYPOTONY AND WOUND LEAKAGE. NO ADDITIONAL INFORMATION REGARDING PATIENT'S CURRENT STATUS IS KNOWN AT THIS TIME. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER STATED, "THE LEFT EYE HAS IMPROVED AND IS NOW 20/30-2 WITH CONTACT LENS. NO CELLS WERE SEEN IN EXAM WITH POSTERIOR SYNECHAIE WITH SLIGHT PIGMENT OVER ANTERIOR LENS SURFACE NOTED ON (B)(6) 218 EXAM." THE REPORTER STATED WHEN ASKED ABOUT ANY IDEA AS TO WHY WOUND LEAKAGE AND HYPOTONY HAD OCCURRED THAT, "THE ISSUE THE PATIENT HAD WAS DUE TO WOUND LEAKAGE AND THE PATIENT BEING A SMOKER AND EYE RUBBER." (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606098&pc=MTA | 2023826-2017-00903 | 2013-03-14 | Injury | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; Intraocular Infection; Pain; Blurred Vision; Loss of Vision; Visual Disturbances; Vitrectomy; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.0MM TICM120V4, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. PATIENT STATES SURGERY WAS PAINFUL AND THE SURGEON CAUSED "SURGICAL TRAUMA" WHICH LED TO INFECTION AND BLURRED VISION. ON (B)(6) 2013 PATIENT'S VISION WAS "6/6" BUT BY THE NEXT DAY PATIENT WAS EXPERIENCING PAIN, "HUGE FLASHES," AND YELLOWNESS IN VISION. ON (B)(6) 2013 PATIENT WAS DIAGNOSED WITH PVD (POSTERIOR VITREOUS DETACHMENT). PATIENT STATES HE WAS TOLD HE WAS HAVING "AC REACTION" AND WAS GIVEN "MEDICINAL" DROPS. ON (B)(6) 2013 THE PATIENT WAS DIAGNOSED WITH AN INFECTION IN THE EYE. PATIENT WAS TREATED WITH ANTIBIOTIC AND ANTI-FUNGAL MEDICATION. BOTH TIMES PATIENT TESTED NEGATIVE FOR INFECTION. ON (B)(6) 2013 THE LENS WAS EXPLANTED. LENS WAS NOT SENT TO STAAR (MANUFACTURER) FOR EVALUATION BUT TO AN ALTERNATE LAB, IN WHICH ONE OUT OF THREE PIECES TESTED POSITIVE FOR FUNGAL ELEMENTS. PATIENT HAD SEVERAL MORE SURGERIES, INCLUDING VITRECTOMIES, AND SEVERAL MORE MEDICINAL TREATMENTS. PATIENT STATES THAT AFTER THE 1ST IMPLANT (REFERENCE MDR 2023826-2017-00902) HE HAD DEVELOPED ENDOPHTHALMITIS AND WAS LEFT UNINFORMED AND UNTREATED BY HIS PHYSICIAN. CURRENTLY, PATIENT HAS NO VISION IN THE RIGHT (OD) EYE. Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, IT HAS BEEN DETERMINED THAT NOTHING IN THE MANUFACTURING OF THE LENS COULD HAVE CAUSED THE REPORTED EVENT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826223&pc=MTA | 2023826-2016-00987 | 2016-06-24 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Endophthalmitis; Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINS IMPLANTED. THE PATIENT EXPERIENCED AN ELEVATED IOP, SUSPECTED ENDOPHTHALMITIS, PIGMENT DISPERSION, AND BLURRED VISION. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 AND EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. (B)(4). CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT HAD LOW CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY AND AN ANTERIOR CHAMBER DEPTH LOWER THAN 3.00MM. CORRECTED DATA: THERE WAS NO ENDOPHTHALMITIS OR ELEVATED INTRAOCULAR PRESSURE RISE OBSERVED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9278780&pc=MTA | 2023826-2019-02089 | 2019-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Intraocular Pressure Increased; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019. ELEVATED IOP (29.5 MMHG) WAS ASSESSED ON (B)(6) 2019. TREATMENT OF STEROID, PREDONIN, DIAMOX, EYE DROPS, AND INTERNAL MEDICINE. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, 'AFTER 3 DAYS FROM SURGERY, PATIENT CLAIMED CLOUDY VISION. FLARE VALUE HIGH. REVOFLO, AND SANBETA PRESCRIBED. AFTER 4 DAYS FROM THE SURGERY, CONFIRMED TOXIC ANTERIOR SEGMENT SYNDROME (TASS) SYMPTOMS AND IOP WAS HIGH (29.5). ENFORCED STEROID AND STARTED INTERNAL USE OF PREDONIN, AND DIAMOX. AS OF (B)(6), THE INFLAMMATION HAS GOTTEN BETTER." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- "ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019." SHOULD BE REMOVED FROM THE INITIAL MDR. H6- PATIENT CODE 1835 IS NOT APPLICABLE IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- IT WAS REPORTED THAT AFTER 12-15 DAYS FROM SURGERY FIBRIN WAS RECOVERED, DEPOSIT WAS OBSERVED ON THE SURFACE OF ICL LENS, HYPEREMIA WAS RECOVERED AND AN IOP 14 MMHG WAS OBSERVED. AFTER A MONTH FROM SURGERY FIBRIN LIKE TURBIDITY APPEARED AGAIN AND BETAMETHASONE AND STEROID APPLIED ON PATIENT. AFTER 1.5 MONTH FROM SURGERY FIBRIN LIKE TURBIDITY GOT BETTER. THE PATIENT WAS CHANGED TO FLUOROMETHOLONE 0.1% X 4 TIMES PER DAY FROM BETAMETHASONE X4 TIMES PER DAY. AFTER 3 MONTH FROM SURGERY SLIGHT PIGMENT OBSERVED ON ICL BUT LESS INFLAMMATION. THE PATIENT WAS PRESCRIBED A DECREASED DOSE FREQUENCY TO FLUOROMETHOLONE 0.1% X 2 TIMES PER DAY. AFTER 4 MONTH FROM SURGERY THE PIGMENT OBSERVED ON ICL DECREASED AND THER WAS NO INFLAMMATION. THE PATIENT'S DOSE FREQUENCY WAS DECREASED TO FLUOROMETHOLONE 0.1% X 1 TIME PER DAY. THE PATIENT WENT THROUGH 4 MONTHS OF MEDICAL TREATMENT AND THEN FULLY RECOVERED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8315399&pc=MTA | 2023826-2019-00232 | 2018-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Iritis; Retinal Detachment; Uveitis; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -13.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. LOSS OF BCVA (POST-OP BCVA HM+), BLURRED VISION, UVEITIS/IRITIS AND TASS WAS REPORTED. AQUEOUS TEST REPORT WAS NEGATIVE FOR ENDOPHTHALMITIS. AC WASH, "INTENSIVE MEDICAL TREATMENT" (ANTIBIOTICS AND STEROIDS), AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WERE PERFORMED. IT WAS REPORTED THAT WITH INTENSIVE STEROIDS, THERE IS VERY GRADUAL DECREASE IN VITREOUS INFLAMMATORY EXUDATES. HOWEVER, THERE IS NOT MUCH OF VISION IMPROVEMENT AS OF NOW. B SCAN SHOWS RESOLVING VITREOUS EXUDATES. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2019; OUR INITIAL IMPRESSION WAS TASS. BUT BY THE NEXT COUPLE OF DAYS,WE REALIZED IT WAS ENDOPHTHALMITIS OF THE LEFT EYE. THE PATIENT WAS TREATED AGGRESSIVELY IN THE EYE. THE INFECTION CAME UNDER CONTROL BUT UNFORTUNATELY,BY THE TIME IT WAS CONTROLLED, SHE DEVELOPED ENDOPHTHALMITIS DUE TO CONTRACTING VITREOUS. RETINAL DETACHMENT SURGERY WAS DONE. HER VISION NOW IS HAND MOVEMENTS. ADDITIONAL INFORMATION RECEIVED ON (B)(6) 2019; PATIENT HAS BEEN REFERRED TO A POSTERIOR SEGMENT SURGEON AT SANKARA EYE HOSPITAL, BANGALORE. RETINAL DETACHMENT SURGERY WAS PERFORMED. ORIGINAL ICL HAD BEEN REMOVED. PRESENT PATIENT VISION IS HAND MOVEMENTS ONLY. FURTHER INFORMATION WAS REQUESTED BUT HAS NOT YET BEEN PROVIDED. PATIENT CODE ADDITION INFORMATION - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXPLANT, RETINAL DETACHMENT SURGERY. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS, AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365002&pc=MTA | 2023826-2022-01477 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Iritis; Uveitis; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A CASE REPORT FROM THE JOURNAL OF VITREORETINAL DISEASES WAS RECEIVED IN REGARD TO "CULTURE-NEGATIVE C ACNES ENDOPHTHALMITIS FOLLOWING IMPLANTATION OF A PHAKIC IMPLANTABLE COLLAMER LENS". THE CASE REPORT WAS ABOUT A 34-YEAR-OLD WOMAN WHO PRESENTED WITH CHRONIC UNILATERAL IRITIS, WITH PRECIPITATES 5 MONTHS AFTER BILATERAL IMPLANTABLE COLLAMER (ICL) IMPLANTATION IN THE LEFT EYE (OS). TREATMENT WITH TOPICAL PREDNISOLONE ACETATE WAS INITIALLY SUCCESSFUL BUT THE IRITIS RECURRED. THE PATIENT WAS REFERRED TO A UVEITIS SPECIALIST AT EIGHT MONTHS POSTOP, AND TESTS FOR UVEITIS HAD NEGATIVE FINDINGS. ULTRASOUND BIOMICROSCOPY SHOWED POSSIBLE CHAFFING OF THE IRIS BY THE ICL. AT 10 MONTHS AFTER THE INITIAL SURGERY, THE ICL WAS REMOVED. PATIENT CONTINUED TO DEVELOP WORSENING FLOATERS AND VA. ONE YEAR AFTER ICL PLACEMENT, PATIENT HAD 2+ VITRITIS WITH SMALL DEPOSITS ON POSTERIOR SURFACE OF NATURAL LENS. A DIAGNOSTIC VITRECTOMY WAS PERFORMED THAT WAS CULTURE NEGATIVE. TRIAL OF ORAL PREDNISONE RESULTED IN ONLY TRANSIENT IMPROVEMENT OF UVEITIS. A SECOND VITRECTOMY COMBINED WITH LENSECTOMY WAS PERFORMED. CULTURES HAD NEGATIVE FINDINGS, BUT BACTERIAL PCR TESTING WAS POSITIVE FOR C ACNES. TREATMENT WITH INTRAVITREAL VANCOMYCIN AND CEFTAZIDIME WAS ADMINISTERED. AT 3 MONTHS FOLLOW-UP, PATIENT HAD COMPLETE RESOLUTION OF OCULAR INFLAMMATION AND TREATMENTS DISCONTINUED. AT ONE-YEAR FOLLOW-UP PATIENT'S VISION WAS 20/20 WITH CORRECTION FOR APHAKIA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. HEALTH IMPACT- CLINICAL CODE: (B)(4) PRECIPITATES; IRIS CHAFFING. HEALTH EFFECT IMPACT CODE: (B)(4) TOPICAL PREDNISOLONE ACETATE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552493&pc=MTA | 2023826-2019-00664 | 2018-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON THE SAME DAY, CORNEAL EDEMA AND INFLAMMATORY CELLS DEVELOPED IN THE ANTERIOR CHAMBER. DIAGNOSTIC TESTS WERE PERFORMED AND THE SURGEON REPORTS ENDOPHTHALMITIS. NO AQUEOUS SAMPLES WERE TAKEN. STEROID EYE DROPS, ORAL MEDICINE, AND EYE DROPS WITH HYPOTENSIVE EFFECT WERE ADMINISTERED. CURRENTLY, THE PATIENT IS RECOVERED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437905&pc=MTA | 2023826-2019-00446 | 2019-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -2.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO CULTURE-NEGATIVE ENDOPHTHALMITIS. INTRA-CAMERAL ANTIBIOTICS WERE ADMINISTERED TO THE PATIENT. THE PATIENT IS "OK RIGHT NOW" AND DID NOT WANT ADDITIONAL REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON INDICATES THIS IS A CASE OF TASS. REPORTEDLY, THE SURGEON SUSPECTS THE VISCOELASTIC USED (REPORTED AS 2% HPMC) MAY HAVE BEEN A CONTRIBUTING FACTOR. IN ADDITION, THE SURGEON FOUND A FISSURE IN THE AUTOCLAVE AND BELIEVES THERE WAS A PROBLEM WITH THE STERILIZATION PROCESS OF THE SURGICAL TOOLS USED. THE AUTOCLAVE INSTRUMENTS USED COULD HAVE BEEN THE CAUSE OF THE EVENT. THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. PATIENT CODE 1835 PREVIOUSLY REPORTED DOES NOT APPLY. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (TASS) METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437911&pc=MTA | 2023826-2019-00455 | 2019-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM, VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO CULTURE-NEGATIVE ENDOPHTHALMITIS. INTRAOCULAR INJECTION, INTRA-CAMERAL ANTIBIOTICS WERE ADMINISTERED TO THE PATIENT. THE PATIENT IS "OK RIGHT NOW." THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE SURGEON INDICATES THIS IS A CASE OF TASS. REPORTEDLY, THE SURGEON SUSPECTS THE VISCOELASTIC USED (REPORTED AS 2% HPMC) MAY HAVE BEEN A CONTRIBUTING FACTOR. IN ADDITION, THE SURGEON FOUND A FISSURE IN THE AUTOCLAVE AND BELIEVES THERE WAS A PROBLEM WITH THE STERILIZATION PROCESS OF THE SURGICAL TOOLS USED. THE AUTOCLAVE INSTRUMENTS USED COULD HAVE BEEN THE CAUSE OF THE EVENT. THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WERE FIBERS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 1835 PREVIOUSLY REPORTED DOES NOT APPLY. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (TASS) METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518998&pc=MTA | 2023826-2019-00617 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -16.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ONE DAY POST-OP THE PATIENT HAD EXAGGERATED INFLAMMATION AND EXUDATES IN THE EYE. THE PATIENT WAS DIAGNOSED WITH ACUTE ENDOPHTHALMITIS AND IMMEDIATELY WAS INJECTED WITH INTRAOCULAR ANTIBIOTICS. VITREOUS BIOPSY SAMPLE REVEALED PSEUDOMONAS INFECTION. AUTOCLAVE INDICATORS PASSED. OT AND AIR CULTURES REPORTED NEGATIVE. INSTRUMENTS AND CONSUMABLES REPORTED NEGATIVE FOR ANY MICROBIOLOGICAL GROWTH. VITREO-RETINAL INTERVENTION WAS "NOT POSSIBLE DUE TO CORNEAL INVOLVEMENT." VISION IS FC 2MT. INFECTION IS CURRENTLY INACTIVE. IT IS UNCLEAR AS TO WHETHER THE LENS REMAINS IMPLANTED, REQUESTS FOR ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595430&pc=MTA | 2023826-2019-00772 | 2019-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -9.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE DAY AFTER SURGERY, THE SURGEON REPORTS THAT THERE WAS FIBRIN DEPOSITION; ENDOPHTHALMITIS WAS DIAGNOSED VIA EXAMINATION. THE PATIENT WAS TREATED WITH A STEROID INJECTION AND EYE DROPS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE DOCTOR HAD THE PATIENT WEAR A THERAPEUTIC CONTACT LENS AFTER SURGERY. THE CONTACT LENS CAME OFF IN THE EVENING OF THE DAY OF SURGERY AND THE PATIENT WAHED THE CONTACT LENS WITH A CLEANING SOLUTION AND REPLACED THE LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9280687&pc=MTA | 2023826-2019-02090 | 2019-10-05 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VICMO13.2, -7.50 DIOPTER, INTRAOCULAR LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. SUSPECTED ENDOPHTHALMITIS WAS OBSERVED ON (B)(6) 2019 DURING FOLLOW UP VISIT. ON (B)(6) 2019 THE LENS WAS EXPLANTED. ADDITIONAL MEDICAL INTERVENTION ( COMPOUND TOBRAMYCIN, LEVOFLOXACIN EYE DROPS, PREDNISONE EYE DROPS AND TOBRAMYCIN EYE DROP) WERE PRESCRIBED PRIOR TO ICL EXPLANTATION. AQUEOUS HUMOR WAS SAMPLED FOR BOTH BACTERIAL CULTURE AND SMEAR, RESULTS WERE REPORTED AS BACTERIAL CULTURE- NEGATIVE AND SMEAR- POSITIVE, GRAM-POSITIVE COCCAI INFECTION. FOR STATUS OF THE EYE THE REPORTER INDICATED THAT POST EXPLANT THE PATIENT WAS TREATED WITH: VANCOMYCIN, CEFTAZIDIM, LEVOFLOXACIN EYE DROPS, FUXILIC ACID EYE DROPS, PREDNISONE EYE DROPS AND GATIFLOXACIN GEL. THE CDVA WAS 20/20 AND IOP WAS 15.5 MMHG. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTE TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4618716&pc=MTA | 2023826-2015-00362 | 2015-01-25 | Injury | STAAR SURGICAL COMPANY | 2015-03-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Endophthalmitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SUSPECTED ENDOPHTHALMITIS AND THE PROBLEM WAS RESOLVED. PATIENT WAS ADMINISTERED INTRAOCULAR INJECTION (MEDICATION). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (INCISION OPENED); (SECONDARY SURGERY); ENDOPHTHALMITIS; (LENS EXPLANTED); , EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DHR, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION, AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INFORMATION PROVIDED BY THE FACILITY, THE LENS WAS NOT INDICATED TO HAVE BEEN RECEIVED DAMAGED. THE RETURNED PRODUCT MET SPECIFICATION AND NO DAMAGE WAS REPORTED. IT IS UNCLEAR IF THE LENS ITSELF CAUSED OR CONTRIBUTED TO THE ONSET OF "SUSPECTED ENDOPHTHALMITIS". THEREFORE, NO ROOT CAUSE CAN BE IDENTIFIED. PROBABLE CAUSES FOR THIS EVENT COULD INCLUDE PRE-EXISTING PATIENT CONDITION OR SURGERY FACILITY OR CONDITION. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6370439&pc=MTA | 2023826-2017-00296 | 2016-04-29 | Injury | STAAR SURGICAL COMPANY | 2017-03-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6 IMPLANTABLE COLLAMER LENS, 8/2.5/090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO SUSPECTED ENDOPHTALMITIS AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (CLEAR AND REDDISH/BROWN DRY RESIDUE) AND A PIECE OF HAPTIC AND OPTIC ARE MISSING. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CONCLUSION CODE: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7425870&pc=MTA | 2023826-2018-00546 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. AT ONE DAY POST-OP, THE PATIENT PRESENTED WITH SUSPECTED ENDOPHTHALMITIS. THE LENS WAS EXPLANTED ON (B)(6) 2018 BY A DIFFERENT SURGEON AT A GENERAL HOSPITAL. A CELL CULTURE WAS PERFORMED AND STREPTOCOCCUS PNEUMONIA WAS FOUND. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7426046&pc=MTA | 2023826-2018-00547 | 2018-02-27 | Injury | STAAR SURGICAL COMPANY | 2018-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICM5_12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. AT ONE DAY POST-OP, THE PATIENT PRESENTED WITH SUSPECTED ENDOPHTHALMITIS. THE LENS WAS EXPLANTED ON (B)(6) 2018 BY A DIFFERENT SURGEON AT A GENERAL HOSPITAL. A CELL CULTURE WAS PERFORMED AND STREPTOCOCCUS PNEUMONIA WAS FOUND. ON (B)(6) 2018 IT WAS REPORTED THAT MILD INFLAMMATION CONTINUED IN THE EYE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6030472&pc=MTA | 2023826-2016-01442 | 2015-07-08 | Injury | STAAR SURGICAL COMPANY | 2016-10-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Endophthalmitis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12,00/+1,5/155 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE SUSPECTED ENDOPHTALMITIS AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: PREVIOUSLY, IT WAS REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO SUSPECTED ENDOPTHALMITIS AND PIGMENT DISPERSION. THIS IS INCORRECT. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THIS RESOLVED THE PROBLEM. PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY AS OF (B)(6) 2016 WAS 20/20. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505318&pc=MTA | 2023826-2016-01573 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-19 04:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Endophthalmitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -9.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, SUSPECTED ENDOPHTALMITIS AND PIGMENT DISPERSION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PMA/510(K): NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7144221&pc=MTA | 2023826-2017-01946 | 2017-11-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Endophthalmitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+5.5/145 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER INDICATED AT 8 DAYS POST-OP, THE PATIENT HAD SUSPECTED ENDOPHTHALMITIS. THE PATIENT LIVED FAR AWAY FROM THE SURGERY CENTER, SO ALL CONTROLS AND TREATMENTS (INTRAOCULAR INJECTION (MEDICATION)) WERE ADMINISTERED BY THEIR LOCAL OPHTHALMOLOGIST. THE REPORTER INDICATED THE PATIENT IS GETTING BETTER AND THE LENS REMAINS IMPLANTED. THE PATIENT WILL BE MONITORED. Manufacturer Narrative: PATIENT EYE WAS IMPECCABLE WITH NO SIGNS OF INFLAMMATION AND VISUAL ACUITY WAS BACK TO NORMAL. CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH KERATOCONUS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275467&pc=MTA | 2023826-2018-00236 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Endophthalmitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V4C IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENTS EYE 1 WEEK, DUE TO ACUTE ENDOPHTHALMITIS. THE REPORTER INDICATED THE PATIENT WAS TREATED WITH INTRAVITREAL ANTIBIOTICS. ANOTHER ICL LENS WAS IMPLANTED 9 MONTHS AFTER THE SUCCESSFUL RESOLUTION OF ENDOPHTHALMITIS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) - SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476571&pc=MTA | 2023826-2020-01936 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Endophthalmitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO A CASE OF ENDOPHTHALMITIS WAS FOUND. A DIAGNOSTIC WAS PERFORMED AND (B)(6) WAS DETECTED IN THE AQUEOUS HUMOR AFTER THE INCIDENT. THE PATIENT WAS GIVEN INTRAVENOUS INFUSION OF CEFTAZIDIME AND VANCOMYCIN ANTIBIOTICS, AND TOPICAL USE OF MOXIFLOXACIN EYE DROPS, LEVOFLOXACIN EYE DROPS, TOBRAMYCIN EYE DROPS FOR ANTIBACTERIAL TREATMENT, ANTI-INFLAMMATORY TREATMENT OF PREDNISOLONE EYE DROPS, ATROPINE AND COMPOUND TROPIKAMIDE EYE DROPS, DILATED PUPILS, INJECTION OF CEFTAZIDIME NEXT TO THE EYEBALLS, AND B-ULTRASOUND EXAMINATION OF THE EYE SHOWED VITREOUS TURBIDITY. ON (B)(6) 2020, THE INFLAMMATION HAD SUBSIDED. LATER THAT SAME DAY, THE INFLAMMATION OF THE AFFECTED EYE RECURRED, ROOM FLASH++++, CELL ++++, INCREASING THE FREQUENCY OF MEDICATION, CLOSELY OBSERVATION. ON (B)(6) 2020, A B-ULTRASOUND EXAMINATION OF THE EYE SHOWED THAT THE VITREOUS TURBIDITY WAS SIGNIFICANTLY WORSE THAN BEFORE, SO THE "LEFT EYE VITREOUS PUNCTURE AND INJECTION + ICL REMOVAL" WAS URGENTLY PERFORMED. THE AQUEOUS HUMOR AND ICL WERE TAKEN OUT DURING THE OPERATION. ANTIBACTERIAL TREATMENT OF SIFLOXACIN EYE DROPS, TOBRAMYCIN EYE DROPS, ANTI-INFLAMMATORY TREATMENT OF PREDNISOLONE EYE DROPS, COMPOUND TROPICAMIDE EYE DROPS TO DILATE PUPILS, GATIFLOXACIN EYE GEL AND TOBRAMYCIN PLAIN EYE OINTMENT 1 TIME/NIGHT TREATMENT. FROM (B)(6) 2020 THROUGH TO (B)(6) 2020, THE CONDITION OF THE PATIENT'S EYE IS STABLE, THE PATIENT'S ATRIAL FLASH +, CELL 0, CORRECTED VISION RESTORED TO 1.0 -, ICL BACTERIAL CULTURE 0, AQUEOUS HUMOR TEST: HUMAN ALPHA (B)(6) TYPE 1, ANTIVIRAL TREATMENT IS GIVEN TO THE PATIENT, AND THE REST IS THE SAME AS BEFORE. CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR, PATIENT SUFFERED FROM ORAL ULCERS, CAUSING (B)(6). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725776&pc=MTA | 2023826-2023-01250 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | Endophthalmitis; Red Eye(s); Uveitis; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO ENDOPHTHALMITIS WAS OBSERVED. PATIENT HAD RED EYE AND BLURRED VISION. DIAGNOSTICS WERE PERFORMED FOR UVEITIS AND INFECTION. TREATMENT WAS PERFORMED - A/C IRRIGATION, PREDNILONE EYE DROPS (6.5ML) AND VIGAMOX 0.5% (5MIL). THE PATIENT WAS TREATED AND THE EYE HAS RECOVERED. THE LENS WAS SENT TO A LAB FOR BACTERIAL CULTURE TEST AND AND NO PRESENT OF BACTERIA WAS FOUND ON THE LENS. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - PREDNILINE EYE DROPS (6.5ML); VIGAMOX 0.5% (5ML). H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - A/C IRRIGATION. (B)(4). H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360972&pc=MTA | 2023826-2023-05562 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Retinal Detachment; Toxic Anterior Segment Syndrome (TASS) | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO ENDOPHTHALMITIS WAS OBSERVED. DIAGNOSTICS OF BLOOD SAMPLING WAS PERFORMED. ANTERIOR CHAMBER CLEANING WAS PERFORMED. MEDICATION WAS ADMINISTERED (MELOTHENE CONJUNCTIVALIZATION INJECTION/ANTERIOR CHAMBER LAVAGE). VITREOUS SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. POST-OP, VITREOUS RETINAL DETACHMENT TREATMENT, OIL INJECTION SURGERY, LASER TREATMENT GAS AND ANTERIOR CHAMBER LAVAGE. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENTS CURRENT CONDITION WAS VISUAL ACUITY AT HAND MOTION LEVEL, TREATMENT IS ON-GOING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ANTERIOR CHAMBER CLEANING/LAVAGE; VITREOUS SURGERY, IOL IMPLANTED. H6: HEALTH EFFECT IMPACT CODE: 4644 - MELOTHENE CONJUNCTIVALIZATION INJECTION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207622&pc=MTA | 2023826-2021-02772 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Toxic Anterior Segment Syndrome (TASS) | NA | 1 | Event Description: THE REPORTER INDICATED THERE HAS BEEN REPORTS OF ENDOPHTHALMITIS AND TASS IN PHAKIC IOL PATIENTS AT ONE OF THEIR CLINICS OVER THE LAST COUPLE OF WEEKS. THE REPORTER INDICATED A CULTURE WAS TAKEN AND THE RESULT WAS STREP MITIS. A POST-OP ANTIBIOTIC WAS USED - VIGAMOX QID. THE REPORTER INDICATED THEY WERE ABLE TO IDENTIFY AN ISSUE INTERNALLY WITH THE CLEANING OF THE MST MICRO INSTRUMENT AND THE RE-USABLE INJECTOR THAT MAY HAVE CONTRIBUTED TO THE ISSUE. THE INJECTORS WERE ALSO SWABBED AND CULTURE TAKEN AND THEY WERE UNABLE TO CULTURE IT ON THE ITEMS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007849&pc=MTA | 2023826-2023-04813 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Endophthalmitis; Uveitis | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OS). THE SURGEON REPORTS THE PATIENT HAS UVEITIS 2 MONTHS AFTER IMPLANTATION, AND MAY HAVE CHRONIC ENDOPHTHALMITIS DUE TO THE PRESENCE OF LARGE KP'S AND A SMALL PLAQUE IN THE POSTERIOR ASPECT OF THE EVO ICL. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6562081&pc=MTA | 2023826-2017-00789 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Endophthalmitis; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -8.00 DIOPTER IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT IS EXPERIENCING PHOTOPHOBIA AND "SEEING THROUGH A DIRTY LENS." UPON EXAMINATION, THE SURGEON NOTES WHITISH DEPOSITS ON THE ANTERIOR AND POSTERIOR FACE OF THE LENS. THERE IS SUSPECTED ENDOPHTALMITIS. CULTURES ARE PENDING LAB ANALYSIS RESULTS. AS OF THE DATE OF MDR SUBMISSION, LENS REMAINS IMPLANTED. Manufacturer Narrative: SAMPLES WERE TAKEN FROM THE EYE. AN ANTIBIOTIC INJECTION WAS ADMINISTERED. CULTURES AND PCR TEST RESULTS WERE NEGATIVE. THE INFILTRATES HAVE DISAPPEARED, THOUGH IT IS POSSIBLE THAT SOME RESIDUAL REMAINS, ACCORDING TO CLAIM REPORTER. THE PATIENT IS SATISFIED AND THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7676112&pc=MTA | 2023826-2018-01030 | 2016-03-22 | Injury | STAAR SURGICAL COMPANY | 2018-07-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Endophthalmitis; Visual Disturbances; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, SUSPECTED ENDOPHTHALMITIS, AND PIGMENT DISPERSION. THE PROBLEM WAS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PREVIOUSLY STATED THAT THE LENS WAS EXCHANGED DUE TO ENDOPHTHALMITIS, EXCESSIVE VAULT, AND PIGMENT DISPERSION. CORRECTED-THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT AND PIGMENT DISPERSION ONLY. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: OFF-LABEL USE [PRE-EXISTING CONDITION-NARROW ANTERIOR CHAMBER ANGLES (I.E. GRADE II OR LESS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7081602&pc=MTA | 2023826-2017-01730 |  | Injury | STAAR SURGICAL COMPANY | 2017-12-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Endophthalmitis; Visual Disturbances; Vitrectomy; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, THE PATIENT DEVELOPED ENDOPHTHALMITIS. THE PATIENT HAS BEEN TREATED WITH MULTIPLE INTRAOCULAR INJECTIONS AND A VITRECTOMY BUT "IS NOT DOING WELL VISION-WISE." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12104313&pc=MTA | 2023826-2021-02224 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Injury; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT , TRANSIENT LOSS OF BCVA, AND BLEEDING ON THE BACK OF THE ICL WAS OBSERVED. MEDICAL INTERVENTION WAS PRESCRIBED (LINDERON 0.1%, EYE DROPS 7 TO 8 TIMES, ADNA TABLETS 30 MG TAKEN 3 TIMES). IT WAS REPORTED THAT THE PROBLEM RESOLVED. FOR ADDITIONAL INFORMATION THE REPORTER STATED " ACCORDING TO PATIENT COMMENTS, AFTER WAKING UP ON APR 19, HER VISION APPEARED TO BE VEILED, BUT THE SYMPTOMS DISAPPEARED TWO HOURS LATER. AFTER THAT, HER EYESIGHT BECAME HARD TO SEE FROM THE EVENING". IT WAS REPORTED THAT THERE ARE NO PLANS TO REPLACE THE LENSES BECAUSE THERE IS NO PROBLEM WITH THE EYESIGHT. CAUSE OF EVENT WAS UNKNOWN " DOCTORS SUSPECT THAT THE PATIENT SUFFERED A BRUISE WHILE SLEEPING." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048066&pc=MTA | 2023826-2023-04715 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ORBITAL PAIN DUE TO PATIENT PSYCHOLOGICAL EFFECTS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR: PSYCHOLOGICAL EFFECTS. REPORTEDLY, PATIENT HAS REVERTED TO GLASSES. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048064&pc=MTA | 2023826-2023-04714 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ORBITAL PAIN DUE TO PATIENT PSYCHOLOGICAL EFFECTS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR: PSYCHOLOGICAL EFFECTS. REPORTEDLY, PATIENT HAS REVERTED TO GLASSES. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10529812&pc=MTA | 2023826-2020-02074 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ANGLE CLOSURE. A NEW PERIPHERAL IRIDOTOMY WAS ADDED (LASER) BUT NOTHING CHANGED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS REMAINS IMPLANTED. THE PATIENT IS SEEING WELL BUT EYE HURTS. Manufacturer Narrative: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5190315&pc=MTA | 2023826-2015-01404 | 2015-09-16 | Malfunction | STAAR SURGICAL COMPANY | 2015-10-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Crack; Sticking | Failure of Implant | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS ADHERED TO THE FORCEPS AND CRACKED. THERE WAS NO PATIENT CONTACT. THE SURGERY WAS POSTPONED UNTIL THE NEXT DAY AND ANOTHER LENS WAS SUCCESSFULLY IMPLANTED. THE SURGEON FELT THERE MAY HAVE BEEN A PROBLEM WITH THE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16706816&pc=MTA | MW5116465 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-07 04:00 | MTA | EVO ICL (IMPLANTABLE COLLAMER LENS), LEFT EYE | Product Quality Problem | Failure of Implant; Loss of Vision |  | 1 | Event Description: HAD STAAR MEDICAL'S IMPLANTABLE COLLAMER LENS, EVO ICL IMPLANTED INTO BOTH RIGHT AND LEFT EYE. NOTICED NEXT DAY A LOSS OF PERIPHERAL VISION. RETURNED TO DOCTOR, WAS TOLD TO WAIT A WEEK. RETURNED TO HAVE RETINA CHECKED, RETINA WAS OKAY. TODAY, (B)(6), ISSUE STILL PRESENT. ATTEMPTED TO CONTACT COMPANY (B)(6), WAS REFUSED INFORMATION. THIS SIDE EFFECT WAS NOT LISTED AS POSSIBLE SIDE EFFECT. I AM NOW TOLD THE LENS, EVEN WHEN ADVERTISED AS REMOVABLE, IS INDEED NOT REMOVABLE, SO SIDE EFFECTS ARE NOT REVERSIBLE. I WOULD HAVE NEVER HAD THIS PROCEDURE DONE IF I KNEW THIS WAS A POSSIBLE SIDE EFFECT. REFERENCE REPORT: MW5116464.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983342&pc=MTA | 2023826-2022-02208 | 2022-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Fatigue; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS REFRACTIVE SURPRISE, BLURRED VISION AND VISUAL FATIGUE. IT NOTES "PATIENT PRESBYOPIA +1.25". ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11506684&pc=MTA | MW5100045 | 2017-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-03-16 | MTA | VISIAN ICLS | Patient-Device Incompatibility | Fatigue; Headache; Discomfort |  | 1 | Event Description: I GOT VISIAN ICLS IN 2017 AND EVER SINCE THEN, I HAVE HAD HEADACHES, PULLING OF EYE AND FACE, BAD INTERACTION WITH HEAD, TIREDNESS. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14046187&pc=MTA | 2023826-2022-01093 | 2022-03-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Fatigue; Headache; Inflammation; Intraocular Pressure Increased; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND AC INFLAMMATION CELLS. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- CORRECTED TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLNATABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), AC INFLAMMATION CELLS, PAIN, HEADACHES, NAUSEA, NUMBNESS OF LEG, GENERALIZED BODY WEAKNESS, INFLAMMATION, PIGMENT CELLS, MEDICATION WAS PRESCRIBED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; NUMBNESS OF LEG. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 MED WATCH REPORT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10969908&pc=MTA | 2023826-2020-02886 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Flashers | NA#S-P#NA | 1 | Event Description: RECEIVED A COMPLAINT FROM (B)(6), THE PATIENT REPORTED EYE FLASHES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279013&pc=MTA | 2023826-2023-02619 | 2022-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Foreign Body Sensation in Eye | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO PATIENT DISCOMFORT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT HAD CONTINUOUSLY COMPLAINED OF FOREIGN BODY SENSATION OF EYES. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279021&pc=MTA | 2023826-2023-02618 | 2022-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Foreign Body Sensation in Eye | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO PATIENT DISCOMFORT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT HAD CONTINUOUSLY COMPLAINED OF FOREIGN BODY SENSATION OF EYES. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO : " THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO PATIENT DISCOMFORT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT HAD CONTINUOUSLY COMPLAINED OF FOREIGN BODY SENSATION OF EYES." D6B - CORRECTED TO: 10-FEB-2023 IN THE INITIAL MDR. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227502&pc=MTA | 2023826-2019-02011 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT FELT BAD BECAUSE HER VISUAL ACUITY HAD IMPROVED TOO MUCH, AND HAD A FEELING OF A FOREIGN BODY. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227507&pc=MTA | 2023826-2019-02012 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PATIENT FELT BAD BECAUSE HER VISUAL ACUITY HAD IMPROVED TOO MUCH, AND HAD A FEELING OF A FOREIGN BODY. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9939900&pc=MTA | 2023826-2020-00816 | 2020-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT FELT A FOREIGN BODY SENSATION IN HIS EYE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941222&pc=MTA | 2023826-2020-00817 | 2020-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT FELT A FOREIGN BODY SENSATION IN HIS EYE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8259220&pc=MTA | 2023826-2019-00123 | 2018-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Foreign Body Sensation in Eye; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ABOUT AN HOUR AFTER THE LENS WAS REMOVED DUE TO THE PATIENT COMPLAINING OF FOREIGN BODY SENSATION AND NOT COOPERATING WITH THE DOCTOR FOR EYE INSPECTION. IT WAS REPORTED THAT IN SURGERY THE DOCTOR DISCOVERED THAT THE PATIENT HAS IFIS WITH IRIS PROLAPSE. ADDITIONAL SUTURE WAS REPORTED. IN THE SURGEONS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. AT PATIENTS LAST VISIT BCVA WAS 20/2, IOP 15, THE EYE IS WELL. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6487472&pc=MTA | 2023826-2017-00620 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Foreign Body Sensation in Eye; Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT REPORTED HAD SEVERE PAIN BEHIND THE EYE AND FOREIGN BODY SENSATION. THE LENS REMAINS IMPLANTED BUT THE PLAN IS TO EXPLANT AND EXCHANGE FOR A SHORTER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590903&pc=MTA | 2023826-2020-02255 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPLANT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10590933&pc=MTA | 2023826-2020-02256 | 2020-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Foreign Body Sensation in Eye; Visual Disturbances; Eye Pain | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VICMO12.6, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED, IT WAS REPORTED THAT AFTER LENS IMPLANT THE PATIENT EXPERIENCE FOREIGN BODY SENSATION, PAIN AND DOUBLE IMAGES AND REQUIRED REMOVAL. LENS EXPALNT RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE ON THE LENS SURFACE AND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: D10 - INJECTOR MODEL CORRECTED TO MSI-PF IN THE INITIAL MDR. H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734856&pc=MTA | 2023826-2021-01173 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS MALIGNANT GLAUCOMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT: UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR UNK ENTRY, IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11734862&pc=MTA | 2023826-2021-01174 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma | N/A | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE SURGEON REPORTS MALIGNANT GLAUCOMA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986326&pc=MTA | 2023826-2021-01847 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Glaucoma | P030016 | 1 | Event Description: REPORTEDLY AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED DUE TO A GLAUCOMA COMPLICATION. A TUBE SHUNT WAS THEN PLACED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: DATE OF EVENT: UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16230473&pc=MTA | 2023826-2023-00271 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -13 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND GLAUCOMA. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND IT'S UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED; ADDITIONALLY IT NOTED "PATIENT GOT GLAUCOMA BECAUSE OF HIGH VAULT". Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16230029&pc=MTA | 2023826-2023-00272 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -12.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND GLAUCOMA. ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND IT'S UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED; ADDITIONALLY IT NOTED "PATIENT GOT GLAUCOMA BECAUSE OF HIGH VAULT". Manufacturer Narrative: B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. B4 - CORRECTED TO: 15-JAN-2023 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048692&pc=MTA | 2023826-2022-04678 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; No Apparent Adverse Event | Glaucoma; Blurred Vision; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT REPORTED EYE PROBLEMS AND POSSIBLE LOSS OF VISION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY, "AS SOON AS I WOKE UP FROM ANESTHESIA I HAD UNBEARABLE PAIN IN THE LEFT EYE WITH VOMITING AND NAUSEA. IT TURNED OUT THAT I HAD AN ATTACK OF OCULAR HYPERTONIA. THE OPHTHALMOLOGIST PUT ME ON A MANNITOL INFUSION AND GAVE ME DIAMOX TABLETS WITH DROPS TO LOWER THE OCULAR TENSION WITHOUT RESULTS." ON (B)(6)2022, AN OPTICAL COHERENCE TOMOGRAPHY FOUND, "IRRIDOCORNEAN ANGLE IS CLOSED (2 DEGREES). THE LENS WAS EXPLANTED AND THE REPORTER STATED, "I HAD CORNEAL EDEMA, A LOSS OF ENDOTHELIAL CELLS (I ONLY HAVE 1000 AT THE MOMENT) AN OPACITY OF THE LENS (GLAUCOMAFLECKEN BECAUSE OF THE OCULAR HYPERTENSION CRISIS) PUPILLARY MYDRIASIS AND GLAUCOMA WITH SIGNIFICANT LOSS OF OPTIC NERVE FIBERS IN THE EYE." THE RIGHT LENS REMAINS IMPLANTED AND IS CAUSING CHRONIC PAIN AND DOUBLE VISION, AND WITH THE LEFT EYE DAMAGE THE REPORTER STATED, "SUFFERING ANISOMETROPIA DUE TO THE DIFFERENCE OF VISION." REPORTEDLY, "WE HAVE TWO HYPOTHESES TO KNOW WHAT HAPPENED TO MY LEFT EYE EITHER THE IMPLANT WAS BIGGER THAN MY EYE IN THIS CASE THE OPHTHALMOLOGIST HAD TO CHANGE THE IMPLANT IMMEDIATELY OR IT'S DUE TO OBSTRUCTION OF THE CENTRAL AQUAPORT WITH VISCOELASTIC AND INFLAMMATORY DEBRIS." H6: WORK ORDER SEARCH FOUND ONE SIMILAR COMPLAINT - 727444 FOR FELLOW RIGHT EYE (OD). CLAIM#: (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13417556&pc=MTA | 2023826-2022-00338 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Glaucoma; Eye Pain | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT, LENS ROTATION, PAIN, AND GLAUCOMA. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671676&pc=MTA | 2023826-2019-01013 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Glaucoma; Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: AN ARTICLE PUBLISHED IN CLINICAL PRACTICE AND CASES IN EMERGENCY MEDICINE WAS RECEIVED ON (B)(6) 2019, ORIGINALLY PUBLISHED IN (B)(6) 2019, ENTITLED 'ACUTE ANGLE CLOSURE GLAUCOMA SECONDARY TO A PHAKIC INTRAOCULAR LENS, AN OPHTHALMIC EMERGENCY.' THE ARTICLE DESCRIBES A CASE IN WHICH A (B)(6) WOMAN WHO'D HAD BILATERAL ICL IMPLANTATION PERFORMED 5 YEARS PREVIOUSLY PRESENTED WITH HEADACHE, BLURRY VISION, ANISOCORIA, AND LIGHT SENSITIVITY IN HER RIGHT (OD) EYE. SHE WAS ALSO DIAGNOSED WITH ELEVATED IOP, A ROUND, FIXED PUPIL, +1 INJECTION, DIFFUSE MICROCYSTS, A SHALLOW ANTERIOR CHAMBER, FIXED, MINOR IRIS BOMBE, AND PUPILLARY BLOCK. SHE WAS ADMINISTERED TIMOLOL, ACETAZOLAMIDE, AND BRIMONIDINE AND UNDERWENT A YAG LASER PERIPHERAL IRIDOTOMY. POST PROCEDURE THE PATIENT WAS GIVEN TOPICAL BRIMONIDINE AND DIFLUPREDNATE. SUBSEQUENTLY SHE REQUIRED AN ADDITIONAL PERIPHERAL IRIDOTOMY AND REPORTEDLY, CONTINUES IN TREATMENT. SHE WAS DIAGNOSED WITH DELAYED PUPILLARY BLOCK AND ACUTE ANGLE CLOSURE GLAUCOMA. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743980&pc=MTA | 2023826-2015-00510 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER RIGHT EYE (OD). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY INDICATED THEY AGREED WITH THE PATIENT REPORT. THE PATIENT HAD HIGH PRESSURE (11 - 36MM HG) AND WAS DIAGNOSED WITH GLAUCOMA (OCULAR HYPERTENSION - GLAUCOMA SUSPECT) FOR THE PERIOD OF (B)(6) 2012. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE HIGH PRESSURE AND GLAUCOMA. THE FACILITY INDICATED THE CAUSE OF THE HIGH PRESSURE AND GLAUCOMA WAS FAMILIAL VERSUS OUTFLOW PROBLEM, THE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT'S POST-OP VA WAS 20/20. THE LENS REMAINS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PER MEDICAL REVIEW - MANY PATHOPHYSIOLOGICAL CHANGES COULD CAUSE OCULAR HYPERTENSION SUCH AS ANTERIOR CHAMBER DISTORTION, VISCOELASTICS, INFLAMMATION, HEMORRHAGE, PIGMENT DISPERSION, PUPILLARY BLOCK AND VITREOUS IN THE ANTERIOR CHAMBER. EXCESSIVE VAULTING AS A FACTOR COULD CAUSE PUPILLARY BLOCK, NARROW ANTERIOR CHAMBER ANGLE OR CLOSURE LEADING TO OCULAR HYPERTENSION. MOST LIKELY VISCOELASTIC SUBSTANCE COULD CAUSE THE TEMPORARY IOP RISE POSTOPERATIVELY WHICH CAN BE CONTROLLED MEDICALLY OR RESUME TO NORMAL PRESSURE AUTOMATICALLY. ALSO MYOPIA IS HIGHLY ASSOCIATED WITH OCULAR HYPERTENSION ESPECIALLY HIGH DEGREE MYOPIA. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743982&pc=MTA | 2023826-2015-00509 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS FOR HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE LENS WAS IMPLANTED ON (B)(6) 2012. THE FACILITY INDICATED THEY AGREED WITH THE PATIENT REPORT. THE PATIENT HAD HIGH PRESSURE (10 - 21MM HG) AND WAS DIAGNOSED WITH GLAUCOMA (OCULAR HYPERTENSION - GLAUCOMA SUSPECT) FOR THE PERIOD OF (B)(6) 2012. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE HIGH PRESSURE AND GLAUCOMA. THE FACILITY INDICATED THE CAUSE OF THE HIGH PRESSURE AND GLAUCOMA WAS FAMILIAL VERSUS OUTFLOW PROBLEM, THE EVENT WAS NOT RELATED TO THE DEVICE. THE PATIENT'S POST-OP VA WAS 20/20-2. THE LENS REMAINS IMPLANTED AND THE PATIENT'S PROGNOSIS IS GOOD. Manufacturer Narrative: (B)(4) IMPLANTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - MANY PATHOPHYSIOLOGICAL CHANGES COULD CAUSE OCULAR HYPERTENSION SUCH AS ANTERIOR CHAMBER DISTORTION, VISCOELASTICS, INFLAMMATION, HEMORRHAGE, PIGMENT DISPERSION, PUPILLARY BLOCK AND VITREOUS IN THE ANTERIOR CHAMBER. EXCESSIVE VAULTING AS A FACTOR COULD CAUSE PUPILLARY BLOCK, NARROW ANTERIOR CHAMBER ANGLE OR CLOSURE LEADING TO OCULAR HYPERTENSION. MOST LIKELY VISCOELASTIC SUBSTANCE COULD CAUSE THE TEMPORARY IOP RISE POSTOPERATIVELY WHICH CAN BE CONTROLLED MEDICALLY OR RESUME TO NORMAL PRESSURE AUTOMATICALLY. ALSO MYOPIA IS HIGHLY ASSOCIATED WITH OCULAR HYPERTENSION ESPECIALLY HIGH DEGREE MYOPIA. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513709&pc=MTA | 2023826-2015-00142 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Glaucoma; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS). THE PATIENT REPORTED WAS PRESCRIBED EYE DROPS FOR AT LEAST THREE CONSECUTIVE MONTHS DUE TO HIGH PRESSURE OR GLAUCOMA INSIDE THE EYE. THE ICL WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316432&pc=MTA | 2023826-2021-03114 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision | NA | 1 | Event Description: ON 13 JUL 2021 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN BMC OPHTHALMOLOGY ENTITLED, 'DIODE LASER TRANSSCLERAL CYCLOPHOTOCOAGULATION CAUSES INTRAOCULAR COLLAMER LENS DISPLACEMENT IN PSEUDOPHAKIC EYE: A CASE REPORT'. THE ARTICLE REPORTS ELEVATED IOP AND LOSS OF VISUAL ACUITY IN OD AFTER ICL IMPLANTATION. A REFRACTORY GLAUCOMA WAS DIAGNOSED AND THE IOP WAS CONTROLLED WITH MEDICATION. PATIENT RECEIVED LASER TRANSSCLERAL CYCLOPHOTOCOAGULATION (TSCP) TO TREAT GLAUCOMA CAUSING ICL TO MOVE FORWARD WHICH INCREASED IOP AGAIN. THE IOP WAS FINALLY CONTROLLED BY IMPLANTING AN AHMED GLAUCOMA VALVE. BCVA REMAINED 20/40 AND NO PROGRESSION GLAUCOMA OCCURRED IN THE NEXT 10 MONTHS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - AN ARTICLE WAS RECEIVED ON 13/07/2021 ENTITLED 'DIODE LASER TRANSSCLERAL CYCLOPHOTOCOAGULATION CAUSES INTRAOCULAR COLLAMER LENS DISPLACEMENT IN PSEUDOPHAKIC EYE: A CASE REPORT'. THE ARTICLE REPORTS ELEVATED IOP; PIGMENT DISPERSION; AND LOSS OF VISUAL ACUITY ASSOCIATED WITH AN ICL IMPLANTATION IN A 39 YEAR OLD PATIENT'S RIGHT (OD) EYE. REFRACTORY GLAUCOMA WAS DIAGNOSED WHICH WAS CONTROLLED WITH MEDICATION AND TREATED WITH LASER TSCP. AFTER SURGERY THE ANTERIOR CHAMBER BECAME SHALLOW, THE ICL WAS INCARCERATED IN THE PUPIL EVEN MOVING THE ICL FORWARD. FOR THIS AN AHMED GLAUCOMA VALVE (AGV) WAS IMPLANTED TO CONTROL IOP AND NO PROGRESSION OF GLAUCOMA OCCURRED IN THE NEXT 10 MONTHS. LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11334358&pc=MTA | 2023826-2021-00364 |  | Injury | STAAR SURGICAL COMPANY | 2021-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: AN ARTICLE WAS RECEIVED FROM BMC OPHTHALMOLOGY (2021) THAT CITED A CASE STUDY OF "EX-PRESS SHUNT IMPLANTATION FOR INTRACTABLE GLAUCOMA WITH POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS". A (B)(6) YEAR OLD MALE PRESENTED WITH PROGRESSIVE BLURRED VISION AND MEDICALLY UNCONTROLLED INTRAOCULAR PRESSURE (IOP) IN THE RIGHT EYE (OD). THE PATIENT WAS PRESCRIBED MEDICATION. THE PATIENT HAD AN IMPLANTABLE COLLAMER LENS (ICL) IMPLANTED IN 2009. ON OPHTHALMIC EXAMINATION, THE ICL WAS WELL PLACED WITH A VAULT HEIGHT OF 456UM. THE ANTERIOR CHAMBER ANGLES WERE OPEN BUT NARROW, AND MILD TO MODERATE TRABECULAR PIGMENTATION WAS NOTED. EX-PRESS GLAUCOMA FILTRATION SURGERY WITHOUT ICL REMOVAL WAS PERFORMED TO CONTROL IOP. AN EX-PRESS P50 SHUNT WAS INSERTED INTO THE ANTERIOR CHAMBER. TRANSIENT HYPOTONY AND SHALLOW ANTERIOR CHAMBER OCCURRED IN THE FIRST WEEK AFTER SURGERY, ALONG WITH AN ICL TILT TOWARDS THE CORNEA WITH REDUCED VAULT HEIGHT. NO OTHER COMPLICATIONS RELATED TO EITHER THE ICL OR THE EX-PRESS SHUNT WERE NOTED. IOP REMAINED STABLE AT 12-14 MMHG AT THE FIRST 3 MONTH FOLLOW-UP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119687&pc=MTA | 2023826-2023-02156 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2017. THE LENS WAS EXPLANTED DUE TO PERMANENT LOSS OF BCVA, ELEVATED IOP 38MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON (B)(6)2023), PIGMENT DISPERSION, BLURRED VISION, IRIS ATROPHY, AND EXCESSIVE VAULT ON (B)(6)2023. ADDITION OF NEW PI, YAG, PERFORMED ON (B)(6)2017. THIS DID NOT RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS POSSIBLY TOO LARGE. IRIDOTOMY WAS NOT DONE PER CHARGE REVIEW. LPI DONE 9 MONTHS POST-OP. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED AN 13.7MM MICL 13.7 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED GLAUCOMA, EXCESSIVE VAULT, PERMANENT LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, BLURRED VISION, IRIS ATROPHY, ZONULAR LAXITY, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, YAG WAS PERFORMED, AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXPLANTED BUT THE PROBLEM WAS NOT RESOLVED. STATUS OF EYE: "SOME ZONULAR LAXITY NOTED AT SURGERY. HAS GLAUCOMATOUS OPTIC NEUROPATHY WITH IOP CONTROLLED ON TIMILOL." ADDITIONAL INFORMATION PROVIDED: "ICL SAVED FOR ANALYSIS BY STAAR AND CONFIRMATION OF SIZE/POWER. RT VISIAN ICL SUBSEQUENTLY REMOVED AND KDB GONIOTOMY PERFORMED ON (B)(6) 2023". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: " LENS POSSIBLY TOO LARGE, W EXCESSIVE VAULT, PIGMENT DISPERSION. PREOP IRIDOCTOMY WAS NOT DONE PER CHART REVIEW. HAD LPI DONE 9 MONTHS POST OP". H6- HEALTH EFFECT- CLINICAL CODE: 4581- "SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES" AND "ZONULAR LAXITY" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "1875" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4644" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. H4 - UNK. H6- HEALTH EFFECT IMPACT CODE 4581: PIGMENT DISPERSION, IRIS ATROPHY. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419929&pc=MTA | 2023826-2018-00537 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/+2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (STEROID TREATMENT). CONCLUSION CODE: OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.89MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7420014&pc=MTA | 2023826-2018-00535 | 2018-03-07 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/+2.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS IRITIS. REPORTEDLY, THE PATIENT WAS DIAGNOSED WITH AN ALLERGY TO BRIMONIDINE ON (B)(6) 2018. ON (B)(6) 2018 SHE WAS DIAGNOSED WITH REBOUND IRITIS. AFTER BEGINNING STEROID TREATMENT THE PATIENT IS IMPROVING. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL/CORRECTED INFORMATION: REPORTER STATES THAT, WHILE BEING TREATED FOR IRITIS, THE PATIENT DEVELOPED ELEVATED IOP. THIS WAS TREATED WITH TRAVATAN Z. THE PATIENT WAS ALLERGIC TO THIS MEDICATION, WHICH WAS DISCONTINUED AND WAS GIVEN COSOPT PF. THE PATIENT HAS NOW REPORTEDLY DEVELOPED, "SIGNIFICANT ONGOING PIGMENT DISPERSION, PIGMENTARY GLAUCOMA, WITH IOPS IN THE 30'S DESPITE MAXIMALLY TOLERATED MEDICAL THERAPY." THE PATIENT IS ALSO CURRENTLY PHOTOPHOBIC. THE ICL WAS EXPLANTED ON (B)(6) 2018; STEROID INCREASED TO EVERY 2 HOURS AND CYCLOGEL GIVEN FOR PHOTOPHOBIA. (B)(4). METHOD AND RESULTS CODES CORRECTED. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419890&pc=MTA | 2023826-2018-00531 | 2018-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Iritis; Visual Impairment; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.5/+2.5/166 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, REPORTEDLY DUE TO PIGMENT DISPERSION, UNREACTIVE (MID-DILATED FIXED) PUPIL, AND IRITIS. A YAG PI WAS PERFORMED ON (B)(6) 2018. THE SURGEON STATED HE THOUGHT THE PATIENT MAY HAVE HAD AN ASYMPTOMATIC PUPIL BLOCK DUE TO THE VISCOELASTIC USED IN SURGERY, BUT IS UNABLE TO THEN EXPLAIN WHY IT TOOK 2-3 WEEKS FOR IRITIS AND PIGMENT DISPERSION TO DEVELOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: AFTER THE LENS EXPLANT, THE PATIENT WAS ADVISED TO TAKE PRED/GATI 3 TIMES DAILY. AFTER DISCONTINUING, THE PATIENT DEVELOPED AN INCREASE IN IOP AND INFLAMMATION IN THE ANTERIOR CHAMBER. CURRENTLY, THE PATIENT IS BEING TREATED WITH PREDNISOLONE AND COMBIGAN, AZOPT, AND TRAVATAN Z FOR PRESSURE SPIKES. IT IS BELIEVED THAT THE SPIKES ARE DUE TO "STEROID USE VS INFLAMMATORY CAUSES." REPORTEDLY, NO INFECTION HAS OCCURRED WITHIN THE EYE. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT REPORTS "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA, AND LOSS OF VISION. ALSO STATED IS THAT THE IRIS IS PERMANENTLY FIXED AND DILATED. GLAUCOMA SURGERY PERFORMED. G3-REPORT SOURCE ADDED. H6-ADDITIONAL 3191: (BLOWN IRIS, FIXED IRIS, GLAUCOMA SURGERY). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, LENS EXPLANT, FIXED PUPIL) OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT (46YRS), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.8MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16326950&pc=MTA | 2023826-2023-00266 | 2022-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON(B)(6)2022. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00265. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - ENLARGE PERIPHERAL IRIDOTOMY. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ELEVATED INTRAOCULAR PRESSURE AND GLAUCOMA. AN IRIDOTOMY WAS PERFORMED BUT THE TENSION DID NOT GO DOWN. THE PATIENT HAD A CATARACT LENS IMPLANTED AND THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00265. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: MEDICAL DEVICE CODE: 2993. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031623&pc=MTA | 2023826-2019-01698 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT & REFRACTIVE SURGERY IN JUNE 2019 TITLED, '10-YEAR CLINICAL OUTCOMES AFTER IMPLANTATION OF A POSTERIOR CHAMBER PHAKIC LENS FOR MYOPIA.' THE ARTICLE STATES THAT GLAUCOMA FILTERING SURGERY WAS PERFORMED ON ONE EYE DUE TO PIGMENTARY GLAUCOMA. THE ICL WAS NOT REMOVED. AFTER FILTERING SURGERY, THE IOP WAS STABILIZED TO NORMAL RANGE WITHOUT ANY IOP-LOWERING MEDICATIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INITIAL MDR SHOULD HAVE ALSO STATED: 'INITIALLY, TOPICAL MEDICATIONS WERE USED TO REDUCE IOP, BUT GLAUCOMATOUS CHANGES PROGRESSED, SO FURTHER INTERVENTION WAS REQUIRED.' PATIENT CODE 3191 (SECONDARY SURGICAL INTERVENTION - TREATMENT WITH MEDICATIONS, GLAUCOMA FILTERING SURGERY) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8853040&pc=MTA | 2023826-2019-01329 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - EXPLANTATION WAS PERFORMED IN 2 EYES OF ONE PATIENT DUE TO DE-CENTRATION OF THE ICL AND DEVELOPMENT OF ACUTE GLAUCOMA ATTACK WITH INSUFFICIENT EFFECT OF CONSERVATIVE AND SURGICAL SOLUTION OF THE CONDITION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 2993 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR #1. CLAIM#: (B)(4). Manufacturer Narrative: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - ONE EYE (OS) WITH AN IOP GREATER THAN 50MMHG, GLAUCOMA, AND ADDITIONAL ADVERSE EVENTS WAS OBSERVED SOON AFTER SURGERY. IN THE PATIENT'S LEFT EYE (OS), AN ACUTE GLAUCOMA ATTACK DEVELOPED WITH INADEQUATE RESPONSE TO TREATMENT. THE ICL WAS EXPLANTED AND THE PATIENT WENT ON TO HAVE LASER REFRACTIVE SURGERY. AT LAST FOLLOW UP SEVEN YEARS AFTER SURGERY THE IOP REMAINED STABLE AT 11 MMHG. THE AUTHOR STATED THAT THE 'POSTOPERATIVE COURSE WAS CONDITIONED BY THE IMPLANTATION OF THE INCORRECT SIZE OF ICL' AND THAT 'SOON AFTER THE PROCEDURE, EXCESSIVE VENTRAL BUCKLING OCCURRED DUE TO THE INFLUENCE OF THE INAPPROPRIATE LENGTH OF THE ICL, WITH THE DEVELOPMENT OF ACUTE PUPILLARY BLOCK'. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.' COMMON DEVICE NAME CORRECTED TO PHAKIC INTRAOCULAR LENS & PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: MTA & QCB (COHORT INCLUDED SEVERAL LENS MODELS). DEVICE INFORMATION CORRECTED TO ICM V4, VICMO, TICM V4, VTICM, VTICMO (COHORT INCLUDED SEVERAL LENS MODELS). REPORT SOURCE SHOULD HAVE INCLUDED LITERATURE IN INITIAL MDR SHOULD HAVE INCLUDED PATIENT CODE 3191 (EXCESSIVE VENTRAL BUCKLING, ACUTE PUPILLARY BLOCK; SECONDARY INTERVENTION, TREATED WITH MEDICATION, EXPLANT, LASER REFRACTIVE SURGERY) IN INITIAL MDR. DEVICE CODE 2923 IS NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6855474&pc=MTA | 2023826-2017-01419 | 2016-09-30 | Injury | STAAR SURGICAL COMPANY | 2017-09-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2014. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP, PIGMENT DISPERSION, AND MILD VISUAL FIELD LOSS. SMALL SIGNS OF PIGMENT DISPERSION WAS SEEN AT PREOP EXAM, BUT (CONSIDERING DEVELOPMENT) WAS MIS-ASSESSED. SURGICAL ENLARGEMENT OF PI WAS PERFORMED AND IOP LOWERING EYE DROPS WERE USED. GLAUCOMATOUS DISC WAS SEEN IN THE EYE, BUT OTHERWISE THE EYE IS NORMAL. PATIENT'S POST-OP BCVA ON (B)(6) 2017 WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS -8.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. SURGICAL ENLARGEMENT OF PI WAS NOT PERFORMED. (B)(4). Manufacturer Narrative: EXPLANTED (B)(6) 2017. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6110077&pc=MTA | 2023826-2016-01663 | 2016-01-02 | Injury | STAAR SURGICAL COMPANY | 2016-11-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, UNKNOWN DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE PATIENT UNDERWENT AN ENLARGEMENT OF PERIPHERAL IRIDECTOMY OR ADDITION OF NEW PERIPHERAL IRIDECTOMY USING YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS OF A DIFFERENT MODEL AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND LENS UNABLE TO EVALUATE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6249746&pc=MTA | 2023826-2017-00051 | 2016-03-20 | Injury | STAAR SURGICAL COMPANY | 2017-01-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/098 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), AND DEVELOPMENT OF GLAUCOMA. AS OF THE DATE OF THIS SUBMISSION, THE LENS HAS BEEN RETURNED, BUT THE EVALUATION RESULTS ARE NOT YET AVAILABLE. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS PERFORMED. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6054107&pc=MTA | 2023826-2016-01522 | 2016-10-10 | Injury | STAAR SURGICAL COMPANY | 2016-10-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 LENS, -3.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. FOLLOWING THE SURGERY, THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ACD. ON (B)(6) 2016, THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE (IOP) AND CLOSED ANGLE GLAUCOMA DUE TO EXCESSIVE VAULTING OF THE LENS. THE PATIENT RECEIVED A YAG PERIPHERAL IRIDOTOMY, AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8451848&pc=MTA | 2023826-2019-00491 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED PATIENT STARTED TO GET SECONDARY GLAUCOMA IN RIGHT EYE (OD). LENS REMAINS IMPLANTED. REPORTEDLY PATIENT IS ON ANTIGLAUCOMA MEDICATIONS. PRESENT IOP IS UNDER CONTROL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF 14.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD). THE PATIENT DEVELOPED SECONDARY GLAUCOMA AND THE VAULT WAS FOUND TO BE 1197 MICRONS (DETERMINED TO BE EXCESSIVE VAULT). THE PATIENT IS ON ANTI-GLAUCOMA MEDICATIONS AND PRESENT IOP IS UNDER CONTROL. THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8547014&pc=MTA | 2023826-2019-00667 | 2018-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED 12.1MM, VTICM5_12.1,-12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION, BLURRED VISION AND IRIS ASTROPHY. PHACO-EMULSIFICATION AND GLAUCOMA FILTRATION SURGERY WERE REPORTED AS SECONDARY SURGICAL INTERVENTIONS. THE REPORTER STATED "THE PATIENT DEVELOPED PIGMENT DISPERSION GLAUCOMA, BEYOND ANTI-GLAUCOMA CONTROLLED AND THE SURGEON PLANNED FOR GLAUCOMA FILTRATION SURGERY WITH PHACO". POST ICL EXPLANT THE IOP WAS STILL HIGH AND REQUIRED GLAUCOMA SURGERY. PHACO TRABECULECTOMY WAS PERFORMED. THE GLAUCOMA WAS REPORTED AS "CONTROLLED", PATIENTS BCVA 6/20-2. PHACO SURGERY WAS PERFORMED AND IOL WAS IMPLANTED TO REPLACE THE NATURAL CRYSTALLINE LENS. Manufacturer Narrative: BCVA AT LAST POST-OP VISIT WAS 20/20-2. ADDITIONALLY, IT WAS REPORTED THAT THERE WAS NO ASCO (ANTERIOR SUBCAPSULAR LENS OPACITY). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE OPTIC OF THE LENS WAS BENT AND THE HAPTIC WAS BENT AND DEFORMED. H3- DIMENSIONAL INSPECTION: LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703234&pc=MTA | 2023826-2020-00267 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | Glaucoma; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -01.75 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION AND UNREACTIVE (FIXED) PUPIL. THE SURGEON DECIDED NOT TO IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. THE PATIENTS POST-OP CONDITION WAS GLAUCOMA ASSOCIATED WITH PIGMENTARY DISPERSION. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (SAFETY AND EFFECTIVENESS OF THE LENS HAS NOT BEEN ESTABLISHED IN PATIENTS WITH HISTORY OF PREVIOUS OCULAR SURGERY INCLUDING REFRACTIVE CORNEAL SURGERY). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - METHOD CODE: 3331 SHOULD BE CORRECTED TO 4110 IN SUPPLEMENTAL MEDWATCH REPORT #2. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7419948&pc=MTA | 2023826-2018-00533 | 2018-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2018-04-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/+1.0/007 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS PUPIL BLOCK, ANGLE CLOSURE, ELEVATED IOP, AND PIGMENT DISPERSION. THE SURGEON STATED HE THOUGHT THE HIGH IOP WAS DUE TO THE VISCOELASTIC. REPORTEDLY, HE ASSUMES THE AQUA PORT WAS BLOCKED BY THE VISCOELASTIC. A YAG WAS PERFORMED AND THEN IOP RETURNED TO "NORMAL." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THE PATIENT REPORTS THE LENS WAS EXPLANTED. ALSO REPORTED: "PRESSURE ISSUES," "BLOWN IRIS," ONGOING GLAUCOMA AND VISION LOSS, GLAUCOMA SURGERY PERFORMED, IRIS PERMANENTLY FIXED AND DILATED. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (PIGMENT DISPERSION, SECONDARY SURGERY, ANGLE CLOSURE, YAG) CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6)), ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.67MM)] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7294520&pc=MTA | 2023826-2018-00267 | 2018-01-23 | Injury | STAAR SURGICAL COMPANY | 2018-02-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Glaucoma; Intraocular Pressure Increased; Pupillary Block; No Code Available | NA | 1 | Event Description: MALIGNANT GLACOMA WITH EXPLANT WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 VTICM5_13.2, -13.5/+1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED 5 HOURS LATER DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL PROCEDURES (ENLARGEMENT OF PI OR ADDITION OF NEW PI AND YAG) WERE REPORTED TO BE PERFORMED. ENTERED STAAR (B)(4)'S ADDRESS DEVICE EVALUATION: LENS WAS RETURNED IN A CONTACT CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6262587&pc=MTA | 2023826-2017-00092 | 2015-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Glaucoma; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO COMBINED MECHANISM GLAUCOMA. THE INCREASED INTRAOCULAR PRESSURE WAS NOT CAUSED BY THE ICL BUT WAS DUE TO AN UNKNOWN SICKLE CELL TRAIT AND YAG PI. THE PATIENT HAD REFRACTIVE LENS EXCHANGE AS RECOMMENDED BY A GLAUCOMA PROVIDER AND WAS PERFORMED AT THE TIME OF EXPLANTATION. THE PATIENTS POST-OP BSCVA WAS 20/25. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH A PIECE OF THE HAPTIC MISSING. THERE WAS DRY, CLEAR RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221710&pc=MTA | 2023826-2015-01525 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Glaucoma; Loss of Vision | NA | 1 | Event Description: THE REPORTER INDICATED A PATIENT HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED. THE PATIENT DEVELOPED ACUTE GLAUCOMA AND LOST MOST OF HER VISION IN ONE EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACUTE GLAUCOMA AFTER ICL IMPLANTATION MAY BE DUE TO: REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, NON-PATENT/FUNCTIONING IRIDECTOMIES (IN THE V4 AND V4B MODELS), OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, BLOCKAGE OF THE PORT (IN THE V4C MODEL), IRIDO-CICLITIS, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ETC). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10032334&pc=MTA | MW5094406 | 2018-02-16 | Injury | STAAR SURGICAL COMPANY | 2020-05-06 | MTA | STAAR EVO+ VISIAN ICL/ IOL INTRAOCULAR LENSE | Pressure Problem | Glaucoma; Loss of Vision; No Code Available |  | 1 | Event Description: INSTALL ICL LED TO PRESSURE ISSUES, BLOWN IRIS, ONGOING GLAUCOMA AND LOSS OF VISION. LENS WAS IMPLANTED AND AFTER A FEW WEEKS EXPERIENCED PRESSURE ISSUES, IRIS PERMANENTLY FIXED A DILATED AND ONGOING GLAUCOMA ISSUES.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8799310&pc=MTA | 2023826-2019-01274 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 DUE TO GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: RETURN DATE SHOULD BE CORRECTED TO 21-JUN-2019 IN THE FOLLOW UP #1 MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8799452&pc=MTA | 2023826-2019-01273 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED FROM THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 DUE TO GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RETURN DATE SHOULD BE CORRECTED TO 21-JUN-2019 IN PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5193790&pc=MTA | 2023826-2015-01396 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ANGLE CLOSURE GLAUCOMA. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: RESULTS (EVALUATION RESULT): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A DENT ON A LENS HAPTIC. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, VISUAL INSPECTION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW - REPORTEDLY ICL WAS EXPLANTED FOUR DAYS POSTOPERATIVELY TO ADDRESS ELEVATED IOP DUE TO ANGLE CLOSURE GLAUCOMA. NO POSTOPERATIVE SEQUELAE WERE REPORTED. ACCORDING TO THE REPORT, DATED ON (B)(6) 2015, THE SURGEON STATED THAT HE BELIEVED THAT THERE WAS RETAINED VISCOELASTIC (OCUCOAT) BEHIND IMPLANTED ICL. IT SHOULD BE NOTED THAT ACUTE RISE IN IOP COULD BE A CONSEQUENCE OF RETAINED VISCOELASTIC AND PER DFU PERCAUTIONS: "3. INADEQUATE FLUSHING OF THE VISCOELASTIC FROM THE EYE MAY LEAD TO INTRAOCULAR PRESSURE (IOP) SPIKES. IOP SHOULD BE CHECKED 24 HOURS POSTOPERATIVELY." GIVEN THIS INFORMATION IT IS VERY LIKELY THAT PROCEDURE RELATED FACTOR (LEFT VISCOELASTIC) HAS CAUSED THE REPORTED EVENT. CONCLUSION: (EVALUATION CONCLUSION): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, IT IS VERY LIKELY THAT PROCEDURE RELATED FACTOR (LEFT VISCOELASTIC) HAS CAUSED THE REPORTED EVENT. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641729&pc=MTA | 2023826-2017-00977 |  | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | INPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Glaucoma; No Code Available | N/A | 1 | Event Description: ON (B)(6) 2017 THE REPORTER STATED THAT A PATIENT WITH PIGMENTATION GLAUCOMA AFTER 18 MONTHS OF ICL IMPLANTATION HAD THE LENS EXPLANTED. IT IS UNKNOWN IF THERE IS ANY CORELATION BETWEEN THE GLAUCOMA AND ICL IMPLANTATION. THE PATIENT ALSO HAS PIGMENTATION PROBLEMS AND A DAMAGED OPTICAL NERVE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: INITIAL REPORTER DETAILS IN THE INTIAL HAVE BEEN CORRECTED. (B)(4). Manufacturer Narrative: THE PATIENT ALREADY HAD PIGMENTATION GLAUCOMA PRIOR TO ICL SURGERY. CONCLUSION CODE PER DFU, VTICL'S ARE CONTRAINDICATED FOR IMPLANTATIONS IN PATIENTS WITH PRE-EXISTING CONDITION SUCH AS PRIMARY OPEN ANGLE OR NARROW ANGLE GLAUCOMA. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372539&pc=MTA | 2023826-2019-00285 | 2018-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Glaucoma; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. ON (B)(6) 2019 THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND GLAUCOMA. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "DIAMETER OF LENS TOO LONG." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIP LOCK BAG. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9244518&pc=MTA | 2023826-2019-01974 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Glaucoma; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -02.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. REPORTEDLY, CELL COUNT DECREASE AND ACUTE GLAUCOMA WERE OBSERVED. PER THE REPORTER ON (B)(6) 2019, LENS WAS REMOVED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. DEVICE CODE 1494: OFF LABEL USE, ACD < 3.0MM. DEVICE CODE 2993 NOT APPLICABLE. EVENT: THE LENS WAS IMPLANTED ON (B)(6) 2019. A VAULT OF 2000 MICRONS WAS ALSO REPORTED. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS DUE TO THE DEVICE. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM# (B)(4). Manufacturer Narrative: A LASER PI WAS PERFORMED 1 WEEK PRIOR TO IMPLANTATION. EXCESSIVE VAULT WAS REPORTED. CLAIM# (B)(4). Manufacturer Narrative: COUNT DECREASE AND ACUTE GLAUCOMA WERE OBSERVED.' CORRECTED TO ' CELL COUNT DECREASE AND CAUTION OF ACUTE GLAUCOMA WERE OBSERVED.' CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. WEIGHT: UNK. DATE OF EVENT: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5807434&pc=MTA | 2023826-2016-00939 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Pupillary Block; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES BILATERALLY IN THE PATIENT'S EYES. THE REPORTER INDICATED POST-OP THE PATIENT HAD PUPILLARY BLOCK GLAUCOMA DUE TO SHALLOW CHAMBER. BOTH LENSES WERE EXPLANTED, ONE LENS HAD A LENS EXCHANGE AND THE OTHER LENS WAS RE-IMPLANTED. THE PATIENT IS STILL HAVING ISSUES WITH THE LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR. REPORT #2023826-2016-00940 FOR THE FELLOW EYE. Manufacturer Narrative: DATE OF EVENT: UNK. EXPIRATION DATE - UNK. DATE OF IMPLANT: UNK. DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810884&pc=MTA | 2023826-2016-00940 |  | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Glaucoma; Pupillary Block; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENSES BILATERALLY IN THE PATIENT'S EYES. THE REPORTER INDICATED POST-OP THE PATIENT HAD PUPILLARY BLOCK GLAUCOMA DUE TO SHALLOW CHAMBER. BOTH LENSES WERE EXPLANTED, ONE LENS HAD A LENS EXCHANGE AND THE OTHER LENS WAS RE-IMPLANTED. THE PATIENT IS STILL HAVING ISSUES WITH THE LENSES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR. REPORT #2023826-2016-00939 FOR THE FELLOW EYE. Manufacturer Narrative: DATE OF EVENT: UNK. EXPIRATION DATE - UNK. DATE OF IMPLANT: UNK. DATE OF EXPLANT: UNK. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179666&pc=MTA | 2023826-2016-01773 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED A HALO FOGGY RING AROUND LIGHTS WHEN IT IS DARK, CAUSING ISSUES WHEN DRIVING. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. THE PATIENT HAS NOT SEEN A DOCTOR IN OVER THREE YEARS AND NO MORE INFORMATION WAS AVAILABLE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6179667&pc=MTA | 2023826-2016-01774 |  | Injury | STAAR SURGICAL COMPANY | 2016-12-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem | Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN HER LEFT EYE (OS). THE PATIENT REPORTED A HALO FOGGY RING AROUND LIGHTS WHEN IT IS DARK, CAUSING ISSUES WHEN DRIVING. THE LENS WAS IMPLANTED ON (B)(6) 2011 AND REMAINS IMPLANTED. THE PATIENT HAS NOT SEEN A DOCTOR IN OVER THREE YEARS AND NO MORE INFORMATION WAS AVAILABLE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4522031&pc=MTA | 2023826-2015-00164 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD) IN (B)(6) 2013. THE PATIENT REPORTED HE HAS BEEN EXPERIENCING SYMPTOMS OF HALOS, PREDOMINANTLY IN THE LEFT OPTIC. THE PATIENT REPORTED THE HALOS ARE PROMINENT AT NIGHT WHEN HE IS DRIVING. DURING THE DAY HE OCCASIONALLY HAS DOUBLE VISION. AFTER CONSULTING WITH HIS SURGEON, HE WAS INFORMED THAT HE HAS THE OPTION OF REMOVING THE LENS BUT HAS NOT DECIDED. THE PATIENT IS ALSO SCHEDULED FOR PRK SURGERY. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4) - HALO; (DOUBLE VISION). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THERE ARE INDIVIDUAL VARIANCES OF EYE BALL AND PUPIL SIZE. THE DECENTRATION OF ICL LENS AND /OR THE OPTIC DIAMETER OF ICL LENS THAT IS TOO SMALL RELATIVE TO PUPIL SIZE CONTRIBUTED TO THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4863342&pc=MTA | 2023826-2015-00708 | 2015-04-27 | Injury | STAAR SURGICAL COMPANY | 2015-06-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAS ANGLES ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE PATIENT ALSO EXPERIENCED BLURRED VISION AND GLARES/HALOS THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. THE PATIENT'S LAST BCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375453&pc=MTA | 2023826-2023-05764 | 2023-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -7.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT HAS LOW VAULTING WITH ROTATION, THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE, AND THE PATIENT HAS GLARE/HALOES. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077720&pc=MTA | 2023826-2022-02310 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VISIAN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD A HIGH VAULT AND THE PATIENT COMPLAINED OF HALOES POST-OP. THE PATIENT IS USING BRIMONIDINE WITH SOME HELP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT WAS ADMINISTERED 0.5% PILOCARPINE. THE PATIENT IS HESITANT FOR ANY LENS REMOVAL/REPLACMENT. H6: HEALTH EFFECT IMPACT CODE: 4644 - PILOCARPINE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. HEALTH EFFECT IMPACT CODE: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077724&pc=MTA | 2023826-2022-02309 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A VISIAN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD A HIGH VAULT AND THE PATIENT COMPLAINED OF HALOES POST-OP. THE PATIENT IS USING BRIMONIDINE WITH SOME HELP. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5: THE PATIENT WAS ADMINISTERED 0.5% PILOCARPINE. THE PATIENT IS HESITANT FOR ANY LENS REMOVAL/REPLACEMENT. H6: HEALTH EFFECT IMPACT CODE: 4644 - PILOCARPINE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7405119&pc=MTA | 2023826-2018-00528 |  | Malfunction | Unknown Manufacturer | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Halo | N/A | 1 | Event Description: PATIENT REPORTED THAT THEY ARE STILL EXPERIENCING HALOS FROM LIGHT SOURCES AT NIGHT 11 MONTHS AFTER VISIAN ICL SURGERY. PATIENT HAS HAD MULTIPLE OPTOMETRISTS EXAMINE HER EYES SINCE THEN AND THEY CONFIRM THAT VISION IS PERFECT AND EYES ARE OTHERWISE HEALTHY. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17664984&pc=MTA | 2023826-2023-03784 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Halo | UNK | 1 | Event Description: IN THE ARTICLE, "CYSTOID MACULAR EDEMA ASSOCIATED WITH OMIDENEPAG ISOPROPYL IN A PHAKIC EYE WITH AN IMPLANTABLE COLLAMER LENS: A CASE REPORT." AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S LEFT EYE (OS). 12 YEARS LATER, GLAUCOMA WAS OBSERVED. OMIDENEPAG ISOPROPYL WAS PRESCRIBED AND AFTER 9 MONTHS, BLURRED VISION WAS OBSERVED. SWEFT SOURCE OPTICAL COHERENCE TOMOGRAPY REVEALED CYSTOID MACULAR EDEMA IN THE EYE AND OMIDENEPAG ISOPROPHYL USAGE WAS DISCONTINUED AND BROMFENAC SODIUM HYDRATE WAS ADMINISTERED. THIS HELPED THIS ISSUE AND REPORTEDLY, "OPTICAL COHERENCE TOMOGRAPHY EXAMINATION CONFIRMED THAT THE MACULA HAD NORMALIZED." Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11200456&pc=MTA | 2023826-2021-00185 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-01-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Halo | NA | 1 | Event Description: THE REPORTER STATED " TOO MUCH HALOES", UNSATISFIED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ESUBMITTER DOES NOT ALLOW YOU TO PUT UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15468496&pc=MTA | 2023826-2022-03254 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Halo | P030016 | 1 | Event Description: HALOS IN A PATIENT THAT WAS IMPLANTED WITH A PAIR OF EVO LENSES IN (B)(6) 2022 WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: NO SERIAL NUMBERS REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347362&pc=MTA | 2023826-2023-03025 | 2023-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -10.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE PATIENT HAD GLARE/HALOES. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. IT WAS ADDITIONALLY INDICATED THERE WILL BE NO REPLACEMENT LENS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17273280&pc=MTA | 2023826-2023-02720 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -18.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. IT WAS INDICATED THE NEXT DAY AFTER IMPLANT THE PATIENT SAW A SECTOR OF HALOS, HIS INITIAL LENS IMPLANT OF DOUBLE VISION DISAPPEARED, AND HE FELT MUCH BETTER. THE REPORTER INDICATED HIS HALO SYMPTOM ALMOST DISAPPEARED AND IT'S UNLIKELY AN EXCHANGE WILL OCCUR. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8932871&pc=MTA | 2023826-2019-01547 | 2019-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 AND THE SURGERY WENT WITHOUT ANY COMPLICATION. ON (B)(6) 2019, THE PATIENT COMPLAINED OF HALOS AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: OTHER SERIOUS, SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT. SERIOUS INJURY SHOULD BE CORRECTED TO MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241368&pc=MTA | 2023826-2021-02899 | 2021-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT COMPLAINED OF MILD HALOS AND THE LENS REMAINED IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02043. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538491&pc=MTA | 2023826-2022-01436 | 2022-04-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL 13.2 IMPLANTABLE COLLAMER LENS OF -14.5 DIOPTER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS. "PATIENT IS BOTHERED BY NIGHTTIME HALOES." THE LENS REMAINS IMPLANTED. SEE MFR# 2023826-2022-01426 FOR THE INITIAL IMPLANT LENS. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584545&pc=MTA | 2023826-2022-03620 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -8.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED GLARE/HALOS. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "DYSHOTOPSIA FOR AQUAPORT". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748577&pc=MTA | 2023826-2023-01351 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -3.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO HALO AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. FOR CAUSE DETAILS THE REPORTER INDICATED "THE PATIENT WAS NOT COMFORTABLE WITH HIS VISION". Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253088&pc=MTA | 2023826-2023-02690 | 2023-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -3.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO DISSATISFACTION AND GLARE/HALOES. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583784&pc=MTA | 2023826-2020-00093 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE PATIENT HAD ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN RIGHT EYE (OD) TO CORRECT MYOPIA AND ASTIGMATISM. THE PATIENT SEES VERY WELL AND CLEAR BUT HAS COMPLAINED OF HALOS AROUND LIGHTS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583785&pc=MTA | 2023826-2020-00092 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE PATIENT HAD ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN LEFT EYE (OS) TO CORRECT MYOPIA AND ASTIGMATISM. THE PATIENT SEES VERY WELL AND CLEAR BUT HAS COMPLAINED OF HALOS AROUND LIGHTS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11478868&pc=MTA | 2023826-2021-00637 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS. IT WAS REPORTED THE PATIENT POSSIBLY EXPERIENCED HALOS DUE TO AQUAPORT. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE PATIENT IS DOING WELL, AND IS UNDER OBSERVATION BY THE DOCTOR. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11632493&pc=MTA | 2023826-2021-00941 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-04-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | NA | 1 | Event Description: ON (B)(6) 2021 WE RECEIVED NOTIFICATION OF AN ARTICLE ENTITLED, 'VISUAL OUTCOMES AFTER SMALL INCISION LENTICULE EXTRACTION AND IMPLANTABLE COLLAMER LENS V4C FOR MODERATE MYOPIA: 1-YEAR RESULTS.' THE ARTICLE REPORTS THAT THIRTY-TWO EYES OF 20 PATIENTS UNDERWENT EVO-ICL IMPLANTATION, LEADING COMPLAINTS AFTER ICL IMPLANTATION WERE HALOS. THE ARTICLE STATED "THE MOST PREVALENT VISUAL DISTURBANCE AFTER ICL WERE HALOS WITH A PREVALENCE OF 84.4%. HOWEVER, THIS COMPLAINT WAS CONSIDERED TO BE DISTRESSING IN LESS THAN 10% AND SECONDARY INTERVENTION WAS NEVER REQUIRED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378655&pc=MTA | 2023826-2021-03368 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS HALOS AROUND LIGHT IN THE EVENING AND NIGHT IN THE CENTRAL VISUAL FIELD. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE-UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466238&pc=MTA | 2023826-2022-01633 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS; -14.5 DIOPTER INTO THE PATIENT'S EYE. 6 WEEKS POST OP REPORTS A HALO EFFECT IN ENVIRONMENTS THAT ARE NOT BRIGHTLY ILLUMINATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16080994&pc=MTA | 2023826-2022-04797 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo |  | 1 | Event Description: THE REPORTER (CONSUMER) INDICATED THAT AN EVO ICL WAS IMPLANTED INTO THE PATIENTS EYE AND "HALOS ARE SO BAD AT NIGHT."NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND MEDICATION WAS PRESCRIBED. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." IT WAS ALSO REPORTED "THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: ADDIITONAL DATA: B5- "THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." THE LENS REMAINS IMPLANTED." H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE CONSUMER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED HALOES AND MEDICATION WAS PRESCRIBED. REPORTEDLY "I HAVE THE EVO ICL BUT THE HALOS ARE SO BAD AT NIGHT WOULD A REPLACEMENT WITH AN EVO + ICL HAVE BETTER RESULTS AND HELP WITH REDUCING HALOS." IT WAS ALSO REPORTED "THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". IT WAS LATER REPORTED " I HAD AN APPOINTMENT TODAY AND THEY SAID THE LENS' LOOK GOOD AND I COULD GET ANOTHER OPINION FROM ONE OF THE SURGEONS BUT THEY SUGGESTED THAT I KEEP THE LENS' AS IS AND USE GLAUCOMA DROPS TO SHRINK MY PUPIL TO HELP GET RID OF THE HALOS." THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152675&pc=MTA | 2023826-2023-00105 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO HALOS. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOES. REPORTEDLY "NO REPLACEMENT LENS WAS IMPLANTED". CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152716&pc=MTA | 2023826-2023-00104 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO HALOS. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.00 INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO HALOES. REPORTEDLY "NO REPLACEMENT LENS WAS IMPLANTED". CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16579605&pc=MTA | 2023826-2023-00975 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE SURGEON REPORTED THAT A PATIENT HE AND A CO-MANAGED DOCTOR TREATED WITH ICL BACK IN (B)(6) 2022 COMPLAINS OF GOLDEN HALO IN BRIGHT LIGHT AND IN HEADLIGHTS. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748170&pc=MTA | 2023826-2023-01388 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER (CONSUMER) INDICATED THAT A 13.2MM VICM5_13.2 -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 29-JUN-2022. REPORTER (CONSUMER) EXPERIENCED HALOS, HE STATES " ITS REALLY BAD AT NIGHT". LENS REMAINS IMPLANTED. REPORTER STATED " RIGHT NOW THEY GAVE ME GLAUCOMA DROPS AND THAT SEEMS TO HELP". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: IT WAS LATER REPORTED THAT THE PATIENT HAD AN APPOINTMENT ON (B)(6) 2023. THE REPORTER STATED " THEY SAID THE LENS' LOOK GOOD AND I COULD GET ANOTHER OPINION FROM ONE OF THE SURGEONS BUT THEY SUGGESTED THAT I KEEP THE LENS' AS IS AND USE GLAUCOMA DROPS TO SHRINK MY PUPIL TO HELP GET RID OF THE HALOS". LENS REMAINS IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16899774&pc=MTA | 2023826-2023-01718 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "PATIENT WITH ONGOING COMPLAINTS OF GLARE AND HALO POST-EVO," ALSO, "THE PATIENT IS 20/15." Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17794635&pc=MTA | 2023826-2023-04099 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED DUE TO HALOES. Manufacturer Narrative: H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT HAS HALOES. LENS REMAINS IMPLANTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17795124&pc=MTA | 2023826-2023-04098 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REMOVED DUE TO HALOES. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN PARAMETERS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT THE LENS REMAINS IMPLANTED, AND THAT THE PATIENT IS PLANNING FOR REMOVAL SURGERY DUE TO INTOLERANCE TO HALOS. CLAIM # (B)(4). Manufacturer Narrative: H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17980256&pc=MTA | 2023826-2023-04711 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO 10 PATIENT'S EYES. IT WAS REPORTED THERE WAS PERFECT VA, PERFECT VAULT 600-800 MICROS BUT PATIENT'S INSIST OF AN EXPLANT DUE TO THEY CANNOT TOLERATE THE GLARES AND HALOES ANYMORE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4522032&pc=MTA | 2023826-2015-00165 |  | Injury | STAAR SURGICAL COMPANY | 2015-02-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Halo | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN HIS LEFT EYE (OS) IN (B)(6) 2013. THE PATIENT REPORTED HE HAS BEEN EXPERIENCING SYMPTOMS OF HALOS, PREDOMINANTLY IN THE LEFT OPTIC. THE PATIENT REPORTED THE HALOS ARE PROMINENT AT NIGHT WHEN HE IS DRIVING. DURING THE DAY HE OCCASIONALLY HAS DOUBLE VISION. THE PATIENT REPORTED THE ICL ROTATED 70 DEGREES IN THE LEFT EYE. AFTER CONSULTING WITH HIS SURGEON, HE WAS INFORMED THAT HE HAS THE OPTION OF REMOVING THE LENS BUT HAS NOT DECIDED. THE PATIENT IS ALSO SCHEDULED FOR PRK SURGERY. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THERE ARE INDIVIDUAL VARIANCES OF EYE BALL AND PUPIL SIZE. THE DECENTRATION OF ICL LENS AND /OR THE OPTIC DIAMETER OF ICL LENS THAT IS TOO SMALL RELATIVE TO PUPIL SIZE CONTRIBUTED TO THE EVENT. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802022&pc=MTA | 2023826-2023-01515 | 2023-02-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE VAULT IS STABLE BUT PATIENT IS FEELING MYOPIC. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT- CLINICAL CODE: 4581- PHOTOPHOBIA. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.0D INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. IT WAS REPORTED THAT THE PATIENT WAS FEELING MYOPIC, THE PATIENT ALSO EXPERIENCED GLARES/HALOES AND PHOTOPHOBIA. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - MYOPIC. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802019&pc=MTA | 2023826-2023-01489 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -4.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THAT THE PATIENT HAD A SLIGHT INCREASE IN VAULT OVER A FEW DAYS AND HYPEROPIA. THE LENS WAS REPOSITIONED IN THE VERTICAL POSITION; FROM 30 DEGREES TO 90 AND VAULT IS AT 480UM. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. H6: HEALTH EFFECT- CLINICAL CODE: 4581- PHOTOPHOBIA. B5: THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER +4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, HYPEROPIA, GLARES/HALOES, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND PHOTOPHOBIA. ON (B)(6) 2023 THE LENS WAS REPOSITIONED VERTICALLY. REPORTEDLY "ICL STILL TILTED. STILL VERY MUCH BOTHERED BY PHOTOPHOBIA". THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA, SLIGHTLY TILTED ICL". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, HYPEROPIA, GLARES/HALOES, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND PHOTOPHOBIA. ON (B)(6) 2023 THE LENS WAS REPOSITIONED VERTICALLY. REPORTEDLY "ICL STILL TILTED. STILL VERY MUCH BOTHERED BY PHOTOPHOBIA". THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "SEVERE PHOTOPHOBIA, SLIGHTLY TILTED ICL". THE CAUSE OF THE EVENT WAS THE DEVICE. CLAIM# (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - HYPEROPIA. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13430293&pc=MTA | 2023826-2022-00285 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING A HIGH VAULT AND THE PATIENT COMPLAINED OF HALO IMAGES NEXT TO OBJECTS PERIPHERALLY. ALSO HAD DOUBLE IMAGES INFERIORLY AND DIAGONALLY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. OTHER #: EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770483&pc=MTA | 2023826-2023-04112 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LENS ROTATION, HALOS, STARBURST, DIPLOPIA, THE LENS WAS REPOSITIONED ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED AND IT WAS NOTED "VAULT IS HIGHER BUT OK, LENS IS CENTERED". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR WITH NO DEVICE CAUSALITY AND NOTED "DESCENTRATION CAUSED SHORT LENGTH". Manufacturer Narrative: CORRECTED INFORMATION: B5 - DECENTRATION SHOULD HAVE BEEN INCLUDED. H6 - HEALTH EFFECT CLINICAL CODE: 4581 - DECENTRATION CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: 4581 - STARBURST, DIPLOPIA, CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000599&pc=MTA | 2023826-2022-04621 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH_12.6; +9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 05-DEC-2019. THE PATIENT EXPERIENCED GLARE/HALOES AND ASTHENOPIA. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH ANOTHER IOL. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000618&pc=MTA | 2023826-2022-04622 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH_12.6; +9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 05-DEC-2019. THE PATIENT EXPERIENCED GLARE/HALOES AND ASTHENOPIA. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH ANOTHER IOL. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15670354&pc=MTA | 2023826-2022-03828 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE PATIENT INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THEIR EYE. IT WAS INDICATED THAT THEY WERE SEEKING A SOLUTION FOR HALOES AND STARBURST. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082712&pc=MTA | 2023826-2021-04925 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem; No Apparent Adverse Event | Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE PATIENT REPORTED HAD AN EVO VISIAN ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN HER RIGHT EYE (OD) IN (B)(6) 2021 TO CORRECT MYOPIA. SINCE THEN, THE PATIENT HAD EXPERIENCED MULTIPLE BRIGHT CONCENTRIC CIRCLES AROUND THE FOCUS POINT WHEN LIGHT IS SHINING. THE PATIENT IS NOT ABLE TO DRIVE OR PERFORM HER WORK ON THE PC. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PATIENT STILL REPORTS HALOS AND IT WAS PROBABLY DUE TO THE CENTRAL HOLE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579377&pc=MTA | 2023826-2016-00469 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; Discomfort; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT HAVING DISCOMFORT WITH GLARES AND HALOS. THE REPORTER STATED THE VAULT AND OUTCOME WAS GOOD BUT THE PATIENT REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579747&pc=MTA | 2023826-2016-00470 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; Discomfort; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PATIENT HAVING DISCOMFORT WITH GLARES AND HALOS. THE REPORTER STATED THE VAULT AND OUTCOME WAS GOOD BUT THE PATIENT REQUESTED THE LENS BE REMOVED. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW RESULT: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16690916&pc=MTA | 2023826-2023-01210 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO_12.1 IMPLANTABLE COLLAMER LENS -15.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAUL WITH ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTICS TORN/BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- CLINICAL CODE: "2227 SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3001" AND "2923" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15176916&pc=MTA | 2023826-2022-02618 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Halo; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S LEFT EYE (OS). THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND HALOS. THE PROBLEM WAS RESOLVED. REPORTEDLY, "DR. (B)(6) HAS REPORTED PATIENT 45271 OS WITH A VUALT OF 1200 MICRONS AND BOTHERSOME HALOS. HE HAS PERFORMED AN EXCHANGE FOR A SMALLER LENS WHICH HAS RESOLVED BOTH ISSUES - NEW VAULT 280 MICRONS AND DYSPHOTPSIA RESOLVED." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831218&pc=MTA | 2023826-2023-04210 | 2024-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Optical Problem | Halo; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS EYE ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO OVER CORRECTION. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6734732&pc=MTA | 2023826-2017-01234 |  | Injury | STAAR SURGICAL COMPANY | 2017-07-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO THE PATIENT WAS UNHAPPY WITH THE LENS DUE TO PERSISTENT GLARE AND HALOS. THE LENS WAS EXPLANTED WITH NO PATIENT INJURY, NO INCISION ENLARGEMENT OR SUTURES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAS RETURNED TO WEARING CONTACT LENSES AND IS DOING WELL. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6556157&pc=MTA | 2023826-2017-00720 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HALOS. THERE WAS NO REPORTED PATIENT INJURY OR SURGICAL INTERVENTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THEY IMPLANTED BILATERAL LENSES, WITH THE EXPECTATION OF THE LENSES BEING UNDERCORRECTED AND THEY PLANNED TO DO LASIK TO CORRECT THE VISION. IT WAS A 2 STEP PROCESS, BUT DUE TO A LANGUAGE BARRIER, THE PATIENT WAS UNCOOPERATIVE AND COMPLAINED OF HALOS. HE WANTED THE LENSES REMOVED. THE LENSES WERE EXPLANTED, WITH NO PATIENT INJURY AND THE PATIENT IS DOING WELL, VISION IS RECOVERING. THE LENSES WERE NOT EXCHANGED FOR OTHER LENSES. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6556158&pc=MTA | 2023826-2017-00719 | 2017-04-13 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Insufficient Information | Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO HALOS. THERE WAS NO REPORTED PATIENT INJURY OR SURGICAL INTERVENTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE REPORTER STATED THEY IMPLANTED BILATERAL LENSES, WITH THE EXPECTATION OF THE LENSES BEING UNDERCORRECTED AND THEY PLANNED TO DO LASIK TO CORRECT THE VISION. IT WAS A 2 STEP PROCESS, BUT DUE TO A LANGUAGE BARRIER, THE PATIENT WAS UNCOOPERATIVE AND COMPLAINED OF HALOS. HE WANTED THE LENSES REMOVED. THE LENSES WERE EXPLANTED, WITH NO PATIENT INJURY AND THE PATIENT IS DOING WELL, VISION IS RECOVERING. THE LENSES WERE NOT EXCHANGED FOR OTHER LENSES. THE REPORTER STATED THE EVENT WAS NOT PRODUCT RELATED, IT WAS A PATIENT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A CONTACT LENS CASE. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5304844&pc=MTA | 2023826-2015-01611 | 2012-02-08 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855106&pc=MTA | 2023826-2015-00696 | 2014-05-01 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO GLARE/HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. LENS EXPLANTED. (B)(4). Manufacturer Narrative: METHOD: - WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5440040&pc=MTA | 2023826-2016-00152 | 2014-07-02 | Injury | STAAR SURGICAL COMPANY AG | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -16.0/5.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO GLARES/HALOS AND THE PATIENT WAS NOT SATISFIED WITH THE BEST-CORRECTED VISUAL ACUITY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC TORN AND BENT. FOREIGN MATERIAL IN THE FORM OF FIBERS WERE PRESENT ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4702503&pc=MTA | 2023826-2015-00447 | 2014-09-09 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-17 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE PATIENT REPORTED GLARE AND HALOES. AN ICL EXCHANGE IS PLANNED FOR (B)(6) 2015. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT:UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED. Manufacturer Narrative: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW-REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4870113&pc=MTA | 2023826-2015-00712 | 2015-04-02 | Injury | STAAR SURGICAL COMPANY | 2015-06-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536611&pc=MTA | 2023826-2016-00357 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARES AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PROCESS EVALUATION: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536840&pc=MTA | 2023826-2016-00358 | 2015-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO GLARES AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. SECONDARY SURGERY. GLARE. EXPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # 708676
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043822&pc=MTA | 2023826-2016-01111 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/088 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT OPTIC WAS TORN/ SPLIT, HAPTIC WAS TORN/ BROKEN/ BENT/ DEFORMED AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6043825&pc=MTA | 2023826-2016-01112 | 2016-02-20 | Injury | STAAR SURGICAL COMPANY | 2016-10-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/010 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN, BUT THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS SPOT. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5909288&pc=MTA | 2023826-2016-01176 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/005 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767214&pc=MTA | 2023826-2015-00519 |  | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN HIS RIGHT EYE (OD). THE PATIENT REPORTED HE HAS BAD NIGHT VISION (GLARES/HALOS). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5009187&pc=MTA | 2023826-2015-01053 | 2014-12-30 | Malfunction | STAAR SURGICAL COMPANY AG | 2015-08-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Break; Torn Material | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) AND THE LENS TORE/BROKE. THE PATIENT EXPERIENCED GLARES/HALOS. THE LENS WAS REMOVED AND REPLACED WITH ANOTHER SAME MODEL LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST BCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -14.50. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5481997&pc=MTA | 2023826-2016-00234 |  | Injury | STAAR SURGICAL COMPANY AG | 2016-03-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, DIOPTER - UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD) MORE THAN FIVE YEARS AGO. THE PATIENT REPORTED INCREASE IN HALOS AND GLARE AT NIGHT. THE ICL WAS FOUND TO HAVE ROTATED AND THE LENS WAS RE-POSITIONED. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: THE PATIENT HAD MORE ASTIGMATISM OVER TIME AND THE SURGEON NOTICED MILD PIGMENT DISPERSION AND WHITE DEPOSITS ON ICL. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6404331&pc=MTA | 2023826-2017-00389 | 2016-10-28 | Injury | STAAR SURGICAL COMPANY | 2017-03-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/35 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARE/HALOES. ON (B)(6) 2016 THE LENS WAS EXPLANTED. AS THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: LENS WAS RETURNED IN A CENTRIFUGE VIAL. VISUAL INSPECTION FOUND CLEAR RESIDUE AND DEBRIS ON LENS SURFACE. LENS IS FOLDED AT THE HAPTICS AND OPTIC IS TWISTED DUE TO POSITION IN VIAL. NO VISIBLE TEARS OR SCRATCHES. UNABLE TO FURTHER EVALUATE DUE TO LENS SHAPE. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7248713&pc=MTA | 2023826-2018-00206 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM MICL12.6, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THAT HALOES AND INCREASE IN FLOATERS WERE NOTICED. ATTEMPT TO OBTAIN ADDITION INFORMATION WAS UNSUCCESSFUL. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THAT HALOS AND INCREASE IN FLOATERS WERE NOTICED. THE PATIENT ALSO USED "EYE DROPS TO CONSTRICT THE PUPIL" FOR IMPROVING VISION. (B)(4) Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7378433&pc=MTA | 2023826-2018-00469 |  | Injury | STAAR SURGICAL COMPANY | 2018-03-28 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) APPROXIMATELY 2 YEARS AGO. THE PATIENT COMPLAINED OF HALOS AND STATED SHE SEES THE AQUAPORT. SHE THEN "DISAPPEARED" FROM HIS CONSULTATION. SHE RETURNED STATING SHE CAN NO LONGER TOLERATE THE HALOS, ESPECIALLY AT NIGHT, AND ON (B)(6) 2018 THE LENS WAS EXPLANTED. REPORTEDLY, THE PROBLEM IS RESOLVED. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, ABLE TO SEE AQUAPORT). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6286637&pc=MTA | 2023826-2017-00151 | 2016-11-04 | Injury | STAAR SURGICAL COMPANY AG | 2017-01-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5 MM ICM125V4 IMPLANTABLE COLLAMER LENS, DIOPTER -12.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A BENT CONDITION IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOES. AS OF (B)(6) 2017, THE PATIENT IS WELL ON TOWARDS RECOVERY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MANUFACTURED OR MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5870494&pc=MTA | 2023826-2016-01060 | 2015-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-08-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED GLARE AND HALOS AND THE PLAN IS TO EXPLANT THE LENS. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2016. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY IS 20/20 AND UNCORRECTED VISUAL ACUITY IS 20/25. THE PATIENT PLANS TO EXCHANGE THE LENS AT A LATER DATE. DEVICE EVALUATED BY MANUFACTURER? LENS EXPLANTED BUT NOT RETURNED. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: THE REPORT SOURCE WAS INDICATED TO BE DISTRIBUTOR, NOT USER FACILITY. (B)(4). CLAIM #: (B)(4). LENS EXPLANTED BUT NOT RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923615&pc=MTA | 2023826-2016-01217 | 2015-06-16 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE AND HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. PERFORMED VISUAL INSPECTION AND FOUND NO VISIBLE DAMAGE ON LENS AND FOREIGN MATERIAL ON LENS SURFACE (SPOT). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577529&pc=MTA | 2023826-2016-00440 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.5/+2.50/X095 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00441 FOR LEFT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH FOREIGN MATERIAL ON LENS SURFACE (FIBERS) AND A SCRATCH. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.5/+2.50/X095. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577591&pc=MTA | 2023826-2016-00441 | 2016-01-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.00/+3.00/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED GLARES/HALOES. THE ICL WAS EXPLANTED ON (B)(6) 2016 AND WAS NOT REPLACED. NO OTHER LENS WAS IMPLANTED. LAST VISIT, BCVA WAS 20/20. SEE MFR # 2023826-2016-00440 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CHANGE "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." TO "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -12.00/+3.00/X91 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016." CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.00/+3.00/X091. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5926685&pc=MTA | 2023826-2016-00681 | 2016-04-02 | Injury | STAAR SURGICAL COMPANY | 2016-09-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND THAT THERE WERE SPOTS ON LENS SURFACE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.83MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5938281&pc=MTA | 2023826-2016-01247 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-09-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/005 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO PIGMENT DISPERSION, GLARE AND HALOS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. A VISUAL INSPECTION WAS PERFORMED, NO DAMAGE TO THE LENS. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6091629&pc=MTA | 2023826-2016-01605 | 2016-07-06 | Injury | STAAR SURGICAL COMPANY | 2016-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: REPORT SOURCE: USER FACILITY WAS INCORRECT; REPLACED BY DISTRIBUTOR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; SURGICAL RESIDUE WAS CLEAR; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). METHOD: WORK ORDER SEARCH PERFORMED AND NO SIMILAR COMPLAINTS FOUND. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4933380&pc=MTA | 2023826-2015-00904 | 2014-04-04 | Injury | STAAR SURGICAL COMPANY | 2015-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO GLARE/HALOS. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. PER MEDICAL REVIEW - GLARES AND HALOS MAY BE NOTED BY PATIENTS EVEN IF THEY NEVER HAD ANY OCULAR SURGERY. THIS IS COMMONLY DUE TO ABERRATIONS OR IRREGULARITIES IN THE CORNEA. IT MAY BE NOTED MORE AT LOW LIGHT CONDITIONS WHEN THE PUPIL IS DILATED. IN ICL SURGERY, THE CENTRAL AND MID PERIPHERAL CORNEA IS NOT TOUCHED, THEREFORE PATIENTS WHO HAVE GLARES AND HALOS BEFORE THE SURGERY, WILL COMMONLY RETAIN THESE SYMPTOMS BUT LITTLE OR NO INCREASE IN SEVERITY SHOULD BE EXPERIENCED SINCE THE CENTRAL AND MID PERIPHERAL REGION REMAINS UNABLATED AND ONLY A CORNEAL INCISION WITH A MAXIMUM LENGTH OF 3.2MM AT THE TEMPORAL PERIPHERAL REGION IS MADE. GLARE AND HALOS MAY ALSO BE PERCEIVED MORE DUE TO THE INCREASED CLARITY OF VISION AFTER ICL SURGERY OR IF THE PATIENT HAS RETAINED ASTIGMATISM. ANOTHER FACTOR COULD ALSO BE THE EFFECTIVE OPTICAL CORNEAL ZONES (EOCZ) OF THE ICLS, WHICH HAVE A MAXIMUM DIAMETER OF 6.17 TO 7.30 MM DEPENDING ON THE ICL POWERS. THIS IS A DESIGN LIMITATION WHICH MAY CREATE SIMILAR SYMPTOMS OF GLARE AND HALOS DUE TO THE INTERFACE OF THE OPTICAL ZONE AND THE PATIENT'S OPTICAL FIELD OF VISION, WHICH MAY BE NOTED WHEN THE PUPIL SIZE IS GREATER THAN THE EOCZ. THE 0.36MM HOLE AT THE CENTER OF THE ICL HAS BEEN COMPARED TO THE ICL WITHOUT A HOLE WHICH SHOWED THAT MODULATION-TRANSFER FUNCTIONS (MTF) ARE SMALL AND CLINICALLY NEGLIGIBLE. GLARES AND HALOS ARE COMMONLY TEMPORARY, LASTING 1-6 MONTHS, BUT MAY ON VERY RARE OCCASIONS BECOME PERMANENT. THIS CONDITION IS CONSIDERED TRIVIAL AND TEMPORARY AND SHOULD NOT BE MDR REPORTABLE UNLESS MEDICAL OR SURGICAL INTERVENTION WAS DONE TO ELIMINATE SEVERE SYMPTOMS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6085625&pc=MTA | 2023826-2016-00882 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-11-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/073 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING OF THE LENS, GLARE/HALOES, AND A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN/ BENT. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701679&pc=MTA | 2023826-2017-01154 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICM5_13.7, DIOPTER -5.5/+1.0/089 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, HALOES, AND WIDE PUPIL. THE SURGEON COMMENTED THAT THE PATIENT'S PUPIL > 8 MM ON (B)(6) 2017, AND PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (DEBRIS). (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP INDICATED NORMAL PUPIL SIZE, BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6701680&pc=MTA | 2023826-2017-01155 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-07-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7, DIOPTER -5.5 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, HALOES, AND WIDE PUPIL. THE SURGEON COMMENTED THAT THE PATIENT'S PUPIL > 8 MM ON (B)(6) 2017, AND PLANS TO EXCHANGE THE LENS FOR THE PATIENT. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP INDICATED NORMAL PUPIL SIZE, BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) TO BE 20/20. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND CLEAR RESIDUE) AND THE LENS HAS A CURVED SHAPE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836431&pc=MTA | 2023826-2017-01376 | 2017-07-24 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/140 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS. THE PATIENT IS OK. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN MATERIAL ON LENS SURFACE (DEBRIS AND RESIDUE ON LENS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4613548&pc=MTA | 2023826-2015-00328 | 2014-02-16 | Injury | STAAR SURGICAL COMPANY | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED GLARES/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - (NEW INCISION); HALO; (GLARE); SIZE INCORRECT FOR PATIENT; (VAULTING). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS : A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4724366&pc=MTA | 2023826-2015-00456 | 2015-02-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-23 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2007. LOW VAULT WAS NOTED, AS WELL AS REFRACTIVE CHANGE OVER TIME (INCREASED MYOPIA), AND GLARE/HALOS. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: N/A. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO - 81(OTHER): LENS REMAINS IMPLANTED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788339&pc=MTA | 2023826-2016-00883 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GLARE/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/13. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5676444&pc=MTA | 2023826-2016-00658 | 2016-04-11 | Injury | STAAR SURGICAL COMPANY | 2016-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS REPORTED AS HAVING AN EXCESSIVE VAULT AND IS PLANNED TO EXCHANGE THE LENS. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREATIVE PUPIL (FIXED), AND GLARE/HALOES. THE PATIENT ALSO HAD YAG AND IS ON PILOCARPINE TREATMENT. THE PATIENT'S BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/40. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5371940&pc=MTA | 2023826-2016-00065 |  | Injury | STAAR SURGICAL COMPANY | 2016-01-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ANOTHER SIMILAR COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5462012&pc=MTA | 2023826-2016-00199 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND LENS DISLOCATION/SUBLUXATION. THE PATIENT EXPERIENCED GLARE/HALOS. THE LENS WAS EXPLANTED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9751864&pc=MTA | 2023826-2020-00384 | 2020-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VICM5_13.7, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6)2020 THE LENS WAS EXPLANTED DUE TO GLARE/HALOES AND THE PROBLEM RESOLVED. REPORTER STATED "PATIENT HAD MASSIVE HALOS. CANCELLED SURGERY FROM 2ND EYE. COMES TO RLE WITHIN NEXT YEARS." THE REPORTER STATES THAT THE CAUSE OF THIS EVENT WAS DUE TO LIGHT SENSATION FROM AQUAPORT. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794397&pc=MTA | 2023826-2020-00433 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO A PATIENT RELATED FACTOR, THE REPORTER STATED " PT. HAS A HUGE PUPIL (8.5MM) AND COULD NOT LIVE WITH THE HALOS. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2227 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794396&pc=MTA | 2023826-2020-00432 | 2019-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO A PATIENT RELATED FACTOR, THE REPORTER STATED " PT. HAS A HUGE PUPIL (8.5MM) AND COULD NOT LIVE WITH THE HALOS. LENS REMOVAL RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2227 SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5011063&pc=MTA | 2023826-2015-01052 | 2014-12-09 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS ROTATED AND WAS REPOSITIONED. THE PATIENT EXPERIENCED GLARE AND HALOS. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND THIS RESOLVE THE PROBLEM. PATIENT LAST UCVA WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -15.0/3.0/118. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8997601&pc=MTA | 2023826-2019-01704 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Halo; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING LOW VAULTING. THE LENS DECENTERED, WITH LENS ROTATION OF 160 DEGREES. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON DOES NOT PLAN TO EXPLANT/EXCHANGE THE LENS AT THIS TIME, THE HALOS ARE IMPROVING AND THE SURGEON WANTS TO WAIT. DATE OF REPORT: 20-MAR-2019. CLAIM #: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9886971&pc=MTA | 2023826-2020-00697 | 2020-03-05 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Halo; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VICM_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS AND IT IS REPORTED THAT "PATIENT STILL HAVE HALOES." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. REFERENCE MFR# 2023826-2019-02021 FOR INITIAL LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12168068&pc=MTA | 2023826-2021-02624 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED HEADACHES. IT WAS NOTED THERE IS NO PLAN TO REPLACE THE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION : LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170268&pc=MTA | 2023826-2021-02625 | 2021-03-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6; -12.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCED HEADACHES. IT WAS NOTED THERE IS NO PLAN TO REPLACE THE LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6997961&pc=MTA | 2023826-2017-01639 |  | Injury | STAAR SURGICAL COMPANY | 2017-11-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Headache | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN HER RIGHT EYE (OD). THE PATIENT REPORTED HAS EXPERIENCED "MIGRAINE AURA". THE LENS WAS IMPLANTED ON (B)(6) 2013 AND REMAINS IMPLANTED. THE PATIENT REPORTED THE EVENT STARTED IN (B)(6) 2017 AND OCCURS EVERY COUPLE OF WEEKS. THE DOCTOR'S OFFICE CONFIRMED THE EVENT AND INDICATED THE CAUSE WAS UNKNOWN. PATIENT HAD LASIK ON (B)(6) 2013. THE PATIENT HAS NOT BEEN PRESCRIBED ANY MEDICATION AND THE DATE OF LAST VISIT WAS (B)(6) 2017. THE PATIENT'S VISUAL ACUITY POST-OP WAS 20/25. Manufacturer Narrative: (B)(4) Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10950703&pc=MTA | 2023826-2020-02777 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Headache | UNK | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT "A PRESSURE ISSUE THE DAY/NIGHT OF SURGERY AND INTO THE NEXT DAY" WAS OBSERVED. ACUTE ANGULAR CLOSURE GLAUCOMA WAS DIAGNOSED. REPORTEDLY, "PRESSURE THING WAS ONE TIME." VISION LOSE, THROWING UP, TERRIBLE HEADACHE (THROBBING/PULSING FEELING) WAS OBSERVED. DROPS TO PREVENT PRESSURE RISING AND MORE ORAL MEDICATION WAS PRESCRIBED. "THEY DETERMINED THAT MY PIS WERE PLUGGED IN ONE EYE, AND THEY HAD TO ADD AN ADDITIONAL IRIDOTOMY. RELIEF STARTED TO HAPPEN SHORTLY AFTER. SO NO ONGOING PROBLEM," LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11019222&pc=MTA | 2023826-2020-02778 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Headache | UNK | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT "A PRESSURE ISSUE THE DAY/NIGHT OF SURGERY AND INTO THE NEXT DAY" WAS OBSERVED. ACUTE ANGULAR CLOSURE GLAUCOMA WAS DIAGNOSED. REPORTEDLY, "PRESSURE THING WAS ONE TIME." VISION LOSE, THROWING UP, TERRIBLE HEADACHE (THROBBING/PULSING FEELING) WAS OBSERVED. DROPS TO PREVENT PRESSURE RISING AND MORE ORAL MEDICATION WAS PRESCRIBED. "THEY DETERMINED THAT MY PIS WERE PLUGGED IN ONE EYE, AND THEY HAD TO ADD AN ADDITIONAL IRIDOTOMY. RELIEF STARTED TO HAPPEN SHORTLY AFTER. SO NO ONGOING PROBLEM," LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957436&pc=MTA | 2023826-2015-00990 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Anxiety; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2014 AND ON (B)(6) 2015 THE PATIENT WAS EXPERIENCING HEADACHES, ANXIETY AND INSOMIA. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS NOT REPLACED. SEE MFR REPORT# 2023826-2015-00991 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4957478&pc=MTA | 2023826-2015-00991 | 2015-04-13 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-31 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Anxiety; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2014 AND ON (B)(6) 2015 THE PATIENT WAS EXPERIENCING HEADACHES, ANXIETY AND INSOMNIA. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND WAS NOT REPLACED. SEE MFR REPORT# 2023826-2015-00990 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT INFORMATION AND WORK ORDER SEARCH, WE ARE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15142576&pc=MTA | 2023826-2022-02548 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Headache; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: A PATIENT REPORTED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THEIR EYE. THE PATIENT INDICATES THEY HAVE LOSS OF NIGHT VISION; SEVERE HEADACHES AND VERTIGO. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6560720&pc=MTA | 2023826-2017-00791 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED THERE WAS OFF-TARGET REFRACTION (REFRACTIVE SURPRISE) AND LOW VAULTING WITH THE LENS. THE PATIENT WAS EXPERIENCING HEADACHES. THE LENS REMAINS IMPLANTED BUT WILL BE EXPLANTED AT A LATER DATE. Manufacturer Narrative: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11510275&pc=MTA | 2023826-2021-00681 | 2021-01-11 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Headache; Blurred Vision | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VICM5_12.1, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER LENS DUE TO REFRACTIVE SURPRISE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED FOR CAUSE OF EVENT, THE REPORTER STATED "PATIENT EXPERIENCED HEADACHE." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440020&pc=MTA | 2023826-2022-01618 | 2022-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. PIGMENT DISPERSION, FIXED PUPIL AND IRIS ATROPHY WAS OBSERVED. ON (B)(6) 2022 PI YAG WAS PERFORMED. LENS REMAINS IMPLANTED. FOR CAUSE DETAILS THE REPORTER STATED, "GOOD SURGICAL OUTCOME, HAPPY VISION, BUT HAS HAD EPISODES OF HEADACHE AND BLURRED VISION ABOUT 2-3 WEEKS POST-OP. IOP ON EXAM WAS NORMAL, VAULT NORMAL." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "PERSISTENT, MILD DILATED, FIXED PUPIL WITH DIFFUSE TID OD ONLY. ACUITY GOOD, BUT FLUCTUATING." FOR ADDITIONAL INFORMATION THE REPORTER STATED "THIS IS THE THIRD CASE OF FIXED PUPIL WITH DIFFUSE TIDS. ALL WITHIN LAST MONTH, ALL RIGHT EYES, NO COMMON THEMES. HAPPY WITH VISION, SO NOT LIKELY TO REMOVE LENS, BUT CONCERNED ABOUT LONG TERM ISSUES (IOP,GLARE,NIGHT VISION). AFTER DISCUSSION WITH COLLEAGUES, FEEL THIS MAY BE UNUSUAL CASE OF MOXIFLOXACIN TOXICITY." Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 - SHOULD BE CORRECTED TO IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4600863&pc=MTA | 2023826-2015-00311 | 2015-02-10 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, HEADACHE AND BLURRED VISION. THE LENS REMAINS IMPLANTED. THE PATIENT WAS TREATED FOR CHOROIDAL EFFUSION (SUSPECT) WITH TOPIC CORTICOSTEROIDS, THE ANGLES WERE COMING BACK TO PREVIOUS AND THERE WAS GOOD EVOLUTION OF THE PATIENT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. EXPLANT DATE: NA. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587683&pc=MTA | 2023826-2017-00874 | 2016-09-22 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -4.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND SEVERE HEADACHE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: LENS HAS NOT BEEN RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN A SMALL VIAL, WITH CLEAR DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. METHOD CODE: ADDITIONAL CODE ADDED. RESULT CODE: UPDATED BASED ON ADDITIONAL NFORMATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052958&pc=MTA | 2023826-2021-02041 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES OR VA LOSS. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052969&pc=MTA | 2023826-2021-02042 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Blurred Vision; Visual Disturbances; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO TRANSIENT LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA), GLARE/HALOS, HEADACHE AND BLURRED VISION. THE PIS WERE ENLARGED BY YAG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE ICL IS IN A GOOD POSITION AND NO MORE HEADACHES. Manufacturer Narrative: ENLARGE PIS. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441206&pc=MTA | 2023826-2021-00619 | 2021-01-02 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -02.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT HAD A HEADACHE AND DISCOMFORT. THE SYMPTOM WAS RESOLVED AFTER THE LENS WAS REMOVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THERE WAS NO ABNORMAL ISSUE IN POST-OP, THE PATIENT FELT DISCOMFORT BECAUSE OF HER SUBJECTIVE SYMPTOMS AND PERSONALITY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11441242&pc=MTA | 2023826-2021-00618 | 2021-01-02 | Injury | STAAR SURGICAL COMPANY | 2021-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -02.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT HAD A HEADACHE AND DISCOMFORT. THE SYMPTOM WAS RESOLVED AFTER THE LENS WAS REMOVED. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THERE WAS NO ABNORMAL ISSUE IN POST-OP, THE PATIENT FELT DISCOMFORT BECAUSE OF HER SUBJECTIVE SYMPTOMS AND PERSONALITY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073174&pc=MTA | 2023826-2021-02030 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT IN SEEING THINGS, SEVERE HEADACHE, AFFECTING NORMAL LIFE AND WORK. ON (B)(6) 2021 THE LENS WAS EXPLANTED PER PATIENTS REQUEST AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED " THE PATIENT WAS (B)(6) YEARS OLD WITH PRESBYOPIA. IN ORDER TO DELAY THE TIME OF WEARING PRESBYOPIA AFTER OPERATION, IT IS SUGGESTED TO KEEP A CERTAIN DEGREE AFTER OPERATION. THE PRESBYOPIA PROBLEM WAS COMMUNICATED TO THE PATIENT MANY TIMES BEFORE OPERATION, BUT THE PATIENT WAS UNWILLING TO KEEP THE PRESBYOPIA DEGREE AND INSISTED ON COMPREHENSIVE CORRECTION." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073149&pc=MTA | 2023826-2021-02029 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED THAT THE PATIENT COMPLAINED OF DISCOMFORT IN SEEING THINGS, SEVERE HEADACHE, AFFECTING NORMAL LIFE AND WORK. ON (B)(6) 2021 THE LENS WAS EXPLANTED PER PATIENTS REQUEST AND THE PROBLEM RESOLVED. IN THE REPORTER'S OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR ADDITIONAL INFORMATION THE REPORTER STATED "THE PATIENT WAS (B)(6) WITH PRESBYOPIA. IN ORDER TO DELAY THE TIME OF WEARING PRESBYOPIA AFTER OPERATION, IT IS SUGGESTED TO KEEP A CERTAIN DEGREE AFTER OPERATION. THE PRESBYOPIA PROBLEM WAS COMMUNICATED TO THE PATIENT MANY TIMES BEFORE OPERATION, BUT THE PATIENT WAS UNWILLING TO KEEP THE PRESBYOPIA DEGREE AND INSISTED ON COMPREHENSIVE CORRECTION." Manufacturer Narrative: B5: LENS WAS IMPALNTED ON (B)(6) 2021. CLAIM (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442519&pc=MTA | 2023826-2022-03148 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO DISCOMFORT AND HEADACHE. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442553&pc=MTA | 2023826-2022-03149 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO DISCOMFORT AND HEADACHE. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477607&pc=MTA | 2023826-2023-03351 | 2023-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Discomfort | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS WAS EXPLANTED DUE TO PATIENT DISSATISFACTION- HEADACHE AND DISCOMFORT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766106&pc=MTA | 2023826-2023-01417 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Headache; Discomfort; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "THE PATIENT UNDERWENT BILATERAL ICL LENS IMPLANTATION IN OUR HOSPITAL ON (B)(6) 202 AND (B)(6) 2022 RESPECTIVELY. AFTER THE OPERATION, HE COMPLAINED OF A HEADACHE AND DISCOMFORT ." THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "1880" AND "2330" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. PATIENT RELATED FACTOR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766108&pc=MTA | 2023826-2023-01418 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | Headache; Discomfort; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "THE PATIENT UNDERWENT BILATERAL ICL LENS IMPLANTATION IN OUR HOSPITAL ON (B)(6) 2023. AFTER THE OPERATION, HE COMPLAINED OF A HEADACHE AND DISCOMFORT ." THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: H6-HEALTH EFFECT- IMPACT CODE: "1880" AND "2330" SHOULD BE ADDED. D.6A SHOULD BE CORRECTED TO 02/18/2023. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0 INTO THE PATIENT'S LEFT EYE (OS) ON 18/FEB/2023. THE PATIENT EXPERIENCED HEADACHES AND DISCOMFORT. ON 04/MARCH/2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839502&pc=MTA | 2023826-2015-00631 | 2014-01-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO PATIENT DISSATISFACTION AND HEADACHE/DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, MEDICAL REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - OCULAR DISCOMFORT IS A VERY SUBJECTIVE AND BROAD SYMPTOM THAT COULD BE MAINLY RELATED TO INFLAMMATION OR IRRITATION OF SENSITIVE STRUCTURES. THE ICL IS A PLATE-HAPTIC LENS THAT SITS ON THE POSTERIOR SULCUS-CILIARY BODY AREA. ALTHOUGH HIGHLY UNLIKELY IN THE PRESENCE OF OPTIMALLY VAULTING ICL, THE ICL FOOTPLATES MIGHT COMPRESS THE CILIARY BODY STRUCTURE LEADING TO OCULAR DISCOMFORT. UNUSUALLY LARGE CILIARY PROCESSES (ANATOMICAL VARIANCE) COULD ALSO BE THE SOURCE OF THE SYMPTOM, AS WELL AS THE PRESENCE/RUPTURE OF AN UNDETECTED CYST. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839503&pc=MTA | 2023826-2015-00632 | 2014-01-05 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Discomfort; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO PATIENT DISSATISFACTION AND HEADACHE/DISCOMFORT. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - OCULAR DISCOMFORT IS A VERY SUBJECTIVE AND BROAD SYMPTOM THAT COULD BE MAINLY RELATED TO INFLAMMATION OR IRRITATION OF SENSITIVE STRUCTURES. THE ICL IS A PLATE-HAPTIC LENS THAT SITS ON THE POSTERIOR SULCUS-CILIARY BODY AREA. ALTHOUGH HIGHLY UNLIKELY IN THE PRESENCE OF OPTIMALLY VAULTING ICL, THE ICL FOOTPLATES MIGHT COMPRESS THE CILIARY BODY STRUCTURE LEADING TO OCULAR DISCOMFORT. UNUSUALLY LARGE CILIARY PROCESSES (ANATOMICAL VARIANCE) COULD ALSO BE THE SOURCE OF THE SYMPTOM, AS WELL AS THE PRESENCE/RUPTURE OF AN UNDETECTED CYST. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030225&pc=MTA | 2023826-2022-01092 | 2022-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Optical Problem | Headache; Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP), PIGMENT DISPERSION, GLARE/HALOES, BLURRED VISION, AC INFLAMMATION CELLS AND IRREGULAR PUPIL. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022. AT 3 DAY POST-OP, FIBRIN CONTRACTING AS PATIENT DEVELOPED SEVERE INFLAMMATION POST EXPLANT. EARLY CATARACT FORMATION NOTED. THE PATIENT IS BEING MONITORED. CLAIM # (B)(4). Manufacturer Narrative: B5- PATIENT ALSO EXPERIENCED HEADACHES; NAUSEA; WEAKNESS; MEDICATION REQUIRED. SHOULD HAVE BEEN INCLUDED IN INITIAL MED WATCH REPORT. H6- HEALTH IMPACT: CLINICAL CODE: 4581- NAUSEA; WEAKNESS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5199928&pc=MTA | 2023826-2015-01357 | 2015-09-17 | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED INTRAOCULAR PRESSURE (UP TO 25MMHG) STARTED BY THE 4TH DAY. THE PATIENT WAS GIVEN PRESSURE DROPS, THE PRESSURE WENT DOWN TO 22 BUT THEN IT WENT UP TO BETWEEN 32MMHG AND 50MMHG. THE PATIENT'S LAST UCVA WAS 20/20 AND IS SATISFIED WITH THE ICL BUT WAS COMPLAINING OF HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. RECENT IMAGES SHOWS GOOD VAULT AND NORMAL IOP BUT THE PATIENT IS STILL ON MEDICATION. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+2.0/X093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. LOW VAULT WAS OBSERVED AND ELEVATED INTRAOCULAR PRESSURE (IOP) STARTED BY THE 4TH DAY. THE SURGEON ESTABLISHED MEDICAL MANAGEMENT AND IOP WAS CONTROLLED TO NORMAL LEVELS WITH ONE GLAUCOMA MEDICATION. PATIENT ALSO REPORTED HEADACHES. THE PHYSICIAN STATED THE PATIENT HAS A HEAVILY PIGMENTED IRIS. ON LAST VISIT, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, VAULT WAS NORMAL, AND IOP WAS NORMAL. THE LENS REMAINS IMPLANTED. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT IS MARKETED IN THE U.S. DIOPTER: -17.0/+2.0/X093. DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15964027&pc=MTA | 2023826-2022-04484 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965745&pc=MTA | 2023826-2022-04483 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Blurred Vision; Dizziness | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -8.5 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED. REPORTEDLY, "AFTER THE PATIENT'S SURGERY, VISION, REFRACTIVE STATE, INTRAOCULAR PRESSURE, AND ARCH HIGH ARE NORMAL. THE PATIENT CONSCIOUSLY DIZZY AND HEADACHE." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1880 ADDED FOR HEADACHE. 2194 ADDED FOR DIZZINESS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7404820&pc=MTA | 2023826-2018-00508 | 2018-03-06 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE SURGEON INDICATED THAT A 13.2MM MICL13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. PATIENT REPORTED TO HAVE BLURRY VISION AND HEADACHE. ON (B)(6) 2018 THE SURGEON OBSERVED EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ACD (ANTERIOR CHAMBER DEPTH). PATIENT WAS PUT ON GLAUCOMA MEDICATION AND AN ANTERIOR CHAMBER WASH OUT WAS PERFORMED. THE LENS WAS EXPLANTED ON (B)(6) 2018. GLAUCOMA DROPS WERE GIVEN AND IOP (INTRAOCULAR PRESSURE) RETURNED TO NORMAL. SURGEON REPORTED THAT THE CAUSE OF EVENT IS UNKNOWN- COMBINATION OF TRAUMA (PATIENT ACCIDENTLY HIT HER EYE WITH HAND ON (B)(6) 2018), LENS LENGTH, SMALL 2ND LPI (LASER PERIPHERAL IRIDOTOMY). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE: SECONDARY SURGICAL INTERVENTION, EXPLANT. NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ACD. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444214&pc=MTA | 2023826-2017-00528 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT AND HEADACHE. THE PROBLEM IS NOT RESOLVED AND A LENS EXCHANGE IS PLANNED. REFERENCE 2023826-2017-00527 FOR ASSOCIATED EYE. Manufacturer Narrative: PREVIOUSLY STATED: THE LENS EXCHANGE WAS PLANNED. CURRENT: LENS WAS EXCHANGED. BOTH PATIENT AND PHYSICIAN ARE HAPPY WITH THE RESULTS. UNABLE TO OBTAIN ANY FURTHER DETAILS. "THE LENS REMAINS IMPLANTED." (B)(4). Manufacturer Narrative: PREVIOUS STATED: SYMPTOMS OF EXCESSIVE VAULT AND HEADACHES; NO DETAILS FOR THE LENS EXCHANGE. ADDITIONAL/CORRECTED INFORMATION: PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRED VISION. LENS WAS REPOSITIONED ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. PROBLEM IS RESOLVED. TO DATE, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. IN FOLLOW-UP #2, IT WAS STATED THAT THE LENS WAS REPOSITIONED ON (B)(6) 2017 AND WAS EXCHANGED ON (B)(6) 2017. IT SHOULD BE THAT THE LENS WAS EXCHANGED ON (B)(6) 2017, AND THE REPLACEMENT LENS WAS REPOSITIONED ON (B)(6) 2017. ADDITIONALLY, IT WAS STATED THAT THE PROBLEM WAS RESOLVED. IT SHOULD BE THAT THE PROBLEM WAS NOT RESOLVED. (B)(4). Manufacturer Narrative: DEVICE SERIAL NUMBER AND EXPIRATION DATE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6444673&pc=MTA | 2023826-2017-00527 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-03-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.00/+6.0/085 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED (B)(6) 2016 AND THE SURGEON NOTICES A GRADUAL INCREASE IN IOP. THE PROBLEM IS NOT RESOLVED AND A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+6.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS ROTATION, ELEVATED IOP, POOR VISION, AND HEADACHES. REPOSITIONING WAS PERFORMED ON (B)(6) 2016 AND THE SURGEON NOTICED A GRADUAL INCREASE IN IOP. THE LENS WAS EXPLANTED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781720&pc=MTA | 2023826-2020-00391 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Headache; Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 LENS, -10.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, BLURRED VISION, AND HEADACHES. THE PROBLEM IS RESOLVED. REPORTEDLY THE PATIENT IS "SEEING WELL, NO HEADACHES, PATIENT REPORTS HAVING SLIGHT TWITCH IN THIS EYE NO PAIN." THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6324404&pc=MTA | 2023826-2017-00233 | 2016-12-28 | Injury | STAAR SURGICAL COMPANY | 2017-02-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pain; Pupillary Block; Blurred Vision; Vomiting; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXTREME VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED: THE REPORTER INDICATED THE PATIENT CALLED POST-SURGERY WITH EXTREME PAIN, HEADACHE, NAUSEA AND VOMITING. THE PATIENTS PRESSURE WAS 50. THE WOUND WAS BURPED BUT WOULD NOT GO DOWN BELOW 37. A SECONDARY YAG PI WAS PERFORMED BUT WITH NO PRESSURE RELIEF. THERE WAS NARROWING OF THE ANGLE AND MECHANICAL ANGLE CLOSURE. THE LENS WAS EXPLANTED ON (B)(6) 2016 AND A SHORTER LENS WAS IMPLANTED ON (B)(6) 2016. THE PATIENTS CURRENT VA IS 20/20 AND IS DOING WELL. THE REPORTER STATED THE EVENT WAS PRODUCT-RELATED. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID AND THERE WAS CLEAR RESIDUE ON THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065520&pc=MTA | 2023826-2017-01781 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK) AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH (ACD). A REVISION OF THE PERIPHERAL IRIDECTOMIES WAS PERFORMED. THE PATIENT EXPERIENCED BLURRY VISUAL ACUITY, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 21. Manufacturer Narrative: (B)(4) . DATE OF EVENT: IT IS CORRECTED FROM UNK TO (B)(6) 2017. (B)(4) Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7065648&pc=MTA | 2023826-2017-01782 | 2017-10-13 | Injury | STAAR SURGICAL COMPANY | 2017-11-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; Nausea; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (PUPILLARY BLOCK). THE PATIENT EXPERIENCED BLURRY VISION, HEADACHE AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25 AND IOP WAS 18. Manufacturer Narrative: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND PIECES OF LENS HAPTIC TORN OFF AND MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6853814&pc=MTA | 2023826-2017-01415 |  | Injury | STAAR SURGICAL COMPANY | 2017-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED ADDITIONAL SYMPTOMS - HEADACHE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S CURRENT CONDITION POST-OP - BSCVA - 20/30, INTRAOCULAR PRESSURE - 18MM HG. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PIGMENT DISPERSION/ANATOMICAL ANOMALY (ABNORMAL SULCUS ANATOMY). THE PATIENT HAS PLANNED THERAPY OF LASER VISION CORRECTION (LVC). Manufacturer Narrative: LENS WAS RETURNED IN LENS CASE. VISUAL INSPECTION FOUND HAPTIC DEFORMED. CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5728010&pc=MTA | 2023826-2016-00823 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2016-06-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED HEADACHES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO ABNORMAL POSTERIOR IRIS ANATOMY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5728011&pc=MTA | 2023826-2016-00822 | 2015-10-02 | Injury | STAAR SURGICAL COMPANY | 2016-06-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED HEADACHES. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO ABNORMAL POSTERIOR IRIS ANATOMY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS DISCARDED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6074744&pc=MTA | 2023826-2016-01574 |  | Injury | STAAR SURGICAL COMPANY | 2016-11-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THROUGH SOCIAL MEDIA HE HAD AN MICL IMPLANTABLE COLLAMER LENS IMPLANTED ON (B)(6) 2016. APPROXIMATELY ONE HOUR AFTER THE SURGERY THE EYE PRESSURE HAD INCREASED AND THE PATIENT COMPLAINED OF A HEADACHE. THE SURGEON WAS ABLE TO REDUCE THE PRESSURE. THE PATIENT HAS COMPLAINED OF GLARE BUT IT HAS REDUCED OVER TIME. THE PATIENT'S VISION IS 20/15. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6765864&pc=MTA | 2023826-2017-01265 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SMALLER SIZE LENS ON (B)(6) 2017 DUE TO HIGH VAULT, ELEVATED IOP, AND NARROWING OF THE ANGLE. HEADACHE WAS ALSO REPORTED AS AN ADDITIONAL SYMPTOM. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENTS POST-OP BCVA WAS 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED, IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN A LIQUID PRESENT IN A VIAL. VISUAL INSPECTION FOUND A PIECE OF HAPTIC MISSING FROM THE LENS. (B)(4). Manufacturer Narrative: THE PATIENT'S POST-OP BCVA WAS 20/20" IS INCORRECT. THE PATIENT'S POST-OP REFRACTION BCVA WAS FOUND TO BE 20/20 ON (B)(6) 2017 BEFORE THE EXCHANGE. CONCOMITANT MEDICAL PRODUCTS: INJECTOR, CARTRIDGE AND INJECTOR SYSTEM MODEL ARE KNOWN. (B)(4). Manufacturer Narrative: (B)(4). A WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320744&pc=MTA | 2023826-2018-00327 | 2018-02-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.00, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE SURGEON NOTES THE FOLLOWING: EXCESSIVE VAULT, ELEVATED IOP, IRIS PIGMENTS ON ICL, "MASSIVE" HEADACHE, APPOSITIONAL ANGLE CLOSURE. ADDITIONAL INFORMATION PROVIDED: 1 WEEK POST-OP THE IOP BECAME ELEVATED, PATIENT WAS GIVE IOP DROPS (COSOPT+ALPHAGAN). ON (B)(6) 2018 IT WAS REPORTED THAT "THE CENTRAL HOLE IS OPEN." THERE IS "THE PRESENCE OF APPOSITIONAL ANGLE CLOSURE...ACETAZOLAMIDE WAS ADDED WHEN THE PRESSURE REACHED HIGH THIRTIES ON FULL TOPICAL MEDICATIONS. STEROIDS WERE STOPPED TO RULE OUT THE POSSIBILITY OF STEROID RESPONSE. NOW THE IOP IS 19MMHG ON COSOPT ONLY WITH A HIGH VAULT...AND LOTS OF IRIS PIGMENTS ON THE ICL..." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: TWO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668555&pc=MTA | 2023826-2016-00732 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT AND PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT COMPLAINED OF A HEADACHE. RESULTS: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF BOTH HAPTICS TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL. CONCLUSIONS: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453292&pc=MTA | 2023826-2017-00534 | 2017-01-30 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Manufacturing, Packaging or Shipping Problem | Headache; Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED HEADACHE, ELEVATED IOP. THE CUSTOMER STATES THERE WAS NO CENTER HOLE ON THE LENS. ON (B)(6) 2017 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN SMALL VIAL, THERE WAS MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, NO HOLE ON THE LENS, CLEAR SURGICAL RESIDUE ON LENS SURFACE, AND THERE WAS INK STAINING ON THE SURFACE OF THE LENS. THE LENS WAS REMEASURED, AND PASSED THE LENS VERIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6918173&pc=MTA | 2023826-2017-01503 | 2017-07-18 | Injury | STAAR SURGICAL COMPANY | 2017-10-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Shelf Life Exceeded; Inadequacy of Device Shape and/or Size | Headache; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO INTERMITTENT ELEVATED IOP, HEADACHES, EXCESSIVE VAULTING, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED WITH A NEW LENS. Manufacturer Narrative: PATIENT CODE: (B)(4) NARROWING OF THE ANGLE, SHALLOWING OF THE ACD, SECONDARY SURGERY, LENS EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6606309&pc=MTA | 2023826-2017-00914 | 2012-12-10 | Injury | STAAR SURGICAL COMPANY | 2017-06-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Headache; Intraocular Pressure Increased; Pain; Blurred Vision; Visual Disturbances; Vitreous Detachment | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 11.5MM TICM115V4, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2012. BEGINNING ON (B)(6) 2012 THE PATIENT BEGAN TO EXPERIENCE BLURRED VISION. PATIENT ALSO EXPERIENCES FLASHES DUE TO POSTERIOR VITREOUS DETACHMENT (PVD). ADDITIONAL SYMPTOMS INCLUDE PAIN, OBSERVANCE OF MULTIPLE IMAGES, SEVERE HEADACHE, INABILITY TO FOCUS ON OBJECTS OR WORDS, POOR, VISION, ELEVATED IOP DUE TO STEROIED RSPONSE, AND LENS ROTATION. ON (B)(6) 2013 THE PAITENT WAS PUT ON ANTIGLAUCOMA THERAPY AS HE IS A "STEROID RESPONDER." THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. REFERENCE MDR 2023826-2017-00903 FOR FELLOW EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). DEVICE REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138562&pc=MTA | 2023826-2020-01184 | 2019-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (29MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138567&pc=MTA | 2023826-2020-01185 | 2019-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL | Patient-Device Incompatibility; Optical Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -04.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (27MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020, GLARE/HALOES WERE OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR (ACCOMMODATION SPASM). REPORTEDLY, "GLARE/HALOES, HEAD PAIN DIMINISHED, PATIENT STILL CANNOT DRIVE BY NIGHT. PATIENT HAS HAD SOME STOMATOLOGY CAL PROBLEMS. AFTER RESOLVING THOSE, THE SYMPTOMATOLOGY AMELIORATED SIGNIFICANTLY." PER THE REPORTER ON (B)(6) 2020, "PATIENT WAS SCHEDULED FOR LENS REMOVAL BUT CANCELED BECAUSE OF THE SIGNIFICANT AMELIORATION OF THE SX AND NORMAL IOP WITHOUT TREATMENT." Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1 - CORRECTED TO ADVERSE EVENT B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT B5 - PHOTOPHOBIA WAS ALSO REPORTED. SECONDARY YAG WAS PERFORMED. H6 - HEALTH EFFECT CLINICAL CODE 1880 H6 - HEALTH EFFECT IMPACT CODE 4644 H6 - HEALTH EFFECT IMPACT CODE 4625: ADDITIONAL SURGERY (SECONDARY YAG) ADDITIONAL INFORMATION: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510(K) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582873&pc=MTA | 2023826-2022-00595 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -12.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL# (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13582877&pc=MTA | 2023826-2022-00594 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Visual Disturbances; Halo; Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM512.6, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE PATIENT EXPERIENCED ELEVATED IOP (INTRAOCULAR PRESSURE), HEADACHE AND NAUSEA, MEDICAL INTERVENTION (PILOCARPINE/SAMPIRO 3-4/DAY) WAS PRESCRIBED. ON (B)(6) 2022 THE PATIENT VISITED AN EMERGENCY CLINIC DUE TO EYE PAIN, HEADACHE AND NAUSEA, IT WAS REPORTED THAT IOP (INTRAOCULAR PRESSURE) WAS NORMAL. PATIENT COMPLAINED OF FEELING SICK FROM HALO AND GLARE (LIGHTS ON THE STREET AT NIGHT) AND INCREASED EYE PAIN IN THE EVENING. THE PATIENT ALSO COMPLAINED OF FEELING SICK UNDER FLUORESCENT LIGHTS AND NOT BEING ABLE TO SLEEP DUE TO PAIN DURING SLEEP. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT ELEVATED IOP (INTRAOCULAR PRESSURE) IS TRANSIENT, CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION HALO, GLARE AND NIGHT BLINDNESS ARE BEING MONITORED. THERE IS NO PLAN TO REPLACE THE LENS NOW. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 - SERIAL # (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4767216&pc=MTA | 2023826-2015-00531 | 2015-03-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. AT THE SAME DAY POST-OPERATIVE EXAMINATION, THE PATIENT HAD 20/60 UNCORRECTED VISUAL ACUITY, INTRAOCULAR PRESSURE BY APPLANATION WAS 41MM HG AND LENS HAD A 200% VAULT. THE PATIENT COMPLAINED OF SLIGHT HEADACHE AND SHE WAS EXPERIENCING "STARBURSTS SIMILAR TO MIGRAINE AURA". LPI ENHANCEMENT WAS NECESSARY AND "BURPING" OF THE WOUND WAS PERFORMED. THE INTRAOCULAR PRESSURE WAS LOWERED TO SATISFACTORY LEVELS AND MEDICATION WAS PRESCRIBED. AT ONE DAY POST-OPERATIVE VISIT, THE SURGEON DECIDED THE LENS WAS TOO LARGE AND THE LENS WAS EXPLANTED ON (B)(6) 2015. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT SAME DAY POST-OPERATIVE VISIT, THE INTRAOCULAR PRESSURE WAS 44MM HG, UNCORRECTED VISUAL ACUITY WAS 20/50, WITH 1+ CORNEAL EDEMA AND 75% VAULTING OF LENS. "BURPING" OF THE WOUND LOWERED THE INTRAOCULAR PRESSURE AND PATIENT WAS PRESCRIBED MEDICATION. AT POST-OPERATIVE VISIT ON (B)(6) 2015, PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/30, INTRAOCULAR PRESSURE WAS 18MM HG WITH CLEAR CORNEA AND LENS HAD 75% VAULT. Manufacturer Narrative: (B)(4). INCISION OPENED. SECONDARY SURGERY. NEW IRIDECTOMY. EXPLANTED. VAULTING. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, MOST LIKELY THE EXCESSIVE VAULTING CAUSED SHALLOW ANTERIOR CHAMBER, ANGLE NARROWING OR CLOSURE, AND/OR PUPILLARY BLOCK WHICH WERE ASSOCIATED WITH INCREASED IOP. THE SECOND IOP INCREASE AFTER THE LENS EXCHANGE COULD BE CAUSED BY SURGICAL RELATED ANTERIOR CHAMBER ANGLE DISTORTION, PRIOR SUSPECT GLAUCOMA, PRIOR OPEN ANGLE GLAUCOMA OR VISCOELASTIC SUBSTANCE USED IN THE SURGERY. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: EVALUATION: RESULTS - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF DARK SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7794679&pc=MTA | 2023826-2018-01199 | 2018-01-25 | Injury | STAAR SURGICAL COMPANY | 2018-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. TWO DAYS AFTER THE SURGERY, THE PATIENT SUFFERED FROM SEVERE VOMITING, HEADACHE, AC LOSS AND TENSION OF MORE THAN 40. THE SURGEON ADMINISTERED MANNITOL BY IV AND THE SYMPTOMS DECREASED. AFTER 1 WEEK, THE LENS WAS EXPLANTED ON (B)(6) 2018. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4542564&pc=MTA | 2023826-2015-00176 | 2015-01-15 | Injury | STAAR SURGICAL COMPANY | 2015-02-24 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Intraocular Pressure Increased; Vomiting; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THAT EVENING AFTER SURGERY THE PATIENT HAD HEADACHE AND VOMITING, SUGGESTIVE OF INCREASED INTRAOCULAR PRESSURE. THE PATIENT HAS ALSO COMPLAINED OF GLARE. AT THE LAST POST-OP VISIT, THE PATIENT'S SYMPTOMS HAVE SUBSIDED BUT INTRAOCULAR PRESSURE WAS 24. THE LENS REMAINS IMPLANTED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - INTRAOCULAR PRESSURE RISE; VOMITING; HEADACHE; (GLARE). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373025&pc=MTA | 2023826-2017-00299 | 2017-01-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Headache; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS, -9.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE PATIENT EXPERIENCED PIGMENT DISPERSION, HEADCHES, AND OCULAR PRESSURE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6381321&pc=MTA | 2023826-2017-00325 | 2017-01-05 | Injury | STAAR SURGICAL COMPANY | 2017-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Intraocular Pressure, Delayed, Uncontrolled; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2017 THE PATIENT EXPERIENCED PIGMENT DISPERSION, HEADACHES, AND OCULAR PRESSURE. ON (B)(6) 2017 IT WAS DOCUMENTED THAT THERE WAS NO FREE PIGMENT IN ANTERIOR CHAMBER OF THE PATIENT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). INTRAOCULAR PRESSURE, DELAYED UNCONTROLLED NOT REPORTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITH IN THE SAME LOT. CLAIM#(B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614654&pc=MTA | 2023826-2015-00344 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4614667&pc=MTA | 2023826-2015-00345 | 2014-05-30 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Nausea; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO THE PATIENT EXPERIENCED GLARE/HALOS, HEADACHE AND WAS NAUSEOUS. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/24. Manufacturer Narrative: (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - GLARE/HALOS ARE A SUBJECTIVE PERCEPTION OF STARBURSTS OF LIGHT (LIGHT AROUND OR COMING OFF OBJECTS/LIGHT SOURCES). USUALLY MORE DETECTABLE AT NIGHT TIME, BUT IT CAN ALSO BE DETECTED DURING DAYTIME. POTENTIAL CAUSES ARE: RESIDUAL REFRACTIVE ERRORS, OPTICAL ZONE OF THE TREATMENT/IMPLANT SMALLER THAN PRE-OPERATIVE PUPIL SIZE IN MESOPIC CONDITIONS, DECENTERED TREATMENT/IMPLANT, LIGHT FROM THE IRIDOTOMY SITES, ETC. USUALLY THE PERCEPTION OF GLARE IS TRANSIENT, ONLY RELATED TO SOME LIGHTING CONDITIONS AND DIMINISHES/DISAPPEARS OVER A COURSE OF WEEKS/MONTHS. IN SOME CASES THE PERCEPTION MAY PERSIST AND BE SIGNIFICANT. IF THE PATIENT IS HAVING SIGNIFICANT DISTURBANCES AND IF THE IMPLANT IS THE SOURCE OF GLARE, IT CAN BE EXPLANTED THROUGH THE SAME INCISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335509&pc=MTA | 2023826-2020-01564 | 2020-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; Nausea; Pain; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS HAD EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF ICA AND PIGMENT DISPERSION. THE PATIENT EXPERIENCED EYE PAIN, BLURRING OF VISION AND NAUSEA. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS EXPLANTED ON (B)(6) 2020. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT DOES NOT FEEL HEADACHE ANYMORE. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: UNK. EXPIRATION DATE UNK. NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671658&pc=MTA | 2023826-2019-01019 | 2015-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Headache; Nausea; Pain; Red Eye(s); Blurred Vision; Burning Sensation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2015. THE SURGEON REPORTS PIGMENT DISPERSION, RED EYE, AND PERMANENT EYE PAIN SINCE IMPLANTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE SURGEON THINKS THE FOOTPLATES OF THE ICL SEEM TO BE LOW, PLACED ON/IN BETWEEN ZONULAS OR ON/IN THE CILIARY BODY-WHICH COULD CAUSE THE PAIN. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PATIENT CODE1880: HEADACHE. CLAIM#: (B)(4). Manufacturer Narrative: PER EMAIL ON 2019/07/25. OD IMPLANT SURGERY WAS ON "(B)(6)" OR (B)(6) 2015. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: DATE IS APPROXIMATE. INITIAL OPERATION OCCURRED EITHER (B)(6) 2015. AFTER IMPLANT PATIENT REPORTS HYPERMETROPIA OF +1 AND ASTIGMATISM OF 0.75. GLASSES WERE WORN AND PARTIALLY IMPROVED THE SITUATION. REPORTEDLY LASIK WAS PERFORMED IN (B)(6) 2017 WHICH IMPROVED THE PROBLEM. HOWEVER, IN (B)(6) 2018 THE RIGHT EYE BECAME WORSE AND THE PATIENT STATES PAIN IS CONSTANT ESPECIALLY WHEN READING. THERE IS ALSO A SLIGHT BURNING BEHIND THE EYE AND A FEELING OF NAUSEA ON THE RIGHT SIDE, EVEN AT NIGHT. IMPLANT DATE: DATE IS APPROXIMATE. ADDITIONAL PATIENT CODE ADDED, 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LASIK), LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: PRODUCT IS MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257758&pc=MTA | 2023826-2019-02014 | 2018-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT COMPLAINED OF A HEADACHE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9266062&pc=MTA | 2023826-2019-02013 | 2018-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT COMPLAINED OF A HEADACHE. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5553054&pc=MTA | 2023826-2016-00415 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Adverse Event Without Identified Device or Use Problem; Appropriate Term/Code Not Available | Headache; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT REPORTED IS WEARING GLASSES AT NIGHT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED. THE PATIENT REPORTED SHE HAD HEADACHES AT NIGHT AND WENT TO AN OPTOMETRIST. THE PATIENT WAS PRESCRIBED "LOW PRESCRIPTION" GLASSES AND NO LONGER HAS HEADACHES OR VISION PROBLEMS. THE PATIENT CONFIRMED NO OCULAR CONDITIONS/COMPLICATIONS. THE PATIENT IS FINE AND BETTER WITH GLASSES. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS REMAINS IMPLANTED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112582&pc=MTA | 2023826-2015-01188 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.30/+1.5/075 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT SUFFERED FROM HEADACHE. NO OTHER LENS WAS IMPLANTED. SEE MFR #2023826-2015-01187 FOR RIGHT EYE. Manufacturer Narrative: BASED ON THE DHR REVIEW, THE MANUFACTURING AND PACKAGING PROCESSES WAS PERFORMED ACCORDING TO PROCEDURES AND NO FINDINGS OR DEVIATIONS WERE REPORTED. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED, HOWEVER, THE COMPLAINT IS NOT DEVICE RELATED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THAT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND TWO OF THE FOOTPLATES WERE MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.30/+1.5/X075. (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5112607&pc=MTA | 2023826-2015-01187 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-09-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT SUFFERED FROM HEADACHE. NO OTHER LENS WAS IMPLANTED. SEE MFR #2023826-2015-01188 FOR LEFT EYE. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: BASED ON THE DHR REVIEW, THE MANUFACTURING AND PACKAGING PROCESSES WAS PERFORMED ACCORDING TO PROCEDURES AND NO FINDINGS OR DEVIATIONS WERE REPORTED. THE PRODUCT EVALUATION CONFIRMED THE LENS WAS IN SPECIFICATION. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINES, HOWEVER, THE COMPLAINT IS NOT DEVICE RELATED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THAT PIECES OF THE LENS HAPTICS WERE TORN OFF. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.00 (B)(4). METHOD CODE: EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE: EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146314&pc=MTA | 2023826-2015-01251 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/X087 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO THE PATIENT EXPERIENCING DISCOMFORT, A HEADACHE. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR A FOLLOW-UP. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -17.5/+2.0/X087. LENS NOT RETURNED. (B)(4). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151588&pc=MTA | 2023826-2015-01250 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.50/+2.50/X061 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT EXPERIENCED DISCOMFORT (HEADACHE) AND THE LENS WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. THE PATIENT DID NOT GO FOR FOLLOW UP VISIT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -14.50/+2.50/X061. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6688151&pc=MTA | 2023826-2017-01109 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING HEADACHES. PROBLEM WAS RESOLVED. THERE IS NO PLAN TO IMPLANT AN ALTERNATE LENS. SEE MFR REPORT# 2023826-2017-01110 FOR THE OTHER EYE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6688220&pc=MTA | 2023826-2017-01110 | 2017-03-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING HEADACHES. PROBLEM WAS RESOLVED. THERE IS NO PLAN TO IMPLANT AN ALTERNATE LENS. SEE MFR REPORT# 2023826-2017-01109 FOR THE OTHER EYE. THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTICS SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6048731&pc=MTA | 2023826-2016-01083 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCING HEADACHE. Manufacturer Narrative: CORRECTED INFORMATION: (B)(4) SHOULD BE UNDER PATIENT CODE. ADDITIONAL INFORMATION: CONCLUSION CODE: PATIENT IS UNDER THE AGE OF 21 AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENT'S UNDER THE AGE OF 21 YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6052593&pc=MTA | 2023826-2016-01084 | 2016-01-26 | Injury | STAAR SURGICAL COMPANY | 2016-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCING HEADACHE. Manufacturer Narrative: (B)(4). CONCLUSION CODE: PATIENT IS UNDER THE AGE OF (B)(6) AND PER DFU FOR THIS MODEL, THIS LENS MODEL IS CONTRADICTED FOR THE PATIENTS UNDER THE AGE OF (B)(6) YEARS. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE HAPTIC BENT AND PRESENCE OF FOREIGN MATERIAL (SPOT) ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587675&pc=MTA | 2023826-2017-00873 | 2016-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.0/+0.5/104 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, IRIS ATROPHY, AND HEADACHE. THE SURGEON COMMENTED THE VAULT ONE MONTH AFTER THE SURGERY IS 600 MICRONS, AND 3 MONTHS AFTER IS 902 MICRONS. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, SIGNIFICANT REDUCTION OF ICA, IRIS ATROPHY). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8298844&pc=MTA | 2023826-2019-00191 | 2018-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+2.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE PATIENT COMPLAINED OF A HEADACHE ON LEFT SIDE AFTER SURGERY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2018. ANOTHER LENS WAS NOT IMPLANTED. (B)(4). Manufacturer Narrative: DEVICE EVALUATION-THE LENS WAS RETURNED DRY IN LENS CASE/VIAL, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310562&pc=MTA | 2023826-2020-01501 |  | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Headache; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED INTERMITTENT HEADACHES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE HEADACHES HAVE BEEN RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A STERILE POUCH, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5788401&pc=MTA | 2023826-2016-00898 | 2016-06-07 | Injury | STAAR SURGICAL COMPANY | 2016-07-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Headache; Pain; Blurred Vision; Vomiting; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE LENS BEING TOO LARGE AND HYPOTONY (LOW IOP). THE PATIENT EXPERIENCED BLURRED VISUAL ACUITY, EYE PAIN, HEADACHE, AND VOMITING. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: (B)(4). RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND PIECES OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN A SMALL CLEAR CONTAINER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057261&pc=MTA | 2023826-2021-04915 | 2021-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Pain; Dizziness; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057173&pc=MTA | 2023826-2021-04916 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Pain; Dizziness; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IN GOOD CONDITION AFTER SURGERY.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT COMPLAINED OF EYE DISTENSION, EYE PAIN, DIZZINESS AND HEADACHE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT IN GOOD CONDITION AFTER SURGERY. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047711&pc=MTA | 2023826-2015-01091 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT DISSATISFACTION WITH HEADACHE AND OPHTHALMODYNIA. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DIOPTER: -09.00. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047815&pc=MTA | 2023826-2015-01092 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Headache; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO PATIENT DISSATISFACTION WITH HEADACHE AND OPHTHALMODYNIA. NO OTHER LENS WAS IMPLANTED. SEE MFR. # 2023826-2015-01091 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -09.50 (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS WAS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14661750&pc=MTA | 2023826-2022-01788 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Headache; Pupillary Block; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: ON 13-MAY-2022 WE RECEIVED NOTIFICATION OF A BOOK CHAPTER FROM 2015 ENTITLED, 'ICL AND SEQUENTIAL BILATERAL PUPILLARY BLOCK'. THE CHAPTER REPORTS A CASE OF BILATERAL PUPILLARY BLOCK AFTER ICL IMPLANTATION. THE ARTICLE STATES " ON THE FIRST DAY AFTER SURGERY, THE PATIENT COMPLAINED OF HEADACHE, NAUSEA, AND LOW VISION IN THE LEFT EYE. THE INTRAOCULAR PRESSURE (IOP) IN THE LE WAS 32 MMHG. A SLIT LAMP EXAMINATION REVEALED SHALLOW ANTERIOR CHAMBER AND A PARTIAL THICKNESS IRIDOTOMY WHICH DID NOT INCLUDE THE PIGMENTARY EPITHELIUM. THIS PEAK OF IOP WAS SUCCESSFULLY TREATED, AS DESCRIBED LATER. THE RIGHT EYE HAD A WIDE PHARMACOLOGIC MYDRIASIS AND IOP OF 8 MMHG THE FIRST DAY, BUT THE NEXT DAY THE IOP WAS 30 MMHG. SLIT LAMP EXAMINATION OF THE RE THE SECOND DAY WAS COMPARABLE TO THAT OF THE LE THE PREVIOUS DAY. SLIT LAMP EXAMINATION REVEALED THAT THE PERIPHERAL IRIDOTOMY WAS NOT APPROPRIATELY THICK; EACH IRIDOTOMY WAS COMPLETED WITH A NEODYMIUM:YAG (ND:YAG) LASER UNDER HIGH POWER. YAG LASER TREATMENT WAS APPLIED OVER THE PIGMENTARY EPITHELIUM. A SINGLE BURST OF 3.0 MJ WAS ENOUGH TO PERFORATE THE PIGMENTARY EPITHELIUM. A FLOW OF AQUEOUS FROM THE POSTERIOR TO ANTERIOR CHAMBER OCCURRED INSTANTLY UPON PENETRATION OF THE IRIS PIGMENT EPITHELIUM, WITH IMMEDIATE DEEPENING OF THE PERIPHERAL ANTERIOR CHAMBER AND NORMALIZATION OF THE ICL VAULT. ORAL ACETAZOLAMIDE AND BRIMONIDINE TARTRATE/TIMOLOL MALEATE DROPS 2 TIMES A DAY WERE PRESCRIBED. THREE MONTHS AFTER SURGERY, THE PATIENT'S VISUAL ACUITY WAS 20/20 IN BOTH EYES. THE SLIT LAMP EXAMINATION REVEALED A NORMAL ANTERIOR SEGMENT." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986867&pc=MTA | 2023826-2023-04600 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Headache; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO GLARE/HALOS, HEADACHE, AND PATIENT DISSATISFACTION. PROBLEM RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694745&pc=MTA | 2023826-2017-01119 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/95 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED IN MICRO CENTRIFUGE VIAL WITH LITTLE MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694750&pc=MTA | 2023826-2017-01120 | 2017-01-04 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -5.0/+2.0/90 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT IS NO LONGER HEADACHE AFTER THE LENS EXPLANT, AND DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WITH A CURVED SHAPE. THERE WAS EVIDENCE OF CLEAR RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6694800&pc=MTA | 2023826-2017-01121 | 2017-05-10 | Injury | STAAR SURGICAL COMPANY | 2017-07-07 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Headache; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED HEADACHE, GLARE/HALOES. ON (B)(6) 2017 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CUSTOMER STATED THAT THE PATIENT DOES NOT HAVE PLAN TO EXCHANGE THE LENS. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A LENS HAPTIC SLIGHTLY BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221698&pc=MTA | 2023826-2015-01522 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Hematoma | P030016 | 1 | Event Description: A PATIENT REPORTED HAD AN ICL IMPLANTABLE COLLAMER LENS IMPLANTED IN HIS RIGHT EYE (OD). SOON AFTER, THE PATIENT HAD A HEMATOMA FORM ON HIS EYE AND IT LASTED SEVERAL DAYS. THE PATIENT REPORTED THE EYE HAS SINCE HEALED AND IS SEEING FAIRLY WELL OUT OF THAT EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - THE PATIENT REPORTED THAT HE HAD EXPERIENCED "HEMATOMA FORMED OVER EYE" THAT LASTED SEVERAL DAYS FOLLOWING ICL IMPLANTATION. NO MEDICAL OR SURGICAL INTERVENTION, AND NO LOSS OF BCVA WERE REPORTED. THE ISSUE RESOLVED ("HEMATOMA HEALED"). THE PATIENT WEARS CONTACT LENS IN THE FELLOW EYE. IT SHOULD BE NOTED THAT HEMATOMA REFERS TO BRUISING OF THE TISSUE UNDER SKIN AROUND THE EYE THAT USUALLY RESOLVES SPONTANEOUSLY. BECAUSE ALL EYE PROCEDURES INVOLVE SOME FORM OF MANIPULATION ON EYELID (NO MATTER HOW SLIGHT), THERE IS A RISK OF FORMING A HEMATOMA. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7535087&pc=MTA | 2023826-2018-00787 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-22 | MTA | UNK | Device Operates Differently Than Expected | Hemorrhage/Bleeding | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN PLOS ONE WAS RECEIVED ON 27-APR-2018 THAT CITED A PROSPECTIVE COHORT STUDY, ARTICLE TITLED 'IRIDOCILIARY CYSTS DO NOT IMPACT ON POSTERIOR PHAKIC INTRAOCULAR LENS IMPLANTATION FOR HIGH MYOPIA CORRECTION: A PROSPECTIVE COHORT STUDY IN 1569 EYES.' THE ARTICLE REPORTS THAT AT 20 MONTHS AFTER INTIAL SURGERY, ONE EYE EXPERIENCED MACULAR HEMORRHAGE. REPORTEDLY, THIS DID NOT CAUSE A SIGNIFICANT DECREASE IN BCVA. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005344&pc=MTA | 2023826-2021-01925 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Hyphema | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE LEFT EYE (OS) OF A FEMALE PATIENT. THE REPORTER INDICATED THE PATIENT WAS WORKING OUT AND DEVELOPED A HYPHEMA IN HER EYE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361184&pc=MTA | 2023826-2017-00270 | 2016-07-22 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Hyphema; Discomfort; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO DISCOMFORT AND HYPHEMA THAT OCCURRED IN THE RIGHT EYE. ON (B)(6) 2017, THE PATIENT'S BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION-LENS RETURNED IN A CENTRIFUGE VIAL AND HAD MOISTURE ON IT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119129&pc=MTA | 2023826-2015-01262 | 2015-09-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Hyphema; Hypopyon; Inflammation; Irritation; Pain; Visual Impairment | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE REPORTER STATED AT ONE DAY POST-OP ((B)(6) 2015), THE PATIENT'S EYE HAD HYPHEMA AND HYPOPHON. MEDICATIONS WERE ADMINISTERED FOR INFLAMMATION BUT THE REPORTER STATED IT LOOKED MORE LIKE AN IRRITATED IRIS. THE PATIENT WAS IN PAIN AND CULTURES WERE TAKEN. THE LENS VAULT IS GOOD. THE PATIENT RETURNED FOR A FOLLOW-UP VISIT ON (B)(6) 2015 AND THE VA WAS 20/20, THERE IS STILL SOME CELL IN THE EYE AND THE VISION IS NOT AS CLEAR. THE REPORTER STATED THE TEST ON THE CULTURE CAME BACK AS TOXIC ANTERIOR SEGMENT SYNDROME. THE REPORTER STATED THE PATIENT'S EYE HAS SETTLED DOWN AND IS RECOVERING. THE PATIENT'S VA WAS 20/20. Manufacturer Narrative: (B)(4). RESULTS: PER MEDICAL REVIEW - ENDOPHTHALMITIS IS A PAINFUL INFECTION OF OCULAR TISSUES THAT MAY OCCUR POSTOPERATIVELY. UNLESS IT IS VIRAL IN ORIGIN, A CULTURE WILL USUALLY BE POSITIVE. TOXIC ANTERIOR SEGMENT SYNDROME (TASS) IS A PAINLESS INFLAMMATORY POSTOPERATIVE CONDITION THAT HAS BEEN ASSOCIATED WITH DRUGS, DEVICES, AND SURGICAL ISSUES. CULTURE IS USUALLY NEGATIVE. IT IS NOT CLEAR WHICH OF THESE CONDITIONS DEVELOPED IN THE PATIENT'S EYE, AND THE CAUSE OF EITHER CONDITION IN THIS CASE IS UNKNOWN. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). METHOD - (PROCESS EVALUATION): LENS WORK ORDER SEARCH, CARTRIDGE LOT NUMBER SEARCH, FOAM TIP PLUNGER LOT NUMBER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. A CARTRIDGE LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A FOAM TIP PLUNGER LOT NUMBER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS, INJECTOR AND CARTRIDGE THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, LENS WORK ORDER SEARCH, CARTRIDGE AND FOAM TIP PLUNGER LOT NUMBER SEARCHES AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6258251&pc=MTA | 2023826-2017-00083 | 2016-10-05 | Injury | STAAR SURGICAL COMPANY | 2017-01-17 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Hyphema; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED HYPHEMA. THE LENS WAS EXPLANTED ON (B)(6) 2016. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED, THEREFORE NO EVALUATION HAS BEEN PERFORMED. Manufacturer Narrative: PATIENT DOB-UNK IS INCORRECT, PATIENT'S DATE OF BIRTH WAS NOTED ON COMPLAINT QUESTIONAIRE. PRODUCT RETURNED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4704669&pc=MTA | 2023826-2015-00431 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Hypopyon; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE PATIENT HAS BEEN DIAGNOSED WITH ACUTE IRIDOCYCLITIS AND HYPOPYON. THE PATIENT IS BEING PRESCRIBED MEDICATION AND THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PT WEIGHT: UNK. DATE EVENT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS IMPLANTED. EVALUATION, METHOD : WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE INFORMATION PROVIDED BY THE FACILITY, THE LENS WAS NOT INDICATED TO HAVE BEEN RECEIVED DAMAGED OR DEFECTIVE. IT IS UNCLEAR IF THE LENS ITSELF CAUSED OR CONTRIBUTED TO THE ONSET OF IRIDOCYCLITIS AND HYPOPYON. THEREFORE, NO ROOT CAUSE CAN BE IDENTIFIED. PROBABLE CAUSES FOR THIS EVENT COULD INCLUDE PRE-EXISTING PATIENT ISSUES OR CONDITIONS. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762426&pc=MTA | 2023826-2017-01232 |  | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | UNK | Device Operates Differently Than Expected | Increased Sensitivity; No Code Available | N/A | 1 | Event Description: ON (B)(6) 2017, STAAR BECAME AWARE OF A FACEBOOK REVIEW OF A STAAR LENS. THE REPORTER STATED THAT THEY HAD PERMANENT DAMAGE DONE TO ONE PUPIL, MAKING THE PUPIL OF THE LEFT EYE (OS) PERMANENTLY LARGER THAN THE PUPIL OF THE RIGHT EYE (OD). REPORTER ALSO STATES GOOD DISTANCE VISION BUT IS HYPER-SENSITIVE TO LIGHT IN THE LEFT EYE AND "SEES COLORS DIFFERENTLY." REPORTER STATES THE DAMAGE IS PERMANENT. IT IS UNKNOWN WHETHER THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: INFORMATION ENTERED INTO DATA FIELDS DUE TO REQUIREMENT OF ESUBMITTER, HOWEVER INFORMATION IS NOT KNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7186676&pc=MTA | 2023826-2018-00014 | 2014-03-18 | Injury | STAAR SURGICAL COMPANY | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2014. AT ONE DAY POST-OP, INFLAMMATION WAS NOTED. THE SURGEON ADMINISTERED ANTIBIOTICS AND THE EYE RETURNED TO NORMAL. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS NOT LENS RELATED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT WAS GIVEN STEROIDS AFTER INFLAMMATION. NOW EVERYTHING IS FINE AND THE LENS IS STILL IN THE EYE. AT THE LAST VISIT ON (B)(6) 2017, THE PATIENT UCVA WAS 20/20. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2014. IMPLANTED DATE: (B)(6) 2014 CORRECTED TO (B)(6) 2014. (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED FROM (B)(6) 2014 TO (B)(6) 2014. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7186699&pc=MTA | 2023826-2018-00015 | 2014-03-18 | Injury | STAAR SURGICAL COMPANY AG | 2018-01-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2014. AT ONE DAY POST-OP, INFLAMMATION WAS NOTED. THE SURGEON ADMINISTERED ANTIBIOTICS AND THE EYE RETURNED TO NORMAL. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE EVENT WAS NOT LENS RELATED. Manufacturer Narrative: THE REPORTER INDICATED THE PATIENT WAS GIVEN STEROIDS AFTER INFLAMMATION. NOW EVERYTHING IS FINE AND THE LENS IS STILL IN THE EYE. AT THE LAST VISIT ON (B)(6) 2017, THE PATIENT UCVA WAS 20/20. (B)(4) Manufacturer Narrative: DATE OF EVENT: CORRECTED FROM (B)(6) 2014. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. IMPLANTED DATE: CORRECTED FROM (B)(6) 2014 TO (B)(6) 2014. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12361943&pc=MTA | 2023826-2021-03295 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS INFLAMMATION. THE PATIENT WAS TREATED WITH DEXAMETHASONE, ALMAGATE, CEFIXIME, LOXOPROFEN SODIUM, AND METHYLPREDNISOLONE. REPORTEDLY THE PATIENT IS GRADUALLY GETTING BETTER AND WILL CONTINUE TO BE OBSERVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363342&pc=MTA | 2023826-2021-03296 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS INFLAMMATION. THE PATIENT WAS TREATED WITH DEXAMETHASONE, ALMAGATE, CEFIXIME, LOXOPROFEN SODIUM, AND METHYLPREDNISOLONE. REPORTEDLY THE PATIENT IS GRADUALLY GETTING BETTER AND WILL CONTINUE TO BE OBSERVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15134238&pc=MTA | 2023826-2022-02451 | 2022-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 -14.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO IRIDOCYCLITIS AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18278965&pc=MTA | 2023826-2023-05523 | 2023-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REBOUND POST-OP INFLAMMATION WAS REPORTED. PRE-MOXIBROM WAS RESTARTED AND THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: H6-HEALTH IMPACT:4644: STEROID DROPS, PRE-MOXI-BROM. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16479534&pc=MTA | 2023826-2023-00727 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF A -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). PATIENT DEVELOPED UNCONTROLLED INFLAMMATION AND THE LENS WAS EXPLANTED. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14836996&pc=MTA | 2023826-2022-01969 | 2022-05-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Inflammation |  | 1 | Event Description: THE REPORTER INDICATED THAT A MESSAGE WAS RECEIVED FROM A DOCTOR THAT "EXPERIENCED INFLAMMATION CASES FOR HER ICL PATIENTS ABOUT 3 WEEKS AGO." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B1: MALFUNCTION. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575012&pc=MTA | 2023826-2019-00732 | 2019-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Inflammation | P030016 | 1 | Event Description: IT WAS REPORTED THAT A 12.1MM, MICL12.1, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM BUT MILD INFLAMMATION WAS OBSERVED AT PATIENTS LAST FOLLOW UP VISIT, BCVA:20/40.2, UCVA: 30/40.2, VAULT:500 AND IOP 12MMHG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: IT WAS REPORTED THAT THE INFLAMMATION HAS RESOLVED AND PATIENT IS SEEING WELL. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138812&pc=MTA | 2023826-2020-01188 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-06-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Inflammation | N/A | 1 | Event Description: THE SURGEON REPORTED "LATE-ONSET INFLAMMATION" ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943255&pc=MTA | 2023826-2023-01838 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Inflammation | UNK | 1 | Event Description: IN THE ARTICLE, "SUCCESSFUL ANTIBIOTIC MANAGEMENT OF STAPHYLOCOCCUS EPIDERMIDIS ENDOPHTHALMITIS WITHOUT ICL EXPLANTATION OR VITRECTOMY," ENDOPHTHALMITIS OBSERVED 20 DAYS POST ICL. REPORTEDLY, "CONSIDERED FOR POSTOPERATIVE ENDOPHTHALMITIS OS, A VITREOUS TAP AND INJECTION OF VANCOMYCIN 1MG/0.1ML AND CEFTAZIDIME 2MG/0.1 ML WAS PERFORMED," AND "THE PATIENT WAS ALSO TREATED WITH SYSTEMIC (CENTRAZIDME1GLVQ24H) AND TOPICAL ANTIBIOTICS (TOBRAMYCIN DEXAMETHASONE EYE Q15M 15MG:5MG/5ML, LEVOFLOXACIN 0.5% EYE DROP Q2H), 1% PREDNISOLONE ACETATE EYE DROPS THREE TIMES A DAY AND 0.5% TROPICAMIDE QD." THE TOPICAL TREATMENT WAS TAPERED GRADUALLY OVER ONE MONTH. REPORTEDLY, "THE UNCORRECTED DISTANCE VISUAL ACUITY OF20/20 WAS ACHIEVED THREE MONTHS FOLLOWING THE ANTIBIOTIC THERAPY." Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634622&pc=MTA | 2023826-2020-02290 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED INFLAMMATION. THE PATIENT RESPONDED WELL TO TREATMENT WITH STEROIDS (ANTI-INFLAMMATORY), AND BY THE END OF (B)(6) 2020, THE INFLAMMATION HAD DISAPPEARED. THE PATIENT IS OK. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: B5: "WHITE DEPOSITS DUE TO THE INFLAMMATION WERE OBSERVED." SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 4581: WHITE DEPOSITS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634631&pc=MTA | 2023826-2020-02289 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED INFLAMMATION. THE PATIENT RESPONDED WELL TO TREATMENT WITH STEROIDS (ANTI-INFLAMMATORY), AND BY THE END OF (B)(6) 2020, THE INFLAMMATION HAD DISAPPEARED. THE PATIENT IS OK. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: B5 - "WHITE DEPOSITS DUE TO THE INFLAMMATION WERE OBSERVED." SHOULD BE ADDED TO THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 4581: WHITE DEPOSITS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919174&pc=MTA | 2023826-2015-00841 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Device Contamination with Chemical or Other Material | Inflammation; Blurred Vision; Loss of Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.0MM ICM120 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2009. ON (B)(6) 2015, THE SURGEON NOTED LOSS OF BCVA, INFLAMMATION AND DEPOSITS ON THE LENS. THE LENS WAS REMOVED AND REPLACED BUT THIS DID NOT RESOLVE THE ISSUE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: P030016 TO N/A - THIS PRODUCT IS NOT MARKETED IN THE U.S. MFR DATE - 03/01/2013 TO 12/16/2008, PREVIOUS DATE IS INCORRECT. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC WAS TORN AND THE HAPTIC WAS BROKEN. WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7987323&pc=MTA | 2023826-2018-01529 | 2018-09-11 | Injury | STAAR SURGICAL COMPANY | 2018-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE REPORTER INDICATED LOSS OF BEST-CORRECTED VISUAL ACUITY (BCVA), PIGMENT DISPERSION, CELL DEPOSIT AND INFLAMMATION. THE LENS REMAINS IMPLANTED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322791&pc=MTA | 2023826-2019-00188 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE INFLAMMATION WAS TREATED BUT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED.. THE PATIENT IS IMPROVING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, 11.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY AND UNREACTIVE (FIXED) PUPIL. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS ON (B)(6) 2019. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THERE WAS SIGNIFICANT RESOLUTION UPON REMOVAL AND THE PATIENT IS IMPROVING. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLOUDY RESIDUE ON THE LENS. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. PATIENT CODE: 3191 - UNREACTIVE (FIXED) PUPIL. DATE OF BIRTH: (B)(6) 1990. EXPLANTED DATE: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322830&pc=MTA | 2023826-2019-00187 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THE PATIENT IS IMPROVING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO CHRONIC INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST CORRECTED VISUAL ACUITY AND UNREACTIVE (FIXED) PUPIL. THE ANTERIOR CHAMBER WAS IRRIGATED/EVACUATED OF VISCO/FLUIDS ON 01/02/2019. THE INFLAMMATION WAS TREATED BUT IT WAS UNSUCCESSFUL AND THE LENS WAS EXPLANTED. THERE WAS SIGNIFICANT RESOLUTION UPON REMOVAL AND THE PATIENT IS IMPROVING. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLOUDY RESIDUE ON THE LENS. CORRECTED DATA: DATE OF BIRTH - (B)(6) 1990 EXPLANTED DATE - (B)(6) 2019 CLAIM # (B)(4). Manufacturer Narrative: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) NCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180447&pc=MTA | 2023826-2019-01903 | 2019-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | Inflammation; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. IT WAS REPORTED THAT THE PATIENT HAD BLURRED VISION AND TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED IN THE PATIENT'S EYE. THE PATIENT WAS GIVEN STEROIDS AND THE SYMPTOMS GRADUALLY IMPROVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 2993 CORRECTED TO DEVICE CODE 1401 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER APPROXIMATELY 3 MONTHS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983668&pc=MTA | 2023826-2022-01035 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Inflammation; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), DISTANCE AND NEAR VISION SLIGHTLY WORSE THAN THE OD EYE, BLEPHARITIS, IRIS BOMBE WITH PREVIOUS APPARENT TOUCH POINT SUPERIOR TO PUPIL, "VH 1/0," NIGHTTIME HALOS WHEN DRIVING AND ANY CEILING LIGHTS, AND IRIS THINNED WITH PIGMENT ON ICL. THE SURGEON STATES, "VERY NARROW ANGLE...ANGLE SHUT OFF." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO BIG." Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8504763&pc=MTA | 2023826-2019-00588 | 2019-02-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM. MICL12.6, -6.0 DIOPTER, IMPLANTABLE COLLAMER LENS IS TOO LARGE FOR THE PATIENT, INCREASED EYE PRESSURE. WILL BE REPLACING WITH A SHORTER LENGTH LENS. NO PATIENT INJURY. HOWEVER, THE PATIENT LATER STATED " IF THERE WAS SOMETHING WRONG THE IMPLANT WOULD HAVE BEEN REMOVED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DOCTOR IS NOT REMOVING OR REPLACING THE CURRENT IMPLANT. THE DOCTOR FELT BECAUSE THE VIEW WAS HAZY DUE TO INFLAMMATION AND PRESSURE THE VAULT SIZE HAD TO BE TOO BIG, WHICH MUST BE LEADING TO THE INCREASE IN EYE PRESSURE. THE PATIENT WAS MONITORED FOR APPROXIMATELY 2-3 MONTHS. THE PATIENTS PRESSURE WENT DOWN TO NORMAL RANGE AND VISION IS 20/20. THE VAULT SIZE IS 100%. PT IS VERY HAPPY WITH OUTCOME. PATIENT CODE:1932 ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14028818&pc=MTA | 2023826-2022-01108 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. THE LENS REMAINS IMPLANTED. CHRONIC INFLAMMATION IS REPORTED AND DOCTOR STILL CONTINUES TO MONITOR PATIENT WITH COMBIGAN LE. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389438&pc=MTA | 2023826-2022-00352 | 2021-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -5.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (52MMHG) WHICH COULD NOT BE CONTROLLED WITH MAXIMAL TOPICAL ANTIGLAUCOMA +ORAL DIAMOX MEDICATION. PATIENT'S DRAINAGE ANGLE WAS HEAVILY PIGMENTED WITH PIGMENTS AND EXCESSIVE PIGMENTS WERE NOTED ANTERIOR CHAMBER. THE MEDICATION REDUCED THE IOP TO 9MMHG WITH ADDITIONAL ORAL GLYCEROL AT A LATER DATE. IT WAS ALSO INDICATED AT A LATER DATE INFLAMMATION OCCURRED AND ANTI-INFLAMMATORY EYEDROPS AND PREDNISOLONE WAS GIVEN. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS REPORTED THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14030262&pc=MTA | 2023826-2022-01091 | 2022-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND PIGMENT DISPERSION. THE VAULT WAS NORMAL. POST-OP THERE WAS MILD INFLAMMATION IN ANTERIOR CHAMBER. IOP WAS 15MMHG AND MEDICATION WAS PRESCRIBED. THE PATIENT IS BEING MONITORED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506509&pc=MTA | 2023826-2022-03375 | 2022-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | Inflammation; Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED IOP WITHOUT PUPIL BLOCK; PIGMENT DISPERSION AND PRO-LONGED INFLAMMATION; "11 O'CLOCK IRIS TAGS". ON (B)(6) 2022 THE LENS WAS REPOSITIONED TO FREE UP THE LENS FROM THE IRIS TOUCH AND GONIO WASH WAS PERFORMED TO REDUCE THE PIGMENTS IN AC. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873864&pc=MTA | 2023826-2016-01066 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, GLARE, HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED IOP AND VERY SHALLOW AC. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY. THE LENS WAS EXPLANTED AND THERE WAS ENLARGEMENT OF INCISION. THE PATIENTS CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: (B)(4). CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -2.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON 06/20/2016 DUE TO EXCESSIVE VAULTING, GLARE/HALOES, BLURRED VISION, INFLAMMATION, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE, AND VERY SHALLOW ANTERIOR CHAMBER. THE PATIENT HAD A LOSS OF BEST CORRECTED VISUAL ACUITY (BCVA). THE LENS WAS EXPLANTED ON (B)(6) 2016 AND THERE WAS ENLARGEMENT OF INCISION. THE SURGEON PERFORMED AN ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS AND A PARS PLANA VITRECTOMY TO DRAIN THE ACQUEOUS MISDIRECTION. THE PATIENT'S CONDITION ON THE LAST VISIT WAS HYPERAEMIA AND SMALL DISC HEMORRHAGE INFERIORLY. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30. USER FACILITY IS THE INCORRECT REPORT SOURCE, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873698&pc=MTA | 2023826-2016-01062 | 2016-06-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; Blurred Vision; Loss of Vision; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -02.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE AND INFLAMMATION. THE PATIENT EXPERIENCED LOSS OF BEST-CORRECTED VISUAL ACUITY, GLARE/HALOS AND BLURRED VISION. THE INCISION WAS ENLARGED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE PATIENT'S EYE IS IN GOOD CONDITION AFTER EXPLANTATION AND THE INFLAMMATION WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422330&pc=MTA | 2023826-2017-00444 | 2017-01-31 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDE: ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. ON (B)(6) 2017, THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR THICK AND REDDISH BROWN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.5/38 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), SIGNIFICANT REDUCTION OF ICA, ACCOMMODATIVE SPASM, AND BLURRED VISION. INTERVENTIONS INCLUDED ENLARGEMENT OF PI (OR ADDITION OF NEW PI), IRIDECTOMY, AND PILOCARPINE TREATMENT. AFTER LENS EXPLANT, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/100 AND THE PATIENT EXPERIENCED A LITTLE INFLAMMATION OF THE ANTERIOR SEGMENT. DEVICE EVALUATION: THE DEVICE WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH-NO SIMILAR COMPLAINTS FOUND. CONCLUSION: OFF-LABEL USE-KERATOCONUS EYE. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6213394&pc=MTA | 2023826-2016-01834 | 2016-11-02 | Injury | STAAR SURGICAL COMPANY | 2016-12-29 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Insufficient Information | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICH12.1 IMPLANTABLE COLLAMER LENS, +3.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. AFTER IMPLANTATION, THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, LENS OPACITY (ASCO), INFLAMMATION, AND DAY 1 PUPILLARY MEMBRANE. IN ADDITION, A SECONDARY INTERVENTION OF ENLARGEMENT/ADDITION OF PI USING YAG WAS PERFORMED. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: AFTER IMPLANTATION, THE PATIENT ALSO EXPERIENCED HAZY VISION, HIGH IOP, AND SYNECHIA. (B)(4). INITIAL MDR DATE IS (B)(6) 2016. (B)(4). Manufacturer Narrative: "THE LENS WAS EXPLANTED" IS CORRECTED TO " THE LENS WAS EXCHANGED FOR A SIMILAR LENS OF DIFFERENT DIOPTER ON (B)(6) 2017. REPORTER STATED THAT THE PATIENT WAS HAPPY. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). CORRECTED DATA: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: IT IS CORRECTED FROM "OTHER SERIOUS" IN THE PREVIOUS SUPPLEMENTAL TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/ DAMAGE". (B)(4). Manufacturer Narrative: THE LENS WAS EXPLANTED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7721249&pc=MTA | 2023826-2018-01084 | 2018-06-13 | Injury | STAAR SURGICAL COMPANY | 2018-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP, ANGLE CLOSURE, INFLAMMATION, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE, THE PATIENT IS "DOING WELL." AZETAZOLAMIDE 250MG GIVEN TO MANAGE SYMPTOMS. THE SURGEON LISTS THE CAUSE OF THE EVENT AS DEVICE. Manufacturer Narrative: DATE OF THIS REPORT: IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DATE RECEIVED BY MFR: -IT IS CORRECTED TO 01-JUL-2018 IN THE INITIAL MDR. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7714154&pc=MTA | 2023826-2018-01016 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Iritis; Uveitis; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2, -13.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, IRIS ATROPHY, INCREASED CELL, AND FLARE/INFLAMMATION. PATIENT WAS ON ANTIBIOTICS AND ANTI-INFLAMMATORIES TO ADDRESS THE INFLAMMATION. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP THERE WAS NO PAIN, NO ACHE, DECREASED INFLAMMATION AND PATIENT IS DOING WELL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A SMALL PLASTIC CONTAINER. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4529095&pc=MTA | 2023826-2015-00185 | 2014-12-12 | Injury | STAAR SURGICAL COMPANY | 2015-02-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Inflammation; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP), INFLAMMATION AND POSSIBLE IRIDOCYCLITIS. THE LENS WAS NOT REPLACED AND THE PATIENT'S LAST BCVA WAS 20/15. Manufacturer Narrative: DATE OF BIRTH - UNK. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THIS LENS THAT WAS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: REFRACTIVE LENS EXCHANGE (PHACO-IOL) WAS PERFORMED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). CORRECTED DATA: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7714241&pc=MTA | 2023826-2018-01017 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-07-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, INCREASED CELL, AND FLARE/INFLAMMATION. PATIENT WAS ON ANTIBIOTICS AND ANTI-INFLAMMATORIES TO ADDRESS THE INFLAMMATION. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP THERE WAS NO PAIN, NO ACHE, DECREASED INFLAMMATION AND PATIENT IS DOING WELL. Manufacturer Narrative: (B)(4) Manufacturer Narrative: DEVICE EVALUATION:LENS WAS RETURNED IN LIQUID, IN A SMALL PLASTIC CONTAINER. VISUAL INSPECTION FOUND A HAPTIC TEAR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13830688&pc=MTA | 2023826-2022-00880 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Intraocular Pressure Increased; Pupillary Block; Red Eye(s); Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1; -10.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AC CELL++; CHRONIC INFLAMMATION; PUPIL BLOCK; DISCOMFORT AND RED EYE; OVER A PERIOD OF TIME AFTER IMPLANTATION; MULTIPLE MEDICATIONS WERE ALSO PRESCRIBED. THE SURGEON THEN PERFORMED A SECONDARY YAG ON (B)(6) 2021 WHICH THEN RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710489&pc=MTA | 2023826-2022-01917 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Inflammation; Intraocular Pressure Increased; Red Eye(s) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO INFLAMMATION AND PIGMENT DISPERSION OCCURRED.; THEN STEROID EYE DROPS WERE PRESCRIBED AND INDUCED IOP. A TRABECULECTOMY WAS PERFORMED. CAUSE IS DUE TO A PATIENT RELATED FACTOR OF PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14710500&pc=MTA | 2023826-2022-01918 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | Inflammation; Intraocular Pressure Increased; Red Eye(s) | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(4) 2022. ON (B)(4) 2022 THE LENS WAS EXPLANTED DUE TO INFLAMMATION AND PIGMENT DISPERSION OCCURRED.; THEN STEROID EYE DROPS WERE PRESCRIBED AND INDUCED IOP. A TRABECULECTOMY WAS PERFORMED. CAUSE IS DUE TO A PATIENT RELATED FACTOR OF PIGMENT DISPERSION SYNDROME. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7201655&pc=MTA | 2023826-2018-00090 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Inflammation; Intraocular Pressure Increased; Uveitis | NA | 1 | Event Description: AFTER ICL IMPLANT AT THE END OF 2016, THE SURGEON NOTES UVEITIS "WITH SEVERE OCULAR HYPERTENSION THAT ONLY RESPONDS TO MAXIMUM TREATMENT" IN THE PATIENT'S EYE. THE SURGEON STATES THAT SYMPTOMS STARTED AFTER THE IMPLANT OCCURRED AND HE BELIEVES THERE ARE "SIGNS OF REJECTION" TO THE IMPLANTED LENS. NO ADDITIONAL INFORMATION IS CURRENTLY AVAILABLE. Manufacturer Narrative: UPDATED INFORMATION: REPORTEDLY, THE PATIENT'S SYMPTOMS ARE UNDER CONTROL. THE PATIENT HAS MAINTAINED "GOOD VISUAL ACUITY." THE SURGEON DOES NOT PLAN TO TAKE ANY SURGICAL ACTION. (B)(4) Manufacturer Narrative: PATIENT INFORMATION: UNK. ESTABLISHMENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8107278&pc=MTA | 2023826-2018-01723 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Contamination /Decontamination Problem; Unintended Movement | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/114 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE SURGEON REPORTS ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UVEITIS, SLIGHTLY DISTORTED PUPIL, CELL DEPOSITS ON THE ICL, AND INFLAMMATION. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. PER THE SURGEON, THE LENS "HAS BEEN SLIGHTLY TILTED FROM THE BEGINNING (PATHOLOGY IN THE SULCUS?)." CELL DEPOSITS DISAPPEAR WHEN GIVEN PRED FORTE; REAPPEAR WHEN THE "STERIOD" IS "DISCONSTINUED". REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: EVENT-ADDITIONAL DATA RECEIVED: "THE EXCHANGE RESOLVED THE INFLAMMATION, THE EYE IS CALM. VA SC 0.8, CC 1.0 VAULT 350 MICRON. ICL IN VERTICAL POSITION (AWAY FROM THE NASAL PROBLEM AREA). PUPIL SLIGHTLY NASALLY DISTORTED. THE MD ATTRIBUTES THE INFLAMMATION CLEARLY TO THE OVER VAULTED ICL. EVERYONE IS HAPPY". IMPLANT DATE: (B)(6)2018 DEVICE EVALUATED BY MFR:DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT-CORRECTED IMPLANT DATE INTO THE RIGHT EYE IS: "(B)(6) 2015" IN THE INITIAL MDR. RELEVANT TESTS/LABORATORY DATA-CORRECTED TO: "NA" IN THE SUPPLEMENTAL MDR#1 CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8742035&pc=MTA | 2023826-2019-01155 | 2019-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION. THE REPORTER INDICATED THE EXCESSIVE VAULTING CAUSED PIGMENT DISPERSION (INCREASING INTRAOCULAR PRESSURE CONTROLLED ON DROPS AND HEAVY PIGMENT ON ICL), AND ANTERIOR CHAMBER INFLAMMATION SPIKES AS STEROIDS WERE TAPERED. THE PATIENT HAD CONSTANT ACHING SENSATION BUT THE PATIENT IS DOING WELL AT ONE-DAY POST-OP EXPLANT EXAM. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A HAPTIC TEAR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8742046&pc=MTA | 2023826-2019-01154 | 2019-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Intraocular Pressure Increased; Uveitis; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, PIGMENT DISPERSION AND ANTERIOR CHAMBER INFLAMMATION. THE REPORTER INDICATED THE EXCESSIVE VAULTING CAUSED PIGMENT DISPERSION (INCREASING INTRAOCULAR PRESSURE CONTROLLED ON DROPS AND HEAVY PIGMENT ON ICL), AND ANTERIOR CHAMBER INFLAMMATION SPIKES AS STEROIDS WERE TAPERED. THE PATIENT HAD CONSTANT ACHING SENSATION BUT THE PATIENT IS DOING WELL AT ONE-DAY POST-OP EXPLANT EXAM. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DAVE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7152546&pc=MTA | 2023826-2017-01967 | 2017-08-29 | Injury | STAAR SURGICAL COMPANY | 2017-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Inflammation; Iritis; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, DIOPTER -6.50 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED, THE SURGEON REPORTING INFLAMMATION AND IRITIS. THE CLAIMANT STATES THAT THE PATIENT HAD AN INFLAMMATORY RESPONSE AFTER THE SURGERY. THE SURGEON TREATED THE IRITIS BY PRESCRIBING ANTIBIOTICS AND ANTI-INFLAMMATORY MEDICINE. REPORTEDLY, THE SURGEON LEFT THE LENS IMPLANTED SINCE VA WAS 20/20. UPON CHECK-UP THE "VA WAS FAILING" AND THE LENS WAS EXPLANTED. PROBLEM RESOLVED WITH LENS EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS AND CLEAR RESIDUE ON LENS. (B)(4) Manufacturer Narrative: PMA 510(K): THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8159612&pc=MTA | 2023826-2018-01822 | 2018-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Iritis; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1, -03.00 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT, PIGMENT DISPERSION, AND PERSISTENT INFLAMMATION WAS OBSERVED. ON (B)(6) 2018 THE LENS WAS REMOVED. IT WAS REPORTED THAT DURING INFLAMMATION DOCTOR PRESCRIBED G.PRED FORTE, G. VIGAMOX, G.FML, AND AFTER EXPLANT G PRED FORTE AND G TIMOCOMOD. PATIENT WILL HAVE 5 WEEK POST-OP ON (B)(6) 2018. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS REPLACED ON (B)(6) 2019 WITH A LONGER LENGTH LENS OF DIFFERENT POWER AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A LENS CASE/ VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. EVENT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.2, -03.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT, PIGMENT DISPERSION, AND PERSISTENT INFLAMMATION (IRITIS) WAS OBSERVED. ON (B)(6) 2018 THE LENS WAS REMOVED. IT WAS REPORTED THAT DURING INFLAMMATION DOCTOR PRESCRIBED G. PRED FORTE, G. VIGAMOX, G. FML, AND AFTER EXPLANT G. PRED FORTE AND G. TIMOCOMOD. PATIENT UNDER OBSERVATION; AC CELLS 2+. PATIENT WILL HAVE 5 WEEK POST-OP ON (B)(6) 2018. OTHER RELEVANT HISTORY- PATIENT HAD TUBERCULOSIS SIX YEARS AGO. RECENT BLOOD TEST SHOWS NORMAL. PATIENT CODE: "1940" NOT INCLUDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORD ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180440&pc=MTA | 2023826-2019-01902 | 2019-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. IT WAS REPORTED THAT TOXIC ANTERIOR SEGMENT SYNDROME (TASS) WAS OBSERVED IN THE PATIENT'S EYE. THE PATIENT WAS GIVEN STEROIDS AND THE SYMPTOMS GRADUALLY IMPROVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE INFLAMMATION HAD SUBSIDED AFTER APPROXIMATELY 3 MONTHS OF TREATMENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S) INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512618&pc=MTA | 2023826-2019-02433 | 2019-11-12 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS THAT TWO MONTHS AFTER SURGERY INFLAMMATION DEVELOPED. STEROIDS WERE ADMINISTERED AND CURRENTLY THE INFLAMMATION HAS SUBSIDED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512620&pc=MTA | 2023826-2019-02434 | 2019-11-12 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -5.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS THAT TWO MONTHS AFTER SURGERY INFLAMMATION DEVELOPED. STEROIDS WERE ADMINISTERED AND CURRENTLY THE INFLAMMATION HAS SUBSIDED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602192&pc=MTA | 2023826-2016-00522 | 2015-09-25 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO INFLAMMATION. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS HAIR. THERE WAS RESIDUE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559303&pc=MTA | 2023826-2017-00795 | 2017-04-10 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/089 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INFLAMMATION. UPON EXPLANTATION OF THE LENS, THE PROBLEM WAS RESOLVED, THE PATIENTS BCVA WAS 20/20 AND NO INFLAMMATION WAS OBSERVED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A ALCON VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (RED RESIDUE). DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577091&pc=MTA | 2023826-2017-00845 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Insufficient Information | Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE SURGEON IRRIGATED THE EYE ON (B)(6) 2017 DUE TO INFLAMMATION BUT WAS UNSUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN SMALL VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (CLEAR RESIDUE) AND THE LENS MISSING A PIECE OF HAPTIC. CORRECTED DATA: USER FACILITY IS THE INCORRECT REPORT SOURCE, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6559302&pc=MTA | 2023826-2017-00796 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Inflammation; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/093 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INFLAMMATION. UPON EXPLANTATION OF THE LENS, THE PROBLEM WAS RESOLVED, THE PATIENTS BCVA WAS 20/28 AND NO INFLAMMATION WAS OBSERVED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595429&pc=MTA | 2023826-2019-00769 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Inflammation; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO SUSPECTED INFLAMMATION. IN THE REPORTER'S OPINION, THE CAUSE OF THE EVENT IS THE DEVICE. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: "ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO SUSPECTED INFLAMMATION." CORRECTED TO "ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO SUSPECTED INFLAMMATION". (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389173&pc=MTA | 2023826-2017-00331 | 2016-11-01 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Inflammation; No Information; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/085 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED LOW VAULT. MULTIPLE ATTEMPTS HAVE BEEN MADE TO DETERMINE WHETHER THE LENS HAS BEEN EXPLANTED/EXCHANGED, STILL UNABLE TO CLARIFY. ACCORDING TO THE SURGEON, THE EYE IS "QUIET". Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) . Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) . Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016,THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS THEN EXCHANGED FOR A LENS, TWO SIZES LARGER, ON (B)(6) 2016. THE SURGEON HAD DIFFICULTY EXPLANTING THE LENS DUE TO PATIENT'S POOR PUPIL DILATION. SURGEON ALSO NOTED THAT IMPLANT OF THE LARGER LENS WAS ESPECIALLY CHALLENGING DUE TO THE PATIENT'S IRIS EDGE WAS SO WIDE IT "APPEARED LIKE A CAT'S EYE." (B)(4). CLAIM # (B)(4) .
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586135&pc=MTA | MW5092202 | 2019-11-03 | Injury | STAAR SURGICAL AG | 2020-01-13 | MTA | IMPLANTABLE COLLAMER LENS (NON-TORIC) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Increased Sensitivity; Visual Impairment |  | 1 | Event Description: BELOW IS AN EMAIL SENT TO STAAR SURGICAL ON 12/12/2019 ABOUT THEIR MEDICAL DEVICE, AFTER THE IMPLANTATION THE NON-TORIC ICL. I HAVE CALLED STAAR SURGICAL MULTIPLE TIMES AND NOT RECEIVED A RESPONSE ABOUT MY CONDITION OR BEEN ABLE TO GET IN TOUCH WITH MEDICAL MONITORING. EMAIL WITH DETAILS BELOW: ON (B)(6), 2019 I HAD STAAR SURGICAL ICL'S IMPLANTED IN BOTH EYES (NON-TORIC) AND AM HAVING MAJOR COMPLICATIONS. I HAVE CLOUDY VISION IN MY LEFT EYE AND MY PUPILS ARE NOW 2 DIFFERENT SIZES. I AM VERY CONCERNED AS MY SURGEON HAS NEVER SEEN THIS BEFORE. I HAD THE SURGERY PERFORMED AT (B)(6) WITH (B)(6). HAVE YOU SEEN THIS CONDITION BEFORE AND CAN IT BE TREATED? SURGERY WAS PERFORMED ON (B)(6) 2019 AND RESULTS WERE GOOD IMMEDIATELY AFTER. MY LAST DAY TAKING POST-OP STEROIDS WAS ON (B)(6) 2019. COMPLICATIONS BEGAN 9 DAYS POST SURGERY ON (B)(6) 2019 IMMEDIATELY AFTER A 3 HOUR FLIGHT. AFTER GETTING OFF THE PLANE, I HAD A RED RING AROUND MY LEFT EYE AND MY PUPILS WERE NOTICEABLY DIFFERENT SIZES. TWO DAYS LATER ON (B)(6) 2019, DEVELOPED PAIN AND LIGHT SENSITIVITY IN THE LEFT EYE SO I WENT TO A DR. THEY SAID I HAD REBOUND INFLAMMATION AND PUT ME BACK ON STEROIDS. AS THINGS WERE SEEMING TO GET BETTER, ON (B)(6) I HAD AN IMMEDIATE ONSET OF CLOUDINESS IN MY LEFT EYES. HALF OF MY VISION WENT GRAY AND I COULD NOT SEE. TO THIS DAY, I STILL HAVE 2 DIFFERENT SIZE PUPILS AND MY VISION IS CONSISTENTLY CLOUDY IN MY LEFT EYE. THE CLOUDINESS GETS MUCH WORSE WHEN SUNLIGHT IS HITTING MY EYE DIRECTLY, BUT SOME LEVEL OF CLOUDINESS IS ALWAYS THERE. I AM VERY WORRIED ABOUT THE HEALTH OF MY EYE AND I HOPE THAT STAAR SURGICAL CAN PROVIDE GUIDANCE. I'VE BEEN TO (B)(6) MANY TIMES SINCE COMPLICATIONS BEGAN, AND WAS TOLD MY EYES ARE HEALTHY AND THEY'VE NEVER SEEN THIS BEFORE IN ANY OTHER PT. I AM VERY WORRIED ABOUT THE CONDITION OF MY EYES WITH YOUR PRODUCT, PLEASE ADVISE NEXT STEPS. I STILL HAVE CLOUDINESS IN THE LEFT EYE AND 2 DIFFERENT SIZED PUPILS. FDA SAFETY REPORT ID# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9528000&pc=MTA | 2023826-2019-02499 | 2019-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Inflammation; Pain; Visual Disturbances; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED HAD AN MICL IMPLANTABLE COLLAMER LENS, IMPLANTED IN HER LEFT EYE (OS) ON (B)(6) 2019 WITH GOOD RESULTS. THE PATIENT REPORTED ON (B)(6) 2019, AFTER A 3 HOUR FLIGHT, THERE WAS A RED RING AROUND HER LEFT EYE AND HER PUPILS WERE DIFFERENT SIZES. TWO DAYS LATER, SHE DEVELOPED PAIN AND LIGHT SENSITIVITY IN THE LEFT EYE. THE PATIENT WENT TO HER DOCTOR AND THEY SAID THE PATIENT HAD REBOUND INFLAMMATION AND WAS PUT BACK ON STEROIDS. THE PATIENT HAS REPORTED CLOUDINESS AND HALF HER VISION WENT GRAY AND SHE COULD NOT SEE. TO THIS DAY, THE PATIENT HAS 2 DIFFERENT SIZE PUPILS AND VISION IN LEFT EYE IS CLOUDY. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED AN 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED "GRAY VISION TRANSIENT". THE CAUSE OF THE EVENT WAS UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14187607&pc=MTA | 2023826-2022-01182 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Red Eye(s); Eye Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICM5 12.1 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS HAD LOW VAULTING AND THERE WAS PIGMENT DISPERSION. THE IOP WAS CONTROLLED WITH TOPICAL ANTIGLAUCOMA EYEDROPS. THE PATIENT HAD PIGMENT DISPERSION SYNDROME 2 WEEKS POST ICL IMPLANTATION THAT COMPLETELY RESOLVED 2 MONTHS LATER. HOWEVER, SHE AGAIN HAD PIGMENTS IN A/C 1 MONTH LATER. PATIENT IS STILL UNDER CLOSE MONITORING. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- PATIENT EXPERIENCED INFLAMMATION, PAIN, AND REDNESS.- SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422283&pc=MTA | 2023826-2017-00442 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/091 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN FORTE (ALLERGEN). UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: UPDATED INFORMATION-ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6422303&pc=MTA | 2023826-2017-00443 | 2017-02-20 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Inflammation; Uveitis; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/109 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT THEN EXPERIENCED BLURRED VISION, INFLAMMATION, CYCLOPLEGIA AND BILATERAL ANTERIOR SEGMENT UVEITIS (TESTS CAME OUT NEGATIVE). PATIENT WAS TREATED WITH PREDNEFRIN COLLYRIUM. UCVA REMAINS AT 20/20. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED IN THE PATIENT AND THEREFORE, HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: ICL REMAINS IMPLANTED. PATIENT FINISHED WITH PREDNEFRIN FORTE, CURRENTLY IS NOT ON MEDICATION. THE EYE IS WHITE WITH NO ACTIVITY OR INFLAMMATION. BOTH SURGEON AND PATIENT ARE HAPPY WITH THE RESULTS. (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. PATIENT CODE (B)(4) NO CODE AVAILABLE (CYCLOPEGIA, TREATMENT WITH PREDNEFRIN COLLIRIUM) WORK ORDER SEARCH PERFORMED-NO SIMILAR COMPLAINTS FOUND. (B)(4). DEVICE REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7275463&pc=MTA | 2023826-2018-00235 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Inflammation; Uveitis; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A V4 IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENTS EYE 1 YEAR AGO, DUE TO SHALLOW VAULT, WITH RECURRENT UVEITIS. THE REPORTER INDICATED IMPLANTING ANOTHER ICL LENS WAS NOT FEASIBLE IN VIEW OF RECURRENT INFLAMMATORY EPISODES, AND THE PATIENT WAS PRESCRIBED CONTACT LENS. Manufacturer Narrative: THE LENS WAS EXPLANTED AFTER BEING IMPLANTED IN THE PATIENT'S EYE FOR ONE YEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983737&pc=MTA | 2023826-2022-01034 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Inflammation; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES (ICA), BLEPHARITIS, IRIS BOMBE WITH NO APPARENT TOUCH POINT, "VH 1," AND NIGHTTIME HALOS IN PERIPHERY WHEN DRIVING AND ANY CEILING LIGHTS. THE SURGEON STATES, "VERY NARROW ANGLE...ANGLE SHUT OFF." THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO BIG." Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED, AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214361&pc=MTA | 2023826-2021-02693 | 2021-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Insufficient Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: B5: PER NEW INFORMATION RECEIVED ON 12-APR-2022: CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215166&pc=MTA | 2023826-2022-01259 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Insufficient Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215058&pc=MTA | 2023826-2022-01258 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Insufficient Information | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808870&pc=MTA | 2023826-2021-04335 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Insufficient Information | P030016 | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON WILL BE DOING A LENS EXCHANGE ON A PATIENT WHO HAS AN IRIS CYST. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PER UPDATED INFORMATION THE DOCTOR STATED THAT THE CYST WAS SOMETHING THE PATIENT HAD AND WAS NOT CAUSED BY THE ICL. LENS WAS REMOVED DUE TO SIZING, 200% VAULT. PATIENT IS DOING GOOD. CLAIM # (B)(4). Manufacturer Narrative: B5 - "A SURGEON EMAIL ON 06-JUN-2022 STATES THE PATIENT WAS RE-OPERATED WITH AN UNKNOWN ICL PLACED HORIZONTALLY. THE PATIENT VISUAL ACUITY IS NOTED AS 'SHE IS DOING VERY WELL WITH 20/20 UNCORRECTED VISION IN BOTH EYES.'" SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: AGE - RACE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11536790&pc=MTA | 2023826-2021-00735 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Insufficient Information | NA | 1 | Event Description: A SPHERICAL LENS WAS RETURNED DAMAGED. PATIENT CONTACT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE ON PRODUCT, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BENT/DEFORMED AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214224&pc=MTA | 2023826-2021-02653 | 2021-01-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Insufficient Information | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15054858&pc=MTA | 2023826-2022-02184 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Insufficient Information | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6: MEDICAL DEVICE PROBLEM CODE: 1494 OFF-LABEL USE, ACD < 3.0MM. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11853686&pc=MTA | 2023826-2021-01570 | 2021-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Insufficient Information; No Clinical Signs, Symptoms or Conditions | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6428219&pc=MTA | 2023826-2017-00467 | 2017-02-22 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Infection; Visual Impairment; Vitrectomy; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. BEGINNING ON (B)(6) 2017 THE PATIENT BEGAN EXPERIENCING THE FOLLOWING: LOW VAULT, BCVA OF "ONLY HAND MOTIONS," DECREASED ACCOMMODATION, AND DIFFICULTY IN NEAR VISION. THE PATIENT UNDERWENT A VITRECTOMY, THEN WAS HOSPITALIZED FOR APPROXIMATELY 1 WEEK. PATIENT HAD AN INFECTION-SUSPECTED ENDOPHTHALMITIS BUT TESTS CAME OUT NEGATIVE. TESTED POSITIVE HOWEVER, FOR "FUNGAL-LIKE" ORGANISMS AND GRAM POSITIVE COCCI. PATIENT WAS TREATED WITH IV ANTIBIOTICS AND DROPS, PPV. LENS WAS EXPLANTED (B)(6) 2017. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4). Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT AND CLEAR RESIDUE AND FIBERS ON LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IS INCORRECT, CORRECT IMPLANTABLE COLLAMER LENS USED WAS 12.6MM VICMO12.6. MODEL VICMO13.2 IS INCORRECT, CORRECT DATA IS VICMO12.6. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317683&pc=MTA | 2023826-2022-00196 | 2020-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2020. THE SURGEON REPORTS ELEVATED IOP; CONCOMITANT MEDICATIONS PRESCRIBED. REPORTEDLY THE PROBLEM IS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS ELEVATED IOP; CONCOMITANT MEDICATIONS PRESCRIBED. REPORTEDLY THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278995&pc=MTA | 2023826-2023-02608 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(51MMHG)WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IOP WAS TREATED WITH TABLETS OVER TIME. PRESSURE KEPT INCREASING WITHOUT THERAPY(RX) SO SURGEON DECIDED TO EXPLANT. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH IMPACT CODE: "4644" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE IOP WAS TREATED WITH TABLETS OVER TIME. BECAUSE THE PRESSURE KEPT INCREASING WITHOUT THERAPY THE SURGEON DECIDED TO FINALLY EXPLANT THE LENS". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17278997&pc=MTA | 2023826-2023-02609 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP(34MMHG)WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6) 2021) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IOP WAS TREATED WITH TABLETS OVER TIME. PRESSURE KEPT INCREASING WITHOUT THERAPY(RX) SO SURGEON DECIDED TO EXPLANT. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH IMPACT CODE: "4644" SHOULD BE ADDED. B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND MEDICATION WAS PRESCRIBED. REPORTEDLY "THE IOP WAS TREATED WITH TABLETS OVER TIME. BECAUSE THE PRESSURE KEPT INCREASING WITHOUT THERAPY THE SURGEON DECIDED TO FINALLY EXPLANT THE LENS". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840569&pc=MTA | 2023826-2021-04398 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4). Manufacturer Narrative: REPORTEDLY THE LENS WAS EXPLANTED (B)(6) 2021. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840585&pc=MTA | 2023826-2021-04399 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS ELEVATED IOP. REPORTEDLY A WASHOUT WAS PERFORMED, DIAMOX, IOPIDINE, AND IATANOPROST WERE PRESCRIBED. STEROID TREATMENT WAS STOPPED AND NSAID (BROMEGENAC) WAS STARTED WHICH LOWERED THE IOP INITIALLY BUT IT WAS RAISED AGAIN. Manufacturer Narrative: B5: UPDATED INFORMATION: REPORTEDLY THE LENS WAS EXPLANTED ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: H6- INVESTIGATION TYPE: 4110- LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200099&pc=MTA | 2023826-2023-00234 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -16.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE OF 39 MMHG. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464622&pc=MTA | 2023826-2023-00764 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6; -16.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS ELEVATED INTRAOCULAR PRESSURE 35 MMHG; WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE SURGEON NOTES "THE PATIENT DID NOT WANT TO UNDERGO A SECONDARY LENS IMPLANTATION PROCEDURE". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672709&pc=MTA | 2023826-2023-03650 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT WAS TREATED WITH INTRAOCULAR PRESSURE LOWERING DRUGS (CAPTOPROLOL EYE SOLUTION+BROMONIDINE EYE SOLUTION). ONE MONTH AFTER SURGERY, THE IOP WAS 25.0MMHG. AFTER ADDING ORAL INTRAOCULAR PRESSURE LOWERING DRUGS, THE IOP REMAINED HIGH BY ABOUT 25-30MMHG. AT LAST POST-OP VISIT ON (B)(6) 2023, THE IOP WAS 20MMHG AND PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - CAPTOPROLOL EYE SOLUTION; BROMONIDINE EYE SOLUTION; ORAL INTRAOCULAR PRESSURE LOWERING DRUGS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17680175&pc=MTA | 2023826-2023-03651 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND WAS TREATED WITH INTRAOCULAR PRESSURE LOWERING DRUGS (CAPTOPROLOL EYE SOLUTION+BROMONIDINE EYE SOLUTION). ONE MONTH AFTER SURGERY, THE IOP WAS 25.0MMHG. AFTER ADDING ORAL INTRAOCULAR PRESSURE LOWERING DRUGS, THE IOP DROPPED TO A NORMAL LEVEL AND ROUTINE FOLLOW-UP EXAMINATIONS WERE CARRIED OUT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: NA. H6: HEALTH EFFECT IMPACT CODE: 4644 - CAPTOPROLOL EYE SOLUTION; BROMONIDINE EYE SOLUTION; ORAL INTRAOCULAR PRESSURE LOWERING DRUGS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581257&pc=MTA | 2023826-2023-03593 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) .ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS ROTATED VERTICALLY AND THE PROBLEM RESOLVED. SURGEON STATES " DOING GREAT, OFF ALL DROPS". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581260&pc=MTA | 2023826-2023-03592 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF _7.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) .ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS ROTATED VERTICALLY AND THE PROBLEM RESOLVED. SURGEON STATES " LOOKS GREAT, BELOW BASELINE, OFF ALL DROPS". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777978&pc=MTA | 2023826-2023-04108 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE. THE REPORTER STATED " PREOPERATIVE BASAL IOP WAS HIGHER THAN NORMAL, AND NO FURTHER EXAMINATION WAS PERFORMED TO EXCLUDE GLAUCOMA. AFTER THE CRYSTAL IS REMOVED, THE INTRAOCULAR PRESSURE IS TEMPORARILY STABILIZED TO NORMAL VALUE AFTER DRUG TREATMENT. THE LENS WAS REMOVED AND TRANSFERRED TO THE GLAUCOMA DEPARTMENT FOR FURTHER CONSULTATION." Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346570&pc=MTA | 2023826-2023-05625 | 2023-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (30MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON 26/OCT/2023) ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437835&pc=MTA | 2023826-2024-00032 | 2023-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, ELEVATED IOP WITHOUT PUPIL BLOCK WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS REMOVED. REPORTEDLY, "LEFT EYE ANTERIOR CHAMBER STILL HAS DENCE AC CELLS AND INV PIGMENTS DESPITE FREQUENT TOMZAL STERIOD." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10979857&pc=MTA | 2023826-2020-02956 |  | Injury | STAAR SUGICAL COMPANY | 2020-12-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENTS EYE. THE PATIENT REPORTS ELEVATED IOP DUE TO STEROID USER AFTER SURGERY. IT WAS REPORTED THAT MEDICAL INTERVENTION (DROPS) WERE PRESCRIBED AND IT QUICKLY STABILIZED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; SUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8678993&pc=MTA | 2023826-2019-01027 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-06-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF ADVANCES IN MEDICINE AND MEDICAL RESEARCH IN MARCH 2019 TITLED, 'TORIC AND PHAKIC IOLS FOR THE TREATMENT OF ASTIGMATISM AND/OR HIGH MYOPIA: OUR EXPERIENCE AT PRINCE HASHEM HOSPITAL ZARQA.' THE ARTICLE STATES THAT ONE PATIENT HAD HIGH INTRAOCULAR PRESSURE OF 32 MMHG AFTER ICL IMPLANTATION IN BOTH EYES AND THIS WAS TREATED WITH ANTI-GLAUCOMA DROPS AND ACETAZOLAMIDE 250 MG TABLETS. THE PATIENT WAS FOUND TO BE STEROID RESPONDER AS HER INTRAOCULAR PRESSURE WENT DOWN AFTER SHIFTING TO PRESSURE SPARING STEROID DROPS LOTEMAX. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846156&pc=MTA | 2023826-2019-01294 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-07-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE INDIAN JOURNAL OF OPHTHALMOLOGY IN JULY 2019 TITLED, 'COMPARATIVE ANALYSIS OF CLINICAL OUTCOMES BETWEEN TWO TYPES OF POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENSES FOR CORRECTION OF MYOPIA AND MYOPIC ASTIGMATISM.' THE ARTICLE STATES THAT "TWO EYES (0.98%) IN THE ICL GROUP DEVELOPED MODERATELY HIGH IOP RISE SECONDARY TO STEROID RESPONSE. THEY WERE MANAGED CONSERVATIVELY WITH ANTIGLAUCOMA MEDICATIONS, FOLLOWING WHICH IOP NORMALIZED." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL #: UNK. IMPLANTED: UNK. EXPLANTED: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476063&pc=MTA | 2023826-2020-01919 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE PATIENT WAS REPORTED WITH HIGH INTRAOCULAR PRESSURE (IOP) AT 2ND DAY POST OP. THE VAULT WAS REASONABLE CENTRALLY BUT WAS SHALLOW PERIPHERALLY. THE IOP HAS DECREASED TO A MORE NORMAL LEVEL. THE SURGEON ATTRIBUTED THE IOP ISSUE TO RETAINED VISCOELASTIC AND PIS THAT WEREN'T PATENT ENOUGH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476081&pc=MTA | 2023826-2020-01920 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE PATIENT WAS REPORTED WITH HIGH INTRAOCULAR PRESSURE (IOP) AT 2ND DAY POST OP. THE VAULT WAS REASONABLE CENTRALLY BUT WAS SHALLOW PERIPHERALLY. THE IOP HAS DECREASED TO A MORE NORMAL LEVEL. THE SURGEON ATTRIBUTED THE IOP ISSUE TO RETAINED VISCOELASTIC AND PIS THAT WEREN'T PATENT ENOUGH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551356&pc=MTA | 2023826-2020-02124 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: AN ARTICLE PUBLISHED IN SPRINGER NATURE WAS RECEIVED ON 21-AUG-2020, ENTITLED 'VISUAL OUTCOMES AFTER REFRACTIVE LASER CORNEAL SURGERY AND PHAKIC IOL IN AMBLYOPIC EYES.' THE ARTICLE STATES THAT AFTER IMPLANTATION OF PHAKIC IOLS, 4 EYES HAD INCREASED INTRAOCULAR PRESSURE DURING THE IMMEDIATE POST-OP PERIOD. REPORTEDLY, THESE WERE "CONSERVATIVELY MANAGED." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5- ADDITIONAL INFORMATION: REPORTEDLY THE IOP RISE IS NOT BEYOND ANY STANDARD GUIDELINE AND WAS MAINLY DUE TO VISCOELASTIC. IOP CAME DOWN WITHIN A DAY OR TWO FOR ALL EYES. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. PRODUCT CODE-UNK. ESUBMITTER DOES NOT ALLOW FOR UNK OR BLANK ENTRY DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10898536&pc=MTA | 2023826-2020-02758 |  | Injury | STAAR SURGICAL COMPANY | 2020-11-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS ELEVATED IOP AND ADDITION OF SURGICAL PI'S. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11620586&pc=MTA | 2023826-2021-00862 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS ELEVATED IOP "SEVERAL" MONTHS POSTOP. CURRENTLY THE PATIENT IS ON AN IOP-LOWERING TOPICAL AGENT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A12.6MM MICL12.6 ICL WAS IMPLANTED ON (B)(6) 2020. THE PATIENT IS CURRENTLY ON LATANOPROST, THE SURGEON WILL CONTINUE TO OBSERVE. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11654123&pc=MTA | 2023826-2021-00954 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS ELEVATED IOP "SEVERAL" MONTHS POSTOP. CURRENTLY THE PATIENT IS ON AN IOP-LOWERING TOPICAL AGENT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE PATIENT IS CURRENTLY ON LATANOPROST. THE SURGEON WILL CONTINUE TO OBSERVE THE PATIENT. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887005&pc=MTA | 2023826-2021-01617 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS EYE. THE REPORTER INDICATED AT ONE DAY POST-OP, THE INTRAOCULAR PRESSURE (IOP) WAS 50MMHG. THE PATIENT WAS PRESCRIBED DIAMOX AND TIMOLOL WITH PRESSURE OF 14 TO 17MMHG AND THE PIS WERE ENLARGED. LATER THAT SAME DAY, THE PRESSURE WAS NORMAL (17MMHG) AND THE ANGLES WERE OPENED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: EXPIRATION DATE UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677889&pc=MTA | 2023826-2021-04042 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | NA | 1 | Event Description: AN ARTICLE (ONE-STEP VISCOELASTIC AGENT TECHNIQUE FOR ICL V4C IMPLANTATION FOR MYOPIA) WAS RECEIVED THAT CITED A CASE STUDY OF 100 EYES (52 PATIENTS) OF PATIENTS IMPLANTED WITH HOLE ICLS. OF THE 100 EYES, FOUR EYES HAD INTRAOCULAR PRESSURE (IOP) EXCEEDING 30MMHG THREE HOURS AFTER SURGERY, AND CORNEAL PARACENTESIS WAS CONDUCTED. THROUGH ROUTINE POSTOPERATIVE MEDICATION, IOP OF ALL THE EYES DECREASED TO BASELINE VALUE AT 1 DAY POST-OP. THE LENSES REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE : UNK. EXPLANT DATE : UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. DEVICE MANUFACTURE DATE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082438&pc=MTA | 2023826-2021-04957 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AFTER IMPLANTING AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE, THE IOP WENT TO 50. REPORTEDLY, THIS WAS "REMEDIED WITH MORE ENERGY IN PIS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT AFTER IMPLANTING AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE, THE IOP WENT TO 50. REPORTEDLY, THIS WAS "REMEDIED WITH MORE ENERGY IN PIS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: SECTION A- UNK. DATE OF EVENT- UNK. DATE OF REPORT- UNK. DEVICE MANUFACTURE DATE- UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488086&pc=MTA | 2023826-2022-00427 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6; -8.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE. THE PATIENT HAD INTRAOCULAR PRESSURE SPIKES AND THE SURGEON HAD TO PERFORMED AN EMERGENCY PERIPHERAL IRIDOTOMY. IT WAS ALSO NOTED THEY BELIEVE THE LENS IS TOO LARGE FOR THE PATIENT'S EYE AND CALCULATIONS WILL BE PERFORMED TO ORDER A NEW LENS SIZE BEFORE IMPLANTING INTO THE PATIENT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14524996&pc=MTA | 2023826-2022-01696 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. ELEVATED INTRAOCULAR PRESSURE WAS REPORTED. THE SURGEON ENLARGED THE PI; AND IT WAS NOTED THE PATIENT WAS MEDICALLY TREATED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15212772&pc=MTA | 2023826-2022-02704 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: ON (B)(6) 2022 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED, ENTITLED, 'ANALYSIS OF MICROCIRCULATION CHANGES IN THE MACULAR AREA AND PARA-OPTIC COLLAMER LENS IMPLANTATION IN PATIENTS WITH HIGH MYOPIA'. THE ARTICLE REPORTS 1 CASE OF IOP > 21MMHG ON THE FIRST DAY POST-OP WHICH WAS TREATED WITH IOP-LOWERING DRUGS AND 1 CASE OF ICL ROTATION TO VERTICAL POSITION DUE TO VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621320&pc=MTA | 2023826-2022-03646 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THE PATIENT HAD SIGNIFICANT PRESSURE SPIKES AND IOP IS -50S. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621416&pc=MTA | 2023826-2022-03647 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THE PATIENT HAD SIGNIFICANT PRESSURE SPIKES AND IOP IS -50S. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15835038&pc=MTA | 2023826-2022-03142 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S LEFT EYE (OS). DROPS WERE GIVEN TO REDUCED EDEMA AND HIGH IOP. GLAUCOMA MEDS WERE ALSO USED AND BURPED THE WOUND. HIGH VAULT WITH NARROW ANGLES AND DISCOMFORT WERE ALSO REPORTED. Manufacturer Narrative: PATIENT WEIGHT AND RACE: UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15835116&pc=MTA | 2023826-2022-03141 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S RIGHT EYE (OD). DROPS WERE GIVEN TO REDUCED EDEMA AND HIGH IOP. GLAUCOMA MEDS WERE ALSO USED AND BURPED THE WOUND. HIGH VAULT WITH NARROW ANGLES WERE ALSO REPORTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965599&pc=MTA | 2023826-2022-04581 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON HAD 28 IMPLANTABLE COLLAMER LENS COMPLAINTS REGARDING PATIENTS WHO HAD ELEVATED INTRAOCULAR PRESSURE; HALF HAD PRESSURE OF 30MMHG OR HIGHER. A PORTION OF THE PATIENTS REQUIRED EXTENSIVE TREATMENT AND OBSERVATION SUCH AS BURPING THE PARACENTESIS OVER A PERIOD OF HOURS; ACETAZOLAMIDE AND DROPS. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16147833&pc=MTA | 2023826-2023-00057 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased |  | 1 | Event Description: IN THE ARTICLE, "EARLY OUTCOMES OF ANTERIOR SEGMENT PARAMETERS AFTER IMPLANTABLE COLLAMER LENS V4C IMPLANTATION," WITHIN THE EYES STUDIED: 3 EYES IN THE POSTOPERATIVE IOP AT 2H WAS 21 TO 30MMHG, WHICH WAS CONTROLLED BY CARTEOLOL EYE DROPS UNTIL NORMAL, 5 EYES WITH IOP LARGER THAN 30 MMHG AND WERE IMMEDIATELY LOWERED TO LESS THAN 15 MMHG VIA PRIMARY INCISON DRAINAGE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102367&pc=MTA | 2023826-2023-02187 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP) 35MMHG. 3 TYPES OF GLAUCOMA MEDS WERE PRESCRIBED. STEROIDS WERE STOPPED ON (B)(6) 2023. IOP WAS 34MMHG ON ALPHAGAN DROPS ON (B)(6) 2023. GANFORT WAS ADDED. AT THE (B)(6) 2023 POST-OP VISIT, THE IOP WAS 11 BE. THE DOCTOR STOPPED ONE OF THE GLAUCOMA MEDS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH EFFECT IMPACT CODE: 4644 - GLAUCOMA MEDS (ALPHAGAN DROPS, GANFORT). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17108733&pc=MTA | 2023826-2023-02188 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE PATIENT HAD PERSISTENT ELEVATED INTRAOCULAR PRESSURE (IOP) 35MMHG. 3 TYPES OF GLAUCOMA MEDS WERE PRESCRIBED. STEROIDS WERE STOPPED ON (B)(6) 2023. IOP WAS 34MMHG ON ALPHAGAN DROPS ON (B)(6) 2023. GANFORT WAS ADDED. AT THE (B)(6) 2023 POST-OP VISIT, THE IOP WAS 11 BE. THE DOCTOR STOPPED ONE OF THE GLAUCOMA MEDS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6B: UNK. H4: UNK. H6: HEALTH EFFECT IMPACT CODE: 4644 - GLAUCOMA MEDS (ALPHAGAN DROPS, GANFORT). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378752&pc=MTA | 2023826-2023-03108 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: A COPY OF A RETROSPECTIVE, OBSERVATIONAL STUDY WAS FOUND BY STAAR SURGICAL COMPANY, DATED (B)(6) 2023. THE STUDY WAS IN REGARD TO "REFRACTIVE OUTCOMES OF POSTERIOR CHAMBER PHAKIC (SPHERIC AND TORIC IMPLANTABLE COLLAMER LENS) INTRAOCULAR LENS IMPLANTATION WITH A CENTRAL PORT DESIGN." A TOTAL OF 254 EYES OF 134 PATIENTS UNDERWENT ICL V4C MODEL AND TORIC ICL V4C MODEL IMPLANTATION. COMPLICATION RATE OF 12.99% BEING ELEVATION OF INTRAOCULAR PRESSURE (IOP) THE MOST COMMON OBSERVED IN NINETEEN EYES, ALL OF WHICH RESPONDED WELL TO TOPICAL ANTIHYPERTENSIVES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK A3: UNK A4: UNK A5: UNK A6: UNK. B3: UNK. D2: COMMON DEVICE NAME: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL ANTIHYPERTENSIVES. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15221421&pc=MTA | 2023826-2022-02751 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 25-MAY-2022. ON 07-JUL-2022: ELEVATED IOP 50MMHG, PIGMENT DISPERSION, AND PRO-LONGED INFLAMMATION WAS RECORDED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON (B)(6) 2022. PROBLEM WAS RESOLVED. REPORTEDLY "MINIMAL PIGMENTS, MORE QUITE EYE." CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426111&pc=MTA | 2023826-2021-03441 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE SURGEON REPORTS ELEVATED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE - UNK. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355865&pc=MTA | 2023826-2016-00029 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -18.00/+1.50/X084 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00030 FOR LEFT EYE. Manufacturer Narrative: CONCLUSION: UNABLE TO CONFIRM COMPLAINT - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -18.00/+1.50/X084. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5202980&pc=MTA | 2023826-2015-01469 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED A SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS AND THE PATIENT WAS EXPERIENCING INCREASED ELEVATED INTRAOCULAR PRESSURE (40MMHG). THE PATIENT WAS PRESCRIBED DROPS AND THE INTRAOCULAR PRESSURE DECREASED TO 18MMHG. THE IRIDOTOMIES WERE PATENT AND THE CHAMBER WAS SHALLOW. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, BY NOT IMPLANTING DOCTOR, PATIENT PRESENTED WITH ELEVATED IOP (40MMHG) FOLLOWING ICL IMPLANTATION. SOME EYE DROPS WERE PRESCRIBED AND INTRAOCULAR PRESSURE DECREASED TO 18 MMHG. THE DOCTOR ALSO REPORTED THAT IRIDOTOMIES WERE PATENT BUT ANTERIOR CHAMBER APPEARED SHALLOW. ICL REMAINS IMPLANTED AND THE PATIENT WAS SCHEDULED FOR ANOTHER VISIT ON A LATER DATE. IT SHOULD BE NOTED THAT INADEQUATE FLUSHING OF VISCOELASTIC MAY LEAD TO IOP SPIKES AS WRITTEN UNDER " PRECAUTIONS" IN THE DFU. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6059231&pc=MTA | 2023826-2016-01500 |  | Injury | STAAR SURGICAL COMPANY | 2016-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED A COMPLICATION WITH PATIENT IMPLANTED WITH AN MICL COLLAMER LENS. THE EVENT MAY HAVE BEEN RELATED TO INCREASED INTRAOCULAR PRESSURE THAT DIDN'T HAVE APPROPRIATE YAG PERIPHERAL IRIDECTOMIES AS THE CAUSE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14092478&pc=MTA | 2023826-2022-01113 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO PRESSURE. "PRESSURE IS CONTROLLED UNDER MAXIMUM MEDS IN ONE EYE AND THE OTHER EYE STILL HAS PRESSURE OF 34 DEGREES UNDER MAXIMUM MEDS." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14092585&pc=MTA | 2023826-2022-01112 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL 12.6 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO PRESSURE. "PRESSURE IS CONTROLLED UNDER MAXIMUM MEDS IN ONE EYE AND THE OTHER EYE STILL HAS PRESSURE OF 34 DEGREES UNDER MAXIMUM MEDS." NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16001450&pc=MTA | 2023826-2022-04427 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-18 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT IN THE IMMEDIATE POST-OP AFTER IMPLANTATION OF AN EVO IMPLANTABLE COLLAMER LENS, A PATIENT'S IOP WAS IN THE 30S. THE SURGEON RELEASED PRESSURE AT A PARACENTESIS INCISION, ONE OF THE ICL HAPTICS MOVED TO THE ANTERIOR CHAMBER. THE SURGEON REPLACED THE INITIAL LENS WITHOUT INCIDENT. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6: HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505274&pc=MTA | 2023826-2017-00666 | 2017-03-08 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE PATIENT BEGAN TO EXPERIENCE ELEVATED IOP. THE SURGEON STATED THAT HE BELIEVED THE HIGH IOP WAS CAUSED BY THE PATIENT'S STEROID (DEXAMETHASONE) TREATMENT. TO DATE, THE LENS REMAINS IMPLANTED. SURGEON STATED THAT "AFTER MEDICATION" THE IOP WENT DOWN TO NORMAL VALUE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641024&pc=MTA | 2023826-2017-00975 | 2017-05-17 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICM5_12.1 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED ELEVATED IOP. THE REPORTER ALSO MENTIONED THAT THE SURGEON CHANGED THE MEDICATIONS EYEDROPS WITHOUT CORTISONE TO FIND OUT IF THE HIGH IOP IS THE OUTCOME OF GIVING CORTISONE EYEDROPS. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. REFERENCE MDR NUMBER 2023826-2017-00701 FOR THE SAME PATIENT'S EYE. Manufacturer Narrative: SINCE THE PHYSICIAN MODIFIED THE MEDICATION, THE IOP IS REPORTED AS NORMAL. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/20. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT, LENS IMPLANT IN A PSEUDO PHAKIC EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6632436&pc=MTA | 2023826-2017-00940 |  | Injury | STAAR SURGICAL COMPANY | 2017-06-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL 13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, WAS EXPLANTED DUE TO PATIENTS INTRAOCULAR PRESSURE BEING HIGH. THE REPORTER INDICATED THAT THE EVENT CAUSE COULD BE DUE TO THE LENS BEING BIG. ADDITIONAL INFORMATION WAS REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID , IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6952443&pc=MTA | 2023826-2017-01570 |  | Injury | STAAR SURGICAL COMPANY | 2017-10-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A MICL12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THREE WEEKS AFTER THE IMPLANTATION, THE IOP WAS NOTED TO BE 29. THE IOP WAS SLIGHTLY ASYMMETRIC A WEEK PRIOR BY 5 BUT WITHIN THE NORMAL LIMITS. GONIOSCOPY WAS COMPLETELY OPEN. THE VAULT WAS 400. THE PI'S WERE PATENT AND FUNCTIONING. DILATED EXAM WAS NORMAL AND NO DROP IN IOP FOLLOWING DILATION. THE PATIENT WAS ON BID PRED FORTE AND THIS WAS THOUGHT TO BE THE CAUSE ACCORDING TO THE SURGEON. AFTER TWO DAYS OF DISCONTINUING THE PRED FORTE, THE PATIENT'S IOP WAS 17. THE PROBLEM WAS RESOLVED. UCVA IS 20/20, IOP NORMALIZED, AND PATIENT IS HAPPY. Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066047&pc=MTA | 2023826-2015-01174 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/-1.50/089 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01175 FOR LEFT EYE. Manufacturer Narrative: THERE WAS NO EXCESSIVE VAULTING. THE INTRAOCULAR PRESSURE RISE WAS UNDER CONTROL AND THE PROBLEM WAS RESOLVED. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. (B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.50/+1.50/089. (B)(4). METHOD:EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4740585&pc=MTA | 2023826-2015-00486 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY AG | 2015-04-30 | MTA | VISIAN ICL (IMPLANTABLE COLLAME LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST KNOWN UCVA WAS 20/25 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. REDUCTION OF IRIDO-CORNEAL ANGLES. EXPLANTED. VAULTING. DEVICE EVALUATED BY THE MANUFACTURER? NO: LENS WAS NOT RETURNED. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH AND MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT SHOWED THE LENS WAS RETURNED DRY WITH NO VISIBLE DAMAGE. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT THE INADEQUATE VAULT IS A CONSEQUENCE OF WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENTS, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST, ECT). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP, ACT.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSION: BASED ON THE COMPLAINT INFORMATION, WORK ORDER SEARCH, PRODUCT EVALUATION AND MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF THE INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5981810&pc=MTA | 2023826-2016-01321 | 2016-04-18 | Injury | STAAR SURGICAL COMPANY | 2016-09-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VICMO 13.2 IMPLANTABLE COLLAMER LENS - -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6900441&pc=MTA | 2023826-2017-01484 | 2017-08-06 | Injury | STAAR SURGICAL COMPANY | 2017-09-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/082 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (CONTROLLED BY MEDICATION). THE LENS REMAINS IMPLANTED BUT THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SMALLER SIZE LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. EXPLANTED DATE IS CORRECTED TO "N/A" FROM "UNK" IN THE INITIAL; AS THE LENS REMAINS IMPLANTED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824489&pc=MTA | 2023826-2018-01252 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13361443&pc=MTA | 2023826-2022-00301 | 2020-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10634642&pc=MTA | 2023826-2020-02293 | 2020-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-10-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13295627&pc=MTA | 2023826-2022-00213 | 2020-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13941313&pc=MTA | 2023826-2022-00951 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. "HIGH VAULT DESPITE APPROPRIATELY SIZED ICL BASED ON WTW. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822759&pc=MTA | 2023826-2021-01436 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VICM5_13.7, -13.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. YES BECAUSE THE INITAL LENS WAS TOO LONG AND THE VAULT TOO HIGH". Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119262&pc=MTA | 2023826-2021-02299 | 2021-03-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233847&pc=MTA | 2023826-2021-02882 | 2021-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (23MMHG) WITH OUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS INCREASED IOP, MADE AC SHALLOWER. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE LENS WAS EXCHANGED ON (B)(6) 2021. "SURGERY WENT WELL, PATIENT HAPPY!" VISION IS 20/25, IOP WITHIN NORMAL RANGE. H6: INVESTIGATION FINDING 114 WAS REPORTED PREVIOUSLY IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16917178&pc=MTA | 2023826-2023-01761 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939026&pc=MTA | 2023826-2021-01671 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (33MMHG, ASSESSED ON (B)(6) 2021) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, PT SYMPTOMS: FOREIGN BODY SENSATION AND ITCHINESS. OBJECTIVE SLE: CORNEA: 2+ DMF. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A SPECIMEN. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155775&pc=MTA | 2023826-2021-02533 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON PERFORMED ADDITIONAL LASER SURGERY TO THE IRIDOTOMY (ENLARGED) AND THAT RESOLVED THE ISSUE. THE PATIENT IS FINE AND THERE WAS NO NEED TO EXPLANT THE LENS, THE LENS REMAINED IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK.. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. ENLARGED IRIDOTOMY BY LASER. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308409&pc=MTA | 2023826-2021-03077 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5 13.7, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ELEVATED IOP AS REDUCED TO 15MMHG BY TREATING WITH DIAMOX. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12399422&pc=MTA | 2023826-2021-03396 | 2021-07-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. THE PATIENT WILL BE MONITORED STRICTLY AND THE IOP HAS DECREASED SLIGHTLY DURING THE LAST WEEKS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309724&pc=MTA | 2023826-2022-00221 | 2021-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.1 VICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309095&pc=MTA | 2023826-2022-00219 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: D6A- IMPLANT DATE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13309644&pc=MTA | 2023826-2022-00220 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT A 12.6 VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13267274&pc=MTA | 2023826-2022-00120 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13267305&pc=MTA | 2023826-2022-00121 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256669&pc=MTA | 2023826-2022-00118 | 2021-11-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: RACE-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15501957&pc=MTA | 2023826-2022-03342 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -15.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BORDERLINE IOP IMPROVING- STILL ON GTTS." Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15503201&pc=MTA | 2023826-2022-03373 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED "BORDERLINE IOP IMPROVING- STILL ON GTTS." Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16455859&pc=MTA | 2023826-2023-00761 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -14.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE 28 MMHG; WITHOUT PUPIL BLOCK. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/ SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE "ALL FINE!" AND "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16554736&pc=MTA | 2023826-2023-00869 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THE EYE IS FINE AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433266&pc=MTA | 2023826-2023-00524 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741438&pc=MTA | 2023826-2023-03980 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.25 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN WITH FIBRE ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17754907&pc=MTA | 2023826-2023-03981 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND OPTIC TORN/DEFORMED AND HAPTIC BROKEN/STRETCHED OUT WITH FIBRE ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831137&pc=MTA | 2023826-2023-04215 | 2023-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -04.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND HIGH IOP. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DOCTOR CHOSE A DIFFERENT ICL BY ERROR FROM HIS STOCK, SO USER HAD TO REPLACE IT SECOND DAY POST-OP. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360955&pc=MTA | 2023826-2023-05558 | 2023-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7835563&pc=MTA | 2023826-2018-01299 |  | Injury | STAAR SURGICAL COMPANY | 2018-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: RECEIVED A REPORT WHERE THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS. THE LENS WAS REPORTED AS HAVING ELEVATED VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE UNK. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611940&pc=MTA | 2023826-2019-00825 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, MICL 13.7, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. HIGH PRESSURE AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2018. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11616800&pc=MTA | 2023826-2021-00841 |  | Injury | STAAR SURGICAL COMPANY | 2021-04-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HIGH VAULTING AND INCREASED INTRAOCULAR PRESSURE (IOP). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281694&pc=MTA | 2023826-2021-03061 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, PATIENT HAS DONE BETTER WITH EXCHANGE AND HAS BETTER IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281698&pc=MTA | 2023826-2021-03060 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -7.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP AND EXCESSIVE VAULT ON (B)(6) 2021. REPORTEDLY, PATIENT HAS DONE BETTER WITH EXCHANGE AND HAS BETTER IOP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: SEX- UNK, WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE OF EVENT- UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794337&pc=MTA | 2023826-2022-01949 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HIS PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND EXCESSIVE VAULT POST ICL IMPLANTATION. LENS REMAINS IMPLANTED. PER UPDATED INFORMATION PATIENT IS DOING BETTER. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078671&pc=MTA | 2023826-2022-04755 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS EYE. THE REPORTER INDICATED THE LENS HAD A HIGH VAULT AND ELEVATED INTRAOCULAR PRESSURE (IOP). AT ONE DAY POST-OP THE IOP WAS AT 10MMHG. THE VISION IS GREAT AND PATIENT IS HAPPY. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE HIGH VAULT WAS RESOLVED AND THE PATIENT IS DOING FINE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DEVICE EXPIRATION DATE UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16456516&pc=MTA | 2023826-2023-00715 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS OF A +8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND ANGLE CLOSURE WITH ELEVATED IOP. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16456539&pc=MTA | 2023826-2023-00714 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS OF A +8.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694223&pc=MTA | 2023826-2023-01293 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -9.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE; AND MEDICATION WAS USED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5241755&pc=MTA | 2023826-2015-01526 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-11-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -12.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. LENS REMAINS IMPLANTED. Manufacturer Narrative: THE SURGEON DID A REPOSITIONING OF THE NASAL HAPTIC WHICH WAS NOT RIGHTLY LOCATED IN THE SULCUS. THE VAULTING WAS RECORDED AS 400UM. AFTER REPOSITIONING, THE VAULTING INCREASED AGAIN, SO THE SURGEON DECIDED TO EXCHANGE THE LENS FOR A SHORTER LENS, SAME MODEL AND DIOPTER ON (B)(6) 2016. THE VAULTING WAS STABLE AT 400UM AND THE PATIENT WAS FINE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. CONCLUSION: AS PER THE DFU OF THIS LENS MODEL,IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRAINDICATED. THIS PATIENT HAD AN ACD OF 2.89MM AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -12.50. DEVICE EVALUATED BY MANUFACTURER? NO, DEVICE REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4). LENS REMAINS IMPLANTED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5355968&pc=MTA | 2023826-2016-00030 | 2015-11-13 | Injury | STAAR SURGICAL COMPANY | 2016-01-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/+3.00/X083 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS REMAINS IMPLANTED. SEE MFR#2023826-2016-00029 FOR RIGHT EYE. Manufacturer Narrative: CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50/+3.00/X083. DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5066072&pc=MTA | 2023826-2015-01175 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/2.00/103 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS NOTED AS WELL AS ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION IS AVAILABLE. SEE MFR # 2023826-2015-01174 FOR RIGHT EYE. Manufacturer Narrative: SURGEON INDICATED THE IOP WAS UNDER CONTROL. NO SECONDARY SURGICAL INTERVENTION WAS NEEDED. LENS IMPLANTED ON (B)(6) 2015. (B)(6). (B)(4). THE PATIENT WAS UNDER THE AGE OF 21 YEARS OLD, AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -14.00/2.00/103. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS REMAINS IMPLANTED
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5185751&pc=MTA | 2023826-2015-01447 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/X067 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (26MMHG). THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574478&pc=MTA | 2023826-2015-00249 | 2015-02-02 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Insufficient Information | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD AND THE LENS WAS REPOSITIONED DUE TO ELEVATED IOP AND SHALLOW ACD ASSOCIATED WITH HIGH VAULT. LENS IS TILTED NOW AND REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S BUT NOT MARKETED IN THE U.S. (B)(4). SHALLOW ACD. VAULTING. REPOSITIONED. TILTED. DEVICE EVALUATED BY MANUFACTURER? NO: LENS REMAINS IMPLANTED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641098&pc=MTA | 2023826-2022-00644 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. AT 4 HOUR POST-OP, THE INTRAOCULAR PRESSURE WAS 36MMHG BUT THE PRESSURE CAME DOWN TO 14MMHG. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. AT A POST-OP VISIT THE PATIENTS REFRACTION HAD IMPROVED TO 6/6 UNCORRECTED, +0.25 SCRIPT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117846&pc=MTA | 2023826-2022-02425 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THERE WAS ELEVATED INTRAOCULAR PRESSURE (IOP). THE IOP WAS RAISED ON DAY OF SURGERY. DIAMOX WAS USED TO BRING IOP DOWN BUT IOP WAS UP TO 50MMHG ON (B)(6) 2022, SO PATIENT WAS BROUGHT BACK TO THE THEATRE FOR AC WASHOUT. IOP 12 AFTER AND VISION IS 6/12 AFTER HAVING BEEN HM (HAND MOVEMENTS). STATUS OF EYE - PUPIL DILATED STILL, POTENTIAL SPHINCTER DAMAGE FROM HIGH IOP. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: POST-OP, THE PATIENT WAS 6/12, WITH FIXED DILATED PUPIL AND DAMAGED SPHINCTER. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17894898&pc=MTA | 2023826-2023-04406 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED ON RESEARCH SQUARE IN (B)(6) 2023 TITLED, SAFETY OF PURE IMPLANTABLE COLLAMER LENS IMPLANTATION WITHOUT VISCOELASTIC AGENT IN THE EARLY PERIOD AFTER SURGERY. THE ARTICLE STATES THAT ELEVATED IOP IN 4 EYES, 2 REQUIRING MEDICAL MANAGEMENT AND TRANSIENT LOSS OF CDVA IN 7 EYES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823655&pc=MTA | 2023826-2020-02586 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED AND ELEVATED IOP (37MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020. USE OF STEROIDS AND IOP LOWERING MEDICATIONS PRESCRIBED. POST-OP MEDICATIONS STOPPED, AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "STEROID RESPONSE". PER THE REPORTER ON (B)(6) 2020, "THE CASE IS RESOLVED; THE STEROIDS WERE STOPPED AND IOP LOWERING DROPS STOPPED AND THE EYE IS FINE. THE STEROIDS WERE STANDARD POST OP CARE." Manufacturer Narrative: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10823662&pc=MTA | 2023826-2020-02585 | 2020-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS WAS OBSERVED AND ELEVATED IOP (37MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS ASSESSED ON (B)(6) 2020 AND USE OF IOP LOWERING MEDICATIONS PRESCRIBED. POST-OP MEDICATIONS STOPPED, AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "STEROID RESPONSE". PER THE REPORTER ON 15/OCT/ 2020, "THE CASE IS RESOLVED - THE STEROIDS WERE STOPPED AND IOP LOWERING DROPS STOPPED AND THE EYE IS FINE." Manufacturer Narrative: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SERIOUS INJURY. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. IMPLANT DATE: N/A. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773859&pc=MTA | 2023826-2023-01423 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, ELEVATED IOP WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLMMATION AND IOP," AND PATIENT CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND PIGMENT DISPERSION WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLAMMATION AND IOP," AND PATIENT CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." D9: RETURN DATE: 31-MAY-2023. H6: 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773860&pc=MTA | 2023826-2023-01424 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. PIGMENT DISPERSION WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLMMATION AND IOP," AND PATIENT'S CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO : THE REPORTER INDICATED THAT A 12.1MM VICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO PIGMENT DISPERSION. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BETTER CONTROL OF OCULAR INFLAMMATION AND IOP," AND PATIENT'S CURRENT MEDICATION: AZARGA, ALPHAGEN, PRED FORTE, VIGAMOX. PATIENT DIAGNOSED ON BE MYOPIC CHOROIDAL NEOVASCULARIZATION WHICH REQUIRED INTRAVITREAL ANTI-VEGF INJECTION." D4: SERIAL NUMBER: SHOULD BE CORRECTED TO (B)(6). D4: UDI: SHOULD BE CORRECTED TO: (B)(4). D6B: (B)(6) 2023 SHOULD BE ADDED FOR CORRECTION. D9: RETURN DATE: 31-MAY-2023 SHOULD BE ADDED FOR CORRECTION. H6: 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14178946&pc=MTA | 2023826-2022-01224 | 2022-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. POST-OP OBSERVATION AT 1DAY IOP 19; 1WEEK IOP 27; 1 MONTH IOP 46 AND THE DOCTOR GAVE A STEROID EYE DROP; 3 MONTH IOP 33. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(6). CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ELEVATED IOP WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS OF -10.50 DIOPTER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: CORRECTED TO " THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(3) 2021. PATIENT EXPERIENCED ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(3) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED UNKNOWN. H6: HEALTH IMPACT: CLINICAL CODE: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359050&pc=MTA | 2023826-2023-00484 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. ELEVATED IOP WAS REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT WAS TREATED WITH TOPICAL DROPS AND ACETAZOLAMIDE. REPORTEDLY THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359060&pc=MTA | 2023826-2023-00485 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. ELEVATED IOP WAS REPORTED. THE LENS REMAINS IMPLANTED AND THE PATIENT WAS TREATED WITH TOPICAL DROPS AND ACETAZOLAMIDE. REPORTEDLY THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B1 - CORRECTED TO: ADVERSE EVENT IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO: SERIOUS IN THE INITIAL MDR. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515448&pc=MTA | 2023826-2020-02023 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A#S-P#N/A | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. PRESSURE SPIKES WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10755783&pc=MTA | 2023826-2020-02461 |  | Injury | STAAR SURGICAL COMPANY | 2020-10-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT AND REFRACTIVE SURGERY IN 2020 ENTITLED, 'COMPARISON OF CLINICAL OUT COMES OF IMPLANTABLE COLLMER LENS IMPLANTATION WITH AND WITHOUT USE OF AN OPHTHALMIC VISCOSURGICAL DEVICE'. THE ARTICLE STATES THAT, " IN THE IOP GROUP, FIVE EYES (9.26%) SHOWED AN INCREASE IN IPO TO >30 MMHG AT 1 H POST-OPERATIVELY, HOWEVER, THE IOP WAS REDUCED BY A SHORT COURSE OF TOPICAL ANTIGLAUCOMA MEDICATION(DORZOLAMIDE/TIMOLOL: DOTIMOL: HANMI PHARMACEUTICAL." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT - IMPACT CODE 4644 - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT CODE: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE U.S.. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10950665&pc=MTA | 2023826-2020-02775 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT "AN INCREASE IN PRESSURE AFTER SURGERY. AND "EVERYTHING SEEMS TO BE OK NOW. HEADACHES ARE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10950698&pc=MTA | 2023826-2020-02776 |  | Malfunction | STAAR SURGICAL COMPANY | 2020-12-05 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | N/A | 1 | Event Description: FACEBOOK COMPLAINT: THE REPORTER INDICATED THAT "AN INCREASE IN PRESSURE AFTER SURGERY. AND "EVERYTHING SEEMS TO BE OK NOW. HEADACHES ARE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12361974&pc=MTA | 2023826-2021-03282 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER INTO THE PATIENT'S EYE. THE SURGEON REPORTS HIGH PRESSURE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426109&pc=MTA | 2023826-2021-03442 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-03 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE. THE SURGEON REPORTS ELEVATED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK OR UNK ENTRY. IMPLANT DATE: UNK. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14028645&pc=MTA | 2023826-2022-01107 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased |  | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE HAD POTENTIAL ISSUE WITH PATIENT IOP. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16525952&pc=MTA | 2023826-2023-00872 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. THE SURGEON NOTES AN IOP (INTRAOCULAR PRESSURE) INCREASE AT 3 WEEKS POST-OP AND REQUESTED ASSISTANCE FROM A CONSULTATION DOCTOR, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16816167&pc=MTA | 2023826-2023-01483 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: REPORTEDLY, "WE ARE HAVING TROUBLE WITH ICL PATIENTS EXPERIENCING POST OP NAUSEA AND VOMITING, OFTEN IN ASSOCIATION WITH MILDLY ELEVATED OR HIGH SIDE OF NORMAL EYE PRESSURES. SOMETIMES IOPS ARE COMPLETELY NORMAL. TYPICALLY THESE PATIENTS ARE EXPERIENCING THIS SEVERAL HOURS (4 SEEMS TO BE THE AVERAGE) AFTER SURGERY, SOMETIMES IN ASSOCIATION WITH HEADACHE. WE ARE CURRENTLY DOING THE FOLLOWING FOR PROCEDURE: IV VERSED AND KETAMINE SEDATION, RARELY WITH FENTANYL, IV DECADRON, IV ZOFRAN, IV HYDRATION, IV TORADOL FOR MANY, INTRAOCULAR MIOSTAT, PREOP PHENYLEPHRINE FTTS, POSTOP PILOCARPINE, INTRAOCULAR PRESERVATIVE FREE LIDOCAINE, TOBRADEX GTTS AT HOME.". Manufacturer Narrative: A4-A6:UNKNOWN. B3: UNKNOWN. D4: (EXPIRATION DATE); UNK, NO SERIAL NUMBER PROVIDED. D6A: UNKNOWN. H4: (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER PROVIDED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213811&pc=MTA | 2023826-2023-02511 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "THE PATIENT HAD RECEIVED ICL IN EGYPT MANY YEARS AGO, AND CAME INTO THIS DOCTORS CLINIC FOR A REGULAR EYE EXAM. THE PATIENT PRESENTED WITH IOP IN THE 30'S OU. THE PATIENT HAD PATENT PL'S," AND "THE PATIENT WAS PUT ON IOP DROPS IN THE INTERIM." Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213813&pc=MTA | 2023826-2023-02510 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | UNK | 1 | Event Description: AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "THE PATIENT HAD RECEIVED ICL IN EGYPT MANY YEARS AGO, AND CAME INTO THIS DOCTORS CLINIC FOR A REGULAR EYE EXAM. THE PATIENT PRESENTED WITH IOP IN THE 30'S OU. THE PATIENT HAD PATENT PL'S," AND "THE PATIENT WAS PUT ON IOP DROPS IN THE INTERIM." Manufacturer Narrative: A4-A6:UNK B3: UNK D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486824&pc=MTA | 2023826-2023-03267 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM SPHERIAL IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. ELEVATED IOP HAS BEEN OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: A1 - UNK, A2 - UNK, A3 - UNK ,A4 - UNK, A5 - UNK, A6 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, H4 - UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486825&pc=MTA | 2023826-2023-03268 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM SPHERIAL IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON 23/FEB/2023. ELEVATED IOP HAS BEEN OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENTS ADDED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY) CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15502149&pc=MTA | 2023826-2022-03357 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145297&pc=MTA | 2023826-2023-05047 | 2023-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18152256&pc=MTA | 2023826-2023-05048 | 2023-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON 02-SEP-2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924848&pc=MTA | 2023826-2023-01811 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: ON (B)(6) 2023 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN FRONTIERS IN NEUROLOGY TITLED, 'A QUANTITATIVE STUDY OF THE EFFECT OF ICL ORIENTATION SELECTION ON POST-OPERATIVE VAULT AND MODEL-ASSISTED VAULT PREDICTION'. THE ARTICLE REPORTS TRANSIENT IOP ELEVATION OBSERVED IN FOUR PATIENTS, WHICH RETURNED TO NORMAL AT THE NEXT FOLLOW-UP VISIT AFTER TOPICAL APPLICATION OF SINGLE GLAUCOMA MEDICATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF-LABEL USE (ACD<3.0MM) AND OFF LABEL USE (AGE<21) . CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8699719&pc=MTA | 2023826-2019-01050 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED IOP (21MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, EXCESSIVE VAULT WAS OBSERVED. PHACO-EMULSIFICATION (CATARACT SURGERY) AND IOL IMPLANTATION WAS DONE, THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS DEVICE, "ICL TOO LARGE". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186556&pc=MTA | 2023826-2019-01905 | 2019-09-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. THE SURGEON REPORTS VAULTING AT 300%, ELEVATED IOP, AND HIGH RISK OF ANGLE CLOSURE. LENS REMAINS IMPLANTED. NO ADDITIONAL INFORMATION HAS BEEN RECEIVED. Manufacturer Narrative: TYPE OF REPORTABLE EVENT: MALFUNCTION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11813620&pc=MTA | 2023826-2021-01419 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815697&pc=MTA | 2023826-2021-01420 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13252941&pc=MTA | 2023826-2022-00114 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065390&pc=MTA | 2023826-2021-04742 | 2021-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13253002&pc=MTA | 2023826-2022-00115 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13253038&pc=MTA | 2023826-2022-00116 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13184529&pc=MTA | 2023826-2022-00103 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6 VICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172928&pc=MTA | 2023826-2022-00098 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT WITH ELEVATED IOP. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172930&pc=MTA | 2023826-2022-00097 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT WITH ELEVATED IOP. A REPLACEMENT LENS WAS LATER IMPLANTED OF SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728674&pc=MTA | 2023826-2023-01333 | 2022-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased | P030016 | 1 | Event Description: THE REPORTER INDICATED, THAT A 12.6MM, VICMO12.6, IMPLANTABLE COLLAMER LENS OF A -14.0 DIOPTER, WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPANTED AND REMOVED. DUE TO ELEVATED IOP WITHOUT PUPIL BLOCK, AND ANGLE CLOSURE, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704133&pc=MTA | 2023826-2023-03950 | 1998-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT PRESENTED WITH REDNESS IN THE EYE, PIGMENT DISPERSION AND ELEVATED IOP IN THE EARLY POST-OP PERIOD. IT WAS REPORTED THAT THE PROBLEM RESOLVED WITH MEDICAL MANAGEMENT. THE REPORTER STATES THAT THE DISPERSED PIGMENTS DECREASED WITH NORMALIZATION OF IOP AFTER CONSERVATIVE MANAGEMENT WITH NO BLURRING OF VISION. HOWEVER, THE PUPIL REMAINED SLIGHTLY DILATED AND IRREGULAR. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: REDNESS, PIGMENT DISPERSION. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE - 4581 - PUPIL ANOMALY. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: 05-NOV-1998 SHOULD BE REMOVED FROM DATE OF EVENT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138789&pc=MTA | 2023826-2020-01186 | 2019-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. PUPIL BLOCK, ANGLE CLOSURE WITH ELEVATED IOP (38MMHG) WAS ASSESSED ON (B)(6) 2019. UNREACTIVE (FIXED) PUPIL, MID-DILATED IRIS WITH PERIPHERAL IRIDO-CORNEAL TOUCH AND FLAT AC WAS OBSERVED. AN ADDITION OF NEW PI AND YAG WAS PERFORMED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "PROBABLY VISCO ELASTIC TRAP." REPORTEDLY, "HAD TO DO SECOND PI YAG. VERY DIFFICULT BUT MANAGED TO GET PUPIL BLOCK REVERSED AND SETTLED DOWN." Manufacturer Narrative: CORRECTED DATA: B4 - CORRECTED DATE TO (B)(6) 2020. G4 - DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10706946&pc=MTA | 2023826-2020-02421 | 2020-09-07 | Injury | STAAR SURGICAL COMPANY | 2020-10-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA#S-P#NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ELEVATED INTRAOCULAR PRESSURE (IOP), UNREACTIVE (FIXED) PUPIL, ANISOCORIA AND IRIS ATROPHY. THE LENS WAS REPOSITIONED (ROTATED) BUT THE PROBLEM WAS NOT RESOLVED. MEDICATION WAS PRESCRIBED BUT WITH NO EFFECT ON PUPIL CONSTRICTION. THE LENS REMAINED IMPLANTED. AT THE LAST POST-OP VISIT ON (B)(6) 2020, THE PATIENT IS RECOVERING SLOWLY TO NORMAL PUPIL APPEARANCE. THIS CASE SEEMS TO BE AN URRETS-ZAVALIA SYNDROME. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11507466&pc=MTA | 2023826-2021-00660 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-03-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED DUE TO ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12277915&pc=MTA | 2023826-2021-03074 | 2021-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 A 12.1MM MICL12.1; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS. THE SURGEON INDICATED THE PATIENT EXPERIENCED SHALLOW AC, ELEVATED INTRAOCULAR PRESSURE AND WAS TREATED WITH MEDICATIONS LIDOCAINE, MOISTAT AND MIOCHOL ALONG WITH TOPICAL PILOCARPINE. PATIENT'S ANGLES WERE STILL NARROW AND WENT ON ALPHAGAN. IT WAS ALSO NOTED THAT "SHE FEELS GOOD WITH SOME IMPROVEMENT IN HER VISION". POSTOPERATIVE REPORT INDICATED PATIENT'S IOP DECREASED TO AN ACCEPTABLE LEVEL. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H10 - INITIAL CLAIM IS NOT APPLICABLE DUE TO WAS FOUND TO BE A DUPLICATE OF MFR# 2023826-2021-03057. REFERENCE MFR# 2023826-2021-03057 FOR CORRECT CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278129&pc=MTA | 2023826-2021-03075 | 2021-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2021 A 12.1MM MICL12.1; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS. THE SURGEON INDICATED THE PATIENT EXPERIENCED SHALLOW AC, ELEVATED INTRAOCULAR PRESSURE AND WAS TREATED WITH MEDICATIONS LIDOCAINE, MOISTAT AND MIOCHOL ALONG WITH TOPICAL PILOCARPINE. PATIENT'S ANGLES WERE STILL NARROW AND WENT ON ALPHAGAN. IT WAS ALSO NOTED THAT "SHE FEELS GOOD WITH SOME IMPROVEMENT IN HER VISION". POSTOPERATIVE REPORT INDICATED PATIENT'S IOP DECREASED TO AN ACCEPTABLE LEVEL. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: H10: INITIAL CLAIM IS NOT APPLICABLE DUE TO CLAIM WAS FOUND TO BE A DUPLICATE OF MFR# 2023826-2021-03058. SEE MENTIONED MFR# FOR CORRECT CLAIM DETAILS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215323&pc=MTA | 2023826-2022-01346 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215325&pc=MTA | 2023826-2022-01347 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506324&pc=MTA | 2023826-2022-03358 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -4.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, PUPIL BLOCK WITH ELEVATED IOP, ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM #(B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393718&pc=MTA | 2023826-2023-05723 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.5/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ANGLE CLOSURE WITH ELEVATED IOP 60MMHG(ASSESSED ON (B)(6) 2023). LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525773&pc=MTA | 2023826-2024-00171 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/3.0/043 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ELEVATED IOP(32MMHG) WITH OUR PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON 3RD WEEK POST-OP), PIGMENT DISPERSION WERE OBSERVED. LENS REMAINS IMPLANTED. PATIENT GIVEN STEROIDS AND IOP MEDS. THIS HAS NOT RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509050&pc=MTA | 2023826-2024-00217 | 2023-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED DURING POST-OP EXAMINATION. THE REPORTER STATED THAT THE SURGEON ROTATED THE LENS TO VERTICAL POSITION ON (B)(6) 2023 SINCE IT WAS FOUND THE VAULT WAS HIGH DURING RE-EXAMINATION ON (B)(6) 2024, BUT THE PROBLEM DID NOT RESOLVE. THE SURGEON FOUND THE CENTRAL HOLE ON THE LENS WAS NOT PERFORATED DURING THE RE-EXAMINATION ON (B)(6) 2023. PER REPORTER THE LENS WAS PERFORATED BY THE DOCTOR ON (B)(6) 2023 AND THE PATIENTS VISION, IOP RETURNED TO NORMAL. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10936830&pc=MTA | 2023826-2020-02830 |  | Injury | STAAR SURGICAL COMPANY | 2020-12-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: (B)(6) COMPLAINT: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S RIGHT (OD) EYE. THE PATIENT REPORTS ELEVATED IOP, ACUTE BLOCK DUE TO OVERSIZED LENS. CLOSURE OF PI'S, PI'S REDONE. MEDICATION ADMINISTERED. LENS EXCHANGED WITH A SHORTER LENS, AND THE PROBLEM IS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: IT IS DETERMINED THAT THE CURRENT CLAIM IS A DUPLICATE OF MFR REPORT# 2023826-2013-00438. PLEASE DISREGARD THIS REPORT. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; (B)(4); CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677893&pc=MTA | 2023826-2021-04041 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: AN ARTICLE (PREDICTION OF THE TRABECULAR IRIS ANGLE AFTER POSTERIOR CHAMBER PHAKIC INTRAOCULAR LENS IMPLANTATION) WAS RECEIVED THAT CITED A CASE STUDY OF 174 EYES (174 PATIENTS) OF PATIENTS IMPLANTED WITH HOLE ICLS. OF THE 174 EYES, A (B)(6) YEAR OLD FEMALE PATIENT, HAD AT ONE DAY POST-OP, ELEVATED INTRAOCULAR PRESSURE (IOP) OF 46MMHG AND ANGLE CLOSURE. INTRAVENOUS GLYCEROL DRIP INFUSION WAS USED TO REDUCE IOP. IN ADDITION, BRINZOLAMIDE 1% AND PILOCARPINE HYDROCHLORIDE 2% WAS INSTILLED. THE IOP DECREASED AND BY THE 3RD DAY WAS 8.6MMHG. AFTER 3 MONTHS, THE IOP REMAINED NORMAL BUT PARTIAL ANGLE CLOSURE WAS OBSERVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13515967&pc=MTA | 2023826-2022-00420 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: A JOURNAL ARTICLE WAS RECEIVED FROM THE FRONTIERS IN MEDICINE, (B)(6) 2021, FOR "SAFETY OF EVO ICL IMPLANTATION WITH AN OPHTHALMIC VISCOSURGICAL DEVICE-FREE TECHNIQUE IN THE EARLY 24 H AFTER SURGERY". A TOTAL OF 180 EYES OF 90 PATIENTS WERE ENROLLED IN THE STUDY, OF WHICH 100 EYES OF 50 PATIENTS WERE TREATED WITH NON-OVD TECHNIQUE. THE REMAINING 80 EYES OF 40 PATIENTS WERE TREATED WITH MIN-OVD TECHNIQUE. IT WAS REPORTED THE OCCURRENCE RATE OF PARACENTESIS TAP WAS SIGNIFICANTLY LOWER IN THE NON-OVD GROUP THAN IN THE MIN-OVD GROUP (1% [1 OF 100] VS. 5% [4 OF 80], P < 0.01). MOREOVER, THE USE OF IOP CONTROL DRUGS WAS SIGNIFICANTLY LOWER IN THE NON-OVD GROUP THAN IN THE MIN-OVD GROUP (0 [0 OF 100] VS. 3.8% [3 OF 80], P < 0.01). THE LENSES REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNKNOWN, SEX: UNKNOWN, WEIGHT: UNKNOWN, RACE: UNKNOWN, DATE OF EVENT: UNKNOWN, ETHNICITY: UNKNOWN, EXPIRATION DATE: UNKNOWN, IMPLANT DATE: UNKNOWN, EXPLANT DATE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325137&pc=MTA | 2023826-2022-02951 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT HIS PATIENT EXPERIENCED INTRAOCULAR PRESSURE POST ICL IMPLANTATION. THE REPORTER STATES " HIS IOP WAS 34 LAST WEEK IN BOTH EYES AND BOTH CHAMBERS WHERE SHALLOW EVEN AFTER ATROPINE USE. I STARTED HIM ON COMBIGAN BID OU AND DIAMOX TID. TODAY IS IOP WAS 11 AND 12. HIS PERIPHERAL A/C LOOKS SLIGHTLY DEEPER, MAYBE IN EACH EYE. HIS VAULT OD IS APPROX. 250UM AND OS IS 150UM. I STOPPED HIS DIAMOX TODAY AND CONTINUED COMBIGAN BID UNTIL I SEE HIM THURSDAY ". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325138&pc=MTA | 2023826-2022-02952 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT HIS PATIENT EXPERIENCED INTRAOCULAR PRESSURE POST ICL IMPLANTATION. THE REPORTER STATES " HIS IOP WAS 34 LAST WEEK IN BOTH EYES AND BOTH CHAMBERS WHERE SHALLOW EVEN AFTER ATROPINE USE. I STARTED HIM ON COMBIGAN BID OU AND DIAMOX TID. TODAY IS IOP WAS 11 AND 12. HIS PERIPHERAL A/C LOOKS SLIGHTLY DEEPER, MAYBE IN EACH EYE. HIS VAULT OD IS APPROX. 250UM AND OS IS 150UM. I STOPPED HIS DIAMOX TODAY AND CONTINUED COMBIGAN BID UNTIL I SEE HIM THURSDAY ". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. DEVICE MANUFACTURING DATE: UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571958&pc=MTA | 2023826-2022-03586 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -15.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND ELEVATED INTRAOCULARE PRESSURE. IT WAS REPORTED THAT IT RESOLVED PROBLEM BUT FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "MILD PIGMENT DISPERSION AND BORDERLINE IOP IMPROVING-STILL ON GTTS". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577782&pc=MTA | 2023826-2022-03587 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO12.6, -16.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND ELEVATED INTRAOCULARE PRESSURE. IT WAS REPORTED THAT IT RESOLVED THE PROBLEM BUT FOR STATUS OF THE EYE (SIGN/SYMPTOMS) THE REPORTER STATED "MILD PIGMENT DISPERSION AND BORDERLINE IOP IMPROVING-STILL ON GTTS". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17714392&pc=MTA | 2023826-2023-03954 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER STATES " HIGH IOP FOLLOWING ICL SURGERY- VAULT DAY 1 1088, INITIALLY HIGH IOP BUT RELEASED AQUEOUS FROM PARACENTESIS AND IOP NORMAL". PER REPORTER NO FURHTER INTERVENTION IS REQUIRED AT THIS STAGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED ELEVATED IOP AND AQUEOUS RELEASE. REPORTEDLY "HIGH IOP FOLLOWING ICL SURGERY - VAULT DAY 1 1088, INITIALLY HIGH IOP BUT RELEASED AQUEOUS FROM PARACENTESIS AND IOP NORMAL" AND "NO FURTHER INTERVENTION REQUIRED AT THIS STAGE". THE LENS REMAINS IMPLANTED. H6-HEALTH EFFECT- CLINICAL CODE: "4581- AQUEOUS RELEASE" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: D6A:UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18561095&pc=MTA | 2023826-2024-00367 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED NARROW ANTERIOR CHAMBER, ELEVATED IOP WITH PIGMENT DISPERSION. IT WAS REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6- HEALTH EFFECT- CLINICAL CODE 4581:NARROW ANTERIOR CHAMBER, PIGMENT DISPERSION H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713915&pc=MTA | 2023826-2022-03878 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING ROTATED AND THE PATIENTS PRESSURE HAD SPIKED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT WAS "BURPED" SEVERAL DAYS IN A ROW AND THE IOP WAS NORMAL. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE EYE IS FOGGY BUT IS SLIGHTLY IMPROVED. THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM. THE PATIENT IS SEEING A GLAUCOMA SPECIALIST WHICH WAS NO DISCLOSED TO THE FACILITY. THE PATIENT SUFFERED A CILIARY BODY INJURY WHICH CAUSED THE VISION TO CHANGE AND HER IOP TO BECOME UNCONTROLLABLE. A IOL ROTATION WAS PERFORMED BUT DUE TO THE ROTATION HER PRESSURES BEGAN TO SPIKE. IT WAS DECIDED TO REMOVE THE LENS FOR PATIENT SAFETY. THE PATIENT IS CURRENTLY SEEKING GLAUCOMA TREATMENT AND THERE ARE NO PLANS TO REPLACE THE LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - "BURPING THE PRESSURE." CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC INTRAOCULAR LENS, PRODUCT CODE: MTA. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11168648&pc=MTA | 2023826-2021-00133 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10535260&pc=MTA | 2023826-2020-01976 | 2020-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENTS EYE. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND ANGLE CLOSURE. THE LENS WILL BE EXPLANTED AND EXCHANGED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DATE OF EVENT: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - HEALTH EFFECT CLINICAL CODE: 4581- ANGLE CLOSURE. B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS; -05.50 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON 06/AUG/2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING, ELEVATED IOP WITH ANGLE CLOSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE "SUGGESTED LENS WAS TOO LARGE". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10963437&pc=MTA | 2023826-2020-02978 | 2020-08-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT, ELEVATED IOP AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11722535&pc=MTA | 2023826-2021-01208 | 2021-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP AND UNREACTIVE PUPIL (PUPIL IS REACTIVE TO LIGHT BUT VERY DILATED). FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "AWAITING FOLLOW UP" WAS REPORTED." DEVICE" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " SURGEON INTERPRETATION IS THAT THE ANGLE IS OPEN BUT THAT THE CONFIGURATION AND VAULT OF ICL IS SUCH THAT THE PUPIL CANNOT CONSTICT OVER THE LIP OF THE OPTIC RESULTING IN THE PUPIL BEING MECHANICALLY PREVENTED FROM RETURNING TO NORMAL SIZE." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT IT WAS AN IRIS CONFIGURATION ISSUE AND NOT AN ISSUE WITH VAULT OR THE ICL. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11721772&pc=MTA | 2023826-2021-01127 | 2021-03-16 | Injury | STAAR SURGICAL COMPANY | 2021-04-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P990013 | 1 | Event Description: IT WAS REPORTED THAT A 13.2MM, MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT (1.42MM VAULT), MILD DILATED PUPIL AND ELEVATED IOP (REQUIRING DROPS). THIS EXCHANGE RESOLVED THE PROBLEM. DEVICE, PATIENT- RELATED FACTOR, UNKNOWN, ICL TOO LARGE WAS REPORTED FOR CAUSE OF THIS EVENT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " IMPROVED-COMFORTABLE". Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK. RACE: UNK. (B)(4). DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11925869&pc=MTA | 2023826-2021-01746 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. HIGH VAULT AND ANGLE CLOSURE WITH ELEVATED IOP WAS OBSERVED. IT WAS REPORTED THAT MEDICAL INTERVENTION (ANTIGLAUCOMATEOUS TREATMENT) WAS PERFORMED WHILE WAITING FOR SURGERY. IT WAS RECENTLY REPORTED THAT LENS EXCHANGE WAS PERFORMED AND THAT THE PATIENTS CURRENT CONDITION IS VERY WELL. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS " OTHER: WTW VALUE OF THE PATIENT MAY CAUSE AN ERROR IN CALCULATION OF LENS DIAMETER, THE LENS MUST BE EXCHANGED BY A 12.6MM LENS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308408&pc=MTA | 2023826-2021-03089 | 2021-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -15.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED IOP (50MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VICMO 12.6; -15.00 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: ONE SIMILAR COMPLAINT EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (CLAIM# (B)(4) EXCESSIVE VAULT; LENS EXPLANTED; PROBLEM RESOLVED). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178435&pc=MTA | 2023826-2021-02702 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS -16.0 DIOPTER INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "DEVICE BEING TOO LARGE CAUSED ANTERIOR VAULT, RESULTING IN NARROWING ANTERIOR CHAMBER AND INCREASED IOP." THE PROBLEM IS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: B5 - PER EMAIL ON 15/OCT/2021, THE REPORTER STATED THAT "REPLACEMENT LENS WERE IMPLANTED ON (B)(6) 2021. RIGHT EYE: (B)(6) & LEFT EYE: (B)(6). ALL PROBLEMS HAVE BEEN RESOLVED AT THIS TIME." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271650&pc=MTA | 2023826-2021-03011 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS OTHER (WHITE TO WHITE MEASUREMENT). Manufacturer Narrative: PT INFO: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341302&pc=MTA | 2023826-2021-03173 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL 13.2, -13.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(32MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2021) AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, RECOMMENDED LENGTH WAS TOO LONG. REPORTEDLY, EYE IS CLEAR, GOOD VAULT, NORMAL. PATIENT REPORTED "RAINBOW EFFECT IN VISION". Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628814&pc=MTA | 2023826-2021-03987 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE INCISION WAS ENLARGED AND ADDITIONAL SUTURE WAS REQUIRED. THE PROBLEM WAS RESOLVED WITH EXPLANT OF LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606079&pc=MTA | 2023826-2021-03851 | 2021-09-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -12.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(23MMHG) WITHOUT PUPIL BLOCK(ASSESSED ON (B)(6) 2021) AND ANGLE CLOSURE (ASSESED ON (B)(6) 2021), AND EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, ANGLE CLOSURE.REPORTEDLY, CORNEA AND LENS OK. PIO NORMAL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO [OR#] SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256529&pc=MTA | 2023826-2022-00157 | 2021-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -17.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "HIGH VAULT; ELEVATED IOP; NARROW ANGLE". Manufacturer Narrative: B5 - IT WAS LATER CLARIFIED THAT THE LENS WAS EXPLANTED ON (B)(6) 2021. ON (B)(6) 2021 THE REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: B5: IT WAS LATER RE-REPORTED THAT ON 15-OCT-2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS, AND NOT REPLACED THE NEXT DAY IN ANOTHER SURGERY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206192&pc=MTA | 2023826-2022-00079 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (24MMHG ASSESSED ON (B)(6) 2021) ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13362155&pc=MTA | 2023826-2022-00349 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7; -11.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE; AND ON (B)(6) 2021 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#:(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179534&pc=MTA | 2023826-2022-02576 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14929821&pc=MTA | 2023826-2022-02102 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS OF DIFFERENT POWER DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTROCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. THE EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " STABLE AN IOP RETURNED TO NORMAL VALUES. REFRACTION IS ALMOST EMMETROPE." IN THE REPORTER OPINION PATIENT- RELATED FACTOR WAS THE CAUSE OF THIS EVENT. FOR CAUSE DETAILS THE REPORTER STATED " OVER CORRECTION. HIGH VAULT AND ANGLE CLOSURE ACCOMPANIED BY DILATED PUPIL." Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916212&pc=MTA | 2023826-2023-01798 | 2023-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -6.00 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; THE LENS WAS EXPLANTED ON THIS SAME DATE. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "WRONG WTW". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216375&pc=MTA | 2023826-2023-02434 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -14.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ELEVATED IOP(22MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESSED ON (B)(6)2023), UNREACTIVE(FIXED) PUPIL, AND EXCESSIVE VAULT WERE OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, LENS EXCHANGE REQUESTED. Manufacturer Narrative: B5 - LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "SOLVED THE VAULT PROBLEM, IS PERFECT." CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: UNREACTIVE(FIXED) PUPIL. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279000&pc=MTA | 2023826-2023-02611 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM512.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP(41MMHG, ASSESSED ON (B)(6) 2023), AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, DIAMETER WAS TOO LARGE, REPLACED SMALLER LENS FIXING PROBLEM. Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES, ANGLE CLOSURE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567731&pc=MTA | 2023826-2024-00371 | 2023-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ,SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND INCREASED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE "LENS WAS NOT FITTING FOR THE PATIENT". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567793&pc=MTA | 2023826-2024-00372 | 2023-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.00 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ,SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND INCREASED IOP. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE "LENS WAS NOT FITTING FOR THE PATIENT". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18568014&pc=MTA | 2023826-2024-00401 | 2023-12-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) AND ANGLE CLOSURE. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. FOR ADDITIONAL INFORMATION THE REPORTER STATED " NO SYMPTOMS, WILL LIKELY NOT EXCHANGE FOR NOW". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15176338&pc=MTA | 2023826-2022-02578 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. "VAULT WAS HIGHER THAN NORMAL OUTCOME; SLIGHTLY NARROW ANGLES; PATIENT IOP 11." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT,ETHNICITY, AND RACE:UNK. CLINICAL CODE (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16036885&pc=MTA | 2023826-2022-04724 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT HAD TO RETURN TO THE OPERATING ROOM AND THE LENS WAS EXPLANTED DUE TO UNCONTROLLABLE EYE PRESSURE WITH HIGH VAULT. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT(S) EXPERIENCED A HIGH VAULT; ELEVATED IOP; "WOUND BURP"; AND MEDICATION WAS PRESCRIBED (ORAL ACETAZOLAMIDE & IOP LOWERING DROPS). REPORTEDLY "OF MY FIRST 28 EVO LENSES, HALF HAD INTRAOCULAR PRESSURES OF 30MMHG OR HIGHER. 1/4 TO 1/3 OF THOSE REQUIRED FAIRLY EXTENSIVE TREATMENT AND OBSERVATION". ADDITIONAL INFORMATION PROVIDED WAS "BURPING THE PARACENTESIS OVER A PERIOD OF HOURS, ACETAZOLAMIDE AND DROPS". THE LENS WAS EXPLANTED. H6- HEALTH EFFECT- CLINICAL CODE: 4581- "WOUND BURP." CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15912720&pc=MTA | 2023826-2022-04281 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN EUROPEAN JOURNAL OF OPHTHALMOLOGY; ON (B)(6) 2022 ENTITLED, "EFFICACY AND SAFETY AFTER POSTERIOR CHAMBER IMPLANTABLE COLLAMER LENS IMPLANTATION ACCORDING TO PREOPERATIVE ANTERIOR CHAMBER DEPTH: SHORT-TERM COMPARATIVE STUDY." THE ARTICLE REPORTS ADVERSE EVENTS FOLLOWING ICL IMPLANTATION. THE ARTICLE REFERENCED 3 CASES OF INTRAOCULAR HYPERTENSION: 1 OF WHICH REQUIRED SURGICAL INTERVENTION. 3 CASES OF TICL ROTATION REQUIRING SURGICAL CORRECTION. 3 CASES OF EXCESSIVE VAULT_2 EYES REQUIRING LENS ROTATION AND 1 REQUIRING EXCHANGE. 1 CASE OF IOL POWER MISCALCULATION REQUIRING EXCHANGE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16127557&pc=MTA | 2023826-2023-00115 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6: INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. B5: "THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "INITIALLY SOME CORNEAL ODEMA AND ENDOTHELIAL CELL LOSS DUE TO TRICKY EXCHANGE (CONSTRICTION OF PUPIL DURING PROCEDURE) WILL COMPLETE ANOTHER COMPLAINT FORM BUT VISION EXCELLENT." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR." CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16128162&pc=MTA | 2023826-2023-00116 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT HAD EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSER; IT WAS LATER ADDITIONALLY NOTED THE PATIENT PRESSURES WERE STABLE ON MEDICATION AND ANGLES NARROW. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT; ANGLE CLOSURE WITH ELEVATED IOP AND MEDICATION WAS PRESCRIBED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "SITUATION RESOLVED AND PATIENT HAS EXCELLENT VISION." THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. CORRECTED DATA: G2: OTHER: "UNITED KINGDOM" SHOULD BE CORRECTED TO "IRELAND". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12276620&pc=MTA | 2023826-2021-03058 | 2021-07-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; RETAINED VISCOELASTIC/OCUCOAT; AND ENLARGEMENT OF PI (SURGICAL) WAS PERFORMED ALONG WITH PARACENTESIS TO REDUCE IOP. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DAY REPLACED WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED AND "PATIENT'S VAULT HAS SLOWLY DECREASED TO ACCEPTABLE LEVEL WITH DECREASING IOP". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "RETAINED VISCOELASTIC IN SMALL EYE & THICK BROWN IRIS". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271376&pc=MTA | 2023826-2021-03057 | 2021-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6; -11.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING; ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; RETAINED VISCOATELASTIC/OCUCOAT; AND ENLARGEMENT OF PI (SURGICAL) WAS PERFORMED ALONG WITH PARACENTESIS TO REDUCE IOP. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DAY REPLACED WITH A SHORTER LENGTH LENS. IT WAS NOTED POSTOPERATIVELY THE PATIENT HAD SHALLOW AC 1000MM VAULT BUT INDICATED THE PROBLEM WAS RESOLVED AND "PATIENT'S VAULT HAS SLOWLY DECREASED TO ACCEPTABLE LEVEL WITH DECREASING IOP". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "RETAINED VISELASTIC IN SMALL EYE & THICK BROWN IRIS". Manufacturer Narrative: CORRECTED DATA: B5 - (B)(6) 2021 SHOULD BE CORRECTED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, RETAINED OCUCOAT, SHALLOW AC. HEALTH EFFECT IMPACT CODE: (B)(4) ENLARGEMENT OF PI (SURGICAL), BURPED PARACENTESIS, CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082457&pc=MTA | 2023826-2021-04967 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE."Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS -18.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "EXCESSIVE VAULT, SHALLOW ANTERIOR CHAMBER, WTW IMPRECISE." CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10537758&pc=MTA | 2023826-2020-02068 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO HIGH PRESSURE. THE LENS WAS EXCHANGED FOR A DIFFERENT SIZE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A SPECIMEN CUP, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6: HEALTH EFFECT IMPACT CODE: 4625 - ADDITIONAL SURGERY; YAG - ADDITION OF NEW PI. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL SURGERY WAS PERFORMED, ADDITION OF NEW PI BY YAG AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2020-02119 FOR EXCHANGED LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR. CLAIM # (B)(4). Manufacturer Narrative: DEVICE BLA NUMBER: NA. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755748&pc=MTA | 2023826-2021-04202 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON REPORTS 3+ PIGMENT IN THE ANGLE 360 DEGREES MORE PROMINENTLY VISIBLE INFERIOR. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: IOP IN THE RIGHT EYE WAS 18MMHG ON COMBIGAN BID. IT WAS DETERMINED THAT THE INCREASED PIGMENT IN ANGLE WAS LIKELY DUE TO THE REPEAT YAG PI AND PIGMENT WILL LIKELY RESOLVED OVER TIME. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PRODUCT CODE:UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR BLANK/UNK ENTRY. IMPLANT DATE - UNK. DEVICE MANUFACTURE DATE - UNK. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178452&pc=MTA | 2023826-2021-02714 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS ELEVATED IOP. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: ON (B)(6) 2021 A 13.7MM VTICM5_13.7 ICL WAS IMPLANTED INTO THE PATIENT'S EYE. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO ANGLE CLOSURE WITH ELEVATED IOP AND THE PROBLEM WAS RESOLVED. REPORTEDLY A SMALL TEMPORAL IRIS DEFECT IS REMAINING DUE TO AN IRIS PROLAPSE THAT OCCURRED DURING THE EXCHANGE. A CORNEAL SUTURE WAS ALSO ADDED. H6: CLINICAL CODE 4581 (ANGLE CLOSURE, IRIS PROLAPSE); IMPACT CODE 4650 (CORNEAL SUTURE); INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: ADDITION OF CORNEAL SUTURE IS NOT THE SURGEON'S STANDARD PRACTICE. IT WAS ADDED BECAUSE THE INCISION SIZE IS "A BIT TOO LARGE" AND HE DIDN'T WANT ANY RISK OF LEAKING. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT -UNK. PRODUCT CODE UNK. EXPLANT DATE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE -UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443376&pc=MTA | 2023826-2021-03535 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE PATIENT EXPERIENCED ANGLE CLOSURE AND ELEVATED IOP AFTER ICL IMPLANTATION. PER UPDATED INFORMATION INTRAOCULAR PRESSURE RETURNED TO NORMAL AND PRESSURE IS BEING MONITORED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341285&pc=MTA | 2023826-2023-05659 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: NARROW ANTERIOR CHAMBER, ELEVATED INTRAOCULAR PRESSURE WITH PIGMENT DISPERSION AFTER ICL IMPLANTATION WAS REPORTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANTERIOR CHAMBER, PIGMENT DISPERSION (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14056490&pc=MTA | 2023826-2022-01147 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6, -5.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED PIGMENT DISPERSION. ON (B)(6) 2022 ADDITIONAL SURGICAL INTERVENTION (YAG PI) AND MEDICAL INTERVENTION ( PRED FORTE, TIMO COMOD, LOTEMAX) WERE PERFORMED/PERSCRIBED PRIOR TO ICL EXPLANTATION. IT WAS REPORTED THAT ELEVATED IOP: 20MMHG WAS ASSESSED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. STATUS OF THE EYE (SIGNS/SYMPTOMS) AT LAST VISIT: BCVA: 20/20; UCVA: 20/20; IOP: 16MMHG; AC QUIET AND DEEP. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846197&pc=MTA | 2023826-2023-04221 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-04220. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # 731544
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846540&pc=MTA | 2023826-2023-04220 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PIGMENT DISPERSION. THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-04221 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: HEALTH IMPACT- CLINICAL CODE: 4581 - PIGMENT DISPERSION. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552675&pc=MTA | 2023826-2022-00438 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15068918&pc=MTA | 2023826-2022-02299 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT CLINICAL CODE (B)(4). DEVICE CODE (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725919&pc=MTA | 2023826-2023-04004 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -13.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. ON THIS SAME DATE THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY NOTED "THE PATIENT UNDERWENT RIGHT EYE ICL IMPLANTATION SURGERY ON (B)(6) 2023. POSTOPERATIVELY ELEVATED INTRAOCULAR PRESSURE WAS OBSERVED EXCEEDING 50MMHG. THE ANTERIOR CHAMBER WAS SHALLOW WITH EXECESSIVE VAULTING. THE LENS WAS REMOVED ON THE SAME AFTERNOON CONSIDERING THE PATIENT'S UNIQUE OCULAR CONDITION, CHARACTERIZED BY A SMALL CORNEAL DIAMETER BUT A SMALL CORNEAL DIAMETER BUT A LARGE POSTERIOR CHAMBER SPACE AND SIGNS OF CILLARY BODY ATROPHERY LIKE CHANGES. AFTER INTERNAL DISCUSSION IT WAS SUGGESTED TO POSTPONE THE SURGERY. HOWEVER DUE TO THE STRONG WILLINGNESS OF THE PATIENT. IT IS RECOMMENDED THE PATIENT SEEK FURTHER TREATMENT AT A MORE EXPERIENCED HIGHER-LEVEL HOSPITAL AFTER A RECOVERY PERIOD OF THREE MONTHS". Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 - ANGLE CLOSURE. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908804&pc=MTA | 2023826-2023-04480 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS EXPLANT RESOLVED THE PROBLEM CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (AGE<21) SHOULD HAVE BEEN INCLUDED. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909001&pc=MTA | 2023826-2023-04481 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE . H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18176646&pc=MTA | 2023826-2023-05183 | 2023-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.0 DIOPTER INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED WITH SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ELEVATED IOP AND ANGLE CLOSURE AFTER SURGERY. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6- CLICICAL CODE4581- SIGNIFICAT REDUCTION OF IRIDOCORNEAL ANGLES; ANGLE CLOSURE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525966&pc=MTA | 2023826-2022-01543 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -5.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT; ELEVATED IOP; AND SIGNIFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED. IOP CONTROLLED DAY AFTER SURGERY. LENS REMAINS IMPLANTED. EYE IS STABLE AND UNDER OBSERVATION. CAUSE IS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140355&pc=MTA | 2023826-2021-02386 | 2021-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS FOR A SHORTER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND PATIENT PLEASED WITH OUTCOME. THE PIS WERE ENLARGED BY (B)(6) ON (B)(6) 2021. SEE MFR. REPORT#: 2023826-2021-04212 FOR REPLACEMENT LENS. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ANGLE CLOSURE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12817115&pc=MTA | 2023826-2021-04361 | 2021-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -7.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14382243&pc=MTA | 2023826-2022-01508 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: REPORTEDLY A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS REMOVED DUE TO HIGH VAULT. IT WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022. Manufacturer Narrative: PATIENT IDENTIFIER: UNKNOWN,AGE OR DATE OF BIRTH:UNKNOWN, SEX::UNKNOWN, WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. DATE OF EVENT: UNK. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS OF -3.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY EXCESSIVE VAULT, ELEVATED IOP, AND NARROW ANGLES WERE FOUND. THE LENS WAS EXPLANTED ON (B)(6) 2022, AND REPLACED ON THE SAME DAY IN A SEPARATE SURGERY WITH A SHORTER LENGTH LENS OF A DIFFERENT MODEL. THIS RESOLVED THE PROBLEM. H6- HEALTH EFFECT- CLINICAL CODE 4581: NARROW ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330098&pc=MTA | 2023826-2023-00399 | 2022-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT, ELEVATED IOP WITHOUT PUPIL BLOCK, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED IOP, AND PIGMENT DISPERSION WERE OBSERVED. CAUSE REPORTED AS UNKNOWN. REPORTEDLY, "HIGHER VAULT THAN EXPECTED PROBABLY DUE TO LONGER LENS." ON (B)(6) 2022, THE LENS WAS REPOSITIONED, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, THE REPORTER SAID THAT THE DOCTOR GAVE UP EXPLANTING THE LENS BECAUSE IN THE LAST CONTROLS, IT COULD BE SEEN THAT THE VAULT HAD DECREASED, AND SHE DECIDED TO CONTINUE EVALUATING THE PATIENT. Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT- CLINICAL CODE: 4581- SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES; ANGLE CLOSURE; PIGMENT DISPERSION. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12052902&pc=MTA | 2023826-2021-02043 | 2021-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 AND EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED DUE TO LARGE VAULT, SHALLOW ANTERIOR CHAMBER, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE (ICL TOO LONG). POST-OP THE PATIENT HAD MILD HALOS AND THE EYE WAS QUIET. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, SHALLOW ANTERIOR CHAMBER. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795690&pc=MTA | 2023826-2017-01321 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT HAS BEEN TREATED BY MEDICATIONS: COSOPT, SPERSADEX, AND DIAMOX. MULTIPLE ATTEMPS TO GET UPDATED INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE PATIENT'S ELEVATED IOP WAS TREATED WITH MEDICATIONS: COSOPT, SPERSADEX AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795695&pc=MTA | 2023826-2017-01322 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE, ELEVATED IOP. IN ONE WEEK POST-OP, THE SURGEON STATED THAT THE PATIENT IOP IS 26.7MMHG, AND HAS BEEN TREATED BY TRUSOPT, AND TOBRADEX MEDICATIONS. THE SURGEON ALSO NOTED THAT THE PATIENT IOP IN TWO WEEKS POST-OP IS 20.6 MMHG. MULTIPLE ATTEMPS TO GET ADDITIONAL INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): THE SURGEON NOTED THAT THE PATIENT'S IOP IN TWO WEEKS POST-OP IS 20.6MM HG AFTER THE PATIENT WAS TREATED WITH MEDICATIONS: TRUSOPT AND TOBRADEX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795682&pc=MTA | 2023826-2017-01320 | 2017-06-28 | Injury | STAAR SURGICAL COMPANY | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.0/159 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, REFRACTIVE SURPRISE, ELEVATED IOP. ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT IOP IS 20.6MMHG, AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN, AND DIAMOX. MULTIPLE ATTEMPS TO GET MORE INFORMATION HAS BEEN UNSUCCESSFUL. Manufacturer Narrative: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES): ON (B)(6) 2017, THE SURGEON COMMENTED THAT THE PATIENT'S IOP IS 20.6MM HG AND HAS BEEN TREATED BY MEDICATIONS: TRUSOPT, EFFLUMIDEX, FENEFRIN AND DIAMOX. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12761919&pc=MTA | 2023826-2021-04212 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE SURGEON ENLARGED THE PIS BY YAG DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (28MMHG) AND SLIGHT VA BLURRINESS PER PATIENT. THE IOP RESOLVED AND WAS 25MMHG SAME DAY. BY (B)(6) 2021 THE IOP WAS 14MMHG. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02386. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK EXPLANT DATE: NA (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343177&pc=MTA | 2023826-2023-02983 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: ANGLE CLOSURE. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17343354&pc=MTA | 2023826-2023-02984 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATES " SULCUS TO SULCUS SMALLER THAN THE ESTIMATE". Manufacturer Narrative: H6- HEALTH EFFECT - CLINICAL CODE: "1937" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF ICA H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135946&pc=MTA | 2023826-2021-00043 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP IS 16/16MMHG AND VAULT IS STILL HIGH AT 1243. CLAIM# (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: AS OF (B)(6) 2021, THE VAULT IS (B)(4). TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATIENT IS OK" POST EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11135969&pc=MTA | 2023826-2021-00044 | 2020-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-01-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM6_13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -1.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE SURGEON REPORTS EXCESSIVE VAULT, ANGLE CLOSURE, AND ELEVATED IOP. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. REPORTEDLY THE IOP AND VAULT ARE BETTER AFTER REPOSITIONING. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF (B)(6) 2020, UCVA DISTANCE IS 0.4, READING DISTANCE IS GOOD, IOP 16/16MMHG AND VAULT IS STILL HIGH AT 1184. CLAIM# (B)(4). Manufacturer Narrative: B5-UPDATED INFORMATION: AS OF 03-FEB-2021 THE VAULT IS 1093. TRANSIENT LOSS OF BCVA REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, "THE PATENT IS OK." POST EXCHANGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195867&pc=MTA | 2023826-2022-02598 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND BLURRED VISION. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. MEDICATION WAS ADMINISTERED. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THERE WAS HIGH PRESSURE, ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATROPHY AND SHALLOWING OF AC. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - ZONULAR INSTABILITY, IRIS PROLAPSE, IRIS ATOPHY, SHALLOWING OF A/C. H6: HEALTH EFFECT IMPACT CODE: 4644 - TOPICAL AND ORAL ANTI-OCULAR HYPERTENSIVE MEDICATION. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750641&pc=MTA | 2023826-2023-01227 | 2023-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (PD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP(60MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE(ASSESED ON 19/FEB/2023), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, BLURRED VISION, AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10544158&pc=MTA | 2023826-2020-02118 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THAT THE LENS IS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED INTO THE OD EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTATION. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY AND POSTERIOR SYNECHIAELYSIS ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, ENCOURAGED TO USE +1.25 READERS TO HELP WITH COMPUTER AND PHONE. CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR, "EYE COULDN'T HANDLE DEVICE." H6: LINICAL CODE 4581 (ANGLE CLOSURE). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10551350&pc=MTA | 2023826-2020-02119 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-17 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S EYE ON (B)(6) 2020. ON (B)(6) 2020 THE SURGEON EXPLANTED THE LENS DUE TO ELEVATED IOP. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: LENS HAD BEEN IMPLANTED IN THE OS EYE. ADDITIONAL SYMPTOMS REPORTED: EXCESSIVE VAULT, ANGLE CLOSURE, BLURRED VISION. ADDITION OF NEW PI-YAG, PHACO-EMULSIFICATION AND IOL IMPLANTED. PERFORMED REFRACTIVE LENS EXCHANGE WITH PUPILLOPLASTY ON (B)(6) 2020. STATUS: PSEUDOPHAKIA, "GAVE SCL'S +0.75 BASE CURVE 8.4 AND WILL SCHEDULE LASIK IN THE NEAR FUTURE." H6-CLINICAL CODE 4581 (ANGLE CLOSURE) CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDU ON THE LENS AND THAT THE LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426116&pc=MTA | 2023826-2021-03454 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, DIOPTER +4.0 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ELEVATED IOP, FIXED PUPIL, BLURRED VISION, AND ENLARGEMENT OF PI-YAG. SYNECHIAE ALSO NOTED, LYSIS PERFORMED BUT NO IMPROVEMENT. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE-"LENGTH ISSUE, TOO LARGE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426103&pc=MTA | 2023826-2021-03455 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, DIOPTER +6.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ELEVATED IOP, BLURRED VISION, AND ENLARGEMENT OF PI-YAG. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE-"TOO LARGE." STATUS OF THE EYE IS REPORTED AS NORMAL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16309075&pc=MTA | 2023826-2023-00445 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6; -9.5 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ELEVATED INTRAOCULAR PRESSURE; GLARE/HALOES; BLURRED VISION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED; IT WAS ADDITIONALLY INDICATED AN INCISION ENLARGEMENT AND ADDITION SUTURE WAS PERFORMED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "SHALLOW VAULT DUE TO LENGTH". Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14057555&pc=MTA | 2023826-2022-01148 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. ON (B)(6)2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTER THAT INTRAOCULAR PRESSURE WAS TAPER DOWN TO 17MMHG. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE: 4581- RAINBOW LIGHT. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -4.50 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE; BLURRED VISION AND HALOS/RAINBOW LIGHT. ON (B)(6) 2021 SECONDARY YAG WAS PERFORMED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT MEDICATION WAS USED TO LOWER IOP. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089564&pc=MTA | 2023826-2015-01252 | 2015-08-19 | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, IN HIS RIGHT EYE (OD) ON (B)(6) 2015. THE PATIENT REPORTED THE SURGERY WENT WELL BUT AT ONE DAY POST-OP, HE COULDN'T SEE OUT OF HIS EYE, HE HAD TROUBLE FOCUSING AND IT WAS LIKE LOOKING THROUGH A WHITE SHEET OF PAPER. HE WENT IN FOR A FOLLOW-UP VISIT THAT SAME DAY AND IT WAS NOTED THE PERIPHERAL IRIDOTOMIES WERE NOT OPEN AND THE PRESSURE WAS UP. THE PERIPHERAL IRIDOTOMIES WERE OPENED AND THE PATIENT WAS SENT HOME. ON (B)(6) 2015 THE PATIENT HAD THE SAME PROBLEM AND HIS VISION WAS BLURRY. HE WENT BACK TO HIS DOCTOR AND THE DOCTOR DECIDED TO EXPLANT THE LENS THAT SAME DAY. HE WAS STILL HAVING PROBLEMS WITH HIS EYE AND HIS DOCTOR GAVE HIM DROPS. THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO PERSISTENT ANGLE CLOSURE. Manufacturer Narrative: PER MEDICAL REVIEW - ANGLE CLOSURE FOLLOWING PIOL IMPLANTATION IS A RARE COMPLICATION. STAAR RECOMMENDS PREVENTION BY PERFORMING TWO LASER PERIPHERAL IRIDOTOMIES BEFORE SURGERY, AND IN THIS CASE, THE PIS HAD BEEN PERFORMED BUT THEY WERE NOT OPEN. EVEN AFTER THE PIS WERE OPENED UP WITH YAG LASER, IOP ROSE AGAIN. ONE POSSIBLE MECHANISM IS INITIAL PUPIL BLOCK DUE TO PUPIL-PIOL CONTACT THAT REMAINED BECAUSE SYNECHIAE DEVELOPED. ANOTHER POSSIBLE MECHANISM IS CILIOLENTICULAR BLOCK, BUT THERE ARE NO UBM OR OCT REPORTS TO DETERMINE THE CAUSE OF PERSISTENT ANGLE CLOSURE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE INCORRECT DIMENSIONS OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED INCORRECT SIZING OF THE LENS RESULTING IN ELEVATED IOP AND PERSISTENT ANGLE CLOSURE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6715656&pc=MTA | 2023826-2017-01142 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO BLURRED VISION AND ELEVATED IOP. AN AHMED VALVE WAS IMPLANTED IN THE PATIENT. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: PATIENT IS FINE. BCVA IS 10/10 AND IOP IS CONTROLLED. PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7465307&pc=MTA | 2023826-2018-00607 | 2018-02-18 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -8.5/+3.0/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA, ELEVATED IOP, PIGMENT DISPERSION, AND BLURRED VISION. THE LENS WAS EXPLANTED ON (B)(6) 2018. TO DATE, THE PATIENT HAS NOT DECIDED WHETHER TO REPLACE WITH ANOTHER LENS OR NOT. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, THE PATIENT DEVELOPED RAISED IOP AFTER TICL IMPLANT, WHICH WAS "REFRACTORY TO TOPICAL ANTI-GLAUCOMA MEDICATIONS." THE TICL WAS EXPLANTED AND A TRABECULECTOMY PERFORMED. THE CORNEA IS CURRENTLY CLEAR. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS AND BLACK SPOTS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7467360&pc=MTA | 2023826-2018-00604 | 2018-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-04-26 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. THE LENS REMAINS IMPLANTED. REPORTEDLY, TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. NO UPDATED INFORMATION IS AVAILABLE AT THIS TIME. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. THE REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED "AROUND (B)(6)" FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: B5 - UPDATED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS LOSS OF BCVA AND ELEVATED IOP. REPORTEDLY TIMOLOL WAS ADMINISTERED TO THE PATIENT IN ORDER TO LOWER IOP. THE SURGEON NOTES FLUCTUATING REFRACTION. REPOSITIONING WAS PERFORMED ON (B)(6) 2018 BUT DID NOT RESOLVE THE PROBLEM. AN EXCHANGE WAS PERFORMED (B)(6) 2018 FOR ANOTHER VICM5_13.2 LENS OF -13.0 DIOPTER. THE VA WAS STILL LOW AT 0.65 AND THE PROBLEM WAS NOT RESOLVED. THE REPORTER STATED THAT "THEY DON'T THINK THAT THE ICL IS CAUSING THE LOW VA. THEY ARE THINKING NOW THAT OPACITIES IN THE VITREOUS ARE CAUSING THE PROBLEMS (THE LENS AND MACULA LOOK GOOD)." D6B - (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876795&pc=MTA | 2023826-2017-01460 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/142 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6876796&pc=MTA | 2023826-2017-01459 | 2016-05-26 | Injury | STAAR SURGICAL COMPANY | 2017-09-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/091 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6661714&pc=MTA | 2023826-2017-01012 | 2017-02-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -12.50 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. SECONDARY INTERVENTION INCLUDES ENLARGEMENT OF INCISION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6641431&pc=MTA | 2023826-2017-00976 | 2017-05-12 | Injury | STAAR SURGICAL COMPANY | 2017-06-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/083 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT YET BEEN RETURNED FOR EVALUATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN LIQID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802614&pc=MTA | 2023826-2017-01350 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. AFTER IMPLANTATION, THE PATIENT EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ANGLE CLOSURE WITH ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND BLURRED VISION. AT A LATER DATE THE SURGEON EXPLANTED THE LENS, CUT IT IN ORDER TO SHORTEN IT, THEN RE-IMPLANTED THE LENS BACK INTO THE EYE. IT WAS REPORTED THAT THE VAULT AND EYE ARE GOOD NOW. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: PMA 50(K): - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL, UNAPPROVED, OR CONTRAINDICATED USE (KERATOCONUS EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7441301&pc=MTA | 2023826-2018-00565 | 2017-10-17 | Injury | STAAR SURGICAL COMPANY | 2018-04-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE SURGEON REPORTS EXCESSIVE VAULT, REFRACTIVE SURPRISE, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID; IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE, ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, SIGNIFICANT REDUCTION OF ICA). (B)(4) OFF-LABEL USE [ANTERIOR ENDOTHELIUM CHAMBER DEPTH < 3.0MM (2.81MM)]. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113305&pc=MTA | 2023826-2017-01883 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO VAULTING AND INCREASED INTRAOCULAR PRESSURE. A SHORTER LENGTH LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED , A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM -9.5 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, ELEVATED IOP, NARROWING OF ANGLE, SHALLOWING OF ACD. THE SURGEON ALSO PERFORMED A SECONDARY YAG ON (B)(6) 2017. THE PATIENT HAD BLURRY VISION DURING THE TIME OF POST-OP. IMPLANTED DATE IS CORRECTED TO (B)(6) 2017. (B)(4). CORRECTED DATA: METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND REPLACED WITH SHORTER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7267534&pc=MTA | 2023826-2018-00209 | 2017-12-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), REFRACTIVE SURPRISE, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTION: LENS WAS IMPLANTED (B)(6) 2017. THE LENS WAS EXCHANGED WITH A SMALLER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL OBSERVATION FOUND THE LENS HAPTIC MISSING A PIECE AND RESIDUE ON LENS. CONCLUSION: OFF-LABEL USE [ANTERIOR CHAMBER DEPTH < 3.00MM (2.84MM)]. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7320724&pc=MTA | 2023826-2018-00319 | 2018-01-03 | Injury | STAAR SURGICAL COMPANY | 2018-03-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. REPORTEDLY, ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, ANGLE CLOSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), BLURRED VISION, AND IRIS ATROPHY. SURGEON ALSO STATES HE PERFORMED IRRIGATION/ASPIRATION, PHACO-EMULSIFICATION, AND IOL IMPLANTATION. PATIENT IS NOW PSEUDOPHAKIC. Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED AND REPLACED WITH A NON-STAAR LENS. AFTER SURGERY DAY 1 THERE WAS NO ANGLE CLOSURE (ONLY ELEVATED IOP), "BUT THEN THE IRIS WAS GETTING CLOSER TO THE CORNEA." DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7504918&pc=MTA | 2023826-2018-00693 | 2018-04-19 | Injury | STAAR SURGICAL COMPANY | 2018-05-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM MICL13.7, -11.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTIONS: ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT WITH BLURRED VISION AND ELEVATED IOP. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF A HAPTIC MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7907389&pc=MTA | 2023826-2018-01423 | 2018-08-15 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, REDUCTION OF BEST-CORRECTED VISUAL ACUITY, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE (FIXED) PUPIL. A YAG WAS PERFORMED. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO DIFFICULTY SIZING WTW/ANGLE LENGTH AND THE ANATOMY OF THE PATIENT'S EYE. SEE MFR. REPORT 2023826-2018-01422 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5037267&pc=MTA | 2023826-2015-01139 | 2015-07-22 | Injury | STAAR SURGICAL COMPANY | 2015-08-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND INCREASED ELEVATED INTRAOCULAR PRESSURE (UP TO 56MMHG). MEDICATIONS WERE ADMINISTERED BUT DID NOT HELP AND A YAG OF THE PERIPHERAL IRIDOTOMIES WAS PERFORMED BUT THE SURGEON DECIDED TO EXPLANT THE LENS. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT HAD ADDITIONAL SYMPTOMS OF BLURRY VISION AND NAUSEA. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. THE SURGEON INDICATED THE CAUSE OF THE EVENT WAS A MISCALCULATION/RECOMMENDED LENS SIZE. Manufacturer Narrative: THE PATIENT HAD SHALLOW ANTERIOR CHAMBER WITH NARROWED ANGLES OS, AND INCREASED IOP RESULTING FROM AN OVER-SIZED LENS. THE SURGEON DETERMINED SIZING USING THE FDA APPROVED LABELING. THE RISE IN IOP RESOLVED WITH LENS EXCHANGE, THEREFORE THE CHANGES IN ANTERIOR CHAMBER CONFIGURATION AND INCREASED IOP WERE RELATED TO THE PRESENCE OF THE OUTSIZE LENS IN THE EYE. BASED ON THE COMPLAINT HISTORY, MEDICAL REVIEW, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW. RESULTS: UPON REVIEW OF THE DEVICE HISTORY RECORD, THERE IS NOTHING IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5731066&pc=MTA | 2023826-2016-00819 | 2016-04-22 | Injury | STAAR SURGICAL COMPANY | 2016-06-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -5.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT, ANGLE CLOSURE (WITH ELEVATED IOP), BLURRED VISION, AND PERMANENTLY DILATED PUPIL. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTIONAL DATA: (B)(4). CODE NOT APPLICABLE, NO OTHER ADVERSE EVENT REPORTED OUTSIDE OF SECONDARY SURGERY. (B)(4). SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5681611&pc=MTA | 2023826-2016-00759 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER . THE PATIENT EXPERIENCED BLURRY VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP CONDITION IS GOOD AND BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL AND DEVICE HISTORY RECORD REVIEWS, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5923610&pc=MTA | 2023826-2016-01211 | 2016-08-01 | Injury | STAAR SURGICAL COMPANY | 2016-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, ELEVATED INTRAOCULAR PRESSURE AND REFRACTIVE SURPRISE. THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: -DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTIC BROKEN AND BENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417003&pc=MTA | 2023826-2019-02307 | 2018-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. BLURRED VISION AND CONSTANT VERTICAL BLACK LINE WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK ONE WAS SEEN." PER REPORTER ON (B)(6) 2019, DOCTOR HAS REQUESTED A REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. BLURRED VISION, ELEVATED IOP, AND CONSTANT VERTICAL BLACK LINE WERE OBSERVED. REPORTEDLY "BLURRING OF VISION SINCE SURGERY ((B)(6) 2018) AND A CONSTANT VERTICAL BLACK LINE WAS SEEN." THE ELEVATED IOP 'WAS BOUGHT UNDER CONTROL WITH DROPS' AND 'PBI WAS ALSO DONE.' THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT THE LINE REMAINED AND BECAME HORIZONTAL. THE LENS REMAINS IMPLANTED WITH PLANNED EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. PATIENT CODE 3191 - SECONDARY INTERVENTION: MEDICAL INTERVENTION (DROPS), LENS REPOSITIONING, PBI. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242379&pc=MTA | 2023826-2019-00102 | 2018-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -14.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER WAS TOO LARGE. Manufacturer Narrative: ADDITONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8242395&pc=MTA | 2023826-2019-00101 | 2018-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -15.0 DIOPTER, IMPLANTABLE COLLMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS LATER REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PUPIL BLOCK, ANGLE CLOSURE, INTRAOCULAR PRESSURE, AND BLURRED VISION. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT WAS STABLE. PATIENT IS CURRENTLY ON ALPHAGAN. IN THE REPORTERS OPINION THE CAUSE OF EVENT IS DUE TO THE RECOMMENDED DIAMETER, WAS TOO BIG. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. PATIENT CODE: "2137" NOT INCLUDED IN THE INITIAL MDR. DEVICE CODE: "1401" NOT INCLUDED IN THE INITIAL MDR. ADDITIONAL DATA: CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8269193&pc=MTA | 2023826-2019-00142 | 2019-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-22 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, LOSS OF BEST CORRRECTED VISUAL ACUITY (BCVA), ELEVATED INTRAOCULAR PRESSURE AND BLURRED VISION. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO RECOMMENDED LENS WAS TOO LARGE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595420&pc=MTA | 2023826-2019-00765 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +5.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS ELEVATED IOP, IRIDOTOMIES FOUND TO NOT BE OPENED, AND RESIDUAL REFRACTION OF -2.25D. THE PATIENT IS DETERMINED TO BE A STEROID RESPONDER AND WAS TREATED WITH DIAMOX TABLETS FOR ELEVATED IOP. YAG WAS PERFORMED ON (B)(6) 2019. PRESSURE WENT DOWN BUT STILL "NO GOOD." NO ADDITIONAL INFORMATION CURRENTLY AVAILABLE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: AT A FOLLOW UP VISIT, THE SURGEON STATES IT'S DIFFICULT TO JUDGE THE OPTIC BUT THAT THE LENS IS SLIGHTLY DECENTERED UP/TEMPORAL, "SOME MM UP AND 1-2MM TEMPORAL CENTERED". THE VAULT IS "A BIT TOO SMALL," O.5CT NASALLY, 0.75CT DOWN AND 1CT TEMPORALLY. IT IS SUSPECTED THAT THE TEMPORAL UPPER HAPTIC EAR IS BENT INWARDS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002877&pc=MTA | 2023826-2020-00983 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE IN A SECOND SURGERY, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND BLURRED VISION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774163&pc=MTA | 2023826-2019-01188 | 2019-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS EXCESSIVE VAULT, LOSS OF BCVA, PUPIL BLOCK, ELEVATED IOP, "FLAT CAMERA," AND ANGLE CLOSURE. ELEVATED IOP WAS TREATED WITH "MEDICAL DROPS" AND PAIN PILLS. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9794960&pc=MTA | 2023826-2020-00430 | 2019-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -12.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE (IOP) AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. H6- DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4725289&pc=MTA | 2023826-2015-00428 |  | Injury | STAAR SURGICAL COMPANY | 2015-04-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Blurred Vision; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE REPORTER INDICATED THE LENS HAD A REFRACTIVE SURPRISE AND THE LENS ROTATED (NOT ASSOCIATED TO A LOW VAULT). THE INCISION WAS ENLARGED AND THE LENS WAS REPOSITIONED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED AT THE FIVE MONTH VISIT, THE SURGEON CONFIRMED THE OFF-AXIS ALIGNMENT AND DECIDED TO REPOSITION THE LENS. THE PATIENT EXPERIENCED BLURRED VISION. AFTER THIS INTERVENTION, THE PATIENT EXPERIENCED AN IOP RISE, WHICH WAS SOLVED BY ANTERIOR CHAMBER ASPIRATION AND IRRIGATION. FOR THIS SECONDARY INTERVENTION A NEW INCISION WAS CREATED, LEADING TO INCREASED CORNEAL ASTIGMATISM. THE LENS REMAINS IMPLANTED AND IS STABLE AFTER REPOSITIONING. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULTS: (EVALUATION RESULT): UPON REVIEW OF THE DEVICE HISTORY RECORD, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING AND PACKAGING PROCESSES IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. BASED ON THE ABOVE INVESTIGATION, NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO - (OTHER): LENS IMPLANTED METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (OTHER): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS IMPLANTED. Manufacturer Narrative: (B)(4). METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - SEVERAL FACTORS MAY BE INVOLVED IN A POST-OPERATIVE REFRACTIVE SURPRISE, INCLUDING INAPPROPRIATE REFRACTIVE SURGICAL PLANNING, A WRONG LENS CALCULATION/SELECTION, PREOPERATIVE INACCURATE DETERMINATION OF THE EYE'S MEASUREMENTS TO DETERMINE THE POWER OF THE ICL, INDUCED REFRACTIVE CHANGE BY INCISIONS, UNSTABLE CORNEA AND/OR REFRACTION PRIOR TO IMPLANTATION, CL-INDUCED WARPAGE, UPSIDE-DOWN LENS, WRONG LENS, MISALIGNED LENS, ROTATION, TILTING, ETC. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE RESIDUAL REFRACTION IS A COMBINATION OF SELECTING A LENS POWER THAT WAS UNDER-CORRECTING THE CYLINDER AND SURGICAL INDUCED ASTIGMATISM SECONDARY TO A NEW INCISION. THE LENS WAS OFF-AXIS AT ONE DAY POSTOPERATIVELY THEREFORE THIS COULD BE DUE TO SURGICAL MISALIGNMENT OR EARLY MOVEMENTS OF THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9338950&pc=MTA | 2023826-2019-02190 | 2019-10-23 | Injury | STAAR SURGICAL COMPANY | 2019-11-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE INTRAOCULAR PRESSURE WAS REPORTED AS BEING HIGH. THE PATIENT COMPLAINED OF POOR VISION AND SEES DISTORTED LIGHTS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE UNK. EXPLANTED: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998035&pc=MTA | 2023826-2022-00825 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS OF -13.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXESSIVE VAULT, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, UNREACTIVE (FIXED) PUPIL OF 5MM, AND SHALLOW ANTERIOR CHAMBER WERE REPORTED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=MTA | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18060314&pc=QCB | 2023826-2023-04944 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED. DOCTOR REQUESTED CONSULTATION FOR PATIENT WITH ICL ROTATION AND PIGMENT DISPERSION. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: 4581-PIGMENT DISPERSION CLAIM#: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11416096&pc=MTA | 2023826-2021-00507 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-03-04 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED IOP, BLURRED VISION, GLARE/HALOES, AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT-RELATED FACTOR, "UNUSUAL ANATOMY". REPORTEDLY, "ALL LOOKS GOOD, MUCH LESS GLARE IN NEW LENS PER PATIENT." D9- CHANGED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6440682&pc=MTA | 2023826-2017-00502 | 2016-12-15 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF ICA, REFRACTIVE SURPRISE, LOSS OF BCVA, PIGMENT DISPERSION, UNREACTIVE PUPIL (FIXED), GLARE/HALOES, BLURRED VISION, AND IRIS ATROPHY. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SAME DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: METHOD, RESULTS CODES: UPDATED: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED TO MANUFACTURER.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6802597&pc=MTA | 2023826-2017-01349 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -6.00 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. AFTER IMPLANTATION, PATIENT EXPERIENCED EXCESSIVE VAULT, ANGLE CLOSURE WITH ELEVATED IOP, GLARE/HALOS, AND BLURRED VISION. A LENS EXCHANGE HAS BEEN PLANNED BUT NOT PERFORMED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6546105&pc=MTA | 2023826-2017-00768 | 2017-01-09 | Injury | STAAR SURGICAL COMPANY | 2017-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN EVO+ COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED A VISIBLE DISTORTION TO THE IRIS. NO FURTHER INFORMATION HAS BEEN PROVIDED. AT THE DATE OF MDR SUBMISSION, MULTIPLE ATTEMPTS TO CONTACT REPORTER HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VICM5_13.2, 13.2MM, -10.00 DIOPTER. PATIENT ALSO EXPERIENCED EXCESSIVE VAULT, LOSS OF BCVA, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND GLARE/HALOES. LENS REMAINS IMPLANTED. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9695111&pc=MTA | 2023826-2020-00291 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-11 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICH13.2, +9.50 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. EXCESSIVE VAULT, LOSS OF BCVA, ANGLE CLOSURE WITH ELEVATED IOP, PIGMENT DISPERSION, GLARE/HALOES AND BLURRED VISION WAS OBSERVED. THE REPORTER STATED " THE PATIENT HAD 2 PI WITH YAG BEFORE UNEVENTFUL SURGERY. HOWEVER, AT THE FIRST POST-OP DAY THE PATIENT HAD HIGH IOP WITH NARROW ANGLE AND ANTERIOR CHAMBER. ICL SHOWED HIGH VAULT THEN AFTER 3 DAYS FROM FIRST SURGERY HE PERFORMED SECOND SURGERY WITH A LITTLE ROTATING OF ICL, ADDITIONAL PI AND AC EVACUATION OF VISCO UNDER THE ICL BUT IT DIDNT SOLVE THE PROBLEM. PATIENT HAS +3D RESIDUAL HYPEROPIA." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H6- PATIENT CODE 2140 - 2227 SHOULD BE IN THE INITIAL MDR. H6- DEVICE CODE 1401 - 3001 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9746415&pc=MTA | 2023826-2020-00379 | 2019-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Blurred Vision; Visual Disturbances; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT, ELEVATED IOP, GLARE/HALOES AND BLURRED VISION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. RETURNED PRODUCT EVALUATION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322382&pc=MTA | 2023826-2021-03131 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Discomfort; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, UNREACTIVE (FIXED) PUPIL, ELEVATED IOP, AND DISCOMFORT. AFTER EXPLANT THE PUPIL AND IOP ARE BACK TO NORMAL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4625074&pc=MTA | 2023826-2015-00385 |  | Injury | STAAR SURGICAL COMPANY | 2015-03-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Discomfort; No Code Available | UNK | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) IN (B)(6), 2014. THE REPORTER INDICATED THE SIZE OF THE LENS WAS TOO BIG AND THE ANTERIOR CHAMBER WAS SHALLOW. THE INTRAOCULAR PRESSURE WAS A PROBLEM AND A 2ND PERIPHERAL IRIDECTOMY WAS PERFORMED. THE PATIENT EXPERIENCED SOME DISCOMFORT. THE LENS WILL BE EXPLANTED AND EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS IMPLANTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE EXCHANGE OF THE LENS TO A SHORTER LENS RESOLVED THE PROBLEM. Manufacturer Narrative: PER MEDICAL REVIEW - MOST LIKELY THE MICL LENS SIZING ERROR CAUSED THE EXCESSIVE LENS VAULTING AND CONSEQUENT IOP SPIKE ASSOCIATED WITH NARROW ANGLE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: PT AGE AND WEIGHT: UNK. EVENT DATE: UNK. EXPIRATION DATE - UNK. EXPLANT DATE: UNK. (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288508&pc=MTA | 2023826-2023-05482 | 2023-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VICM513.7 IMPLANTABLE COLLAMER LENS OF -4.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PUPIL BLOCK WITH ELEVATED IOP AND GLARES/HALOES ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10816077&pc=MTA | 2023826-2020-02662 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-11-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON 06-APR-2020. AT ONE DAY POST-OP, THE VAULT AND VISION WAS IMPROVED BUT WITH MYDRIASIS AND HALOS. THE PATIENTS INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-00945. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE PATIENT IS DOING WELL WITH MILD RESIDUAL ASTIGMATISM, THAT WE HOPE TO ADDRESS WITH POSSIBLY PRK. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10515451&pc=MTA | 2023826-2020-02022 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Optical Problem; No Apparent Adverse Event | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. PRESSURE SPIKES WERE OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONA INFORMATION: B5- THE REPORTER INDICATED THAT A 13.7MM MICL 13.7 OF -7.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS REPORTED. PATIENT STATED "SHE NOTICED HALOES AROUND ANY LIGHTS. THE LENS REMAINS IMPLANTED; HOWEVER THE SURGEON PERFORMED ENLARGEMENT OF PI; AND IT IS REPORTED THAT THE SURGEON HAS GONE BACK AND PERFORMED A SECOND IRIDOTOMY. PATIENT STATES NO PAIN OR VISION CHANGES. THE LEFT EYE PRESSURE IS OK. D6A: IMPLANT DATE: (B)(6) 2020. H1-SERIOUS . H6- CLINICAL CODE:4581-SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. HEALTH EFFECT:4625- ENLARGEMENT OF PI; IRIDOTOMY AND SECOND IRIDOTOMY. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 IN PREVIOUS MDR THE LENS DIOPTER REPORTED -7.0 IS CORRECTED TO -11.5. CLAIM# (B)(4). Manufacturer Narrative: NO APPARENT HARM OCCURRED IN RELATION TO THE ADVERSE EVENT. UNABLE TO ENTER CODE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11795155&pc=MTA | 2023826-2021-01396 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Optical Problem | Intraocular Pressure Increased; Halo; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | 1 | Event Description: IT WAS REPORTED THAT A 12.6MM, VICMO12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA BUT IT DID NOT RESOLVE THE PROBLEM. FOR STATUS OF THE EYE " PRESENCE OF HALOS; UNREACTIVE PUPIL" WAS REPORTED. " DOCTOR BELIEVES THAT HIGH IOP, MAY BE DUE TO CORTICOSTEROIDS OR MANIPULATION" WAS REPORTED FOR ADDITIONAL INFORMATION. ON (B)(6) 2021 THE REPORTER REPORTED "CURRENT CONDITION: PATIENT COMPLAINING OF GLARE; LITTLE REACTIVE PUPILS AND PIO 18. THE USE OF CORTICOSTEROIDS IS PART OF THE POSTOPERATIVE TREATMENT. A PROGRESSIVE REDUCTION WAS MADE AND STOPPED YESTERDAY, (B)(6) 2021. CURRENTLY USING ANTI-GLAUCOMA DORZOLAMIDE EYE DROPS AND BRIMONIDINE. THE DOCTOR WILL TRY, NEXT WEEK, TO ADAPT THE CONTACT LENSES TO REDUCE THE RESIDUAL DEGREE AND CHECK IF THE VISUAL COMPLAINT IMPROVES." ON (B)(6) 2021 THE REPORTER REPORTED " PATIENT IS A LITTLE MORE SATISFIED. LAST REVIEW ON (B)(6) 2021: DOCTOR HAD TO USE PILOCARPINE, TO CLOSE THE PUPIL A LITTLE, WHICH WAS 6MM. 0.73 VAULT. THE PATIENT REMAINS UNDER OBSERVATION AND EVALUATION WEEKLY". IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT "AFTER A WHILE THE GLARE DISAPPEARED." CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955095&pc=MTA | 2023826-2016-01266 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE HAPTIC WAS BROKEN. UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955096&pc=MTA | 2023826-2016-01267 | 2016-03-25 | Injury | STAAR SURGICAL COMPANY AG | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2011. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, GLARE AND HALOS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT WAS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6589772&pc=MTA | 2023826-2017-00878 | 2015-12-01 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+1.5/080 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND GLARE/HALOS. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS. (B)(6). (B)(4). Manufacturer Narrative: WORK ORDER SEARCH PERFORMED: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028643&pc=MTA | 2023826-2016-01460 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. FOREIGN MATERIAL - HAIR ON THE LENS SURFACE PRESENT. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6028650&pc=MTA | 2023826-2016-01459 | 2016-02-06 | Injury | STAAR SURGICAL COMPANY | 2016-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, AND GLARES/HALOES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN CASE/VIAL; VISUAL INSPECTION FOUND FOREIGN MATERIAL (SPOT) ON SURFACE, BUT NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6200082&pc=MTA | 2023826-2016-01824 | 2016-10-13 | Injury | STAAR SURGICAL COMPANY | 2016-12-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, GLARE/HALOES, AND ANGLE CLOSURE WITH ELEVATED IOP. AS OF THE DATE OF MDR FILING, THE LENS HAS NOT YET BEEN EXPLANTED AND REMAINS IMPLANTED IN THE PATIENT'S EYE. Manufacturer Narrative: ADDITIONAL INFORMATION: ADDITIONAL EVENT DESCRIPTION: LENS DIOPTER: -17.0/+1.5/70 (SPHERE/CYLINDER/AXIS). THE SURGEON ADMINISTERED IOP LOWERING MEDICATION. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988375&pc=MTA | 2023826-2020-00945 | 2020-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | Intraocular Pressure Increased; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE SURGEON HAD PERFORMED A LIMBAL RELAXING INCISION BUT IT WAS LEAKING SO THE SURGEON SUTURED THE INCISION CLOSED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS IMPROVED VISION AND VAULT BUT WITH MYDRIASIS AT DAY ONE AND PATIENT COMPLAINED OF HALOS. INTRAOCULAR PRESSURE WAS ELEVATED BUT WAS TREATED AGGRESSIVELY WITH DIAMOX AND TOPICAL DROPS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A ZIP LOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- "(B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE PREVIOUS MDR. H6- PATIENT CODE 1937 SHOULD BE REMOVED FROM INITIAL MDR. H6- PATIENT CODE 3191- (MYDRIASIS) SHOULD BE REMOVED FROM INITIAL MDR. H6- DEVICE CODE 3001 SHOULD BE REMOVED FROM INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5572295&pc=MTA | 2023826-2016-00437 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Unintended Movement | Intraocular Pressure Increased; Halo; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS DISLOCATED/SUBLUXED, WITH THE PATIENT EXPERIENCING ELEVATED INTRAOCULAR PRESSURE AND GLARES/HALOS. THE PATIENT COULD SEE A CRESCENT SHAPE ON THE RIGHT SIDE OF THEIR VISUAL FIELD, WHICH, WITH REFLECTION OF LIGHT, WAS MAKING IT IMPOSSIBLE FOR THEM TO DRIVE AT NIGHT. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7404316&pc=MTA | 2023826-2018-00506 | 2018-02-08 | Injury | STAAR SURGICAL COMPANY | 2018-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; Iritis; No Code Available | P030016 | 1 | Event Description: THE SURGEON INDICATED THAT HE IMPLANTED A 12.6MM MICL12.6, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 IRITIS WAS OBSERVED. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PERSISTENT IRITIS AND HIGH IOP (INTRAOCULAR PRESSURE). THE SURGEON REPORTED TO USE COMBIGAN, DUREZOL, PRED FORTE, AND LOTEMAX OVER THE COURSE OF TREATMENT. IT WAS REPORTED THAT AN IRIDECTOMY WAS PERFORMED. THE SURGEON REPORTED THAT HE BELIEVED THE CAUSE OF THE IRITIS WAS PIGMENT DISPERSION FROM ICL CHAFING OF POSTERIOR IRIS. ON (B)(6) 2018 THE SURGEON REPORTED THAT HE WAS NOT CERTAIN OF WHY THE INFLAMMATION CONTINUED. HE STATED THAT THERE WERE SOME PIGMENTARY CELLS, BUT MOST WERE WHITE CELLS AND THAT THE LENS SIZING WAS CORRECT. HE STATED THAT THE PATIENT IS DOING WELL IN SOFT CONTACT LENSES, INFLAMMATION IS RESOLVING. HE ALSO REPORTED THAT THEY WILL BE PERFORMING A PROPER FITTING FOR THE PATIENT IN THE FUTURE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A CARTRIDGE CASE. LENS WAS STUCK TO TAPE. VISUAL INSPECTION FOUND TAPE RESIDUE ON LENS. (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: PATIENT CODE:(B)(4) SECONDARY SURGICAL INTERVENTION,EXPLANT. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7631046&pc=MTA | 2023826-2018-00931 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-06-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; Iritis; Visual Disturbances; Halo; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM MICL13.2, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFTT EYE (OS) ON (B)(6) 2018. ON (B)(6) /2018 THE LENS WAS REMOVED DUE TO OBSERVATION OF EXCESSIVE VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE), SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE, PIGMENT DISPERSION, UNREACTIVE (FIXED) PUPIL, GLARE/HALOES, AND IRIS ATROPHY. ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT, ADDITION OF NEW PI AND REPOSITIONING) WERE PERFORMED. IT WAS REPORTED THAT AT PATIENTS LAST FOLLOW UP ((B)(6) 2018) THE PATIENT WAS ON FOUR DIFFERENT GLAUCOMA DROPS, IOP 25MMHG. ADDITIONAL INFORMATION REPORTED, PI PATENT, PATIENT REPORTS DECREASED GLARE. EXPLANTED ICL AFTER ROTATION, IRIS CELL AND FLARE INCREASED THE IOP. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A THERMA SCIENTIFIC CONTAINER. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON LENS SURFACE (RESIDUE ON LENS). (B)(4). Manufacturer Narrative: ADDITION OF NEW PI WAS NOT REPORTED. . Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT HAD IRITIS DUE TO THE RUBBING OF THE LENS. THE LENS WAS ROTATED AND THE IRITIS STILL CONTINUED. ANTIBIOTICS AND ANTI-INFLAMMATORIES WERE USED TO ADDRESS THE INFLAMMATION. THE PROBLEM OF THE CELL AND FLARE WAS RESOLVED AFTER THE REMOVAL OF THE ICL, THERE WERE NO CELL AND FLARE POST REMOVAL. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5458648&pc=MTA | 2023826-2016-00197 |  | Injury | STAAR SURGICAL COMPANY | 2016-02-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Loss of Vision; Discomfort; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED AN MICL IMPLANTABLE COLLAMER LENS IN THE PATIENT'S EYE. AT THE SAME DAY OF SURGERY AND ONE DAY POST-OP THE PATIENT WAS DOING WELL AND VA WAS 20/15. THE PATIENT TRAVELLED OUT OF STATE THAT SAME WEEK OF SURGERY AND REPORTED A SLIGHT DECLINE IN VISION AND DISCOMFORT TO THE EYE. THE PATIENT WAS PRESCRIBED A STEROID AND PRESSURE DROP. THE PATIENT WAS ADVISED TO SEE A LOCAL DOCTOR AND PRESENTED WITH ANGLE CLOSURE. THE PATIENT RETURNED TO THE SURGEON THE FOLLOWING WEEK AND PRESENTED WITH A SHALLOW CHAMBER. THE SURGEON HAD DIFFICULTY OPENING AN ADDITIONAL PI, TEMPORIZING HER PRESSURE BUT THE PATIENT WANTED THE ICL REMOVED. THE LENS WAS EXPLANTED AND THE PATIENT HAS RECOVERED HER VISION WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INFORMATION RECEIVED - THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7875111&pc=MTA | 2023826-2018-01395 | 2018-08-15 | Injury | STAAR SURGICAL COMPANY | 2018-09-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Missing Value Reason | MTA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2018 WITH A SHORTER LENGTH LENS DUE TO PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE), TOO LARGE LENS WITH AC SHALLOWING, HIGH VAULTING AND NARROWING OF ANGLE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). LENS WAS RETURNED DRY, IN A SPECIMEN CONTAINER. VISUAL INSPECTION FOUND A TORN HAPTIC AND RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005327&pc=MTA | 2023826-2021-01927 | 2021-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Nausea; Pain; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE SURGEON ENLARGED THE PIS. THE PATIENT COMPLAINED OF EYE PAIN AND NAUSEA. ANOTHER LENS WAS NOT IMPLANTED. PATIENT WAS TAKING KRATOM OR POSSIBLY OTHER MEDS/DRUGS, PT ADMITTED TO USING "EDIBLE" CBD GUMMY AND KRATOM. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A CONTACT LENS CASE. VISUAL INSPECTION FOUND A HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6961181&pc=MTA | MW5072842 | 2016-09-27 | Injury | UNK | 2017-10-19 | MTA | IMPLANTABLE CONTACT LENSES | Fitting Problem | Intraocular Pressure Increased; Nausea; Pain; Loss of Vision; Discomfort |  | 1 | Event Description: I RECEIVED ICL IMPLANTABLE CONTACT LENSES IN BOTH EYES, IN (B)(6) 2016. APPROX. A WEEK AFTER THE SECOND (RIGHT EYE) WAS IMPLANTED, I STARTED HAVING A LOT OF EYE PAIN/ PRESSURE AND REDUCED VISION. I MADE A CALL TO THE DOCTOR'S OFFICE WHERE THEY SAID TO TAKE ANTIINFLAMMATORIES. AFTER ABOUT AN HOUR OF THOSE NOT WORKING, I BECAME VERY NAUSEOUS AND WE WERE ADVISED TO COME INTO THE OFFICE FOR AN EMERGENCY PROCEDURE. UPON ARRIVAL, THE EYE SURGEON ON CALL DETERMINED THE IRIDOTOMY IN MY RIGHT EYE HAD CLOSED, AND I WAS EXPERIENCING ACUTE ANGLE CLOSURE. HE CUT INTO MY CORNEA TO RELIEVE THE PRESSURE AND PROCEEDED TO ZAP MY EYE WITH TWO DIFFERENT KINDS OF LASERS, UPWARDS OF 30 TIMES. MY EYES HAVE NEVER FELT THE SAME SINCE AND ALTHOUGH THE PRESSURE WAS RECEIVED. THE IMPLANTS STILL NEEDED TO COME OUT BECAUSE THE IRIS IN MY RIGHT EYE WAS BOWING FROM THE FIT OF THE LENSES AND CAUSING ADDITIONAL PAIN AND DISCOMFORT. THEY WERE BOTH REMOVED IN (B)(6) 2016. WHY THE FDA HAS APPROVED THESE LENSES, THAT HAVE BEEN SHOWN TO CAUSE CATARACTS IN UP TO 50 PERCENT OF PATIENTS, WITHIN 5 YEARS OF IMPLANT, IS BEYOND ME.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832459&pc=MTA | 2023826-2017-01382 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected | Intraocular Pressure Increased; Nausea; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT RETURNED THE NEXT DAY, (B)(6) 2017, IN PAIN, NAUSEOUS, WITH A PRESSURE SPIKE (60MMHG) AND STEAMY EYE. ANOTHER IRIDOTOMY WAS PERFORMED AND PRESSURE DROPPED DOWN TO 40MMHG. THE CHAMBER WAS SHALLOW, WITH SOME ANGLE CLOSURE AND THE SURGEON DECIDED TO EXPLANT THE LENS THE SAME DAY. A SHORTER LENS (13.2MM) WAS IMPLANTED BUT THE PRESSURE AGAIN SPIKED AND THE LENS WAS EXPLANTED SAME DAY ((B)(6) 2017) - SEE MDR #2023826-2017-01383. A THIRD LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. THE PATIENT IS DOING FINE. Manufacturer Narrative: (B)(4). DEVICE MANUFACTURER DATE: IT IS CORRECTED TO 23-DEC-2016. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777874&pc=MTA | 2023826-2023-04107 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.00 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THE REPORTER STATED " PREOPERATIVE BASAL IOP WAS HIGHER THAN NORMAL, AND NO FURTHER EXAMINATION WAS PERFORMED TO EXCLUDE GLAUCOMA. AFTER THE CRYSTAL IS REMOVED, THE INTRAOCULAR PRESSURE IS TEMPORARILY STABILIZED TO NORMAL VALUE AFTER DRUG TREATMENT. THE LENS WAS REMOVED AND TRANSFERRED TO THE GLAUCOMA DEPARTMENT FOR FURTHER CONSULTATION." Manufacturer Narrative: H6- HEALTH EFFECT- CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939033&pc=MTA | 2023826-2021-01672 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2, -5.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, OBJECTIVE SLE: TRACE DMF. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN SPECIMEN CUP. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. H6: CLINCAL CODE - 1937 CORRECTED TO 4582. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986305&pc=MTA | 2023826-2021-01864 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986314&pc=MTA | 2023826-2021-01865 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), ANGLE CLOSURE, AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4,
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669504&pc=MTA | 2023826-2021-04033 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). AQUEOUS WAS RELEASED FROM THE CHAMBER. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B3- CORRECTED TO (B)(6) 2021. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. AQUEOUS WAS RELEASED FROM THE CHAMBER. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ELEVATED INTRAOCULAR PRESSURE (IOP) WAS REPORTED AS A POST-OPERATIVE CONDITION. THE CAUSE OF THE EVENT WAS UNKNOWN. H6- HEALTH EFFECT- CLINICAL CODE: REMOVED CODE 1937 AND REPLACED WITH CODE 4582. CLAIM# (B)(4). Manufacturer Narrative: B5: PER EMAIL ON 22-FEB-2022, IT WAS CLARIFIED THAT ELEVATED IOP WAS EXPERIENCED AFTER THE INITIAL LENS WAS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082666&pc=MTA | 2023826-2021-04919 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS WAS REPORTED AS HAVING A HIGH VAULT AND POST-OP A PRESSURE SPIKE WAS REPORTED DUE TO NON PATENT IRIDOTOMY. THE PATIENT WAS BROUGHT BACK TO SURGERY TO SURGICALLY OPEN THE PI. THE PRESSURE DROPPED FROM 57MMHG TO 7MMHG. THE LENS REMAINED IMPLANTED. THE PATIENT IS SEEING 20/20 AND IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS WAS REPORTED AS HAVING A HIGH VAULT AND POST-OP A PRESSURE SPIKE WAS REPORTED DUE TO NON PATENT IRIDOTOMY. THE PATIENT WAS BROUGHT BACK TO SURGERY TO SURGICALLY OPEN THE PI. THE PRESSURE DROPPED FROM 57MMHG TO 7MMHG. THE LENS REMAINED IMPLANTED. THE PATIENT IS SEEING 20/20 AND IS VERY HAPPY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, 10.00 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H6: HEALTH EFFECT: CLINICAL CODE: 1937, THERE WAS NO INTRAOCULAR PRESSURE INCREASE. CORRECT CODE IS 4582. H6: HEALTH EFFECT: IMPACT CODE: 4625, THERE WAS NO ADDITIONAL SURGERY. CORRECT CODE IS 4629. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK ETHNICITY: UNK RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANTED: UNK. EXPLANTED: UNK. MANUFACTURE DATE: UNK. (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK ETHNICITY: UNK RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANTED: UNK. EXPLANTED: UNK. MANUFACTURE DATE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14401934&pc=MTA | 2023826-2022-01487 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL13.7 -5.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO HIGH VAULT AND REPLACED WITH A SHORTER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM MICL13.7 -5.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO HIGH VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS LATER REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS "OTHER", FOR CAUSE DETAILS THE REPORTER REPORTED " HIGH VAULT AND NARROW ANGLES". CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099731&pc=MTA | 2023826-2022-02398 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5 DIOPTER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295510&pc=MTA | 2023826-2023-02791 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT." SHOULD BE CORRECTED TO "THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP." IN THE INITIAL MDR. H6 - HEALTH EFFECT CLINICAL CODE 1937: INTRAOCULAR PRESSURE INCREASED. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321642&pc=MTA | 2023826-2021-03144 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER STATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE AND THE PATIENT EXPERIENCED VAULT, ELEVATED IOP AND SHALLOW CHAMBER. THE PATIENT WAS DILATED WITH MYDRIACYL TO SEE IF IT WOULD BREAK OR DEEPENED THE CHAMBER. IT WAS REPORTED THAT ONCE PUPIL DILATED BEYOND EDGE OF ICL THE PRESSURE DROPPED, VAULT INCREASED AND THE CHAMBER DEEPENED. DOCTOR FEELS THE PRESSURE ISSUE IS DUE TO OCUCOAT BLOCKING PIS. (B)(6) 2021 EMAIL FROM DOCTOR STATES "OUR PATIENT IS DOING WELL. HER PRESSURE WAS GREAT TODAY AND HER PUPIL IS ALMOST BACK TO NORMAL. I DON'T THINK SHE WILL NEED AN EXCHANGE". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285150&pc=MTA | 2023826-2022-02891 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO A HIGH VAULT. A REPLACEMENT LENS WAS IMPLANTED OF UNKNOWN MODEL AND SIZE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6-PRODUCT EVALUATION: LENS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: B5: ELEVATED INTRAOCULAR PRESSURE WITH MEDICAL INTERVENTION WAS ALSO REPORTED. H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16669399&pc=MTA | 2023826-2023-01126 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | UNK | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. REPORTEDLY, "IOP SPIKE IN ONE EYE A FEW WEEKS POST OP. IT IS MY UNDERSTANDING IT WAS DUE TO A HIGH VAULT AND WAS RECTIFIED BY A LENS EXCHANGE." Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS OF DIOPTER INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT AND ELEVATED IOP. REPORTEDLY "IOP SPIKE IN ONE EYE A FEW WEEKS POST OP. IT IS MY UNDERSTANDING THAT IT WAS DUE TO A HIGH VAULT AND WAS RECTIFIED BY A LENS EXCHANGE." THE LENS WAS EXPLANTED AND REPLACED. H6-HEALTH EFFECT- IMPACT CODE: "1937" SHOULD BE ADDED. H6- HEALTH EFFECT- CLINICAL CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561598&pc=MTA | 2023826-2023-03510 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A REPLACEMENT 12.6MM VICM512.6 IMPLANTABLE COLLAMER LENS OF A -4.5 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)6) 2023. THE PROBLEM WAS RESOLVED. REPORTEDLY, "A FEW STRIAE ON THE ENDOTHELIUM." Manufacturer Narrative: A4-A6:UNK . H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=MTA | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465175&pc=QCB | 2023826-2023-00767 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:1494-OFF LABEL: PROGRESSIVE MYOPIA. D9: RETURN DATE- 05-APR-2023 CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=MTA | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595289&pc=QCB | 2023826-2023-00985 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. THE REPORTER NOTES A PLANNED LENS EXPLANT AND REPLACEMENT DUE TO EXCESSIVE VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.50 DIOPTER WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494-OFF LABEL- PROGRESSIVE MYOPIA D9: RETURN DATE: 05-APR-2023 (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17684361&pc=MTA | 2023826-2023-03760 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT AND ELEVATED IOP. LATER ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11065610&pc=MTA | 2023826-2020-03104 | 2020-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-12-23 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | Intraocular Pressure Increased; No Clinical Signs, Symptoms or Conditions | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN AND REMOVED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "ANOTHER LENS WILL BE IMPLANTED BUT THE TIMING OF THE OPERATION IS UNCERTAIN." Manufacturer Narrative: ADDITIONAL INFORMATION. D10 - INJECTOR MODEL#MSI-PF,LOT#UNK, CARTRIDGE MODEL#SFC-45,LOT#UNK, FOAMTIPPLUNGER MODEL#FTP,LOT#UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2020. THE LENS WAS IMPLANTED AND REMOVED INTRA-OP WITHIN THE SAME SURGERY DUE TO THE LENS IMPLANTED UPSIDE DOWN. AT POST-OP ELEVATED IOP WAS OBSERVED DUE TO POSTOPERATIVE MEDICATION. "AFTER ADJUSTING THE MEDICATION, THE PATIENT FELT NORMAL AND DID NOT RETURN TO THE HOSPITAL FOR EXAMINATION AGAIN." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6412058&pc=MTA | 2023826-2017-00259 | 2017-01-24 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON PARTIALLY INSERTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/102 DIOPTER, BUT DUE TO INTRAOCULAR POSITIVE PRESSURE, THE SURGEON COULD NOT IMPLANT THE LENS COMPLETELY. THE REPORTER INDICATED LOOSE ZONULES WERE SUSPECTED AND THE ANTERIOR CHAMBER COLLAPSED. THE LENS WAS REMOVED WITHOUT ANY PATIENT INJURY OR COMPLICATIONS AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS RETURNED IN THREE PIECES. THERE WAS RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913311&pc=MTA | 2023826-2020-00744 | 2018-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PROGRESSIVE INCREASE IN INTRAOCULAR PRESSURE (IOP). THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURIGCAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7860761&pc=MTA | 2023826-2018-01247 | 2018-08-10 | Injury | STAAR SURGICAL COMPANY | 2018-09-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. UPON ASSING ON (B)(6) 2018, ELEVATED IOP OF 39 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. LENS WAS EXPLANTED ON (B)(6) 2018 WITH NO PLANS TO EXCHANGE LENS. AFTER REMOVAL OF LENS, THE PATIENT'S INTRAOCULAR PRESSURE RETURNED TO NORMAL. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8077834&pc=MTA | 2023826-2018-01630 | 2018-08-16 | Injury | STAAR SURGICAL COMPANY | 2018-11-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.00/+1.0/X084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) (13MMHG), WHICH CAUSED THE IRIS TO PROLAPSE (POSTERIOR SYNECHIAE) THROUGH THE WOUND. POST-OP, THE PATIENT STILL HAS ELEVATED IOP AND POSTERIOR SYNECHIAE. IT IS UNKNOWN IF ANOTHER ICL LENS WILL BE IMPLANTED. THE REPORTER INDICATED THE SURGEON FELT THE MEASUREMENTS MAY HAVE BEEN INCORRECT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976037&pc=MTA | 2023826-2018-01513 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/025 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 28 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7976051&pc=MTA | 2023826-2018-01512 | 2018-08-20 | Injury | STAAR SURGICAL COMPANY | 2018-10-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/125 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. ON (B)(6) 2018, THE PATIENTS INTRAOCULAR PRESSURE (IOP) WAS ELEVATED AT 21 - 38 MMHG. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES AND IOP WAS 28 MMHG. THE PATIENT WAS PRESCRIBED MEDICATION (SIMBRINZA) AND IOP WAS 23MMHG. THE SURGEON HAS DECIDED TO REDUCE THE MEDICATIONS AND WAIT AND SEE HOW THE IOP DEVELOPS. IF THE IOP INCREASES, HE WILL REPLACE THE LENS. OTHERWISE, THE LENS WILL REMAIN IN THE EYE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337166&pc=MTA | 2023826-2019-00248 | 2018-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND UNREACTIVE PUPIL (FIXED). THE SURGEON ATTEMPTED TO FIX THE PUPIL BLOCK SURGICALLY ON (B)(6) 2018 BUT WAS UNSUCCESSFUL. A SYNECHIA WAS LATER DISCOVERED AT 4 O'CLOCK POSITION. THE PUPIL IS STILL NOT REACTING TO LIGHT AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO FIX THE PUPIL BLOCK SURGICALLY - THE SURGEON WENT IN THROUGH THE TEMPORAL INCISION AND FLUSHED WITH ANAESTHESIA AND MIOCHOL AND LOOSENED A SYNECHIA. THERE WERE CHECKUPS TO SEE IF THERE WERE ANY CHANGES. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215221&pc=MTA | 2023826-2019-00021 | 2018-11-21 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 15.50/+3.5/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED AFTER 1 MONTH POST-OP, THE SURGEON NOTICED RAISED INTRAOCULAR PRESSURE (IOP), WITH OPEN ANGLE BUT SHALLOW CHAMBER. THERE WAS SUSPECTED STEROID RESPONSE, AND THE SURGEON STARTED ANTIGLAUCOMA. AFTER 2 WEEKS, THERE WAS NO IMPROVEMENT TO IOP, SO THE SURGEON PERFORMED LASER PI. THE IOP SHOT UP AFTER 1 WEEK TO 32MMHG AFTER LASER PI. STARTED ANTIGLAUCOMA, BOTH TOPICAL AND SYSTEMIC. AT PRESENT ONLY TOPICAL ANTIGLAUCOMA AND IOP IS 22MMHG. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS IMPLANTED ON (B)(6) 2018 AND EXPLANTED ON (B)(6) 2018 DUE TO ELEVATED INTRAOCULAR PRESSURE AND ANGLE CLOSURE. A YAG WAS PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENTS CURRENT CONDITION WAS ANGLE OPENED AND ANTERIOR CHAMBER WELL FORMED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287616&pc=MTA | 2023826-2019-00180 | 2018-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO INCREASED POST-OP PRESSURE. ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY IS NOTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN GAUZE. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8359916&pc=MTA | 2023826-2019-00291 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-21 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -9.00/1.5/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO ELEVATED IOP (INTRAOCULAR PRESSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, IOP AND IRIDO-CORNEAL ANGLES RETURNED TO NORMAL. THE DOCTOR HAS NO PLANS TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8441979&pc=MTA | 2023826-2019-00438 | 2019-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO PUPIL BLOCK, WITH ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668952&pc=MTA | 2023826-2019-01045 | 2019-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593189&pc=MTA | 2023826-2019-00750 | 2019-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +05.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593164&pc=MTA | 2023826-2019-00749 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +05.00 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON 04/03/2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AND THE LENS HAPTIC HAS A TEAR. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8563071&pc=MTA | 2023826-2019-00706 | 2019-04-03 | Injury | STAAR SURGICAL COMPANY | 2019-04-29 | MTA | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2019. REPORTEDLY, POSTOPERATIVELY THE PATIENT DEVELOPED A PRESSURE SPIKE. THE SURGEON OPENED THE LASER IRIDOTOMY AND THE PRESSURE WENT DOWN TO 18MMHG. THE PUPILLARY BLOCK WAS THEN BROKEN. PUPIL IS NOW REACTIVE AND SURGEON BELIEVES THERE WILL BE MINIMAL DAMAGE FROM ISCHEMIA. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8693232&pc=MTA | 2023826-2019-01073 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | UNK | 1 | Event Description: THE REPORTER STATED " INCREASED INTRAOCULAR PRESSURE IN THE RIGHT EYE FOLLOWING ICL SURGERY ON (B)(6) 2018, PRESUMED BY SURGEON TO BE DUE TO RETAINED OPHTHALMIC VISCOELASTIC DEVICE- REQUIRING MEDICAL INTERVENTION THAT HE MANAGED IN HIS CLINIC REPORTING RESOLUTION ON (B)(6) 2019". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: REPORTER MISTAKENLY STATED THAT THE IMPLANT TOOK PLACE ON (B)(6) 2018 . CORRECT IMPLANT DATE IS (B)(6) 2018 . (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9152716&pc=MTA | 2023826-2019-01856 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED PRESSURE THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI'S AND MEDS.ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -14.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI AND MEDICATION. THE REPORTER STATED " INCREASED IOP D/T EITHER ANGLE CLOSURE FROM ANTERIORLY ROTATED CILIARY BODY VS ATYPICAL AQUEOUS MISDIRECTION." STATUS OF THE EYE: DOING WELL AFTER ICL REMOVAL. IT WAS REPORTED THAT THE ANGLES ARE OPEN. IOP IS NORMAL. NO LOSS OF BEST CORRECTED VISION BUT PATIENT IS BACK TO WEARING GLASSES. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EVENT: "(B)(6) 2019" SHOULD BE IN THE PREVIOUS MDR. CONCOMITANT MEDICAL PRODUCTS: "INJECTOR MODEL: MSIPF - INJ LOT: 1449232. CARTRIDGE MODEL: SFC45 - CART LOT: 1446792. INJECTOR SYSTEM MODEL: FTP - INJ SYSTEM LOT: 1448653" SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9153161&pc=MTA | 2023826-2019-01857 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED PRESSURE THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI'S AND MEDS.ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: EVENT: THE REPORTER INDICATED THAT 13.2MM, MICL13.2, -15.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 05 MAR 2019. ON 21 MAY 2019 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP THAT COULD NOT BE RESOLVED WITH ADDITIONAL PI AND MEDICATION. THE REPORTER STATED "INCREASED IOP D/T EITHER ANGLE CLOSURE FROM ANTERIORLY ROTATED CILIARY BODY VS ATYPICAL AQUEOUS MISDIRECTION." STATUS OF THE EYE: DOING WELL AFTER ICL REMOVAL. IT WAS REPORTED THAT THE ANGLES ARE OPEN. IOP IS NORMAL. NO LOSS OF BEST CORRECTED VISION BUT PATIENT IS BACK TO WEARING GLASSES. PATIENT CODE 3191: ANGLE CLOSURE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- "05/20/2019" SHOULD BE IN THE PREVIOUS MDR. D11- "INJECTOR MODEL: MSIPF - INJ LOT: 1449232; CARTRIDGE MODEL: SFC45 - CART LOT: 1446792; INJ SYSTEM MODEL: FTP - INJ SYSTEM LOT: 1448653" SHOULD BE IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841538&pc=MTA | 2023826-2019-01378 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -7.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THREE WEEKS POST OP THE SURGEON NOTICED DILATED PUPIL. ELEVATED IOP WAS NOTICED ON (B)(6) 2019. SURGICAL INTERVENTION (YAG PIS) AND MEDICAL INTERVENTION (PREDNISOLONE STEROIDS, MIOTICS, MYDRIATICS, MANNITOL, AND ORAL MEDICATION) WERE PERFORMED/PRESCRIBED BUT DID NOT RESOLVE THE ISSUE, THE REPORTER STATED " FROM (B)(6) 2019 NOT RESPONDING, EVEN TO MYDRIATICS, FORM (B)(6) 2019 IOP NOT RESPONDING FOR ANY MEDICATION". FOR LAST STATUS OF PUPIL AND IOP THE REPORTER STATED "PRESSURE IS GETTING UNDER CONTROL AND CONTINUING TOPICAL ANTI GLAUCOMA DROPS AND PATIENT IS STILL UNDER OBSERVATION. PUPIL IS DILATED AND FIXED, NOT RESPONDING TO LIGHT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841649&pc=MTA | 2023826-2019-01379 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5_12.6, -6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019, THREE WEEKS POST OP THE SURGEON NOTICED DILATED PUPIL. ELEVATED IOP WAS NOTICED ON (B)(6) 2019. SURGICAL INTERVENTION (YAG PIS) AND MEDICAL INTERVENTION (PREDNISOLONE STEROIDS, MIOTICS, MYDRIATICS, MANNITOL, AND ORAL MEDICATION) WERE PERFORMED/PRESCRIBED BUT DID NOT RESOLVE THE ISSUE, THE REPORTER STATED " FROM (B)(6) 2019 NOT RESPONDING, EVEN TO MYDRIATICS, FORM (B)(6) 2019 IOP NOT RESPONDING FOR ANY MEDICATION". FOR LAST STATUS OF PUPIL AND IOP THE REPORTER STATED "PRESSURE IS GETTING UNDER CONTROL AND CONTINUING TOPICAL ANTI GLAUCOMA DROPS AND PATIENT IS STILL UNDER OBSERVATION. PUPIL IS DILATED AND FIXED, NOT RESPONDING TO LIGHT". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222176&pc=MTA | 2023826-2020-01405 | 2019-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222185&pc=MTA | 2023826-2020-01406 | 2019-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9507349&pc=MTA | 2023826-2019-02475 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO HIGH PRESSURE. THERE WAS NO PATIENT INJURY. THE REPORTER STATED THAT THE SURGEON WILL BE REPLACING THE LENS WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: THE LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DEVICE WAS RETURNED IN A SPECIMEN CUP, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492475&pc=MTA | 2023826-2019-02445 | 2019-11-19 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -7.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AND THE PATIENT EXPERIENCED ELEVATED IOP. A VAULT VALUE WAS REPORTED AND PUPIL BLOCK WITH ELEVATED IOP (28 MMHG). ADDITIONAL SURGICAL INTERVENTION (YAG, IRIDOTOMIA, IRIS SUTURE) WERE PERFORMED. IT WAS REPORTED THAT THE IOP WAS DECREASED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222180&pc=MTA | 2023826-2020-01410 | 2020-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222182&pc=MTA | 2023826-2020-01409 | 2020-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484156&pc=MTA | 2023826-2020-02011 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -2.25 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO PUPIL BLOCK WITH ELEVATED IOP AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT THE PATIENT FELT UNCOMFORTABLE AND WANTED THE LENS REMOVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6341607&pc=MTA | 2023826-2017-00241 |  | Injury | STAAR SURGICAL COMPANY | 2017-02-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S EYE AND IT WAS REPORTED THE INTRAOCULAR PRESSURE WAS ELEVATED AND THE ANGLE WAS TOO NARROW. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595442&pc=MTA | 2023826-2019-00764 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-09 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS ELEVATED IOP, IRIDOTOMIES FOUND TO NOT BE OPENED. IOP INITIALLY EXCEEDED 50MMHG AFTER TREATMENT WITH DIAMOX. THE PATIENT IS A STEROID RESPONDER. YAG WAS PERFORMED ON (B)(6) 2019, AS ONE OF THE PI'S WAS BLOCKED, POSSIBLY DUE TO IRIS PIGMENTATION MEMBRANE OR "JUST NOT OPEN ENOUGH." PRESSURE WENT DOWN BUT STILL "NO GOOD." NO ADDITIONAL INFORMATION CURRENTLY AVAILABLE. THE LENS CURRENTLY REMAINS IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8853053&pc=MTA | 2023826-2019-01325 |  | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - INTRAOCULAR ELEVATION IN 11 EYES. MEDICATION WAS PROVIDED AND SURGICAL INTERVENTION WAS NOT REQUIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: 'AN ARTICLE WAS RECEIVED THAT CITED A RETROSPECTIVE NON-RANDOMIZED CLINICAL STUDY THAT FOUND PATIENTS IMPLANTED WITH THE IMPLANTABLE COLLAMER LENS (ICL) WITH THE AIM OF CORRECTING MYOPIA OR MYOPIC ASTIGMATISM. POSTOPERATIVELY, THE FOLLOWING COMPLICATIONS WERE OBSERVED - INTRAOCULAR ELEVATION IN 9 EYES. MEDICATION WAS PROVIDED AND SURGICAL INTERVENTION WAS NOT REQUIRED. IN TWO PATIENTS, LOCAL ANTIGLAUCOMATOUS THERAPY IS BEING ADMINISTERED LONG TERM (IN ONE PATIENT THERE IS A POSITIVE FAMILY HISTORY OF GLAUCOMA)'. THE CAUSE OF THE IOP ELEVATION IN THESE CASES WAS ATTRIBUTED TO 'THE EFFECT OF LOCAL CORTICOSTEROIDS IN COMBINATION WITH PRE-OPERATIVE VALUES AROUND THE UPPER LIMIT OF THE NORM (UP TO 21 MMHG)'. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED.' CORRECTED TO PHAKIC INTRAOCULAR LENS & PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: MTA & QCB (COHORT INCLUDED SEVERAL LENS MODELS). DEVICE INFORMATION CORRECTED TO ICM V4, VICMO, TICM V4, VTICM, VTICMO (COHORT INCLUDED SEVERAL LENS MODELS). REPORT SOURCE SHOULD HAVE INCLUDED LITERATURE IN INITIAL MDR. SHOULD HAVE INCLUDED PATIENT CODE 3191 (SECONDARY INTERVENTION, ELEVATED IOP TREATED WITH MEDICATION) IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8998372&pc=MTA | 2023826-2019-01702 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: ON 14 AUG 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN SOCIEDAD ESPANOLA DE OFTALMOLOGIA, ' INTRAOCULAR PRESSURE FLUCTUATIONS IN PATIENTS IMPLANTED WITH AN IMPLANTABLE COLLAMER LENS (ICL V4C). THREE-MONTH FOLLOW-UP.' THE STUDY FOUND THAT THE PERCENTAGE OF OHT WAS 6.6% (2 EYES) OF THE OVERALL SAMPLE, WHICH REQUIRED TREATMENT WITH TOPICAL HYPOTENSORS DURING 8 WEEKS AND RESOLVED IN APPROXIMATELY 10 WEEKS WITHOUT GLAUCOMATOUS STRUCTURAL OR FUNCTIONAL DAMAGES . IT WAS REPORTED THAT AS OF (B)(6) 2019, ALL LENSES REMAINED IMPLANTED AND NO FURTHER PROBLEMS WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 3191 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257786&pc=MTA | 2023826-2019-02033 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF TALMOLOGIA CLINICA Y EXPERIMENTAL IN SEPTEMBER 2019 TITLED, 'IMPLANTE DE LENTE INTRAOCULAR FÁQUICA DE CÁMARA POSTERIOR EN OJOS CON QUERATOCONO: EFECTO A LARGO PLAZO SOBRE LA AGUDEZA VISUAL.' THE ARTICLE REPORTS THREE CASES OF OCULAR HYPERTENSION, IN IMMEDIATE CONTROL AT 2 HOURS OF THE OPERATION. INFORMATION RECEIVED FROM ONE OF THE AUTHORS INDICATED " 3 CASES WITH TRANSIENT OCULAR HYPERTENSION AT 2- HOURS POSTOP CONTROL WAS NO HIGHER THAN 35-40 MMHG. TWO PATIENTS WERE TREATED WITH ORAL ACETAZOLAMIDE AND THE OTHER ONE (IOP UNDER 25 MMHG) WITH DROPS. IN THE 3 CASES ON CONTROL AT POSTOP DAY 1 THE TREATMENT WAS SUSPENDED AS IOP WAS UNDER 18 MMHG. NO CONSEQUENCES FROM HYPERTENSION WERE OBSERVED (IE, URRETZ-ZAVALA'S SYNDROME). THE HYPERTENSION WAS NOT A PROBLEM TO MANAGE NOR IMPLIED ANY VISUAL LOSS FOR THE PATIENTS." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT:UNK. (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. IMPLANT DATE:UNK. EXPLANT DATE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURER DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9485143&pc=MTA | 2023826-2019-02444 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM,VICM5_13.7, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND PI HAD TO BE PERFORMED TO BRING THE PRESSURE UNDER CONTROL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2019. SHOULD BE ADDED IN THE INITIAL MDR. D6- (B)(6) 2019 SHOULD BE IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492665&pc=MTA | 2023826-2019-02446 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: A CASE OF INCREASING IOP IN THE FIRST WEEK AFTER ICL IMPLANTATION WAS REPORTED. ADDITIONAL SURGICAL INTERVENTION (IRIDOTOMIA) WAS PERFORMED AND IOP DECREASED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9507438&pc=MTA | 2023826-2019-02476 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO ACUTE ANGLE CLOSURE AND INCREASED IOP. THE REPORTER STATED THAT THE SURGEON PLANS TO REPLACE THE LENS FOR A SMALLER SIZE (12.6MM). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK .RACE: UNK. DATE OF EVENT: UNK. (EXPIRATION DATE): UNK. NO SERIAL NUMBER REPORTED. EXPLANTED: UNK. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075186&pc=MTA | 2023826-2020-01096 |  | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY OF 49 EYES OF MYOPIC PATIENTS IMPLANTED WITH HOLE ICLS. OF THE 49 EYES, THE SURGERY FOR 21 EYES USED OVD (OPHTHALMIC VISCOSURGICAL DEVICE) AND FOR 28 EYES, OVD WAS NOT USED. A RISE IN INTRAOCULAR PRESSURE (IOP) WAS NOTED IN 14 OF THE 21 EYE GROUP (USED OVD) AND NO RISE IN IOP FOR 28 EYE GROUP. THE INTRAOCULAR PRESSURE WAS LOWERED BY APPROPRIATE MANAGEMENT, INCLUDING INVASIVE INTERVENTION SUCH AS ANTERIOR CHAMBER DRAINAGE. SOME ENDOTHELIAL CELL LOSS WAS NOTED FOR 28 EYE GROUP. NO OTHER COMPLICATIONS WERE FOUND. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ALSO REPORTED SOME SURGICAL INTERVENTION SUCH AS LENS EXCHANGE WAS ALSO PERFORMED. CLAIM# : (B)(4). Manufacturer Narrative: EXPIRATION DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10484825&pc=MTA | 2023826-2020-02012 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: ON 04 AUG 2020 WE RECEIVED NOTIFICATION OF AN ARTICLE ARTICLE ENTITLED,'EVALUATION OF CHANGES IN CHOROIDAL THICKNESS AFTER IMPLANTABLE COLLAMER LENS SURGERY IN HIGH MYOPIA PATIENTS WITH GRAVES' THE ARTICLE REPORTS " THERE WERE FOUR EYES WITH AN IOP GREATER THAN 21MMHG IN THE GROUP A AND SEVEN EYES IN GROUP B. THE IOP RETURNED TO BASELINE WHEN USING THE LOWER INTRAOCULAR PRESSURE DRUGS AFTER SURGERY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBERS REPORTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491988&pc=MTA | 2023826-2020-01958 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S EYE. THE SURGEON REPORTS ONE DAY POSTOPERATIVE ELEVATED IOP AND SHALLOW CHAMBER. REPORTEDLY, THE PROBLEM IS RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR UNKNOWN ENTRY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9169871&pc=MTA | 2023826-2019-01893 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2019. THE REPORTER REPORTED PRE-OP AND POST OP ENDOTHELIAL CELL COUNT, ELEVATED IOP THAT LASTED 4 HOURS AND INCISION ENLARGEMENT. THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT A 12.6MM MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS TORE BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2019. THE REPORTER REPORTED PRE-OP AND POST OP ENDOTHELIAL CELL COUNT, ELEVATED IOP THAT LASTED 4 HOURS AND INCISION ENLARGEMENT. THE LENS WAS REMOVED AND EXCHANGED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED THAT THERE WAS NO SIGNIFICANT LOSS OF ENDOTHELIAL CELLS, PATIENTS CORNEA IS CRYSTAL CLEAR. (B)(4). Manufacturer Narrative: B5 - POST-OP THE REPORTER STATED PATIENT OFFERED NO COMPLAINT OTHER THEN SLIGHTLY BLURRED VISION, WHICH WAS RESOLVED BY RECEIVING A SMALL PRESCRIPTION FOR PRE-EXISTING CORNEAL ASTIGMATISM. H6 - DEVICE CODE: DEVICE CODE: 3189 SHOULD BE CORRECTED TO 2993. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONEOF THE HAPTICS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350566&pc=MTA | 2023826-2019-02188 | 2019-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5 12.6 IMPLANTABLE COLLAMER LENS. -10.00 DIOPTER IN THE PATIENTS RIGHT EYE (OD) ON 06-MAR-2019. ELEVATED IOP; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES; AND PIGMENT DISPERSION WAS NOTICED. ON 03-APR-2019 THE LENS WAS EXPLANTED. THE CAUSE WAS REPORTED AS USER ERROR; LENS IMPLANTED UPSIDE DOWN. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. H6-PATIENT CODE 3191: PIGMENT DISPERSION . CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE (B)(6) IN PREVIOUS MDR IS NOT APPLICABLE. ADDITIONAL INFORMATION: H6- DEVICE CODE (B)(6) FOR LENS FLIPPED UPSIDE DOWN. CLAIM# 714950. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885007&pc=MTA | 2023826-2015-00752 | 2013-07-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM11.5 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2011. ELEVATED IOP WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2015 AND WAS NOT REPLACED. PATIENT'S LAST BCVA WAS 20/19. SEE MFR# 2023826-2015-00753 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4). EVALUATION CODES: CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4885046&pc=MTA | 2023826-2015-00753 | 2013-07-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 12.0MM ICM120 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS)O ON (B)(6) 2011. ELEVATED IOP WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2015 AND WAS NOT REPLACED. PATIENT'S LAST BCVA WAS 20/23. SEE MFR# 2023826-2015-00752 FOR THE OTHER EYE. Manufacturer Narrative: (B)(4). EVALUATION CODES: CONCLUSIONS -(NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889284&pc=MTA | 2023826-2015-00766 | 2013-12-06 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-02 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2013 AND TWO WEEKS POST-OP ANGLE CLOSURE WITH ELEVATED IOP WAS NOTED. AN IRIDOTOMY WAS PERFORMED AND THE LENS WAS EXCHANGE FOR SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR REPORT# 2023826-2015-00767 FOR THE OTHER EYE. Manufacturer Narrative: PREVIOUS MANUFACTURE DATE WAS INCORRECT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (OTHER): BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889274&pc=MTA | 2023826-2015-00761 | 2014-02-17 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013 AND ELEVATED IOP WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT'S POST-OP UCVA WAS 20/18. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5549349&pc=MTA | 2023826-2016-00389 | 2014-04-24 | Injury | STAAR SURGICAL COMPANY | 2016-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 -13.50 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2013. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE RISE WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 ANS NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ON (B)(6) 2016, PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.50 (B)(4). LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578351&pc=MTA | 2023826-2015-00257 | 2014-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW & WORK ORDER SEARCH. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS WAS RETURNED DRY AND NO VISIBLE DAMAGE WAS OBSERVED. THE LENS WAS RE-HYDRATED IN BSS AND BOTH THE LENGTH AND WIDTH WERE RE-MEASURED AND FOUND TO BE WITHIN ORIGINAL DESIGN SPECIFICATIONS. A LENS WORK ORDER SEARCH WAS PERFORMED AND REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND. MEDICAL REVIEW - ELEVATED IOP AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATIENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCO), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZE ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CRODED AC DUE TO SMALL EYE ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, UPSIDE DOWN ICL, ECT. THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE EXACT CAUSE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4578988&pc=MTA | 2023826-2015-00226 | 2014-09-26 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ELEVATED IOP EARLY AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCOELASTIC), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, UPSIDE-DOWN ICL, ETC. PROMPT MANAGEMENT OF IOP RISE AND ITS CAUSE IS CRITICAL TO REDUCE OR ELIMINATE RISKS OF DAMAGE TO THE EYE AND VISION. GENERAL MANAGEMENT INCLUDES IOP-LOWERING DRUGS, BURPING THE INCISION, EVACUATING REMAINING FLUIDS FROM THE AC, ENLARGING THE EXISTING PI OR MAKING A NEW ONE, REMOVING OR REPLACING INADEQUATE SIZE LENSES, ETC. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842429&pc=MTA | 2023826-2015-00657 | 2014-10-14 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. ELEVATED IOP WAS NOTED, THE LENS WAS EXCHANGED ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. CONCLUSION: (NO CONCLUSION CAN BE DRAWN). (B)(4). LENS NOT RETURNED. Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842224&pc=MTA | 2023826-2015-00658 | 2014-10-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS ON (B)(6) 2014. ELEVATED IOP WAS NOTED, THE LENS WAS EXCHANGED ON (B)(6) 2014 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). DEVICE EVALUATED BY THE MANUFACTURER? NO. LENS WAS NOT RETURNED. CONCLUSION: (NO CONCLUSION CAN BE DRAWN). (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: WORK ORDER SEARCH, MEDICAL REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT INADEQUATE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842231&pc=MTA | 2023826-2015-00662 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ELEVATED IOP (INTRAOCULAR PRESSURE) WAS NOTED, THE LENS WAS EXPLANTED ON (B)(6) 2015, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CONCLUSION: (NO CONCLUSION CAN BE DRAWN). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4917149&pc=MTA | 2023826-2015-00839 | 2015-03-19 | Injury | STAAR SURGICAL COMPANY | 2015-07-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015 AND ELEVATED IOP WAS OBSERVED POSTOPERATIVELY. A TRABECULETOMY WAS PERFORMED AND THE LENS WAS EXPLANTED A WEEK LATER. CATARACT SURGERY WAS PERFORMED AND AN IOL WAS IMPLANTED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146056&pc=MTA | 2023826-2015-01247 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | 1 | Event Description: THE REPORTER INDICATED THE PRIMARY LENS WAS EXCHANGED FOR THIS LENS. THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (44MMHG). THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/25. SEE MFR # 2023826-2015-01246 FOR PRIMARY ICL. Manufacturer Narrative: DIOPTER: -13.50, DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD CODE(S): EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S)- (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146356&pc=MTA | 2023826-2015-01245 | 2015-06-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | 1 | Event Description: THE REPORTER INDICATED THE PRIMARY LENS WAS EXCHANGED FOR THIS LENS. THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (44MMHG). THE LENS WAS EXPLANTED AND NO OTHER LENS WAS IMPLANTED. POST-OP VISIT ON (B)(6) 2015 - BCVA WAS 20/25. SEE MFR # 2023826-2015-01244 FOR PRIMARY ICL. Manufacturer Narrative: DIOPTER: -13.50. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5223685&pc=MTA | 2023826-2015-01409 | 2015-07-01 | Injury | STAAR SURGICAL COMPANY | 2015-11-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.50/+0.5/X094 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. UNREACTIVE PUPIL AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION WAS ALSO PERFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -4.50/0.50/X094. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5309171&pc=MTA | 2023826-2015-01693 | 2015-12-03 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED INTRAOCULAR PRESSURE (IOP). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND PUPILLARY BLOCK. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/80. Manufacturer Narrative: RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY CAUSE EXCESSIVE VAULT OF THE IMPLANTED LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY ADVERSE EVENT OR CONDITION WOULD HAVE CAUSED DAMAGE TO THE LENS. THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. PER MEDICAL REVIEW - REPORTEDLY, MICL (13.2 MM) WAS EXPLANTED ON THE FIRST POST-OP DAY TO ADDRESS ELEVATED IOP. NO LENS WAS IMPLANTED AS A REPLACEMENT. ACCORDING TO COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, DATED ON (B)(6) 2016, THE ADVERSE EVENT WAS ATTRIBUTED TO DEVICE (EXCESSIVE VAULTING) THAT HAS CAUSED/CONTRIBUTED TO NARROWING OF THE ANGLE AND SUBSEQUENT IOP ELEVATION. IT SHOULD BE NOTED THAT AT THE TIME OF THE SURGERY PATIENT WAS 19 YEARS OLD AND PER DFU: "INDICATIONS: THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" AND THEREFORE, THERE IS NO SUFFICIENT DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036776&pc=MTA | 2023826-2016-01471 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON(B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THAT THE OPTIC WAS TORN/SPLIT, THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED AND FOREIGN MATERIAL WAS OBSERVED TO BE ON THE LENS SURFACE (BLOOD). UPON DIMENSIONAL INSPECTION, THE LENS WAS FOUND TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579354&pc=MTA | 2023826-2016-00467 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER STATED AT ONE-DAY POST-OP, THE PATIENT RETURNED WITH CLOSED ANGLES AND HIGH PRESSURE (40MMHG). THE PERIPHERAL IRIDECTOMIES (PIS) WERE PATENT, BUT THE SURGEON PERFORMED 2 ADDITIONAL PIS AND DILATED THE PUPILS. A COUPLE OF HOURS LATER, THE VAULT WAS GOOD AND THE PRESSURE WAS IN THE 20S. TO DATE THE VAULT IS STILL A LITTLE HIGH BUT THE PATIENT IS DOING WELL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (NO FAILURE DETECTED) : BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579395&pc=MTA | 2023826-2016-00466 | 2016-03-11 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED AT ONE-DAY POST-OP, THE PATIENT RETURNED WITH CLOSED ANGLES AND HIGH PRESSURE (40MMHG). THE PERIPHERAL IRIDECTOMIES (PIS) WERE PATENT, BUT THE SURGEON PERFORMED 2 ADDITIONAL PIS AND DILATED THE PUPILS. A COUPLE OF HOURS LATER, THE VAULT WAS GOOD AND THE PRESSURE WAS IN THE 20S. TO DATE THE VAULT IS STILL A LITTLE HIGH BUT THE PATIENT IS DOING WELL. THE LENS REMAINS IMPLANTED AND THE PATIENT WILL BE MONITORED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: (B)(4). RESULT CODE(S): (OPERATIONAL PROBLEM) : A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING, INSPECTION AND PACKAGING PROCESS RECORDS TO SUGGEST A CONTRIBUTORY FACTOR TO THE COMPLAINT. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION CODE(S): (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW, AND MEDICAL REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5221692&pc=MTA | 2023826-2015-01523 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT WAS BROUGHT BACK TO SURGERY DUE TO ELEVATED INTRAOCULAR PRESSURE AND HAD A SECONDARY PROCEDURE TO CREATE A NEW PERIPHERAL IRIDOTOMY BUILT IT WAS NOT SUCCESSFUL. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED INTRAOCULAR PRESSURE AND INTERMITTENT ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS (12.6MM) BUT THE PROBLEM WAS NOT RESOLVED. THE PATIENT IS STILL HAVING ISSUES WITH ELEVATED INTRAOCULAR PRESSURE AND THE PLAN IS TO EXPLANT THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): MEDICAL REVIEW, DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): PER MEDICAL REVIEW: REPORTEDLY, VISIAN ICL (13.2 MM) WAS EXPLANTED/EXCHANGED TWO MONTHS POSTOPERATIVELY FOR A SHORTER LENGTH ICL (12.6 MM) TO ADDRESS VAULTING AND SUBSEQUENT ELEVATED IOP AND INTERMITTENT ANGLE CLOSURE. TWO DAYS PRIOR TO LENS EXCHANGE THE SURGEON PERFORMED ANOTHER IRIDOTOMY. ACCORDING TO REPORT FROM THE FACILITY, DATED ON (B)(6) 2015, ELEVATED IOP PERSISTED DESPITE INTERVENTION AND THE SURGEON WAS PLANNING TO EXPLANT/EXCHANGE ICL (12.6 MM) FOR A SHORTER ICL (12.1 MM). NO FURTHER INFORMATION WAS RECEIVED TO DATE. BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, AND WORK ORDER SEARCH, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. NO ROOT CAUSE IS DETERMINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER, HOWEVER, NONE HAS BEEN FORTHCOMING. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596422&pc=MTA | 2023826-2016-00508 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED ON (B)(6) 2016 AND REMOVED FROM THE LEFT EYE (OS) ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND A LITTLE PAIN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/15. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS NOT THE DEVICE BUT WAS DUE TO LENS SIZE SELECTION (TOO LARGE). Manufacturer Narrative: (B)(4). Manufacturer Narrative: IMPLANTED DATE - (B)(6) 2016. EXPLANTED DATE - (B)(6) 2016. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686518&pc=MTA | 2023826-2016-00770 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED MICL IMPLANTABLE COLLAMER LENSES IN BOTH EYES AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THERE WAS AN ANTERIOR CHAMBER REACTION IN THE LEFT EYE AND THE PUPIL IS PERMANENTLY DILATED. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. THE REPORTER INDICATED TO DATE, THE LEFT EYE IS STILL DILATED BUT THE PRESSURE HAS IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR REPORT # 2023826-2016-00771 FOR THE OTHER EYE. Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, WAS IMPLANTED IN THE RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT'S VISUAL ACUITY WAS 20/15 -2 AND HAS BEEN REFERRED TO A GLAUCOMA SPECIALIST. Manufacturer Narrative: EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5686526&pc=MTA | 2023826-2016-00771 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED MICL IMPLANTABLE COLLAMER LENSES IN BOTH EYES AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THERE WAS AN ANTERIOR CHAMBER REACTION IN THE LEFT EYE AND THE PUPIL IS PERMANENTLY DILATED. THE PATIENT WAS ADMINISTERED MEDICATION FOR THE PRESSURE. THE REPORTER INDICATED TO DATE, THE LEFT EYE IS STILL DILATED BUT THE PRESSURE HAS IMPROVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR REPORT # 2023826-2016-00770 FOR THE OTHER EYE.Event Description: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, WAS IMPLANTED IN THE LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT'S VISUAL ACUITY WAS 20/20 AND HAS BEEN REFERRED TO A GLAUCOMA SPECIALIST. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) PUPIL DILATED. IMPLANTED. (B)(4). THE LENS WAS NOT RETURNED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6498299&pc=MTA | 2023826-2017-00548 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. THE SURGEON STATED THAT THE PATIENT MAY NOT BE SUITABLE FOR LENS IMPLANTATION. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CONCLUSION CODE: THE SURGEON STATED THAT THE PATIENT MIGHT NOT HAVE BEEN A SUITABLE CANDIDATE FOR THE LENS. PRE-OPERATIVE INTRAOCULAR PRESSURE WAS VERY HIGH (40 MMHG). (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PRODUCT EVALUATION-FOUND THAT THE LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND THE HAPTIC TORN, BROKEN, BENT, AND DEFORMED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6505320&pc=MTA | 2023826-2017-00549 | 2016-05-21 | Injury | STAAR SURGICAL COMPANY | 2017-04-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE SURGEON STATED THAT EVEN THOUGH THE VAULT VALUE WAS NORMAL AFTER IMPLANT, THE IOP HAD A HIGH VALUE. AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. THE SURGEON STATED THAT THE PATIENT MAY NOT BE SUITABLE FOR LENS IMPLANTATION. Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE:THE SURGEON STATED THAT THE PATIENT MIGHT NOT HAVE BEEN A SUITABLE CANDIDATE FOR THE LENS. PRE-OPERATIVE INTRAOCULAR PRESSURE WAS VERY HIGH (40 MMHG). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND DEFORMED AND THE OPTIC SPLIT. ONLY A PORTION OF THE LENS WAS RETURNED, A LARGE PIECE WAS MISSING. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801085&pc=MTA | 2023826-2017-01285 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTED INDICATED THAT A VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A -9.50 DIOPTER WAS IMPLANTED ON (B)(6)2017 IN THE LEFT EYE (OS). THE PATIENT WAS (B)(6) YEARS OLD AT THE TIME (THIS MODEL IS CONTRAINDICATED FOR USE IN PATIENTS UNDER THE AGE OF 21) THE PATIENT'S IOP ELEVATED WITHOUT PUPIL BLOCK OR ANGLE CLOSURE, SO THE SURGEON EXPLANTED THE LENS ON (B)(6)2017. THIS IS THE SAME PATIENT AS MDR # 2023826-2017-01286. Manufacturer Narrative: LENS WAS RETURNED WITH SOME MOISTURE IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. (B)(4). Manufacturer Narrative: INITIAL LENS WAS IMPLANTED ON (B)(6) 2016. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION CODE: THIS DEVICE IS CONTRAINDICATED FOR USE IN PATIENTS UNDER 21 YEARS ((B)(6) YEAR OLD PATIENT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6801118&pc=MTA | 2023826-2017-01286 | 2016-07-25 | Injury | STAAR SURGICAL COMPANY | 2017-08-17 | MTA | IMPLANTABLE COLLAMER LENS | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTED INDICATED THAT A VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A -9.50 DIOPTER WAS IMPLANTED ON (B)(6)-2017 IN THE RIGHT EYE (OD). THE PATIENT WAS (B)(6) YEARS OLD AT THE TIME (THIS MODEL IS CONTRAINDICATED FOR USE IN PATIENTS UNDER THE AGE OF 21) THE PATIENT'S IOP ELEVATED WITHOUT PUPIL BLOCK OR ANGLE CLOSURE, SO THE SURGEON EXPLANTED THE LENS ON (B)(6)2017. THIS IS THE SAME PATIENT AS MDR # 2023826-2017-01285. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH SOME MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: LENS EXPLANT RESOLVED THE PROBLEM. PATIENT'S POST-OP BCVA WAS 20/20 ON (B)(6) 2017. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CONCLUSION: THIS DEVICE IS CONTRAINDICATED FOR USE IN PATIENTS UNDER 21 YEARS ((B)(6) YEAR OLD PATIENT).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447102&pc=MTA | 2023826-2017-00514 | 2016-12-30 | Injury | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. AFTER EXPLANT, IOP RETURNED TO NORMAL LEVEL. TO DATE, THE SURGEON DOES NOT PLAN TO EXCHANGE THE LENS. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6325597&pc=MTA | 2023826-2017-00117 | 2017-01-13 | Injury | STAAR SURGICAL COMPANY | 2017-02-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2 IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE PATIENT HAD ELEVATED IOP ( WITHOUT PUPIL BLOCK AND ANGLE CLOSURE), UNREACTIVE PUPIL (FIXED). AND REMAINING VISCO IN EYE. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF REMAINING VISCO/FLUIDS WAS PERFORMED. AND AFTER IRRIGATION NO MORE VISCO WAS IN THE EYE. PATIENT'S UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/25 AND THE BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430473&pc=MTA | 2023826-2017-00478 | 2017-02-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, WAS EXPLANTED DUE TO EXTREME HIGH PRESSURE. THE LENS WAS REPLACED ON (B)(6) 2017 WITH A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE PATIENT IS DOING VERY WELL AFTER ICL EXCHANGE. THE LENS EXCHANGE FOR A SHORTER LENS RESOLVED THE PROBLEM. PATIENT CURRENT BSCVA IS 20/20. (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY, IN A POUCH. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLEAR RESIDUE AND LENT WAS OBSERVED ON THE LENS SURFACE. (B)(4). SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6524681&pc=MTA | 2023826-2017-00625 | 2017-03-31 | Injury | STAAR SURGICAL COMPANY | 2017-04-26 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY BECAUSE IT WAS CAUSING PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE, EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION ( LASER PERIPHERAL IRIDECTOMY) AND MEDICAL INTERVENTION (DIAMOX) WERE PERFORMED/PRESCRIBED. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/50 RESIDUAL ASTIGMATISM AT 1 DAY POST -OP OF ICL EXCHANGE. THE REPORTER STATED THAT THE EVENT WAS PRODUCT-RELATED. (B)(4). IMPLANTED DATE WAS CORRECTED TO (B)(6) 2017 , LENS EXCHANGE WAS A SECONDARY SURGERY. CLAIM# (B)(4). Manufacturer Narrative: PREVIOUS IMPLANT DATE, (B)(6) 2017, IS INCORRECT. THE CORRECT IMPLANT DATE IS (B)(6) 2017. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE, EXCESSIVE VAULT, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION ( LASER PERIPHERAL IRIDECTOMY) AND MEDICAL INTERVENTION (DIAMOX) WERE PERFORMED/PRESCRIBED. THE PATIENTS POST-OP UNCORRECTED VISUAL ACUITY WAS 20/50 RESIDUAL ASTIGMATISM AT 1 DAY POST -OP OF ICL EXCHANGE. THE REPORTER STATED THAT THE EVENT WAS PRODUCT-RELATED. (B)(4). IMPLANTED DATE WAS CORRECTED TO (B)(6) 2017 , LENS EXCHANGE WAS A SECONDARY SURGERY. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6812325&pc=MTA | 2023826-2017-01339 | 2017-07-04 | Injury | STAAR SURGICAL COMPANY | 2017-08-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -7.0, INTO THE PATIENT'S LEFT (OS) EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER AND SIMILAR DIOPTER LENS DUE TO ELEVATED IOP. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DATE RECEIVED BY MANUFACTUTRER IN THE INTIAL SHOULD BE (B)(4) 2017. (B)(4). Manufacturer Narrative: DATE OF EVENT: IT IS CORRECTED TO "UNK" FROM "(B)(6) 2017". THE REPORTER STATED THAT THE ADVERSE EVENT (ELEVATED IOP) WAS NOTICED 2-3 WEEKS BEFORE THE EXPLANT DATE; (B)(6) 2017. NO DEFINITE DATE WAS PROVIDED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6867297&pc=MTA | 2023826-2017-01451 | 2017-07-31 | Injury | STAAR SURGICAL COMPANY | 2017-09-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER 13.50/+2.0/106 INTO THE PATIENT'S LEFT (OS) EYE. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR THE SAME LENGTH AND SIMILAR DIOPTER LENS. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN, AND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050888&pc=MTA | 2023826-2017-01706 | 2017-09-21 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.2, -9.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR ANOTHER LENS DUE HIGH PRESSURE AND FIXED DILATED PUPIL. THE BAG WAS ALSO NEEDED TO BE REFORMED. THE PATIENT'S CURRENT CONDITION IS THAT THE EYE HAS A BIT OF EDEMA BUT IS OVERALL FINE WITH THE SURGERY. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS HAPTIC MISSING. METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WERE FOUND WITHIN THE ASSOCIATED LOT. (B)(4). Manufacturer Narrative: PATIENT CODE: (B)(4) FIXED DILATED PUPIL, SECONDARY SURGERY, LENS EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356133&pc=MTA | 2023826-2018-00438 | 2018-01-29 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS (ICL) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "THERE IS AN IOP PROBLEM" AND THE PATIENT "IS ON MEDICATION FOR IT." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: REPORTEDLY, THE ICL THAT THE SURGEON IMPLANTED WAS A 13.2MM VICM5_13.2 ICL, DIOPTER -13.5. THE PATIENT WAS GIVEN MEDICATION ("DROPS") AND THE IOP WENT FROM 47.2MMHG ON (B)(6) 2018 TO 20.8MMHG ON (B)(6) 2018 "WITH 3 DROPS," AND THEN TO 12.4MMHG ON (B)(6) 2018 "WITH NO DROPS." (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6224163&pc=MTA | 2023826-2017-00004 | 2016-11-10 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | INTRAOCULAR LENS | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). AS OF THE DATE OF MDR REPORTING, THE LENS HAS NOT BEEN RETURNED AND THEREFORE NO PRODUCT EVALUATION HAS OCCURRED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978364&pc=MTA | 2023826-2017-01627 | 2017-02-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. USER FACILITY IS INCORRECT, SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978385&pc=MTA | 2023826-2017-01628 | 2017-02-03 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO ELEVATED INTRAOCULAR PRESSURE. A SUTURE WAS REQUIRED. THE PATIENT WAS PRESCRIBED MEDICATION FOR THE PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION FOUND LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS WERE FOUND. (B)(4). Manufacturer Narrative: REPORT SOURCE: USER FACILITY IS INCORRECT; SHOULD BE DISTRIBUTOR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4889289&pc=MTA | 2023826-2015-00767 | 2013-12-06 | Injury | STAAR SURGICAL COMPANY | 2015-07-02 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED THE 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT (OS EYE ON (B)(6) 2013 AND TWO WEEKS POST-OP ANGLE CLOSURE WITH ELEVATED IOP WAS NOTED. AN IRIDOTOMY WAS PERFORMED AND THE LENS WAS EXCHANGE FOR SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. SEE MFR REPORT# 2023826-2015-00766 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EXPLANT. METHOD: WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT INFORMATION AND LENS WORK ORDER SEARCH, WE WERE UNABLE TO DETERMINE A SPECIFIC ROOT CAUSE OF THE EVENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5756735&pc=MTA | 2023826-2016-00792 | 2016-01-21 | Injury | STAAR SURGICAL COMPANY | 2016-06-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -9/+4/167 DIOPTER IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED IOP. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE BUT PRESENCE OF FOREIGN MATERIAL (FIBERS) ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891300&pc=MTA | 2023826-2016-00455 | 2016-02-24 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+0.5/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). TWENTY-FOUR REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21 AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891288&pc=MTA | 2023826-2016-00453 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/094 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A REFRACTIVE SURPRISE AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL BUT DIFFERENT DIOPTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5968179&pc=MTA | 2023826-2016-01310 | 2016-07-03 | Injury | STAAR SURGICAL COMPANY | 2016-09-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE REPORTER INDICATED THE PATIENT'S INTRAOCULAR PRESSURE WAS QUITE HIGH INITIALLY AND THE PATIENT WAS GIVEN AN INJECTION. THE EVENT WAS RESOLVED BUT THE PUPIL IS UNREACTIVE (FIXED) AND FULLY DILATED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5975346&pc=MTA | 2023826-2016-01316 | 2016-07-19 | Injury | STAAR SURGICAL COMPANY | 2016-09-23 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL INFORMATION: - ON (B)(6) 2016, THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724414&pc=MTA | 2023826-2016-00793 |  | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO 13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED IOP . THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE BROKEN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5873701&pc=MTA | 2023826-2016-01061 | 2016-05-16 | Injury | STAAR SURGICAL COMPANY | 2016-08-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected; Insufficient Information | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON WAS UNABLE TO IMPLANT A A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE SULCUS OF THE PATIENT'S RIGHT EYE (OD) DUE TO POSITIVE VITREO PRESSURE. THE LENS WAS INSERTED BUT THE PRESSURE DISPLACED THE LENS OUT OF THE PROPER POSITION IN THE EYE. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND THERE WAS NO APPARENT INJURY. Manufacturer Narrative: DEVICE EVALUATION: EVALUATION FOUND THE LENS RETURNED DRY IN THE LENS CONTAINER, BUT THERE IS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND FIBERS ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6197215&pc=MTA | 2023826-2016-01805 | 2016-01-28 | Injury | STAAR SURGICAL COMPANY | 2016-12-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD CODE(S): NO TESTING METHOD PERFORMED, CLAIM EVALUATED BASED ON REPORTER RESPONSE TO COMPLAINT QUESTIONNAIRE. CONCLUSION CODE: DFU REQUIRES PATIENT'S ANTERIOR CHAMBER DEPTH (ACD) TO BE EQUAL OR GREATER THAN 3.0 MM. PATIENT (ACD) MEASURED AT 2.94 MM. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6036767&pc=MTA | 2023826-2016-01474 | 2016-02-03 | Injury | STAAR SURGICAL COMPANY | 2016-10-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AS OF (B)(6) 2016, PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CONCLUSION: AS PER DFU FOR THIS LENS MODEL, VICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6077547&pc=MTA | 2023826-2016-01562 | 2016-03-10 | Injury | STAAR SURGICAL COMPANY | 2016-11-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -18.0/./+20 DIOPTER IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND PUPIL BLOCK (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: PREEXISTING MEDICAL CONDITION: KERATOCONUS. CONCLUSION CODE: AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL WITH KERATOCONUS IS CONTRAINDICATED. CORRECTED DATA: CORRECT LENS DIOPTER IS -18.0/+2.0/X074 (SPHERE/CYLINDER/AXIS). (B)(4). Manufacturer Narrative: PMA/510(K): N/A THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6795677&pc=MTA | 2023826-2017-01319 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY AG | 2017-08-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION AND NARROWING OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). OFF LABEL USE (THIS LENS MODEL IS CONTRADICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6209910&pc=MTA | 2023826-2016-01829 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-12-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXCHANGED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT UNDERWENT YAG AND THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. THE CORRECT DIOPTER REPORTED FOR VTICMO12.6 IMPLANTABLE COLLAMER LENS WAS -10.00/+1.5/009 DIOPTER. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6410776&pc=MTA | 2023826-2016-01644 | 2016-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-03-16 04:00 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ANGLE CLOSURE (WITH ELEVATED IOP) AND HIGH OCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983954&pc=MTA | 2023826-2016-01325 | 2016-08-31 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. AT A POST-OP VISIT ON (B)(6) 2016, THE SURGEON NOTED AN INCREASE IN THE OCULAR PRESSURE AND LENS WAS CLOUDY. THE SURGEON ATTEMPTED TO RE-POSITION THE LENS BUT WITH NO SUCCESS. THE LENS WAS EXPLANTED AND A SHORTER LENS WAS IMPLANTED. THERE WERE NO COMPLICATIONS. Manufacturer Narrative: CORRECTION - THE REPORTER INDICATED THERE WAS NO INCREASE IN INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED ON (B)(6) 2016 FROM THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT HAD A DISTORTED PUPIL DUE TO CROWDED ANGLE. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING EXCELLENT AND THE POST-OP VISUAL ACUITY WAS 20/25. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460525&pc=MTA | 2023826-2017-00570 | 2016-10-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6460543&pc=MTA | 2023826-2017-00571 | 2016-10-01 | Injury | STAAR SURGICAL COMPANY | 2017-04-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6222127&pc=MTA | 2023826-2017-00012 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/096 DIOPTER, IN THE PATIENT'S LEFT EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: AS PER DFU FOR THIS LENS MODEL, VTICLS ARE CONTRAINDICATED FOR PATIENTS UNDER 21 YEARS OF AGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6224164&pc=MTA | 2023826-2017-00005 | 2016-10-15 | Injury | STAAR SURGICAL COMPANY | 2017-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT REPORTED EXCESSIVE VAULTING AND ELEVATED IOP. ON (B)(6) 2016, THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION NO LENS EVALUATION HAS BEEN PERFORMED, AS THE PRODUCT HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL AND THERE WAS CLEAR SURGICAL RESIDUE DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). DEVICE NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6361169&pc=MTA | 2023826-2017-00266 | 2016-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-02-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -8.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. ON (B)(6) 2016, THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/20. Manufacturer Narrative: LENS RETURNED DRY IN A SMALL CENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND DRY CLEAR RESIDUE AND BLACK DEBRIS ON THE LENS AND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. PATIENT: SECONDARY SURGICAL INTERVENTION, LENS WAS EXCHANGED FOR SHORTER LENS. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172497&pc=MTA | 2023826-2016-01758 | 2016-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. THE PERIPHERAL IRIDOTOMIES WERE ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PATIENT INFO: FEMALE. EXPLANT DATE: (B)(6) 2016. EVENT DATE: UNK. ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT'S CURRENT CONDITION IS BCVA 20/20 +2 AND VAULT IS 3/4 CT. RESULT (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE MISSING FROM THE LENS OPTIC. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6172498&pc=MTA | 2023826-2016-01757 | 2016-11-05 | Injury | STAAR SURGICAL COMPANY | 2016-12-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE. THE PERIPHERAL IRIDOTOMIES WERE ENLARGED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: PATIENT INFO: FEMALE. EXPLANT DATE: (B)(6) 2016. EVENT DATE: UNK. ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT'S CURRENT CONDITION IS BCVA 20/20 +2 AND VAULT IS 3/4 CT. RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6191984&pc=MTA | 2023826-2016-01801 | 2016-11-07 | Injury | STAAR SURGICAL COMPANY | 2016-12-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -3.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGICAL INTERVENTION, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, LENS EXPLANTED. ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6420639&pc=MTA | 2023826-2017-00446 | 2016-12-13 | Injury | STAAR SURGICAL COMPANY | 2017-03-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/105 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIA. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6482395&pc=MTA | 2023826-2017-00601 | 2016-12-14 | Injury | STAAR SURGICAL COMPANY | 2017-04-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND GONIOSYNECHIAE. ON (B)(6) 2016 THE LENS WAS EXCHANGED FOR THE SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE EVALUATION: LENS WAS RETURNED DRY, IN SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DRY, CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. LENS IS TWISTED DUE TO POSITION IN VIAL. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6260641&pc=MTA | 2023826-2017-00101 | 2016-12-21 | Injury | STAAR SURGICAL COMPANY | 2017-01-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/95 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2016, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED TO THE MANUFACTURER BUT NOT YET BEEN EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6281933&pc=MTA | 2023826-2017-00174 | 2017-01-06 | Injury | STAAR SURGICAL COMPANY | 2017-01-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED EVO VISIAN ICL LENSES BILATERALLY IN A PATIENT AND BOTH EYES PRESENT HIGH VAULTING. THE LENSES REMAIN IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE INCISION WAS ENLARGED AND ADDITIONAL SUTURE WAS REQUIRED. THE LENS EXCHANGED RESOLVED THE PROBLEM. THE PATIENTS POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6407568&pc=MTA | 2023826-2017-00412 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-03-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION-LENS WAS RETURNED DRY, WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT IN A SMALL VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6312700&pc=MTA | 2023826-2017-00228 | 2017-01-17 | Injury | STAAR SURGICAL COMPANY | 2017-02-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. WORK ORDER SEARCH: NO SIMILAR CLAIM WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439068&pc=MTA | 2023826-2017-00506 | 2017-01-23 | Injury | STAAR SURGICAL COMPANY | 2017-03-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SMALL VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR/BLACK RESIDUE ON THE PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6424135&pc=MTA | 2023826-2017-00454 | 2017-01-25 | Injury | STAAR SURGICAL COMPANY | 2017-03-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULT, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM12.6 IMPLANTABLE COLLAMER LENS; -5.0 DIOPTER; IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED LOW VAULTING, ELEVATED IOP, AND HAD ADDITIONAL REFRACTIVE TREATMENT. THE LENS WAS EXPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6389631&pc=MTA | 2023826-2017-00348 | 2017-02-07 | Injury | STAAR SURGICAL COMPANY | 2017-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, WAS IMPLANTED ON (B)(6) 2017 AND EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE REPORTER INDICATED THAT THERE WAS NO PATIENT INJURY AND THE EVENT WAS NOT DEVICE RELATED BUT THE EXACT CAUSE WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED STUCK INSIDE CARTRIDGE. CARTRIDGE WAS RETURNED TAPED TO LAF CARD. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE, WHITE/GREY DRY RESIDUE WAS OBSERVED ON CARTRIDGE, AND A STUCK LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6653556&pc=MTA | 2023826-2017-00998 | 2017-03-09 | Injury | STAAR SURGICAL COMPANY | 2017-06-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/072 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULT, ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A VICMO LENS WITH THE SAME SIZE AND DIFFERENT DIOPTER. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR RESIDUE AND DEBRIS ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6711929&pc=MTA | 2023826-2017-01185 | 2017-03-11 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/098 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6)2017, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. UPDATED INFORMATION: THE LENS HAS BEEN RETURNED AND EVALUATED. LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PREVIOUS ADVERSE EVENTS WERE NOT INCLUDED IN INITIAL MDR: PATIENT HAD ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6554053&pc=MTA | 2023826-2017-00781 | 2017-03-15 | Injury | STAAR SURGICAL COMPANY | 2017-05-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/066 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXPLANTED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PREVIOUS: STATED LENS IS NOT RETURNED. UPDATED: LENS HAS BEEN RETURNED AND EVALUATED. DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION REVEALED LENS WAS WITHIN SPECIFICATIONS. METHOD, RESULT CODES: UPDATED. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6587670&pc=MTA | 2023826-2017-00872 | 2017-03-21 | Injury | STAAR SURGICAL COMPANY | 2017-05-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+5/X126 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, THE PUPIL CAN NOT CONTRACT, AND ELEVATED IOP. IN THE SAME DAY, THE LENS WAS EXCHANGED FOR SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED, BUT NOT YET EVALUATED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR AND BROWN RESIDUE ON THE LENS. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6527418&pc=MTA | 2023826-2017-00701 | 2017-03-29 | Injury | STAAR SURGICAL COMPANY | 2017-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS, -14 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ACCORDING TO THE SURGEON THE PATIENT EXPERIENCED EXCESSIVE VAULTING BECAUSE THE LENS WAS IMPLANTED IN A "PSEUDO PHAKIC EYE". THE PATIENT ALSO EXPERIENCED ELEVATED IOP. THE SHORTER LENS WAS ALREADY ORDERED AND SCHEDULED TO EXCHANGE. AS OF THE DATE OF MDR SUBMISSION, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: THE LENS WAS IMPLANTED ON PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION AND CORRECTED: THE LENS WAS IMPLANTED ON PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. ADDITIONAL PATIENT CODE (B)(4) (NEW INCISION): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CONCLUSION: OFF LABEL USE (BELOW 21 YEARS OF AGE AT THE TIME OF IMPLANT, LENS IMPLANT IN A PSEUDO PHAKIC EYE). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6579894&pc=MTA | 2023826-2017-00860 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-05-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, DIOPTER -18.00 IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO THE PATIENT EXPERIENCING EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY (BCVA) WAS 20/30 AND UNCORRECTED VISUAL ACUITY (UCVA) WAS 20/30. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6609924&pc=MTA | 2023826-2017-00921 | 2017-04-11 | Injury | STAAR SURGICAL COMPANY | 2017-06-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2017, THE LENS WAS EXCHANGED WITH A SHORTER AND SIMILAR DIOPTER LENS, DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND PIGMENT DISPERSION. THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS BEEN RETURNED BUT NOT YET EVALUATED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565217&pc=MTA | 2023826-2017-00812 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, CLOSED ANGLE AND SHALLOW ANTERIOR CHAMBER. THE PRESSURE WAS 34MMHG BUT WAS BROUGHT DOWN TO NORMAL WITH DIAMOX. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6565218&pc=MTA | 2023826-2017-00811 | 2017-04-14 | Injury | STAAR SURGICAL COMPANY | 2017-05-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -7.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, CLOSED ANGLE AND SHALLOW ANTERIOR CHAMBER. THE PRESSURE WAS 34MMHG BUT WAS BROUGHT DOWN TO NORMAL WITH DIAMOX. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR CLAIMS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637203&pc=MTA | 2023826-2017-00964 | 2017-04-20 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/88 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION: THE PRODUCT WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RED AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: PREVIOUS CODE NOT APPLICABLE. AS PER DFU, THIS LENS MODEL IS CONTRAINDICATED FOR IMPLANTATIONS IN PATIENTS WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637164&pc=MTA | 2023826-2017-00961 | 2017-04-25 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/085 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. AS OF THE DAY OF MDR SUBMISSION THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS WITH NO VISIBLE DAMAGE AND CLEAR RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6637205&pc=MTA | 2023826-2017-00965 | 2017-04-27 | Injury | STAAR SURGICAL COMPANY | 2017-06-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/88 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PIGMENT DISPERSION. THE PATIENT ALSO UNDERWENT ENLARGEMENT PI OR ADDITION OF NEW PI, AND SECONDARY YAG SURGERY. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS, AND THE PROBLEM WAS RESOLVED. AT THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND A MISSING PIECE OF HAPTIC AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6645493&pc=MTA | 2023826-2017-00983 | 2017-05-16 | Injury | STAAR SURGICAL COMPANY | 2017-06-15 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER INDICATED AT THE INITIAL SURGERY THE PATIENT HAD SLIGHT IRIS PROLAPSE, WHICH WAS MANAGED AT THE TIME. AT ONE DAY POST-OP, THE PATIENT'S INTRAOCULAR PRESSURE WAS HIGH, WITH SHALLOW CHAMBER AND ELEVATED VAULT. THE PATIENT HAD RECURRING IRIS PROLAPSE. THE PATIENT RETURNED TO SURGERY AND THE IRIS PROLAPSE WAS MANAGED. BUT THE SURGEON DECIDED TO EXPLANT THE LENS, WHICH WAS REMOVED ON (B)(6) 2017 AND EXCHANGED FOR A SHORTER LENS. THE PATIENT IS DOING WELL AND THE LENS VAULT IS NORMAL. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6712101&pc=MTA | 2023826-2017-01186 | 2017-06-06 | Injury | STAAR SURGICAL COMPANY | 2017-07-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED FOR EVALUATION. THERE ARE NO PLANS TO EXCHANGE THE LENS. Manufacturer Narrative: PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6728124&pc=MTA | 2023826-2017-01213 | 2017-06-20 | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON EXPLANTED A 13.2 MM MICL13.2 IMPLANTABLE COLLAMER LENS ,-5.0 DIOPTER, THAT WAS IMPLANTED LAST YEAR AND EXCHANGED IT FOR A SHORTER LENGTH LENS. THE REPORTER STATED THAT THE LENS WAS BEING REPLACED DUE TO IT BEING BIG, PROGRESSIVE ANTERIOR SYNECHIAE FROM HIGH VAULT, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD. ADDITIONAL SURGICAL INTERVENTION (LASER IRIDOTOMY) WAS PERFORMED. THE PATIENTS POST- OP VA WAS 20/20. THE REPORTER INDICATED THAT A SURGICAL SYNECHIOLYSIS WAS PERFORMED AT THE SAME TIME OF THE LENS EXCHANGE. Manufacturer Narrative: THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE PATIENT ALSO EXPERIENCED HIGH IOP. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RECEIVED DRY, IN SPECIMEN CONTAINER. VISUAL INSPECTION FOUND A TEAR ON A HAPTIC, DEBRIS, RED,CLEAR RESIDUE ON LENS SURFACE, AND PORTION OF A HAPTIC MISSING. (B)(4). WORK ORDER SEARCH: N O SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6730633&pc=MTA | 2023826-2017-01231 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-20 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, 4.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO LOW VAULTING, ANGLE CLOSURE WITH ELEVATED IOP, ANGLE CLOSURE GLAUCOMA, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS EXCHANGE IS PLANNED AND A NEW ICL WAS ORDERED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A LENS CASE/ VIAL AND HAD CLEAR SURGICAL RESIDUE ON LENS SURFACE. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT AFTER LENS EXPLANT, A NEW SHORTER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THIS EXCHANGE RESOLVED THE PROBLEM. (B)(4). THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT IN THE PATIENT'S RIGHT EYE (OD) AND NOT LOW VAULT AS PREVIOUSLY STATED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6740871&pc=MTA | 2023826-2017-01240 | 2017-06-21 | Injury | STAAR SURGICAL COMPANY | 2017-07-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -2.5/+3.0/162 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED AS AN EMERGENCY IN THE HOSPITAL. THE LENS WAS EXPLANTED BECAUSE OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF PATIENT'S IRIDO-CORNEAL ANGLES (ICA), AND PUPIL BLOCK WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED WITH THE EXPLANT. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6767795&pc=MTA | 2023826-2017-01250 | 2017-07-06 | Injury | STAAR SURGICAL COMPANY | 2017-08-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER. IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS HAD A HIGH VAULT AND DUE TO CONCERNS OVER ANGLE CLOSURE, THE LENS WAS EXPLANTED ON (B)(6) 2017. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE AND NARROWING OF THE ANGLE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S CURRENT CONDITION WAS 20/20 UCVA, WITH GREAT INTRAOCULAR PRESSURE. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS AND UPON IMPLANTATION THERE WAS A RESIDUAL VAULTING. THE SURGEON ENLARGED PI AND RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6890291&pc=MTA | 2023826-2017-01469 | 2017-07-13 | Injury | STAAR SURGICAL COMPANY | 2017-09-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICMO13.2, -15.50 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON THE MORNING OF (B)(6) 2017. EXCESSIVE VAULTING AND HIGH IOP WAS NOTICED AND WAS EXPLANTED LATER THAT EVENING. THE PATIENT'S POST-OP BCVA WAS 20/40. Manufacturer Narrative: DEVICE EVALUATION - THE LENS WAS RETURNED DRY, IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT. DIMENSIONAL EVALUATION FOUND THE LENS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818689&pc=MTA | 2023826-2017-01323 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INCORRECT SIZING. THE LENS WAS EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULT, ELEVATED (INTRAOCULAR PRESSURE) IOP, NARROWING OF ANGLES, ANGLE CLOSURE AND SHALLOWING OF ANTERIOR CHAMBER DEPTH (ACD). AN IRIDECTOMY WAS PERFORMED. THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6818692&pc=MTA | 2023826-2017-01324 | 2017-07-26 | Injury | STAAR SURGICAL COMPANY | 2017-08-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL13.7, -15.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO INCORRECT SIZING. THE LENS WAS EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: FEMALE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP (INTRAOCULAR PRESSURE), NARROWING OF ANGLES, ANGLE CLOSURE AND SHALLOWING OF ACD (ANTERIOR CHAMBER DEPTH). AN IRIDECTOMY WAS PERFORMED. CORRECTED DATA: DATE OF IMPLANTATION WAS CORRECTED TO "(B)(6) 2017". THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL 13.7, -15.5 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE MANUFACTURE DATE WAS CORRECTED TO "10-DEC-2015". (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6850323&pc=MTA | 2023826-2017-01394 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-09-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/088 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT DATE: (B)(6) 2017. DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CONCLUSION CODE: OFF LABEL USE (PATIENT BELOW 21 YEARS OF AGE AT TIME OF IMPLANT). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114677&pc=MTA | 2023826-2017-01841 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -15.0/+3.0/079 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS CURVED AND RED AND CLEAR RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7114680&pc=MTA | 2023826-2017-01859 | 2017-08-10 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -14.0/+1.5/090 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS UNDER REVIEW FOR MORE THAN 3 MONTHS DUE TO THE FLUCTUATION OF IOP. THE DOCTOR DECIDED TO DISCHARGE THE PATIENT ON (B)(6) 2017 AS THE IOP WAS 16 MMHG. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF THE HAPTIC AND THE PRESENCE OF CLEAR AND RED RESIDUE ON THE LENS SURFACE. (B)(4) Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7069472&pc=MTA | 2023826-2017-01802 | 2017-08-17 | Injury | STAAR SURGICAL COMPANY | 2017-11-29 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/10. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6873954&pc=MTA | 2023826-2017-01448 | 2017-08-18 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A 13.7 MM VICMO13.7, -17.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULTING, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS NOTICED. THE SURGEON DECIDED TO OBSERVE THE PATIENT AS THE PATIENT IS ASYMPTOMATIC AND IS HAPPY WITH VISUAL ACUITY. ON (B)(6) 2017, THE PATIENT'S BCVA IS 20/30, UCVA IS 20/40, AND IOP IS 15 MM HG WITHOUT TREATMENT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: PER DFU FOR THIS LENS MODEL, VICL'S ARE CONTRAINDICATED OF IMPLANTATIONS OF A LENS IN A PATIENT WITH PREVIOUS CORNEAL PATHOLOGY. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6920137&pc=MTA | 2023826-2017-01540 | 2017-08-24 | Injury | STAAR SURGICAL COMPANY | 2017-10-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP), NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH (ACD) AND ACUTE ANGLE CLOSURE. THE PERIPHERAL IRIDECTOMIES WERE ENHANCED. THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 -2 AND THE PATIENT IS HAPPY. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER STATED THE PERIPHERAL IRIDECTOMIES WERE "TAPPED/ENHANCED" DUE TO THE ELEVATED INTRAOCULAR PRESSURE BUT WAS UNSUCCESSFUL. THERE WERE ADDITIONAL PATIENT SYMPTOMS; PATIENT'S VISION WENT "WHITE". THE LENS EXCHANGE WAS PERFORMED ON (B)(6) 2017. (B)(4). Manufacturer Narrative: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. LENS WAS RETURNED IN LIQUID IN A CENTRIFUGE VIAL INSIDE A BIOHAZARD BAG. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID IN A CENTRIFUGE VIAL. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6955546&pc=MTA | 2023826-2017-01576 | 2017-09-11 | Injury | STAAR SURGICAL COMPANY | 2017-10-17 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. REPORTEDLY, THE LENS WAS EXPLANTED DUE TO IOP SPIKE, EXCESSIVE VAULT, AND NARROW ANGLES. REPORTER STATES LENS IS PLANNED TO BE EXCHANGED WITH A SMALLER LENS; ORIGINAL LENS HAS BEEN LOST BY PHYSICIAN. ADDITIONAL INFORMATION IS NOT KNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY, ON (B)(6) 2017, A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5/+2.5/108 WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE. THE LENS WAS EXPLANTED ON (B)(6) 2017. THE LENS WAS DAMAGED DURING EXPLANT. THE DOCTOR IS "WAITING FOR PUPIL TO CONSTRICT" BEFORE DECIDING ON FURTHER ACTION. THE DOCTOR SUSPECTS URRETS-ZAVALIA. PREVIOUS REPORT INCLUDED ELEVATED IOP; CURRENT REPORT FROM COMPLAINANT DOES NOT. CURRENT SYMPTOMS REPORTED: EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), UNREACTIVE (FIXED) PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA. ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (FIXED PUPIL, IRIS ATROPHY, SUSPECTED URRETS-ZAVALIA). (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871042&pc=MTA | 2023826-2018-01367 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: DEVICE WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION TOLERANCES. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7871107&pc=MTA | 2023826-2018-01368 | 2017-09-30 | Injury | STAAR SURGICAL COMPANY | 2018-09-12 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE PATIENT DID NOT WANT ANOTHER LENS IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND DEFORMED. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO BROKEN HAPTIC. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7050133&pc=MTA | 2023826-2017-01769 | 2017-10-18 | Injury | STAAR SURGICAL COMPANY | 2017-11-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PRODUCT EVALUATION FOUND LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6978046&pc=MTA | 2023826-2017-01624 | 2017-10-27 | Injury | STAAR SURGICAL COMPANY | 2017-10-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/090 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT AND FIBERS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7113646&pc=MTA | 2023826-2017-01885 | 2017-11-08 | Injury | STAAR SURGICAL COMPANY | 2017-12-13 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTED INDICATED THAT THE SURGEON IMPLANTED A MICL13.2,-16.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED (INTRAOCULAR PRESSURE) IOP, NARROWING OF THE ANGLE AND SHALLOWING OF ANTERIOR CHAMBER DEPTH (ACD). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT POST-OP BCVA IS 20/40. Manufacturer Narrative: SERIAL NUMBER - PREVIOUS SERIAL NUMBER WAS INCORRECT, (B)(4) IS THE CORRECT SERIAL NUMBER. (B)(4) Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7197904&pc=MTA | 2023826-2018-00068 | 2017-12-08 | Injury | STAAR SURGICAL COMPANY | 2018-01-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGERY IMPLANTED A MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXPLANTED DUE TO THE LENS BEING TOO BIG. THE LENS WAS EXCHANGED FOR SHORTER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL TAPED TO PURCHASE RETURNED FORM. VISUAL INSPECTION FOUND THE HAPTIC TORN AND THE LENS MISSING A PIECE OF HAPTIC AND OPTIC. (B)(4) Manufacturer Narrative: ADDITIONAL DATA: DATE OF BIRTH: (B)(6). GENDER: FEMALE. DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A MICL12.6 IMPLANTABLE COLLAMER LENS OF -11.0 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP ASSESSED ON (B)(6) 2017. (B)(4) Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7281453&pc=MTA | 2023826-2018-00262 | 2017-12-22 | Injury | STAAR SURGICAL COMPANY | 2018-02-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RED AND CLEAR RESIDUE ON THE LENS SURFACE. CONCLUSION: THIS LENS IS CONTRAINDICATED FOR USE IN PATIENTS WITH ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0 MM. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356136&pc=MTA | 2023826-2018-00431 | 2018-01-04 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA).THE PROBLEM WAS NOT RESOLVED DUE TO THE REPLACEMENT LENS DID NOT UNFOLD PROPERLY AFTER IMPLANT. SEE MFR REPORT# 2023826-2018-00429 FOR CASE RESOLUTION. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL ON LENS SURFACE (RED AND CLEAR RESIDUE ON LENS) AND THE LENS HAS A CURVED SHAPE. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CONCLUSION CODE: OFF-LABEL USE [OVER 45 YEARS OF AGE AT TIME OF IMPLANT ((B)(6))] (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356130&pc=MTA | 2023826-2018-00437 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC TEAR AND RESIDUE ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356135&pc=MTA | 2023826-2018-00435 | 2018-01-17 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND A "PROTUBERANT" IRIS. ON (B)(6) 2018 A SHORTER LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED INA MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7324462&pc=MTA | 2023826-2018-00359 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HE IMPLANTED A 13.2 MM MICL13.2, -4.0 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT RIGHT EYE (OD) ON (B)(6) 2018. ON POST OP DAY 1 EXCESSIVE VAULT, ELEVATED IOP AND ANGLE CLOSURE WAS DISCOVERED. THE PATIENT WAS PLACED ON COMBIGAN TO LOWER IOP. AT DAY 3 AND 5 THE ELEVATED VAULT PERSISTED WITH NARROWING OF THE ANGLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTED STATED THAT THE PATIENT IS HAPPY WITH GOOD CURRENT VISION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TEARS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7324571&pc=MTA | 2023826-2018-00360 | 2018-01-18 | Injury | STAAR SURGICAL COMPANY | 2018-03-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT HE IMPLANTED A 13.2MM MICL13.2, -4.5 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT LEFT EYE (OS) ON (B)(6) 2018. ON POST OP DAY 1 EXCESSIVE VAULT, ELEVATED IOP AND ANGLE CLOSURE WAS DISCOVERED. THE PATIENT WAS PLACED ON COMBIGAN TO LOWER IOP. AT DAY 3 AND 5 THE ELEVATED VAULT PERSISTED WITH NARROWING OF THE ANGLE. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTED STATED THAT THE PATIENT IS HAPPY WITH GOOD CURRENT VISION. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TEARS. THIS IS A RE-SUBMISSION TO CORRECT MFR # FOR SUPPLEMENTAL #1 IN WHICH THE MFR NUMBER WAS REPORTED INCORRECT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7356143&pc=MTA | 2023826-2018-00434 | 2018-01-24 | Injury | STAAR SURGICAL COMPANY | 2018-03-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SMALLER LENS DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, ELEVATED IOP, AND "PROTUBERANT" IRIS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884850&pc=MTA | 2023826-2018-01266 | 2018-02-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.00/+1.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ASSESMENT ON (B)(6) 2018, ANGLE CLOSURE WITH ELEVATED IOP (26MMHG) AND EXCESSIVE VAULT. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL (RESIDUE) ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7558973&pc=MTA | 2023826-2018-00834 | 2018-02-28 | Injury | STAAR SURGICAL COMPANY | 2018-05-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED IN A 2ND SURGERY THE SAME DAY DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4)- SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621051&pc=MTA | 2023826-2018-00923 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.0/+2.0/089 SPHERE/CYLINDER/AXIS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. SECONDARY SURGERY, LENS EXCHANGE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7621071&pc=MTA | 2023826-2018-00921 | 2018-03-01 | Injury | STAAR SURGICAL COMPANY | 2018-06-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA; DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8275983&pc=MTA | 2023826-2019-00170 | 2018-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -10.00/2.0/85 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE). IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO "FALSE W-TO-W, MEASUREMENT PROVIDED." THE REPORTER STATED THAT THE PATIENTS EYE IS STABLE AND IN GOOD CONDITION. IT WAS REPORTED THAT THE DOCTOR PLANS TO IMPLANT ANOTHER LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7589676&pc=MTA | 2023826-2018-00879 | 2018-05-05 | Injury | STAAR SURGICAL COMPANY | 2018-06-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL AND THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND SHORTER LENS IMPLANTED PATIENT CODE: (B)(4) SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM AND WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172167&pc=MTA | 2023826-2018-01850 | 2018-05-19 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, PUPIL BLOCK, ELEVATED IOP, AND ENLARGEMENT OF OR ADDITION OF PI. IT'S REPORTED THAT ON (B)(6) 2018 THE LENS WAS EXPLANTED. IT IS UNKNOWN AS TO WHETHER A REPLACEMENT LENS WAS IMPLANTED DUE TO CONFLICTING INFORMATION REPORTED. ATTEMPTS TO OBTAIN CLARIFICATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7772418&pc=MTA | 2023826-2018-01180 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-08-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -12.0/1.00/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: PREVIOUSLY STATED, "VISUAL INSPECTION FOUND THE HAPTIC TORN." CORRECTED STATEMENT: "VISUAL INSPECTION FOUND THE HAPTICS BENT." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647578&pc=MTA | 2023826-2018-00984 | 2018-06-06 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO IMPROPER FIT. ANOTHER LENS WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE. THE INCISION WAS ENLARGED AND SUTURE WAS REQUIRED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. THE LENS WAS RETURNED DRY, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: ONE SIMILAR COMPLAINT TYPE EVENT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7683601&pc=MTA | 2023826-2018-01019 | 2018-06-09 | Injury | STAAR SURGICAL COMPANY | 2018-07-12 | MTA | MICL | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE SURGEON INDICATED THAT A12.6MM MICL12.6, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: ANGLE CLOSURE. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: ON (B)(6) 2018 THE LENS WAS EXPLANTED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING AND ANGLE CLOSURE WITH ELEVATED IOP. ANOTHER LENS WAS IMPLANTED ON THE SAME DAY. ON THE PATIENT'S LAST VISIT ON (B)(6) 2018, THE STATUS OF THE EYE WAS COMFORTABLE WITH POOR PUPIL CONSTRICTION AND MINIMAL HALO AND GLARE. (LENS EXCHANGE) TO (LENS EXPLANT) - PREVIOUS CODE IS NOT APPLICABLE, LENS WAS EXPLANTED. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE VAULT WAS IMPROVED, BUT IT WAS STILL HIGH. IOP (INTRAOCULAR PRESSURE) IS WNL (WITHIN NORMAL LIMITS). IT IS BELIEVED THAT THE HIGH IOP (INTRAOCULAR PRESSURE) FROM PUPILARY BLOCK CAUSED THE CONSTRICTION ISSUE. PUPIL CONSTRICTION AND GLARE IMPROVED BUT ARE STILL PRESENT. (B)(4). Manufacturer Narrative: LENS WAS RETURNED DRY IN CARTRIDGE TRAY. VISUAL INSPECTION FOUND A TORN HAPTIC AND FOREIGN MATERIAL ON THE SURFACE OF THE LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903849&pc=MTA | 2023826-2018-01373 | 2018-06-12 | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/6.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 EXCESSIVE VAULT WAS OBSERVED, PATIENT WAS ASSESSED WITH ANGLE CLOSURE WITH ELEVATED IOP (50 MMHG) ON (B)(6) 2018 THE LENS WAS EXPLANTED. SURGEON PLANS TO IMPLANT ANOTHER ICL LENS IN (B)(6) 2018. EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: "ADVERSE EVENT" IN THE INITIAL MDR. CORRECTED TO: "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)" IN THE INITIAL MDR. ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7765108&pc=MTA | 2023826-2018-01169 | 2018-06-19 | Injury | STAAR SURGICAL COMPANY | 2018-08-08 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, PUPIL BLOCK, AND ELEVATED IOP.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED. THE SURGEON REPORTS VAULT VALUES OF 800UM/1.2CT, ELEVATED IOP, AND REFRACTIVE TREATMENT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE (B)(4): NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT, PUPIL BLOCK). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7706773&pc=MTA | 2023826-2018-01067 | 2018-06-20 | Injury | STAAR SURGICAL COMPANY | 2018-07-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM MICL12.6, -16.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED BY A SURGEON THAT WAS ON CALL DUE TO OBSERVATION OF EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL BLOCK AND ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT AND ADDITION OF NEW PI, YAG) WERE PERFORMED. ON (B)(6) 2018 AN EXCHANGED FOR A SHORTER LENGTH LENS WAS PERFORMED. THIS EXCHANGE RESOLVE THE PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7903843&pc=MTA | 2023826-2018-01372 | 2018-06-25 | Injury | STAAR SURGICAL COMPANY | 2018-09-24 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. EXCESSIVE VAULT WAS OBSERVED AND ELEVATED IOP (45MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE ASSESSED ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824081&pc=MTA | 2023826-2018-01272 | 2018-07-21 | Injury | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +10.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND PUPIL BLOCK. ON (B)(6) 2018 THE LENS WAS EXPLANTED. PHACO-EMULSIFICATION AND IOL IMPLANTATION WERE PERFORMED ALONG WITH CLE SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7884841&pc=MTA | 2023826-2018-01392 | 2018-08-21 | Injury | STAAR SURGICAL COMPANY | 2018-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -09.5/6/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: "LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT, SHALLOWING ACD, AND HIGHER IOP" NOT INCLUDED IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090043&pc=MTA | 2023826-2018-01696 | 2018-08-22 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090042&pc=MTA | 2023826-2018-01699 | 2018-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ANGLE CLOSURE, ELEVATED IOP, AND SHALLOW AC. THE PROBLEM WAS RESOLVED. "SURGEON AND PATIENT ARE BOTH HAPPY." THE CAUSE OF THE EVENT IS REPORTED AS "OVERSIZE LENS." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATED BY MFR -DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8277800&pc=MTA | 2023826-2019-00177 | 2018-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-25 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). ADDITIONAL INFORMATION REPORTED: PATIENT CANCELLED THE SURGERY FOR BOTH EYES AFTER THE EXPLANTED OD LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: IMPLANTED DATE: CORRECTED FROM "(B)(6) 2018" TO (B)(6) 2018" FOR INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -2.50/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). PATIENT CANCELLED THE SURGERY AFTER THE EXPLANTED OD LENS AND HAS NOT YET SCHEDULED FOR ANY SURGERIES. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CAS/VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON PRODUCT AND A BENT HAPTIC. DIMENSIONAL INSPECTION WAS WITHIN LENS SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7954052&pc=MTA | 2023826-2018-01504 | 2018-09-06 | Injury | STAAR SURGICAL COMPANY | 2018-10-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -15.5 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. AT ONE DAY POST-OP, THE LENS WAS TRYING TO PROTRUDE THROUGH THE PUPIL AND WAS KINKED (TOO BIG). THE PUPIL WOULD NOT CONSTRICT. THE LENS WAS EXPLANTED ON (B)(6) 2018 BUT THE PUPIL REMAINED DILATED. THE INCISION WAS ENLARGED SLIGHTLY TO REMOVE THE LENS AND SUTURES WERE REQUIRED. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED AT 1 DAY POST-OP, SHOWED HIGH CENTRAL VAULT. THE INTRAOCULAR PRESSURE (IOP) WAS HIGH FOR FIRST FEW DAYS, YET 1 DAY POST-OP, THE EYE WAS SOFT WITH FLAT ANTERIOR CHAMBER. 2 DAY POST-OP, IOP 28MMHG. THE SURGEON FELT THE ORBSCAN WTW MAY HAVE BEEN MEASURED INCORRECTLY. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, TAPED TO A PLASTIC CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE ON THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8176802&pc=MTA | 2023826-2018-01854 | 2018-09-18 | Injury | STAAR SURGICAL COMPANY | 2018-12-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -17.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) (30 MMHG) AND PUPIL DILATION. THE PATIENT WAS TREATED WITH MEDICATION. DUE TO THE PATIENTS SYMPTOMS, THE SURGEON PERFORMED A YAG IRIDOTOMY, WITH ADDITIONAL MEDICAL THERAPY, BUT IT WAS UNSUCCESSFUL. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. AFTER ICL EXCHANGE, THE IOP WAS VERY HIGH (>40MMHG) AND THE SURGEON SUSPECTED SENSITIVITY TO TOPICAL STEROID THERAPY. AFTER TAPERING OFF THE STEROID AND IOP-LOWERING THERAPIES, THE IOP COMPLETELY NORMALIZED. Manufacturer Narrative: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT AND PRESENCE OF RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8048978&pc=MTA | 2023826-2018-01612 | 2018-10-10 | Injury | STAAR SURGICAL COMPANY | 2018-11-07 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VICM5_13.2, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND FIBER ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8113339&pc=MTA | 2023826-2018-01752 | 2018-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-28 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING, WHICH CAUSED AN INCREASE IN PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: PREVIOUS MDR STATED: "EXCESSIVE VAULT WHICH CAUSED AN INCREASE IN PRESSURE". CORRECTED TO : "EXCESSIVE VAULT ONLY" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8217921&pc=MTA | 2023826-2019-00009 | 2018-11-16 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: "TTHE" REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES WERE RETURNED UNIDENTIFIED IN THE SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS; PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00010 FOR CROSS REFERENCE. LENS #2- TWO LENSES WERE RETURNED UNIDENTIFIED IN THE SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00009. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8218075&pc=MTA | 2023826-2019-00010 | 2018-11-16 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: TTHE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -6.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS #1- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A TORN HAPTIC; PRESENCE OF RESIDUE ON LENS SURFACE. MDR # 2023826-2019-00009 FOR CROSS REFERENCE. LENS #2- TWO LENSES RETURNED UNIDENTIFIED IN SAME BAG. LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. MDR# 2023826-2019-00010 FOR CROSS REFERENCE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8768409&pc=MTA | 2023826-2019-01184 | 2018-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13 MM, VICMO13.2, -17.00 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, ELEVATED IOP AND PIGMENT DISPERSION. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER ALSO REPORTER THAT WHEN THE LENS WAS REMOVED THEY NOTICED THAT ONE IF THE HAPTICS HAD A FISSURE (INCOMPLETE BROKE). THE REPORTER STATED " THE LENS WAS REMOVED DURING A SECOND SURGERY, SURGEON TRIED TO MAINTAIN THE IOP USING MEDICATION. BECAUSE THIS TREATMENT WASN'T EFFECTIVE DURING 4-5 MONTHS, FINALLY DECIDED TO REMOVE THE LENS DUE TO POTENTIAL DAMAGE OF RETINA. AN OCT-SCAN WAS PERFORMED PRIOR TO SECOND SURGERY AND RAISED QUESTION ABOUT ICL STABILITY DUE TO A SLIGHTLY ASYMMETRY, INITIALLY INTERPRETED AS A NASAL POSTERIOR SYNECHIAE. THE MOMENT OF SECOND SURGERY WAS POSTPONED FOR A WHILE DUE TO ATTEMPT TO CONTROL IOP WITH MEDICATION, AS REFRACTIVE OUTCOME WAS ACCEPTABLE AND PATIENT LIVE AND WORK ABROAD". Manufacturer Narrative: VISUAL INSPECTION ALSO FOUND RESIDUE ON THE LENS. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, IOP TREATED WITH MEDICATION CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8219260&pc=MTA | 2023826-2019-00044 | 2018-12-05 | Injury | STAAR SURGICAL COMPANY | 2019-01-04 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2018. ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED, DATE ASSESSED: (B)(6) 2018. LENS REMAINS IMPLANTED, SURGEON HAD PLANNED TO EXCHANGE THE LENS ON (B)(6) 2018 FOR A SHORTER LENGTH LENS BUT THE SURGERY WAS POSTPONED BY THE PATIENT AND HAS NOT RESCHEDULED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, MICL13.7, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2018. ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED, DATE ASSESSED: (B)(6) 2018. THE LENS WAS REMOVED ON (B)(6) 2019 AND REPLACED WITH A SHORTER LENGTH LENS. IT WAS REPORTED THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM AND THE PATIENT IS FINE. THE EXPLANTED LENS WAS DISCARDED AND WILL NOT BE RETURNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SECONDARY SURGICAL INTERVENTION, EXPLANT. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356222&pc=MTA | 2023826-2019-00297 | 2018-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A13.2MM, VICMO13.2, -14.50 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP (DATE ASSESSED: (B)(6) 2018). THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PRESENCE OF RESIDUE ON LENS SURFACE WAS ALSO OBSERVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656475&pc=MTA | 2023826-2019-00970 | 2018-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN/BROKEN OFF. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE HAPTIC/OPTIC MISSING. (B)(4). DATE OF EVENT: (B)(6) 2018. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. IMPLANT DATE: (B)(6) 2018. EXPLANT DATE: (B)(6) 2018. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372544&pc=MTA | 2023826-2019-00330 | 2019-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. NO ADDITIONAL INFORMATION IS AVAILABLE. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER STATES ADDITIONAL PATIENT SYMPTOMS OF ANGLE CLOSURE AND ELEVATED IOP AND ALSO NOTES "ANISOCORIA ON OS." REPORTEDLY, ANISOCORIA WAS NOTICED AFTER LENS EXPLANT. BEFORE SURGERY, "PUPIL SIZE WAS FINE." IT IS UNKNOWN AS TO WHETHER THE LENS CAUSED THIS OR NOT. DISTRIBUTOR WAS CHECKED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633167&pc=MTA | 2023826-2019-00877 | 2019-02-06 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND UNREACTIVE (FIXED) PUPIL. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093969&pc=MTA | 2023826-2020-01120 | 2019-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471422&pc=MTA | 2023826-2019-00458 | 2019-03-01 | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -9.5 DIOPTER , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNELA ANGLES, AND ELEVATED IOP (24 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A BIOHAZARD ZIPLOC BAG, IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8488607&pc=MTA | 2023826-2019-00541 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE. ON (B)(6) 2019 A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED AND THEY STATED THAT THE PROBLEM RESOLVED. HOWEVER, FOR STATUS OF THE EYE THEY REPORTED STATED H CELL FLARE. Manufacturer Narrative: IT WAS REPORTED THAT THE CELL FLARE/INFLAMMATION HAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: EXPLANTED: (B)(6) 2019. (B)(4). DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. THERE WAS DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON LENS SURFACE. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633168&pc=MTA | 2023826-2019-00878 | 2019-03-12 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND ELEVATED IOP. CURRENTLY, NO REPLACEMENT LENS HAS BEEN IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279660&pc=MTA | 2023826-2019-02066 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586235&pc=MTA | 2023826-2019-00745 | 2019-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586259&pc=MTA | 2023826-2019-00746 | 2019-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL IN LIQUID. VISUAL INSPECTION FOUND LENS HAPTIC TORN AND BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8588494&pc=MTA | 2023826-2019-00777 | 2019-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. HIGH IOP (INTRAOCULAR PRESSURE) AND A VAULT VALUE WAS REPORTED. PATIENT BCVA IS 20/20, LENS REMAINS IN THE EYE AND SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2, -9.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 21 MAR 2019. HIGH IOP AND A VAULT VALUE WAS REPORTED. ON 25 APR 2019 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM, PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT .(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8591429&pc=MTA | 2023826-2019-00778 | 2019-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. HIGH IOP (INTRAOCULAR PRESSURE) AND A VAULT VALUE WAS REPORTED. PATIENT BCVA IS 20/20, LENS REMAINS IN THE EYE AND SURGEON PLANS TO EXCHANGE THE LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIA WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICM5_13.2, -7.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 21 MAR 2019. HIGH IOP AND A VAULT VALUE WAS REPORTED. ON 25 APR 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317909&pc=MTA | 2023826-2019-02125 | 2019-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -10.5, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH THE SAME SIZE LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957671&pc=MTA | 2023826-2019-01535 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, ELEVATED IOP (27 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WERE OBSERVED ON (B)(6) 2019. LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913404&pc=MTA | 2023826-2020-00721 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8939422&pc=MTA | 2023826-2019-01558 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICM5 12.6, -05.00, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. PUPIL BLOCK WITH ELEVATE IOP (50 MMHG) AND EXCESSIVE VAULT WERE OBSERVED ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019. CAUSE OF THE EVENT IS REPORTED AS "IRIS PROLAPSE OCCURRED DURING SURGERY. OVD REMAINED IN THE EYE AFTER SURGERY. ON THE DAY OF SURGERY, POSTOPERATIVE PUPILLARY MYDRIASIS OCCURRED DUE TO INCREASED INTRAOCULAR PRESSURE." PER THE REPORTER ON (B)(6) 2019, "THE DOCTOR WILL PRESCRIBE AN IRIS IMAGE CONTACT LENS TO THE PATIENT. ON THE DAY OF THE SURGERY THE DOCTOR SUGGESTED TO THE PATIENT THAT RE-OPERATE TO REMOVE OVD, BUT THE PATIENT REFUSED BECAUSE OF SEVERE EYE PAIN." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (MEDICATION: GLYCEOL, DIAMOX). ON THE DAY OF THE SURGERY, MEDICATION WAS PRESCRIBED TO LOWER THE IOP (GLYCEOL 200ML AND DIAMOX) IN THE MORNING AND EVENING. THERE WAS NO POSTOPERATIVE TREATMENT TO MANAGE THE IOP DUE TO THE PATIENT'S IOP LOWERED THE NEXT DAY. THE SURGEON STATED THE CAUSE OF THE IRIS PROLAPSE WAS DUE TO "VITREOUS BODY PRESSURE INCREASED BECAUSE OF THE PATIENT WITH STRONG TENSION DURING THE OPERATION. THERE WAS A POSSIBILITY THAT THE INCISION IS LINER AND SHORT". CLAIM#: (B)(4). Manufacturer Narrative: NO CODE AVAILABLE (IRIS PROLAPSE, PUPILLARY MYDRIASIS, EYE PAIN). (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317874&pc=MTA | 2023826-2019-02139 | 2019-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT, ELEVATED IOP, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA), AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PREVIOUSLY REPORTED ANGLE CLOSURE. UPDATED/CORRECTED INFO: PATIENT DID NOT EXPERIENCE ANGLE CLOSURE. PLEASE DISREGARD. PATIENT CODE 3191: ANGLE CLOSURE WAS REPORTED. THIS SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908374&pc=MTA | 2023826-2019-01488 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE (IOP). AT 4 HOUR IOP CHECK, THE IOP WAS 51 MMHG DUE TO ANGLE CLOSURE FROM EXCESSIVE VAULTING. IOP WAS CONTROLLED BY TAPPING PARACENTESIS AND COMBIGAN DROPS. THE PATIENT WAS TAKEN BACK TO THE OPERATING ROOM TO EXPLANT THE ICL. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: RETURNED IN LIQUID IN SPECIMEN CUP. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910129&pc=MTA | 2023826-2019-01467 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM513.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (24 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2019. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS REPOSITIONED ON (B)(6) 2019 TO 90 DEGREES VERTICAL AND THE PROBLEM RESOLVED. POSTOPERATIVE REPORT STATED "VAULTING IS BETTER, LENS STAYS IN THE EYE, NEXT EXAMINATION IN ONE YEAR". PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910138&pc=MTA | 2023826-2019-01468 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM513.2, -10.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (25 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2019. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS WAS REPOSITIONED 90 DEGREES TO VERTICLE POSITION ON (B)(6) 2019 AND IT RESOLVED THE PROBLEM. POSTOPERATIVE REPORT STATED "VAULTING IS BETTER AFTER ROTATION, NEXT EXAM IN 1 YEAR". H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9588875&pc=MTA | 2023826-2020-00121 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER STATED THAT THE PATIENT EXPERIENCED ELEVATED IOP AND EXCESSIVE VAULT. IOP WAS CONTROLLED WITH ANTI IOP DROPS, BCVA 6/12. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 13.2MM VICM5 13.2 OF -11.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER STATED THE PATIENT EXPERIENCED EXCESSIVE VAULT AND ELEVATED IOP. THE ELEVATED IOP WAS CONTROLLED WITH ANTI IOP DROPS. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308814&pc=MTA | 2023826-2020-01525 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO ELEVATED IOP (40MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE AND EXCESSIVE VAULT. ON (B)(6) 2020 A SHORTER LENGHT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "LENS WAS LARGER THAN IT SHOULD HAVE BEEN." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9702382&pc=MTA | 2023826-2020-00300 | 2019-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THA A 12.6MM, VICH12.6, +6.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS LATER REMOVED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT THE SURGEON PERFOMED FEMTO LASIK TO CORRECT RESIDUAL REFRACTION AFTER LENS EXPLANT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A LENS CASE/VIAL DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9891527&pc=MTA | 2023826-2020-00598 | 2020-03-04 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 26/FEB/2020. PUPIL BLOCK WITH ELEVATED IOP (48MMHG) AND ANGLE CLOSURE WITH ELEVATED IOP (48MMHG) WERE ASSESSED ON (B)(6) 2020, AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE (EXCESSIVE VAULT). REPORTEDLY, "PATIENT HAPPY." Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910966&pc=MTA | 2023826-2020-00675 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -4.5 DIOPTER, INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE SURGEON EXCHANGED THE LENS WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE "(SIZE)." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN LENS CASE/ VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214025&pc=MTA | 2023826-2020-01383 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS EXPLANTED DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (INTRAOCULAR PRESSURE) AND THE PROBLEM RESOLVED. FOR STATUS OF THE EYE THE REPORTER STATED " THE CORNEA IS TRANSPARENT, THE PUPIL IS 3.5X3.0MM, THE PUPIL IS INSENSITIVE TO LIGHT. THE LENS IS TRANSPARENT, AND THER IS A LITTLE PIGMENT ON THE SURFACE OF THE ANTERIOR CAPSULE". ON (B)(6) 2020 THE REPORTER STATED " THE PIGMENTATION IS NOT INCREASED. THIS OBSERVATION IS NOT BEYOND NORMAL". Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR G3 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: SUPPLEMENTAL MDR#2 IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183420&pc=MTA | 2023826-2018-00045 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -15.0/2/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2 IMPLANTABLE COLLAMER LENS, OF -15.0/2/110 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018, DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4) Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7183423&pc=MTA | 2023826-2018-00046 |  | Injury | STAAR SURGICAL COMPANY | 2018-01-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2, -14/1/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. EXCESSIVE VAULT AND ELEVATED IOP WAS NOTICED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, OF -14/1/101 SPHERE/ CYLINDER/ AXIS IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED IOP. THE PROBLEM WAS RESOLVED AND THE SURGEON STATED THAT THE "VAULT WAS NORMAL". (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID INSIDE A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM NUMBER OF THE FILE IN THE PREVIOUS SUPPLEMENTAL IS CORRECTED FROM "CLAIM# (B)(4)" TO " CLAIM# (B)(4)". Manufacturer Narrative: IT IS CORRECTED FROM "OTHER SERIOUS" TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT / DAMAGE" (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7278702&pc=MTA | 2023826-2018-00234 |  | Injury | STAAR SURGICAL COMPANY | 2018-02-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V4C IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED 3 DAYS AFTER BEING IMPLANTED IN THE PATIENTS EYE, DUE TO HIGH VAULT WITH APPOSITIONAL ANGLE CLOSURE. ICL (IMPLANTABLE COLLAMER LENS) WAS EXCHANGED WITH A SMALLER SIZE LENS. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A V4C IMPLANTABLE COLLAMER LENS. THE LENS WAS EXPLANTED 3 DAYS AFTER BEING IMPLANTED IN THE PATIENT'S EYE DUE TO HIGH VAULT WITH APPOSITIONAL ANGLE CLOSURE. THE REPORTER ALSO STATED THAT "EXCESSIVELY HIGH VAULT LED TO RAISED IOP". A SMALLER SIZE LENS WAS IMPLANTED LATER. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED WITH SHORTER LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7527986&pc=MTA | 2023826-2018-00758 |  | Injury | STAAR SURGICAL COMPANY | 2018-05-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE PATIENT REPORTED HAD AN ICL LENS IMPLANTED ON (B)(6) 2018 AND AT ONE DAY POST-OP HAD HIGH EYE PRESSURE. THE FACILITY REPORTED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE INTRAOCULAR PRESSURE WAS STABLE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL (RESIDUE) ON LENS SURFACE. CORRECTED INFORMATION: PATIENT IDENTIFIER SHOULD BE "(B)(6)", (B)(6) IS INCORRECT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. PATIENT CODE: (B)(4) SECONDARY SURGICAL INTERVENTION, LENS EXPLANTED AND EXCHANGED FOR SHORTER LENS CONCLUSION CODE: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7627647&pc=MTA | 2023826-2018-00932 |  | Injury | STAAR SURGICAL COMPANY | 2018-06-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/+3/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT SIZE LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE). THE REPORTED STATED THAT THE PATIENT'S CONDITION IS VERY GOOD NOW. Manufacturer Narrative: CORRECTED DATA: COMPLAINT "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES)". OUTCOME SHOULD HAVE BEEN MENTIONED IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: DATE OF EXCHANGE TOOK PLACE ON (B)(6) 2018. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8194489&pc=MTA | 2023826-2018-01906 |  | Injury | STAAR SURGICAL COMPANY | 2018-12-24 | MTA | UNK | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: JOURNAL OF REFRACTIVE SURGERY ARTICLE WAS RECEIVED ON 30-NOV-2018, ORIGINALLY PUBLISHED IN 2010, WHICH CITED A STUDY TITLED, 'CORRECTION OF RESIDUAL REFRACTIVE ERROR IN PSEUDOPHAKIC EYES WITH THE USE OF A SECONDARY PIGGYBACK TORIC IMPLANTABLE COLLAMER LENS.' THE ARTICLE STATES THAT PUPILLARY BLOCK AND ELEVATED IOP OCCURRED IN ONE EYE ON POSTOPERATIVE DAY 1. LASER IRIDECTOMY USING ARGON AND ND:YAG WAS PERFORMED. IOP DECREASED WITHIN 2 HOURS. IT'S SUSPECTED THAT THE PUPILLARY BLOCK OCCURRED DUE TO INSUFFICIENT PERIPHERAL IRIDECTOMY. IN ADDITION, 6 OUT OF 8 EYES SHOWED HIGH VAULT. HOWEVER, ANGLES REMAINED OPEN AND NO ELEVATED IOP WAS OBSERVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT MANUFACTURED BUT NOT SOLD IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409833&pc=MTA | 2023826-2019-00267 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICH13.2, +07.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, ELEVATED IOP, AND EXCESSIVE VAULT WAS REPORTED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A SHORTER LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8500521&pc=MTA | 2023826-2019-00587 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.50 WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND PUPIL BLOCK WITH ELEVATED IOP (INTRAOCULAR PRESSURE). ADDITIONAL SURGICAL INTERVENTION (ENLARGEMENT OF PI OR ADDITION OF NEW PI/ YAG) WAS REPORTED TO BE PERFORMED. ON (B)(6) 2019 THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9031491&pc=MTA | 2023826-2019-01697 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: AN ARTICLE WAS PUBLISHED IN THE JOURNAL OF CATARACT & REFRACTIVE SURGERY IN (B)(6) 2019 TITLED, '10-YEAR CLINICAL OUTCOMES AFTER IMPLANTATION OF A POSTERIOR CHAMBER PHAKIC LENS FOR MYOPIA.' THE ARTICLE STATES THAT THERE WERE THREE CASES OF TRANSIENT IOP INCREASE AFTER ICL IMPLANTATION. TWO OF 3 EYES WERE DUE TO ANGLE CLOSURE FROM HIGH VAULT, WHICH RETURNED TO NORMAL IOP AFTER ICL EXCHANGE, AND ONE EYE RETURNED TO NORMAL IOP AFTER USING IOP-LOWERING AGENTS FOR ONE MONTH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191 (MEDICAL INTERVENTION - LENS EXCHANGE, TREATMENT WITH MEDICATION) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. MODEL #: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MANUFACTURE DATE: UNK. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013329&pc=MTA | 2023826-2020-00963 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ELEVATED IOP (25 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020, EXCESSIVE VAULT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE CAUSE OF EVENT IS REPORTED AS OTHER - NEW WTW CALCULATION METHOD." SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013330&pc=MTA | 2023826-2020-00964 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ELEVATED IOP (22 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE OBSERVED ON (B)(6) 2020, EXCESSIVE VAULT, AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "THE CAUSE OF EVENT IS REPORTED AS OTHER - NEW WTW CALCULATION METHOD." SHOULD BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689153&pc=MTA | 2023826-2020-00281 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -16.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 "DUE TO EYE PRESSURE TOO HIGH." NO REPLACEMENT LENS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE A12.6MM MICL 12.6 OF -16.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO PRESSURE SPIKE SECONDARY TO HIGH VAULT. NARROW ANGLES, PIGMENT AND BLOOD IN THE ANGLES. PATIENTS IS NOW BACK TO NORMAL. DOCTOR IS PLANNING ON PLACING A SMALLER ICL. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED AT THIS TIME. CORRECTED DATA: DEVICE AVAILABLE FOR EVALUATION DATE IN PREVIOUS MDR SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G3- DATE IN PREVIOUS MDR IS CORRECTED TO (B)(6) 2020. DEVICE CODE- 1583 SHOULD BE INCLUDED IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9648851&pc=MTA | 2023826-2020-00206 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-30 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem; Unintended Movement | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: A DAMAGED (TEAR ON HAPTIC) 13.2MM, MICL 13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS RETURNED TO STAAR. " OS (TANYA MOORE ) REPLACED OLD WITH NEW ICL", WAS HANDWRITTEN ON THE LENS PACKAGING BOX. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -10.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR THE SAME MODEL LENS, LENGTH AND DIOPTER DUE TO LOW VAULT OVERTIME, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PARTIAL ANGLE CLOSURE (FOCAL ANGLE CLOSURE) AND PROGRESSIVE NARROWING OF ANGLE WITH ELEVATED IOP (THAT REQUIRED MEDICAL INTERVENTION). THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THIS EVENT WAS DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. THE REPORTER STATED " THE ORIGINAL ICL WAS INADVERTENTLY INVERTED ON INSERTION. THIS WAS NOT RECOGNIZED UNTIL SEVERAL YEARS AFTER SURGERY, SINCE THE PATIENT HAD GOOD VISION AND WAS TEMPORARILY LOST TO FOLLOW UP. SHE WAS REFERRED BACK TO ME INTERNALLY FROM A COLLEAGUE WHO NOTED FOCAL ANGLE CLOSURE IN ONE EYE ONLY AND INCREASING IOP. UBM REVEALED AN INVERTED ICL IN THAT EYE, WHICH I EXCHANGED SUCCESSFULLY. I BELIEVE THAT THE SMALLER INCISION BUTTERFLY STYLE INJECTOR I WAS USING AT THE TIME FOR SEVERAL MONTHS (ON ADVICE OF THE LOCAL REP AT THAT TIME) CONTRIBUTED TO THE INVERSION OF THE ICL." Manufacturer Narrative: H6- PATIENT CODE 3191: SIGNIFICANT REDUCTION OF ICA, PARTIAL ANGLE CLOSURE (FOCAL ANGLE CLOSURE), NARROWING OF ANGLE, SECONDARY MEDICAL (MEDS) AND SURGICAL (EXCHANGE) INTERVENTION. H6- WORK ORDER SEARCH: A WORK ORDER SEARCH WAS PERFORMED FOR THE CORRECT SERIAL NUMBER AND NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF BIRTH: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS CASE VIAL. VISUAL INSPECTION FOUND A TEAR ON A HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4839520&pc=MTA | 2023826-2015-00627 | 2013-01-14 | Injury | STAAR SURGICAL COMPANY | 2015-06-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2013 DUE TO INADEQUATE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE PERIPHERAL IRIDECTOMY WAS EITHER ENLARGED OR THERE WAS AN ADDITION OF A NEW PERIPHERAL IRIDECTOMY (SURGICAL). THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. PT WEIGHT: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCHS DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENT(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF INADEQUATE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012149&pc=MTA | 2023826-2015-01045 | 2014-06-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR SHORTER ONE AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4554740&pc=MTA | 2023826-2015-00134 | 2014-08-04 | Injury | STAAR SURGICAL COMPANY | 2015-02-27 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED THE NEXT DAY DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. SIGNIFICANT REDUCTION OF IRIDO-CORNEALS ANGLES WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/22. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION, (POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEONS DETERMINES THA THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4599006&pc=MTA | 2023826-2015-00303 | 2014-08-16 | Injury | STAAR SURGICAL COMPANY | 2015-03-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2013. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO ELEVATED IOP ASSOCIATED EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UCVA WAS 20/20. Manufacturer Narrative: (B)(4). EVALUATION CODES RESULTS (OTHER). A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.) TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THA THE LENS IS EXPLANTED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5182255&pc=MTA | 2023826-2015-01326 | 2014-11-05 | Injury | STAAR SURGICAL COMPANY AG | 2015-10-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO 13.7 IMPLANTABLE COLLAMER -13.50 DIOPTER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2013. EXCESSIVE VAULTING ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2014. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE ON (B)(6) 2014. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT IS NOT MARKETED IN THE U.S. DIOPTER: -13.50. (B)(4). CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4629139&pc=MTA | 2023826-2015-00375 | 2014-11-28 | Injury | STAAR SURGICAL COMPANY | 2015-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4): DEVICE EVALUATED BY MANUFACTURER? LENS NOT RETURNED. EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4): LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. MEDICAL REVIEW - ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210642&pc=MTA | 2023826-2015-01434 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -12.00/+2.00/X107 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01433 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.00/+2.00/X107 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5210651&pc=MTA | 2023826-2015-01433 | 2014-12-15 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-09 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/X064 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE ICL WAS EXPLANTED ON (B)(6) 2014. NO OTHER LENS WAS IMPLANTED. SEE MFR. #2023826-2015-01434 FOR RIGHT EYE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT IS NOT MARKETED IN THE U.S. DIOPTER: --11.50/+2.00/X064. (B)(4). CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574078&pc=MTA | 2023826-2015-00240 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER STATED THE EVENT WAS CAUSED BY THE PATIENT TAKING ORAL ANTIHISTAMINE/DECONGESTANT (CAUSING PUPIL DILATION AND ANGLE CROWDING), THIS WAS REVERSED AFTER BEING PRESCRIBED PILOCARPINE DROPS AND EYE MASSAGE. Manufacturer Narrative: (B)(4). EVALUATION, METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD - (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS - (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT WAS THE ROOT CAUSE OF THE COMPLAINT. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USER FAILURE MODE ANALYSIS, IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. EXCESSIVE VAULTING COULD CAUSE ANGLE NARROWING, CLOSURE AND HIGH IOP. ANTIHISTAMINE MEDICATION COULD HAVE THE MYDRASIS EFFECT THAT EXACERBATED THE NARROWING AND CLOSURE OF ANTERIOR ANGLE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4574128&pc=MTA | 2023826-2015-00241 | 2014-12-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-03-05 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER STATED THE EVENT WAS CAUSED BY THE PATIENT TAKING ORAL ANTIHISTAMINE/DECONGESTANT (CAUSING PUPIL DILATION AND ANGLE CROWDING), THIS WAS REVERSED AFTER BEING PRESCRIBED PILOCARPINE DROPS AND EYE MASSAGE. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED. Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - ACCORDING TO USER FAILURE MODE ANALYSIS, IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. EXCESSIVE VAULTING COULD CAUSE ANGLE NARROWING, CLOSURE AND HIGH IOP. ANTIHISTAMINE MEDICATION COULD HAVE THE MYDRASIS EFFECT THAT EXACERBATED THE NARROWING AND CLOSURE OF ANTERIOR ANGLE. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE MEDICAL REVIEW, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4513699&pc=MTA | 2023826-2015-00151 | 2014-12-31 | Injury | STAAR SURGICAL COMPANY | 2015-02-12 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014 AND THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LOW OPACITY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THERE IS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THE EVENT, HOWEVER, CANNOT RULE OUT LOW VAULTING SECONDARY TO A SHORT LENS AS A CONTRIBUTING FACTOR. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). STAAR RECOMMENDS THE LENS TO BE LEFT IMPLANTED IF NO CATARACTS ARE NOTED OR NO PROGRESSION OF ANTERIOR SUB-CAPSULAR CATARACT. THE SERIAL NUMBER WAS NOT PROVIDED SO WE WERE UNABLE TO PERFORMED WORK ORDER SEARCH. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012020&pc=MTA | 2023826-2015-01043 | 2015-01-05 | Injury | STAAR SURGICAL COMPANY AG | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014 AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT'S LAST UCVA WAS 20/20. SEE MFR REPORT #2023626-2015-01042 FOR THE EVENT ON THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DIOPTER: -12.00. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086238&pc=MTA | 2023826-2015-01237 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -23.0/+1.0/075 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086239&pc=MTA | 2023826-2015-01238 | 2015-01-20 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -22.5/+1.5/070 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PATIENT CODE: (B)(4) ANGLE CLOSURE, LENS EXCHANGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4736049&pc=MTA | 2023826-2015-00478 | 2015-01-23 | Injury | STAAR SURGICAL COMPANY | 2015-04-29 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED AS WELL AS ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOW BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). SECONDARY SURGERY. NEW INCISION. ANGLE CLOSURE. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS TORN INTO TWO PIECES. THE LENS COULD NOT BE RE-MEASURED. THE LENS WAS RETURNED DRY. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, EVALUATION OF THE RETURNED PRODUCT AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5012294&pc=MTA | 2023826-2015-01027 | 2015-01-31 | Injury | STAAR SURGICAL COMPANY | 2015-08-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (UP TO 30MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4842228&pc=MTA | 2023826-2015-00660 | 2015-02-09 | Injury | STAAR SURGICAL COMPANY | 2015-06-12 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12MM ICM120V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WERE NOTED, THE LENS WAS EXPLANTED ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016393&pc=MTA | 2023826-2015-01032 | 2015-02-14 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (UP TO 25MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5016398&pc=MTA | 2023826-2015-01033 | 2015-02-24 | Injury | STAAR SURGICAL COMPANY | 2015-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (UP TO 28MMHG). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5086251&pc=MTA | 2023826-2015-01208 | 2015-03-06 | Injury | STAAR SURGICAL COMPANY | 2015-09-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS, -13.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. A VISUAL INSPECTION WAS PERFORMED, OBSERVING THE HAPTIC TO BE TORN/BROKEN/BENT/DEFORMED. (B)(4). CONCLUSION: PER DFU, THIS LENS MODEL IS CONTRAINDICATED IN THE PATIENT'S AGE BEING UNDER 21 YEARS. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4855057&pc=MTA | 2023826-2015-00688 | 2015-03-16 | Injury | STAAR SURGICAL COMPANY | 2015-06-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM ICM125V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119297&pc=MTA | 2023826-2015-01298 | 2015-03-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4759443&pc=MTA | 2023826-2015-00522 | 2015-03-25 | Injury | STAAR SURGICAL COMPANY | 2015-05-08 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT WAS NOTED, AS WELL AS SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ANGLE CLOSURE ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. THE PATIENT'S LAST KNOWN BCVA WAS 20/100 AS OF (B)(6) 2015. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (NO CONCLUSION CAN BE DRAWN) BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743860&pc=MTA | 2023826-2015-00499 | 2015-03-27 | Injury | STAAR SURGICAL COMPANY AG | 2015-05-01 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR REPORT# 2023826-2015-00498 FOR THE OTHER EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN SWITZERLAND AND IS NOT MARKETED IN THE U.S. (B)(4) NEW INCISION. LENS EXPLANTED. VAULTING. METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. (B)(4). Manufacturer Narrative: METHOD: LENS WORK ORDER SEARCH & MEDICAL REVIEW. RESULTS: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A HAPTIC WAS TORN AND THE LENS WAS RETURNED DRY. THE WORK ORDER SEARCH SHOWED NO OTHER COMPLAINT RECORDED RELATED TO THE SAME WORK ORDER. MEDICAL REVIEW - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS (E.G. PATIENT AGE BELOW 21 OR OVER 45 YEARS, ACD BELOW 2.8MM FOR ICL/TICL AND BELOW 3.0MM FOR VICL/VTICL, KERATOCONUS, PREGNANT OR NURSING PATIENTS, PATIENTS WITH LOW/ABNORMAL CORNEAL ENDOTHELIAL CELL DENSITY, FUCH'S DYSTROPHY OR OTHER CORNEAL PATHOLOGY, PATIENTS WHO ARE AMBLYOPIC OR BLIND IN THE FELLOW EYE). OFF-LABEL USE OF THE DEVICE, NO CLINICAL DATA CAN SUPPORT THE COMPLAINT EVENTS(S) OR THE EFFECT(S) ON THE EFFICACY AND SAFETY OF THE DEVICE. THIS INFORMATION HAS BEEN COMMUNICATED TO THE MEDICAL ADVISOR OF STAAR SURGICAL AND TO THE SURGEON IN CASE OF SEVERE DETERIORATION OF PATIENT HEALTH. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). A HIGHER THAN IDEAL VAULTING IN THE ABSENCE OF SIGN AND/OR SYMPTOMS DOES NOT NEED AN EXCHANGE. HOWEVER, THE SURGEON MAY DECIDE TO EXCHANGE IF HE DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119166&pc=MTA | 2023826-2015-01299 | 2015-04-01 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND A SHALLOW CHAMBER. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND ONE OTHER COMPLAINT WAS FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5076256&pc=MTA | 2023826-2015-01128 | 2015-04-03 | Injury | STAAR SURGICAL COMPANY AG | 2015-09-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -10.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. SUBSEQUENTLY, THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4), BUT NOT MARKETED IN THE U.S. DIOPTER: -10.00. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4) NEW INCISION. SECONDARY SURGERY. LENS EXPLANTED. VAULTING. BASED ON THE COMPLAINT HISTORY, A SPECIFIC ROOT CAUSE OF THIS EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106976&pc=MTA | 2023826-2015-01186 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/X101 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON(B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01185 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -9.5/+2.0/X101. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5106991&pc=MTA | 2023826-2015-01185 | 2015-04-18 | Injury | STAAR SURGICAL COMPANY | 2015-09-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+4.0/X091 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ON (B)(6) 2015 EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED ON (B)(6) 2015 AND EXCHANGED FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED. UNCORRECTED VISUAL ACUITY WAS 20/20 ON (B)(6) 2015. SEE MFR #2023826-2015-01186 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -8.0/+4.0/X091. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119168&pc=MTA | 2023826-2015-01291 | 2015-04-20 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5405797&pc=MTA | 2023826-2016-00045 | 2015-04-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -11.5/-4.50/X090 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015 THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. SEE MFR. #2023826-2016-00046 FOR RIGHT EYE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-11.50/+4.50/X090. EVALUATION METHOD: WORK ORDER SEARCH. EVALUATION RESULT: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5178915&pc=MTA | 2023826-2015-01356 | 2015-05-06 | Injury | STAAR SURGICAL COMPANY | 2015-10-26 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER -13.00/+1.0/X133 DIOPTER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -13.00/+1.0/X133 (B)(4). METHOD CODES: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5174763&pc=MTA | 2023826-2015-01390 | 2015-05-16 | Injury | STAAR SURGICAL COMPANY | 2015-10-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4). NO LENS RETURNED, ONLY EMPTY BAG. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5124781&pc=MTA | 2023826-2015-01295 | 2015-05-21 | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO: LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146041&pc=MTA | 2023826-2015-01246 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS DID NOT RESOLVE THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/26. THE SECONDARY ICL WAS EXPLANTED ON (B)(6)2015. SEE MFR # 2023826-2015-01247 SECONDARY ICL. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DIOPTER: -13.50, DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S) : (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5146063&pc=MTA | 2023826-2015-01244 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY | 2015-10-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | PO30016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS DID NOT RESOLVE THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/26. THE SECONDARY ICL WAS EXPLANTED ON (B)(6) 2015. SEE MFR # 2023826-2015-01245 SECONDARY ICL. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THAT LENS DIMENSIONS WERE WITHIN SPECIFICATIONS. CORRECTED DATA: IT WAS INITIALLY REPORTED THAT THE PATIENT'S POST-OP VISIT WAS ON (B)(6) 2015, AND THIS IS INCORRECT. THE PATIENT'S POST-OP VISIT WAS ON (B)(6) 2015. (B)(4); NO OTHER ADVERSE EVENTS REPORTED OUTSIDE OF LENS EXCHANGE AND ELEVATED IOP. (B)(4); CODE NOT APPLICABLE; SHOULD BE IN PATIENT CODE, NOT DEVICE CODE. (B)(4). Manufacturer Narrative: LENS NOT RETURNED. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4919323&pc=MTA | 2023826-2015-00835 | 2015-05-26 | Injury | STAAR SURGICAL COMPANY AG | 2015-07-16 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015 AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS OBSERVED. THE PATIENT'S LEFT PUPIL IS LARGER THAN THE RIGHT. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION CODES: METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4900034&pc=MTA | 2023826-2015-00770 | 2015-05-28 | Injury | STAAR SURGICAL COMPANY | 2015-07-08 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. ANGLE CLOSURE AND ELEVATED IOP ASSOCIATED WITH EXCESSIVE VAULT WAS OBSERVED ON (B)(6) 2015. ADDITIONALLY, SIGNIFICANT REDUCTION OF IRIDO-CORNEALS ANGLES WAS OBSERVED. IT WAS INDICATED THAT THE EYE IS QUIET AND STABLE AND PATIENT IS ON ANTI-GLAUCOMA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). LENS IMPLANTED. EVALUATION CODES: METHOD (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS: (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5119300&pc=MTA | 2023826-2015-01300 | 2015-06-12 | Injury | STAAR SURGICAL COMPANY | 2015-10-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5219006&pc=MTA | 2023826-2015-01410 | 2015-06-25 | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/X094 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL WAS EXPLANTED ON (B)(6) 2015. NO OTHER LENS WAS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONALS INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -11.50/+0.50/X094. DEVICE EVALUATED BY MANUFACTURER? NO- LENS NOT RETURNED. (B)(4). METHOD CODES(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODES(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): NO CONCLUSION CAN BE DRAWN: BASED ON THE COMPLAINT HISTORY AND LENS WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208843&pc=MTA | 2023826-2015-01429 | 2015-07-15 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.00/+1.50/X089 DIOPTER IN PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01428 FOR LEFT EYE. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S.,BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.00/1.50/X089. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206816&pc=MTA | 2023826-2015-01475 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/X082 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5206819&pc=MTA | 2023826-2015-01476 | 2015-07-16 | Injury | STAAR SURGICAL COMPANY AG | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM13.7 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/X075 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC BROKEN AND THERE WAS FOREIGN MATERIAL ON LENS SURFACE SPECIFIED AS FIBERS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(4) AND IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5208884&pc=MTA | 2023826-2015-01428 | 2015-07-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS -16.50/+2.5/X084 DIOPTER IN PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE ICL EXPLANTED AND EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. SEE MFR. #2023826-2015-01429 FOR RIGHT EYE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -16.5/+2.5/X084. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5151143&pc=MTA | 2023826-2015-01249 | 2015-08-03 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00 DIOPTER IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING ASSOCIATED WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. POST-OP VISIT ON (B)(6) 2015 - UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -07.00. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5154950&pc=MTA | 2023826-2015-01354 | 2015-08-08 | Injury | STAAR SURGICAL COMPANY | 2015-10-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287706&pc=MTA | 2023826-2015-01627 | 2015-08-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. METHOD - (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS - (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287692&pc=MTA | 2023826-2015-01638 | 2015-08-22 | Injury | STAAR SURGICAL COMPANY AG | 2015-12-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287680&pc=MTA | 2023826-2015-01622 | 2015-08-25 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 11.5MM ICM115V4 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE NOTED AND THE LENS WAS EXPLANTED ON (B)(6) 2015. PATIENT'S LAST KNOWN BCVA WAS 20/20 AS OF (B)(6) 2015. Manufacturer Narrative: (B)(4). CONCLUSION - REVIEW OF THIS FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. THIS LENS MODEL IS INDICATED FOR USE IN ADULTS 21 - 45 YEARS OF AGE. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN A LENS CONTAINER AND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATED BY THE MANUFACTURER? LENS WAS NOT RETURNED. (B)(4). METHOD CODE (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS CODE (PROCESS RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5409020&pc=MTA | 2023826-2016-00088 | 2015-08-29 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-10.00. PT WEIGHT.: UNK. EVENT PROBLEM AND EVALUATION CODES: (B)(4). METHOD CODE(S): (EVALUATION METHOD): WORK ORDER SEARCH. RESULTS CODE(S): (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5109247&pc=MTA | 2023826-2015-01265 | 2015-08-31 | Injury | STAAR SURGICAL COMPANY | 2015-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VICMO12.1 IMPLANTABLE COLLAMER LENS WITH A 15.50 DIOPER IN THE PATIENT'S RIGHT EYE (OD) ON(B)(6) 2015. ON (B)(6) 2015, THE PATIENT SHOWED INADEQUATE VAULTING, ELEVATED IOP AND PIGMENT DISPERSION. SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED WHEN FURTHER INFORMATION IS AVAILABLE. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2015. LENS WAS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -15.50 DEVICE EVALUATED BY MANUFACTURER? NO. LENS REMAINS IMPLANTED. METHOD: EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION: EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536900&pc=MTA | 2023826-2016-00314 | 2015-09-04 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -12.0/+3.5/91 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.0/+3.5/91. (B)(4). METHOD CODE (EVALUATION METHOD): WORK ORDER SEARCH RESULTS CODE (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5324661&pc=MTA | 2023826-2015-01742 | 2015-09-10 | Injury | STAAR SURGICAL COMPANY | 2015-12-23 | MTA | VISIAN ICL (IMPANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ELEVATED IOP WAS NOTED AND THE LENS WAS REMOVED ON (B)(6) 2015. Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). METHOD: (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS: (EVALUATION RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5287670&pc=MTA | 2023826-2015-01623 | 2015-10-14 | Injury | STAAR SURGICAL COMPANY | 2015-12-10 | MTA | VISIAN ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED AN 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE NOTED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY THE MANUFACTURER? LENS WAS NOT RETURNED. (B)(4). METHOD CODE (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS CODE (PROCESS RESULT): A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION CODE (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5248750&pc=MTA | 2023826-2015-01577 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2015-11-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, IRIS ISCHEMIA, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. THE PATIENT HAS PERSISTENT ANISOCORIA/ IRIS PARESIS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20 AND THE PATIENT IS VERY HAPPY WITH THE IMPROVED VISION. Manufacturer Narrative: METHOD: DEVICE HISTORY RECORD REVIEW, MEDICAL REVIEW. RESULTS: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. IN ADDITION, THE REPORTER STATED, THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. SINCE THE EXCHANGE WITH THE SMALLER SIZE LENS RESOLVED THE ISSUE, HENCE, IT IS VERY LIKELY THE CAUSE OF THE EVENT IS DUE TO LENS SIZING. PER MEDICAL REVIEW - REPORTEDLY, ICL (13.7MM) WAS EXPLANTED/EXCHANGED FOR ANOTHER SHORTER ICL (13.2 MM) EIGHT DAYS POSTOPERATIVELY TO ADDRESS STEEP VAULT AND SUBSEQUENT ELEVATED IOP THAT WAS FIRST NOTED 12 HOURS AFTER IMPLANTATION. IRIDOTOMIES WERE PATENT. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY, SIGNED BY SURGEON ON 11/23/15, THE MOST LIKELY CAUSE OF THE REPORTED EVENTS (VAULTING AND IOP SPIKE) WAS SURGERY RELATED SIZING ISSUE. THE LENS WAS NOT MISLABELED AND THERE WAS NO OTHER LENS PROBLEM. HOWEVER, THE POSTOPERATIVE COMPLICATIONS (ANISOCORIA AND IRIS PARESIS) WERE ATTRIBUTED TO THE DEVICE PER SURGEON`S NOTE. HE ALSO REPORTED THAT IRIS PARESIS WAS IMPROVING. HE ALSO STATED THAT DESPITE ADVERSE EVENT, AFTER LENS EXCHANGE, THE PATIENT WAS VERY HAPPY WITH IMPROVED VISION AND SCHEDULED SURGERY FOR OTHER EYE. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE MEDICAL REVIEW, THE MOST LIKELY CAUSE OF THE REPORTED EVENTS (VAULTING AND IOP SPIKE) WAS SURGERY RELATED SIZING ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5611474&pc=MTA | 2023826-2016-00525 | 2015-10-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, PUPIL DISTORTION AND PERIPHERAL ANTERIOR SYNECHIAE. A SECONDARY YAG WAS PERFORMED TO ENLARGE/ADDITION OF PIS. THE PATIENT IS TAKING MEDICATION FOR THE PRESSURE. Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS: THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND PEN MARKINGS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT PROBLEM AND EVALUATION CODES: PATIENT CODE: (B)(4) INCISION OPENED. SECONDARY SURGERY. ENLARGEMENT/ADDITION IRIDECTOMY. PERIPHERAL ANTERIOR SYNECHIAE. PUPIL DISTORTION. EXPLANTED. VAULTING. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536860&pc=MTA | 2023826-2016-00321 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+3.00/X80 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S LEFT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2015. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -09.5/+3.00/X80. (B)(4). METHOD CODE(S): EVALUATION METHOD: LENS WORK ORDER SEARCH. RESULTS CODE(S): EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. CONCLUSIONS CODE(S): CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5613773&pc=MTA | 2023826-2016-00526 | 2015-10-23 | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE, WITH ELEVATED INTRAOCULAR PRESSURE, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PERIPHERAL ANTERIOR SYNECHIAE. A SECONDARY YAG WAS PERFORMED TO ENLARGE/ADDITION OF PIS. THE PATIENT IS TAKING MEDICATION FOR THE PRESSURE. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS WITH PEN MARKINGS AND FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5643957&pc=MTA | 2023826-2016-00587 | 2015-10-25 | Injury | STAAR SURGICAL COMPANY | 2016-05-10 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -8.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE REPORTED, THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE CASE/VIAL; SURGICAL RESIDUE WAS CLEARED; VISUAL INSPECTION FOUND HAPTIC TORN AND BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5408378&pc=MTA | 2023826-2016-00086 | 2015-11-03 | Injury | STAAR SURGICAL COMPANY | 2016-02-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULTING WITH ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2015 FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER:-12.50 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308661&pc=MTA | 2023826-2015-01659 | 2015-11-09 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/X100 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. Manufacturer Narrative: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS, PATIENT AGE BELOW 21 OR OVER 45 YEARS AND ANTERIOR ENDOTHELIUM CHAMBER DEPTH BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS WAS WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO -LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5439668&pc=MTA | 2023826-2016-00153 | 2015-11-10 | Injury | STAAR SURGICAL COMPANY | 2016-02-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, - 09.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND AN UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT HAD ELEVATED INTRAOCULAR PRESSURE (30MMHG) POST-OP. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (PROCESS EVALUATION): WORK ORDER SEARCH. (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION - BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5399905&pc=MTA | 2023826-2016-00080 | 2015-11-22 | Injury | STAAR SURGICAL COMPANY | 2016-01-29 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5432675&pc=MTA | 2023826-2016-00149 | 2015-11-25 | Injury | STAAR SURGICAL COMPANY | 2016-02-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5506304&pc=MTA | 2023826-2016-00267 | 2015-12-07 | Injury | STAAR SURGICAL COMPANY | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. A SECONDARY YAG WAS PERFORMED ON THE PERIPHERAL IRIDECTOMIES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. THERE WAS LIQUID ON THE LENS SURFACE. CORRECTED DATA: (B)(4). Manufacturer Narrative: (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579290&pc=MTA | 2023826-2016-00474 | 2015-12-21 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND THERE WAS A RISE IN INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.0. PT WEIGHT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5602214&pc=MTA | 2023826-2016-00513 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY | 2016-04-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5505923&pc=MTA | 2023826-2016-00276 | 2015-12-22 | Injury | STAAR SURGICAL COMPANY AG | 2016-03-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICH13.2 IMPLANTABLE COLLAMER LENS, +08.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS CONCLUSION CODE: AS PER DFU OF THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN INDIVIDUAL BELOW THE AGE OF 21 YEARS IS CONTRAINDICATED. THIS PATIENT'S AGE IS (B)(6) YEARS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN (B)(6) AND IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536876&pc=MTA | 2023826-2016-00313 | 2015-12-28 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.5MM ICM125V4 -9.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2015 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DIOPTER: -9.50. (B)(4). RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM #(B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575699&pc=MTA | 2023826-2016-00461 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP ((B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN MATERIAL ON LENS SURFACE (UNKNOWN). (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579567&pc=MTA | 2023826-2016-00475 | 2015-12-30 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND THERE WAS A RISE IN INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2015 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED BENT AND DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CORRECTED DATA: (B)(4). Manufacturer Narrative: DIOPTER: -5.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5535797&pc=MTA | 2023826-2016-00320 | 2015-12-31 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/+2.50/X091 DIOPTER IMPLANTABLE COLLAMER LENS IN PATIENT'S RIGHT EYE ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED ON (B)(6) 2015. THE LENS WAS EXPLANTED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2016. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT, UCVA WAS 20/20. Manufacturer Narrative: THIS PRODUCT WAS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). EVALUATION METHOD: LENS WORK ORDER SEARCH. EVALUATION RESULTS: A LENS WORK ORDER SEARCH REVEALED NO SIMILAR COMPLAINT WITHIN THE WORK ORDER. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5628400&pc=MTA | 2023826-2016-00552 | 2016-01-01 | Injury | STAAR SURGICAL COMPANY | 2016-05-03 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0 DIOPTER IMPLANTABLE COLLAMER LENS IN TO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP WERE REPORTED. THE LENS WAS EXPLANTED ON (B)(6) 2016. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THAT THE LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: COMPLETE DIOPTER VALUE IS: -14.0/+3.0/X092. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5536839&pc=MTA | 2023826-2016-00310 | 2016-01-08 | Injury | STAAR SURGICAL COMPANY | 2016-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 -12.50 DIOPTER IMPLANTABLE COLLAMER LENS ON (B)(6) 2016 IN TO THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THAT THE VAULT OF THE LENS WAS EXCESSIVE AND IOP WAS ELEVATED. THE LENS WAS EXCHANGED FOR A SMALLER LENS ON (B)(6) 2016 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DIOPTER: -12.50 (B)(4) SECONDARY SURGERY. NEW INCISION. EXPLANTED. VAULTING. METHOD: WORK ORDER SEARCH. RESULTS: A WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5577811&pc=MTA | 2023826-2016-00439 | 2016-01-16 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -07.00 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. EXCESSIVE VAULT AND ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/16. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DIOPTER: -07.00 (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5826262&pc=MTA | 2023826-2016-00252 | 2016-01-18 | Injury | STAAR SURGICAL COMPANY | 2016-07-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, PIGMENT DISPERSION AND AN UNREACTIVE PUPIL (FIXED). THE LENS WAS REPOSITIONED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/6/9. Manufacturer Narrative: THE LENS WAS EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. CLEAR SURGICAL RESIDUE/DEBRIS WAS PRESENT ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THE LENS HAD PEN MARKINGS ON IT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525073&pc=MTA | 2023826-2016-00328 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY AND BENT, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5525093&pc=MTA | 2023826-2016-00329 | 2016-01-19 | Injury | STAAR SURGICAL COMPANY | 2016-03-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.5/095 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CONCLUSION: PATIENT ACD BELOW 3.0 MM. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575700&pc=MTA | 2023826-2016-00463 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/29. Manufacturer Narrative: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) (NEW INCISION) IS INCORRECT. NO OTHER ADVERSE EVENTS WERE REPORTED OUTSIDE OF THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5575705&pc=MTA | 2023826-2016-00464 | 2016-01-20 | Injury | STAAR SURGICAL COMPANY | 2016-04-13 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/29. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668556&pc=MTA | 2023826-2016-00734 | 2016-02-10 | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL LASER PERIPHERAL IRIDOTOMIES WERE PERFORMED. THE REPORTER INDICATED THEY FELT THERE WAS POSSIBLE AQUEOUS MISDIRECTION. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: RESULT: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID AND THERE WAS EVIDENCE OF CLEAR SURGICAL RESIDUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891295&pc=MTA | 2023826-2016-01131 | 2016-02-25 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/083 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). CONCLUSION: REVIEW OF THE FILE INDICATES THAT THE LENS WAS NOT IMPLANTED IN ACCORDANCE WITH THE DFU REQUIREMENTS. PATIENT AGE BELOW 21. ANTERIOR ENDOTHELIUM CHAMBER DEPTH IS BELOW 3.0MM FOR VICL/VTICL. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN THE LENS CONTAINER. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/BENT/DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5632888&pc=MTA | 2023826-2016-00535 | 2016-03-06 | Injury | STAAR SURGICAL COMPANY | 2016-05-04 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/097 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, SHALLOW ANTERIOR CHAMBER AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. AT ONE WEEK POST-OP THE LENS HAD A HIGH VAULT WITH CONTROLLED INTRAOCULAR PRESSURE ELEVATION. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670128&pc=MTA | 2023826-2016-00643 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/081 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: (B)(4). DEVICE EVALUATIONS LENS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAVING FOREIGN MATERIAL ON LENS SURFACE (FIBERS). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5670132&pc=MTA | 2023826-2016-00642 | 2016-03-09 | Injury | STAAR SURGICAL COMPANY | 2016-05-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/091 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE (WITH ELEVATED IOP). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND FOREIGN MATERIAL (FIBERS) ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THAT LENS MEASUREMENTS WERE WITHIN SPECIFICATIONS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). THIS PATIENT WAS (B)(6) AT THE TIME OF IMPLANTATION. CORRECTED DATA: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT IS NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5724417&pc=MTA | 2023826-2016-00812 | 2016-03-28 | Injury | STAAR SURGICAL COMPANY | 2016-06-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT, ELEVATED IOP, NARROWING OF THE ANGLE, AND SHALLOWING OF THE ACD. THE LENS WAS EXCHANGED FOR A 12.6MM ICL AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). RESULTS: (OPERATIONAL PROBLEM) - VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THERE WAS EVIDENCE OF DARK COLOR RESIDUE ON THE LENS HAPTIC. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796321&pc=MTA | 2023826-2016-00906 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND SCRATCHED. (B)(4). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5796325&pc=MTA | 2023826-2016-00907 | 2016-04-28 | Injury | STAAR SURGICAL COMPANY | 2016-07-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: (B)(4). PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5955105&pc=MTA | 2023826-2016-01258 | 2016-05-30 | Injury | STAAR SURGICAL COMPANY | 2016-09-15 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/113 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT EVENT WITHIN ASSOCIATED LOTS WAS FOUND. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5810832&pc=MTA | 2023826-2016-00942 | 2016-06-20 | Injury | STAAR SURGICAL COMPANY | 2016-07-20 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.0/+1.0./X180 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. Manufacturer Narrative: CORRECTED INFORMATION: REPORT SOURCE: "USER FACILITY" TO " DISTRIBUTOR". (B)(4) . ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED ON (B)(6) 2016 WITH A SHORTER LENS AND THIS RESOLVED THE PROBLEM. ON (B)(6) 2016, THE PATIENT'S POST-OPERATION BEST-CORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL BUT CLEAR SURGICAL RESIDUE/DEBRIS FOUND ON PRODUCT. VISUAL INSPECTION FOUND THE PRESENCE OF FOREIGN MATERIAL ON THE LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817535&pc=MTA | 2023826-2016-00930 | 2016-06-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: AS PER THE DFU OF THIS LENS MODEL, IMPLANTATION OF LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD), AS MEASURED FROM THE CORNEAL ENDOTHELIUM TO THE ANTERIOR LENS CAPSULE, LESS THAN 3.0 MM IS CONTRADICTED. THIS PATIENT HAD AN ACD OF 2.915 MM AT THE TIME OF IMPLANTATION. CORRECTED INFORMATION: (B)(4) NOT APPLICABLE, SHOULD BE UNDER PATIENT CODE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5817536&pc=MTA | 2023826-2016-00932 | 2016-06-21 | Injury | STAAR SURGICAL COMPANY | 2016-07-22 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CONCLUSION: OFF-LABEL USE: PATIENT ACD BELOW 3.0 MM AND HAD A PRE-EXISTING CONDITION (PREVIOUS CORNEAL OR REFRACTIVE SURGERY). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983960&pc=MTA | 2023826-2016-01326 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPLANTED DATE - (B)(6) 2016. THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 AND VISION IS BETTER. (B)(4). RESULT CODE(S): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THE LENS WAS RETURNED IN LIQUID. CONCLUSIONS CODE(S): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION CODES: (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5983970&pc=MTA | 2023826-2016-01327 | 2016-07-26 | Injury | STAAR SURGICAL COMPANY | 2016-09-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO ANGLE CLOSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20 BUT IS IMPROVING. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5957285&pc=MTA | 2023826-2016-01286 | 2016-08-13 | Injury | STAAR SURGICAL COMPANY | 2016-09-16 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (AT 1 DAY POST-OP) AND ANGLE CLOSURE. THE SURGEON HAD DIFFICULTY "TUCKING" THE FEET OF THE ICL DURING THE INITIAL SURGERY. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/400. Manufacturer Narrative: (B)(4). BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS CUT INTO TWO PIECES, WITH A PIECE OF HAPTIC MISSING. THE LENS WAS RETURNED IN LIQUID. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356345&pc=MTA | 2023826-2019-00279 | 2019-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -11.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO MECHANICAL FAILURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO SIZING. THE LENS EXCHANGE RESOLVED THE PROBLEM. DATE OF EVENT: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT- "(B)(6) 2018" IN THE ORIGINAL MDR SHOULD BE CORRECTED TO "(B)(6) 2018" IMPLANTED DATE SHOULD BE CORRECTED TO (B)(6) 2018 IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY, WRAPPED IN GAUZE IN A ZIPLOCK BAG. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT. THERE WAS DEBRIS ON THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8356327&pc=MTA | 2023826-2019-00280 | 2019-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -13.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO MECHANICAL FAILURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THERE WAS NO PATIENT INJURY. THE LENS WAS DISCARDED BY THE FACILITY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE CAUSE OF THE EVENT WAS DUE TO SIZING. THE LENS EXCHANGE RESOLVED THE PROBLEM. CORRECTED DATA: DATE OF EVENT: (B)(6) 2019. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4822523&pc=MTA | 2023826-2015-00644 |  | Injury | STAAR SURGICAL COMPANY | 2015-06-05 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS THE SAME DAY OF EXPLANT.Event Description: ADDITIONAL INFORMATION RECEIVED - THE PATIENT EXPERIENCED NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH BEFORE EXPLANT OF THE LENS. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/15. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE ROOT CAUSE OF THE COMPLAINT. THE LENS SIZING ERROR MOST LIKELY CAUSED THE EXCESSIVE LENS VAULTING. (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. EVENT DATE: UNK. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. THE INJECTOR, CARTRIDGE AND FOAM TIP PLUNGER WERE NOT RETURNED. CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND EVALUATION OF THE RETURNED PRODUCT, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: MEDICAL REVIEW - IN THIS CASE THE IOP ELEVATION, ANGLE NARROWING AND SHALLOW ANTERIOR CHAMBER WERE THE CONSEQUENCES OF EXCESSIVE VAULTING OF THE LENS. THE TIMELY MANAGEMENT AND LENS EXCHANGE WERE NECESSARY TO PREVENT ANY FURTHER DAMAGE CAUSED BY THE EVENT. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING WAS A CONSEQUENCE OF WRONG LENS USE WHICH INCLUDE THE FOLLOWING: IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENT BASED UPON DIFFERENT TECHNIQUES, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN A INDIVIDUAL CASE, IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST AND IMPROPER LENS LABEL. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5047723&pc=MTA | 2023826-2015-01113 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/106 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6)2015 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, ELEVATED INTRAOCULAR PRESSURE, IRIS ATROPHY AND UNREACTIVE PUPIL (FIXED). THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5089547&pc=MTA | 2023826-2015-01217 |  | Injury | STAAR SURGICAL COMPANY | 2015-09-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -7.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED THE CAUSE OF THE EVENT WAS DUE TO VARIED WHITE-TO-WHITE MEASUREMENTS. THERE WERE NO ADDITIONAL SYMPTOMS RELATED TO THE EVENT. Manufacturer Narrative: A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT WAS DETERMINED. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND THE DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - REPORTEDLY, MICL(13.2 MM) WAS EXPLANTED/EXCHANGED FOR A SHORTER MICL NINE DAYS POSTOPERATIVELY TO ADDRESS EXCESSIVE VAULT AND SUBSEQUENT ELEVATED IOP. ACCORDING TO THE COMPLETED CUSTOMER COMPLAINT QUESTIONNAIRE FOR SURGERY THE EVENT WAS NOT PRODUCT RELATED BUT IT WAS CAUSED BY PROCEDURE RELATED FACTOR (THE SURGEON`S CHOICE OF LONGER LENS BASED ON VARIED WHITE- TO- WHITE MEASUREMENTS). UPON LENS EXCHANGE ALL SYMPTOMS RESOLVED AND UCVA WAS 20/20 CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, DEVICE HISTORY RECORD REVIEW AND THE MEDICAL REVIEW, THE CAUSE OF THE EVENT WAS DUE TO A PROCEDURE RELATED FACTOR (THE SURGEON`S CHOICE OF LONGER LENS BASED ON VARIED WHITE- TO- WHITE MEASUREMENTS). UPON LENS EXCHANGE ALL SYMPTOMS RESOLVED AND UCVA WAS 20/20. THE EVENT WAS NOT PRODUCT RELATED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5125138&pc=MTA | 2023826-2015-01321 |  | Injury | STAAR SURGICAL COMPANY | 2015-10-02 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED DUE TO THE LENS WAS TOO LARGE AND ELEVATED INTRAOCULAR PRESSURE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE FACILITY REPORTED THE LENS WAS EXPLANTED ON (B)(6) 2015, DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. ADDITIONAL PERIPHERAL IRIDECTOMIES WERE PERFORMED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. Manufacturer Narrative: BASED ON THE RESULTS OF THE DHR REVIEW, THE AVAILABLE INFORMATION GIVEN WITHIN THIS COMPLAINT, WORK ORDER SEARCH, AND THE PRODUCT EVALUATION, A CONCLUSION CAN BE DRAWN THAT NOTHING IN THE MANUFACTURING, STERILIZATION AND PACKAGING PROCESSES OF THE LENS IS LIKELY TO HAVE CONTRIBUTED TO THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW THE CAUSE OF THE EVENT IS UNKNOWN. NO DEFINITE ROOT CAUSE FOR THE COMPLAINT IS DETERMINED. PER MEDICAL REVIEW: REPORTEDLY, MICL (12.6 MM) WAS EXPLANTED/EXCHANGED, 24 DAYS POSTOPERATIVELY, FOR A SHORTER LENGTH ICL (12.1MM) TO ADDRESS EXCESSIVE VAULT AND SUBSEQUENT ANGLE CLOSURE AND ELEVATED IOP. SECONDARY SURGICALLY INTERVENTION (ADDITIONAL PERIPHERAL IRIDOTOMY) WAS PERFORMED PRIOR TO LENS EXCHANGE. NO REPORTED POSTOPERATIVE SEQUELAE. ACCORDING TO THE SAME REPORT, CAUSE OF THE EVENT WAS UNKNOWN. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW, AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN. THE LENS WAS RETURNED IN LIQUID. BASED ON THE COMPLAINT HISTORY, EVALUATION OF THE RETURNED PRODUCT AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER - (B)(6) (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5218062&pc=MTA | 2023826-2015-01498 |  | Injury | STAAR SURGICAL COMPANY | 2015-11-11 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/079 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, ANGLE CLOSURE WITH ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THE LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. EVENT DATE: UNK. METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511748&pc=MTA | 2023826-2016-00283 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -09.5 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITH NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT IS DOING GREATEvent Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT IS DOING GREAT Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4). METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION: BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY HAVE CREATED OR NOT DETECTED A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED EXCESSIVE VAULT OF THE LENS. (B)(4). Manufacturer Narrative: METHOD: PROCESS EVALUATION: DEVICE HISTORY RECORD REVIEW. RESULT: (OPERATIONAL PROBLEM): BASED ON THE DHR REVIEW AND ROOT CAUSE ANALYSIS, THE PROBABLE CAUSE COULD NOT BE DETERMINED AT THIS TIME. THERE WERE NO DOCUMENTED ISSUES AND CONCERNS WITH THE MANUFACTURING AND INSPECTION PROCESSES WHICH MAY HAVE CREATED OR NOT DETECTED A NONCONFORMANCE OF THE LENS. PHYSICIAN REPORT DOES NOT INDICATE THAT ANY EXCEPTIONAL EVENT OR CONDITION WOULD HAVE CAUSED EXCESSIVE VAULT OF THE LENS. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5511760&pc=MTA | 2023826-2016-00284 |  | Injury | STAAR SURGICAL COMPANY | 2016-03-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER, IN THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION WAS RECEIVED FROM THE REPORTER, IT WAS INDICATED THAT THE 12.6MM MICL IMPLANTABLE COLLAMER LENS, -11.0 DIOPTER WAS IMPLANTED ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO THE PATIENT EXPERIENCED ELEVATED IOP, NARROWING OF THE ANGLE, ANGLE CLOSURE, AND SHALLOWING OF THE ACD WITH PRE-OP IOP 20MM HG. AN IRIDECTOMY WAS PERFORMED AND THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4). METHOD (PROCESS EVALUATION): WORK ORDER SEARCH. RESULTS (EVALUATION RESULT): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSION (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND BOTH HAPTICS HAD PIECES TORN OFF AND MISSING. THE LENS WAS RETURNED IN A SMALL CLEAR BOTTLE IN LIQUID. PER MEDICAL REVIEW IT SHOULD BE NOTED THAT AT THE TIME OF IMPLANTATION THE PATIENT WAS (B)(6) AND PER DFU, INDICATIONS: "THE VISIAN ICL IS INDICATED FOR USE IN ADULTS 21-45 YEARS OF AGE" AND THEREFORE, THERE IS NO SUFFICIENT SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION IN SUCH PATIENTS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5668553&pc=MTA | 2023826-2016-00731 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -16.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO OVER VAULT AND HIGH PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION RECEIVED - THE REPORTER INDICATED THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT COMPLAINED OF A HEADACHE. THE PUPIL CAME DOWN IN THE RIGHT EYE AND THE SURGEON WAS UNABLE TO REPLACE THE LENS DURING LENS REPLACEMENT SURGERY ON (B)(6) 2016. THE REPLACEMENT LENS SURGERY OCCURRED AT A LATER DATE. RESULTS: (OPERATIONAL PROBLEM): VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE TORN OFF AND MISSING FROM ONE HAPTIC. THE LENS WAS RETURNED DRY AND THERE WAS EVIDENCE OF CLEAR/REDDISH RESIDUE. CONCLUSIONS: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(6).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5680337&pc=MTA | 2023826-2016-00757 |  | Injury | STAAR SURGICAL COMPANY | 2016-05-25 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO INCORRECT SIZE. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE AND SHALLOW ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE IRIS SPHINCTER WAS DAMAGED DUE TO THE PRESSURE. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO MISMEASUREMENT OF THE WHITE-TO-WHITE. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). RESULT: (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH AND PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER? NO - LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5902496&pc=MTA | 2023826-2016-01148 |  | Injury | STAAR SURGICAL COMPANY | 2016-08-24 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, DIOPTER UNKNOWN, IN THE PATIENT'S EYE ON (B)(6) 2016. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT, ELEVATED INTRAOCULAR PRESSURE AND DILATED PUPIL. THE LENS REMAINS IMPLANTED BUT MAY BE EXPLANTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATED BY MANUFACTURER? NO - LENS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6064630&pc=MTA | 2023826-2016-01552 | 2016-10-04 | Injury | STAAR SURGICAL COMPANY | 2016-10-28 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/040 DIOPTER, IN THE PATIENT'S RIGHT EYE ON (B)(6) 2016. ON (B)(6) 2016, THE PATIENT EXPERIENCED EXCESSIVE VAULTING OF THE LENS, ELEVATED INTRAOCULAR PRESSURE (IOP), AND LENS DISLOCATION/SUBLUXATION. THE PATIENT HAD TO UNDERGO AN ENLARGEMENT OF PERIPHERAL IRIDOTOMY (PI) OR ADDITION OF NEW PI USING YAG. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8062409&pc=MTA | 2023826-2018-01634 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE LENS HAD ROTATED AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE (IOP). A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES. THE LENS REMAINS IMPLANTED. THE PUPIL IS DILATED/FIXED AND DOESN'T RESPOND TO TREATMENT. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: EVENT - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE REPORTER INDICATED THE LENS WAS REPOSITIONED AND THE PATIENT EXPERIENCED ELEVATED IOP. A SECONDARY YAG WAS PERFORMED AND THE PUPIL IS DILATED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO AN OVERSIZED ICL. THE LENS REMAINS IMPLANTED WITH AN EXCHANGE PLANNED. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONING DEVICE CODE 2923: PREVIOUS CODE NOT APPLICABLE. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8062402&pc=MTA | 2023826-2018-01633 |  | Injury | STAAR SURGICAL COMPANY | 2018-11-12 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A12.6MM ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS RIGHT EYE (OD). THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE (IOP). THE LENS WAS TORN BEFORE/DURING LOADING. A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMIES. THE LENS ROTATED AND WAS REPOSITIONED VERTICALLY BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE PUPIL IS DILATED AND FIXED AND DOESN'T RESPOND TO TREATMENT. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO OVERSIZED ICL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: EVENT - "THE LENS ROTATED AND WAS REPOSITIONED..." SHOULD BE CORRECTED TO "THE LENS WAS REPOSITIONED..." DEVICE CODE (B)(4). DEVICE CODE(B)(4). CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5891293&pc=MTA | 2023826-2016-00454 | 2016-02-18 | Injury | STAAR SURGICAL COMPANY | 2016-08-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULTING, LENS ROTATION AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THE PATIENT'S POST-OP (ON (B)(6) 2016) UNCORRECTED VISUAL ACUITY WAS 20/20. DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. ELEVATED IOP NOT REPORTED FOR PATIENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5579677&pc=MTA | 2023826-2016-00444 | 2016-02-04 | Injury | STAAR SURGICAL COMPANY | 2016-04-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Operates Differently Than Expected; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2, -11.50 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ELEVATED INTRAOCULAR PRESSURE WAS NOTED. THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2016 FOR A SHORTER LENGTH LENS WITH A SIMILAR DIOPTER SIZE. THIS RESOLVED THE PROBLEM. PATIENT'S LAST VISIT ON (B)(6) 2016, UCVA WAS 20/25. Manufacturer Narrative: DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND THAT THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). AS PER DFU FOR THIS LENS MODEL, IMPLANTATION OF THIS LENS IN AN EYE WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM IS CONTRAINDICATED. THIS PATIENT HAS AN ACD OF 2.89MM. (B)(4). Manufacturer Narrative: DIOPTER: -11.50. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819152&pc=MTA | 2023826-2020-00452 | 2020-01-30 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT WITH ANGLE CLOSURE AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819155&pc=MTA | 2023826-2020-00453 | 2020-01-30 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT WITH ANGLE CLOSURE AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4579225&pc=MTA | 2023826-2015-00223 | 2014-03-17 | Injury | STAAR SURGICAL COMPANY | 2015-03-06 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Unintended Movement; Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM ICM120V4 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. THE LENS WAS EXPLANTED ON (B)(6) 2014 DUE TO EXCESSIVE VAULTING. THE LENS DISLOCATED/SUBLUXED AND THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/40. Manufacturer Narrative: (B)(4): EVALUATION: METHOD - WORK ORDER SEARCH. RESULTS - A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS - (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. LENS NOT RETURNED. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED DRY. THE LENS WAS REHYDRATED IN BSS FOR RE-MEASUREMENT. THE LENS LENGTH WAS MEASURED AND THE RESULT OF THE MEASUREMENT WAS COMPARED AGAINST THE ORIGINAL VALUE AND THE LENS WAS FOUND TO BE IN SPECIFICATION. MEDICAL REVIEW - ELEVATED IOP IN THE PRESENCE OF HIGH VAULT AFTER ICL IMPLANTATION MAY OCCUR DUE TO: NON PATENT/FUNCTIONING IRIDECTOMIES (TOO SMALL, TOO PERIPHERAL AND/OR NOT FULLY PERMEABLE, BLOCKED BY VISCOELASTIC), REMAINING VISCOELASTIC IN THE POSTERIOR CHAMBER, OVERSIZED ICL WITH ANGLE CLOSURE, TOO NARROW ANGLES/CROWDED AC DUE TO SMALL EYE ANATOMY PREOPERATIVELY, UNEXPECTED ABNORMAL ANATOMY OR TISSUE ABNORMALITIES (PRESENCE OF IRIS/CILIARY BODY CYSTS), MALPOSITIONED FOOTPLATES, ETC. ACCORDING TO USE FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). SEVERAL CONDITIONS MAY ARISE FROM THIS EVENT (I.E. PUPILLARY BLOCK, MALIGNANT GLAUCOMA, INCREASED IOP ETC.). TO PREVENT ANY OF THESE COMPLICATIONS FROM OCCURRING, STAAR RECOMMENDS THAT THE LENS BE EXPLANTED AND/OR EXCHANGED WITH THE RIGHT SIZE ONCE THE SURGEON DETERMINES THAT THIS CONDITION MAY AFFECT THE OUTCOME OF THE PATIENT'S VISION. BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, IT WAS FOUND THERE WAS NOT ENOUGH CLINICAL DATA TO DETERMINE THE EXACT CAUSE OF THIS EVENT; HOWEVER, WE CANNOT RULE OUT EXCESSIVE VAULTING SECONDARY TO A LONG LENS AS A CONTRIBUTING FACTOR. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6836423&pc=MTA | 2023826-2017-01377 | 2017-08-01 | Injury | STAAR SURGICAL COMPANY | 2017-08-31 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER. THE LENS WAS EXPLANTED A COUPLE OF HOURS LATER DUE TO ELEVATED INTRAOCULAR PRESSURE. ANOTHER IRIDECTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND REMOVED ON (B)(6) 2017 FROM THE PATIENT'S LEFT EYE (OS). THE INCISION WAS ENLARGED. THE LENS EXPLANT RESOLVED THE PROBLEM. AFTER THE LAST POST-OP VISIT, THE BCVA WAS 20/20 AND THERE WERE NO PATHOLOGICAL SYMPTOMS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A LIQUID INSIDE A CONTACT CASE WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS MISSING A PIECE OF ITS HAPTIC AND PRESENCE OF DEBRIS ON THE PRODUCT. "ADDITIONAL IRIDECTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED" IS CORRECTED TO " ADDITIONAL YAG IRIDOTOMY WAS PERFORMED BUT THE PRESSURE WAS STILL ELEVATED" IN THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633715&pc=MTA | 2023826-2020-00177 | 2020-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, MICL12.6, -10.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. "DOCTOR SAW THAT THE LENS WAS CRACKED." THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020. NO INJURY TO PATIENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIP LOCK BAG WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE DOCTOR NOTICED LOW VAULT, REFRACTIVE SURPRISE AND ELEVATED IOP. THE EXCHANGE RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. H6- MEDICAL DEVICE PROBLEM CODE 1401 - REFRACTIVE SURPRISE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: A2- (B)(6) 1987, SHOULD BE IN THE PREVIOUS MDR. B3- (B)(6) 2020, SHOULD BE IN THE PREVIOUS MDR. H6- CLINICAL CODE 1937 SHOULD BE IN PREVIOUS MDR. H6- DEVICE CODE 1583 SHOULD BE IN PREVIOUS MDR . CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285775&pc=MTA | 2023826-2020-01481 | 2020-06-17 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +2.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ANGLE CLOSURE WITH ELEVATED IOP. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THERE ARE NO PLANS TO IMPLANT AN ALTERNATE LENS. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846146&pc=MTA | 2023826-2019-01278 | 2019-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. REPORTEDLY PATIENT IS "UNDER DORZOLAMIDE AND TIMOLOL EYE DROPS, TWICE DAILY. VAULT IS 350." LENS REMAINS IMPLANTED. PER PATIENTS LAST VISIT, (B)(6) 2019, IOP 20 MMHG, "CONTINUING ANTI GLAUCOMA MEDICATION AND STOPPED STEROIDS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8846151&pc=MTA | 2023826-2019-01279 | 2019-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-31 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VICMO 12.6, -18.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ELEVATED IOP (28 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WERE ASSESSED ON (B)(6) 2019. REPORTEDLY PATIENT IS "UNDER DORZOLAMIDE AND TIMOLOL EYE DROPS, TWICE DAILY. VAULT IS 350." LENS REMAINS IMPLANTED. PER PATIENTS LAST VISIT, (B)(6) 2019, IOP 20 MMHG, "CONTINUING ANTI GLAUCOMA MEDICATION AND STOPPED STEROIDS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF LABEL USE - ACD<3.0MM. PATIENT CODE 3191: NO CODE AVAILABLE - MEDICAL INTERVENTION, USE OF DORZOLAMIDE AND TIMOLOL EYE DROPS TO TREAT ELEVATED IOP. ADVERSE EVENT OR PRODUCT PROBLEM: CORRECTED TO ADVERSE EVENT FROM PRODUCT PROBLEM IN INITIAL MDR. DEVICE CODE 2993 ADDED; DEVICE CODE 3189 NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9983363&pc=MTA | 2023826-2020-00936 | 2020-03-19 | Injury | STAAR SURGICAL COMPANY | 2020-04-21 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICH12.6 IMPLANTABLE COLLAMER LENS, +7.0 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO ELEVATED IOP. THE SURGEON ALSO PERFORMED AN ADDITION OF A NEW PI-YAG. THE PROBLEM WAS RESOLVED AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR-THE PI WASN'T TOTALLY OPEN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966497&pc=MTA | 2023826-2019-01640 | 2018-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -11.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXPLANTED, THE SURGEON REPORTING PUPIL BLOCK, ELEVATED IOP, PIGMENT DISPERSION, AND UNREACTIVE (FIXED) PUPIL. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: "DIOPTER -11.5" SHOULD BE CORRECTED TO "DIOPTER -12.0" IN INITIAL MEDWATCH REPORT. DEVICE CODE:1583 SHOULD BE REMOVED FROM INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. IT WAS REPORTED THAT THE CAUSE WAS DUE TO ACUTE UVEAL REACTION, THE EXUDATION OF ANTERIOR CHAMBER, ANTERIOR CHAMBER PIGMENT LEAD TO PUPILLARY CLOSURE, THE PATIENT WAS ALERGIC. REPORTER INDICATED THERE WAS NO VAULTING AND PROBLEM HAS RESOLVED AFTER THE LENS EXPLANT. CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8432797&pc=MTA | 2023826-2019-00442 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-19 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.7MM MICL13.7, -10.5 DIOPTER , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS EYE ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 "DUE TO MECHANICAL VAULTING, TOO STEEP." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: OTHER RELEVANT HISTORY: ADDED INFORMATION (DENSE NS = NUCLEAR SCLENOTIC CATARACT). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE ON LENS, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8471518&pc=MTA | 2023826-2019-00457 | 2019-03-01 | Injury | STAAR SURGICAL COMPANY | 2019-04-01 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6, -11.0 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP (65 MMHG). THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN BIOHAZARD ZIP LOC BAG, IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279656&pc=MTA | 2023826-2019-02067 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5 DIOPTER, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957673&pc=MTA | 2023826-2019-01536 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICMO13.2, -11.00 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS, ELEVATED IOP (29 MMHG) WITHOUT PUPIL BLOCK AND ANGLE CLOSURE, WITH EXCESSIVE VAULT WERE OBSERVED ON 31/MAY/2019. LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM VICMO 13.2 OF -11.00 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT; ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE WAS OBSERVED. LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PATIENT CODE 3191 FOR SIGNIFICANT REDUCTIO IN IRIDO-CORNEAL ANGLES IS NOT APPLICABLE. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895562&pc=MTA | 2023826-2019-01454 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VICM5_13.2, -8.50 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND PUPIL BLOCK WITH ELEVATED IOP. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DATE RECEIVED BY MANUFACTURER SHOULD HAVE BEEN REPORTED AS 16-JUL-2019 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DEVICE CODE 1494: OFF LABEL USE (AGE OF PATIENT>45 YEARS OLD). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445821&pc=MTA | 2023826-2020-01819 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8403793&pc=MTA | 2023826-2019-00364 |  | Injury | STAAR SURGICAL COMPANY | 2019-03-08 05:00 | MTA | UNK | Optical Problem | Intraocular Pressure Increased; No Code Available | UNK | 1 | Event Description: THE REPORTER INDICATED THAT A LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). PIGMENT DISPERSION, IRIS TRANSILLUMINATION DEFECTS AND HIGH INTRAOCULAR PRESSURE WAS OBSERVED. IT WAS REPORTED THAT IT WAS CONTROLLED AND STABILIZED WITH 1 IOP MEDICATION. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CHANGED FROM (B)(6) 2019 TO (B)(6) 2019 CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. BRAND NAME - UNK. COMMON DEVICE NAME - UNK. SERIAL NUMBER = UNK, EXPIRATION = UNK, UDI - UNK. IMPLANT DATE - UNK. PMA/501K - UNK. DEVICE MANUFACTURE DATE - UNK. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033107&pc=MTA | 2023826-2020-01027 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5 13.2, -10.50 DIOPTER, IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOSS OF BCVA (PRE-OP BCVA = 20/20 POST-OP BCVA = 20/17-18) AND PUPIL BLOCK WITH ELEVATED IOP WAS OBSERVED. LENS WAS REPOSITIONED AND ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF VISCO/ FLUIDS DID NOT RESOLVE THE PROBLEM. REPORTEDLY, PATIENTS LAST VISIT ON (B)(6) 2019 MIDRIASIS (DILATED PUPILS) WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B3 CORRECTED FROM (B)(6) 2019 TO (B)(6) 2019 IN INITIAL MDR. B5 - CORRECTED TO: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.50 D IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTED SAME DAY LOSS OF BCVA; PUPIL BLOCK WITH AN UNSPECIFIED ELEVATED IOP; UNREACTIVE (FIXED) PUPIL AND IRIS ATROPHY. ANTERIOR CHAMBER IRRIGATION/ EVACUATION OF VISCO FLUIDS AND A LENS REPOSITIONING WERE PERFORMED. IT WAS REPORTED THAT THE "ROTATION OF THE LENS" DID NOT RESOLVE THE PROBLEM AND CURRENTLY THE PATIENT HAS PUPIL BLOCK AND UNREACTIVE (FIXED) PUPIL. THE LENS WAS EXPLANTED ON (B)(6) 2019. POST-EXPLANT UCVA AND BCVA WERE REPORTED TO BE 20/18 WITH AND IOP OF 12MM HG AND "MIDRASIS" WAS NOTED. H6 - HEALTH EFFECTS CLINICAL CODE 4581: IRIS ATROPHY, UNREACTIVE PUPIL. H6 - HEALTH EFFECTS IMPACT CODE 4627. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476061&pc=MTA | 2023826-2020-01922 | 2020-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED AND EXCHANGED DUE TO HIGH VAULTING AND PUPILLARY BLOCK. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -15.00 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING, PUPIL BLOCK WITH ELEVATED INTRAOCULAR PRESSURE AND IRIS ATROPHY. ADDITIONAL SURGERY WAS PERFORMED, ENLARGEMENT OF PI BY YAG AND ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. THE LENS WAS REPLACED FOR A SHORTER LENS (DIFFERENT MODEL - TMICL) ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. SOME MILD IRIS ATROPHY BUT PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. H6: HEALTH IMPACT- CLINICAL CODE: 4581 - IRIS ATROPHY. H6: HEALTH EFFECT IMPACT CODE: 4625 - ENLARGEMENT OF PI - YAG; ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7648241&pc=MTA | 2023826-2018-00973 | 2018-04-26 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Positioning Failure; Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Consequences Or Impact To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THAT A 12.1MM, MICL12.1, -12.5 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN, THE LENS WAS REMOVED AND THE BACKUP LENS WAS USED. THE REPORTER STATED THAT THERE WAS NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED INFORMATION: DATE OF EVENT WAS CHANGED TO (B)(6) 2018 PREVIOUS REPORTED DATE WAS INCORRECT. (B)(4). THE LENS WAS NOT IMPLANTED UPSIDE DOWN AS INITIALLY REPORTED. THE LENS WAS IMPLANTED ON (B)(6)) 2018 AND EXCHANGED ON (B)(6) 2018 DUE TO OBSERVATION OF LOW VAULT, ELEVATED IOP (INTRAOCULAR PRESSURE) WITHOUT PUPIL BLOCK AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE REPORTER STATED THAT THE CAUSE OF EVENT WAS DUE TO PATIENT RELATED FACTOR. (B)(4). Manufacturer Narrative: -IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). (B)(4) Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS FOUND. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5308634&pc=MTA | 2023826-2015-01687 | 2015-10-22 | Injury | STAAR SURGICAL COMPANY | 2015-12-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; No Consequences Or Impact To Patient; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.50/+2.50/X076 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2015. EXCESSIVE VAULT AND ELEVATED IOP (WITHOUT PUPIL BLOCK AND ANGLE CLOSURE) WAS NOTED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS WITH A SIMILAR DIOPTER SIZE. Manufacturer Narrative: (B)(4). DEVICE EVALUATION: PRODUCT EVALUATION FOUND THAT THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE. DIMENSION AT INSPECTION FOUND THE LENS WAS WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN US, BUT IS NOT MARKETED IN THE U.S. DIOPTER: -10.5/+2.5/X076. DEVICE EVALUATED BY MANUFACTURER? NO, LENS NOT RETURNED. (B)(4). METHOD: LENS WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH REVEALED THERE WERE NO SIMILAR COMPLAINTS WITHIN THE WORK ORDER. CONCLUSION: CONCLUSION NOT YET AVAILABLE. EVALUATION IS IN PROGRESS. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7824382&pc=MTA | 2023826-2018-01253 | 2018-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2018-08-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/063 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD EXCESSIVE VAULTING. THE LENS REMAINS IMPLANTED. AT ONE WEEK REVIEW, THE PATIENTS INTRAOCULAR PRESSURE WAS 20MMHG AND PUPIL WAS REACTIVE EVEN THOUGH SEEMS SLOWER. THE PATIENT IS BEING MONITORED. Manufacturer Narrative: CORRECTED DATA: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6577254&pc=MTA | 2023826-2017-00849 | 2016-11-20 | Malfunction | STAAR SURGICAL COMPANY | 2017-05-18 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO USE A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS FOR THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE REPORTER STATED THE LENS WAS NOT IMPLANTED NEITHER EXPLANTED, AND PRODUCT RECEIVED DAMAGED. SHE ALSO MENTIONED THE LENS HAS SIZING ISSUE, AND THE PATIENT HAD ELEVATED IOP. MULTIPLE ATTEMPTS TO GET CLARIFY INFORMATION HAS BEEN UNSUCCESSFUL. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BEEN RETURNED. Manufacturer Narrative: THIS REPLACEMENT LENS WAS SUBMITTED IN THE INITIAL MDR, INCORRECTLY. THERE ARE NO ADVERSE EVENTS ASSOCIATED WITH THE REPLACEMENT LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNIT WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6447958&pc=MTA | 2023826-2017-00541 | 2017-01-12 | Malfunction | STAAR SURGICAL COMPANY | 2017-03-30 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Device Operates Differently Than Expected | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2017 THE LENS WAS TORN BEFORE INJECTION INTO THE EYE, AND IT WAS NOT IMPLANTED. ON (B)(6) 2017 THE LENS WAS EXCHANGE, AND THE PROBLEM WAS RESOLVED. AS OF THE DATE OF MDR SUBMISSION, THE LENS HAS NOT BE RETURNED YET. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE AND WAS NOT IMPLANTED. THE LENS WAS EXCHANGED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A SMALL VIAL. VISUAL INSPECTION FOUND THAT THE LENS WAS MISSING A PIECE OF THE HAPTIC. CORRECTED DATA: (B)(4). Manufacturer Narrative: CORRECTED DATA: THE REPORTER STATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2 MM VICMO13.2 IMPLANTABLE COLLAMER LENS, -12 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS TORE/BROKE BEFORE INJECTION INTO THE EYE AND WAS NOT IMPLANTED. THE PATIENT EXPERIENCED POST-OPERATIVE ELEVATED INTRAOCULAR PRESSURE. A DIFFERENT LENS OF THE SAME MODEL AND DIOPTER WAS IMPLANTED ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. RESULT: WORK ORDER SEARCH PERFORMED, NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS IN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7647516&pc=MTA | 2023826-2018-00983 | 2018-05-31 | Injury | STAAR SURGICAL COMPANY | 2018-06-28 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WILL BE EXPLANTED AND EXCHANGED DUE TO BEING TOO HIGH. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE SURGEON ENLARGED THE PERIPHERAL IRIDECTOMIES USING YAG. THE INCISION WAS ENLARGED AND SUTURE WAS REQUIRED. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS EXCHANGE RESOLVED THE PROBLEM. (B)(4) Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8344402&pc=MTA | 2023826-2019-00266 | 2018-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-15 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6, -6.5 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. A LENS VAULT VALUE OF 1150 UM WAS REPORTED AND THE LENS REMAINS IMPLANTED WITH PLANS TO REMOVE AND EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.5 DIOPTER INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND HIGH IOP. THE PATIENT CONTROLLED THE HIGH IOP WITH ANTI-IOP DROPS UNTIL THE LENS WAS EXPLANTED. THE SURGEON PLANS TO IMPLANT A SHORTER LENGTH LENS. DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE ON LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, LENS EXPLANT. PATIENT CODE 2692: CODE IN ORIGINAL MDR NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8998374&pc=MTA | 2023826-2019-01699 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | N/A | 1 | Event Description: ON 14 AUG 2019 WE RECEIVED NOTIFICATION OF AN ARTICLE PUBLISHED IN SOCIEDAD ESPANOLA DE OFTALMOLOGIA, ' INTRAOCULAR PRESSURE FLUCTUATIONS IN PATIENTS IMPLANTED WITH AN IMPLANTABLE COLLAMER LENS (ICL V4C). THREE-MONTH FOLLOW-UP.' THE STUDY FOUND 5 EYES WITH ABNORMAL HIGH VAULT (ABOVE 1MM). IT WAS REPORTED THAT AS OF MARCH 2019, ALL LENSES REMAIN IMPLANTED AND NO FURTHER PROBLEMS WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: INITIAL MDR SHOULD HAVE ALSO STATED: ' THESE 5 EYES AND AN ADDITIONAL TWO EYES OF THE SAME PATIENT WERE COURSED WITH OCULAR HYPERTENSION. THIS REQUIRED TREATMENT WITH TOPICAL HYPOTENSORS, AND THE PROBLEM RESOLVED IN APPROXIMATELY 10 WEEKS.' PATIENT CODE 2692 IS NOT APPLICABLE IN INITIAL MDR, PATIENT CODE 1937 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR, PATIENT CODE 3191 (MEDICAL INTERVENTION WITH TOPICAL HYPOSTENSORS) SHOULD HAVE BEEN INCLUDED IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8172146&pc=MTA | 2023826-2018-01871 | 2018-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-12-17 05:00 | MTA | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; No Apparent Adverse Event | Intraocular Pressure Increased; No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: A VICMO12.6, 12.6MM IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 WAS RETURNED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, DIOPTER -14.0 INTO THE PATIENT'S RIGHT (OD) EYE. ON (B)(6) 2018 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, LENS DISLOCATION/SUBLUXATION, PUPIL BLOCK, AND ELEVATED IOP. THE SURGEON ALSO REPORTS ENLARGEMENT OF PI. REPORTEDLY, THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: TYPE OF REPORTABLE EVENT-CORRECTED. DEVICE EVALUATED BY MFR THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. METHOD CODE ADDED. RESULT CODE UPDATED. CONCLUSION CODE UPDATED. CLAIM # (B)(4). Manufacturer Narrative: PLEASE DISREGARD PREVIOUS STATEMENT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5611457&pc=MTA | 2023826-2016-00544 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-27 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain | P030016 | 1 | Event Description: THE PATIENT REPORTED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN HER LEFT EYE (OS) ON (B)(6) 2013. THE PATIENT REPORTED SHARP PAIN AND PRESSURE. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS REMAINS IMPLANTED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319134&pc=MTA | 2023826-2019-02166 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Optical Problem | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | NA | 1 | Event Description: AN ARTICLE WAS RECEIVED THAT CITED A CASE STUDY THAT FOUND PATIENT IMPLANTED WITH A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00 DIOPTER IN THE PATIENTS LEFT EYE (OS). ON THE 1ST DAY POST-OP, THE PATIENT PRESENTED WITH EYE PAIN AND BLURRY VISION. ANTERIOR DISPLACEMENT OF LENS-IRIS DIAPHRAGM WAS NOTED ON SLIT-LAMP EXAMINATION WITH AN AC DEPTH APPROXIMATING 1 CORNEAL THICKNESS. A CLOSED ANGLE WAS NOTED, IN ADDITION TO HIGH INTRAOCULAR PRESSURE (IOP). THE IRIDOTOMY WAS PATENT AND THERE WERE NO SIGNS OF PUPILLARY BLOCK. A DIAGNOSIS OF MALIGNANT GLAUCOMA WAS MADE. MEDICAL TREATMENT (COMBINATION OF TOPICAL BRIMONIDINE/TIMOLOL, BIMATOPROST, BRINZOLAMIDE AND ATROPINE, IN ADDITION TO ORAL ACETAZOLAMIDE) FAILED OVER THE FOLLOWING 5 DAYS TO NORMALIZE THE PRESSURE AND REFORM THE AC. A POSTERIOR VITRECTOMY WAS PERFORMED, LEADING TO RESOLUTION OF THE AQUEOUS MISDIRECTION AND NORMALIZATION OF IOP. THE PATIENT ACHIEVED 20/20 UNCORRECTED VISUAL ACUITY WITH EXCELLENT PLACEMENT OF ICL. HOWEVER, A SMALL SINGLE AREA OF IRIS ATROPHY SECONDARY TO IRIDOCORNEAL TOUCH ALONG WITH A SLIGHTLY DILATED PUPIL DEVELOPED, CAUSING MILD HALOS AND GLARE UNDER MESOPIC CONDITIONS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE: 2923 IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6430437&pc=MTA | 2023826-2017-00461 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-23 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.0 DIOPTER, IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING AND ELEVATED INTRAOCULAR PRESSURE. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2017 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND THE LENS RETURNED WITH STICKY RESIDUE ON THE LENS AND UNABLE TO EVALUATE DUE TO RESIDUE. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6468284&pc=MTA | 2023826-2017-00598 |  | Injury | STAAR SURGICAL COMPANY | 2017-04-06 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -12.5 DIOPTER, IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED BLURRED VISION AND PAIN. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (THE EVENTS WERE NOT REPORTED ON THE INITIAL MDR). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING, ELEVATED INTRAOCULAR PRESSURE (IOP) AND ANGLE CLOSURE. (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574198&pc=MTA | 2023826-2017-00833 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE; ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF LENS OPTIC AND HAPTIC TORN OFF AND MISSING. A PIECE OF THE OTHER HAPTIC WAS TORN OFF AND MISSING. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6574253&pc=MTA | 2023826-2017-00834 |  | Injury | STAAR SURGICAL COMPANY | 2017-05-17 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -10.0 DIOPTER; IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING; ELEVATED INTRAOCULAR PRESSURE; NARROWING OF THE ANGLE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED BLURRY VISION AND PAIN. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND ONE HAPTIC TORN, WITH A PIECE TORN OFF AND MISSING. THERE WAS EVIDENCE OF CLEAR RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED A SHORTER LENS (13.2MM) WAS IMPLANTED ON (B)(6) 2017 AND THE LENS EXCHANGE RESOLVED THE PROBLEM. THE LENS VAULT WAS GOOD AND THE INTRAOCULAR PRESSURE WAS NORMALIZED. THE PATIENT IS DOING GOOD, WITH A VISUAL ACUITY OF 20/20. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661682&pc=MTA | 2023826-2016-00729 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -6.5 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT CODE(S): (OPERATIONAL PROBLEM) : VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE. THERE WAS EVIDENCE OF DRY RESIDUE. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5661683&pc=MTA | 2023826-2016-00730 | 2016-04-19 | Injury | STAAR SURGICAL COMPANY | 2016-05-17 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 12.6MM MICL12.6 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO HIGH VAULT. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO ELEVATED INTRAOCULAR PRESSURE, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED PAIN AND BLURRY VISION. Manufacturer Narrative: EVENT PROBLEM AND EVALUATION CODES: METHOD CODE(S): (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW RESULT CODE(S): (OPERATIONAL PROBLEM) : THE PRODUCT FOR THIS EVENT WAS RETURNED BUT WAS UNABLE TO EVALUATE DUE TO THE LENS WAS STUCK TO THE LID OF THE CONTAINER THE LENS WAS RETURNED IN. THE LENS WAS RETURNED DRY. (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, THE DEVICES ASSOCIATED WITH THE WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, WAS MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MANUFACTURER: NO, LENS NOT RETURNED. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138530&pc=MTA | 2023826-2022-01231 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICM5_13.2; -13.0 DIOPTER IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "BIGGER LENS". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM #: (B)(4). Manufacturer Narrative: B5 - PATIENT EXPERIENCED PAIN AND ELEVATED INTRAOCULAR PRESSURE - SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7021353&pc=MTA | 2023826-2017-01672 | 2017-01-16 | Injury | STAAR SURGICAL COMPANY | 2017-11-10 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7, -15.5/1.0/105 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO EXCESSIVE VAULTING, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES, AND ELEVATED INTRAOCULAR PRESSURE. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT EXPERIENCED RIGHT EYE PAIN BEFORE THE REPLACEMENT SURGERY. (B)(4). Manufacturer Narrative: PMA 510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6872195&pc=MTA | 2023826-2017-01440 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-09-18 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | N/A | 1 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -18.00/+3.20/100 DIOPTER, IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, EYE PAIN, SHALLOW AC DEPTH, AND HIGH EYE PRESSURE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE PATIENT'S POST-OP ICL EXCHANGE IS STABLE. THE STATUS OF THE EYE IS QUITE. THE PATIENT HAS NO COMPLAINT OF PAIN OR OTHER DISCOMFORTS. THE IOP IS ALSO WITHIN NORMAL LIMITS. THE PATIENT'S BEST CORRECTED VISUAL ACUITY (BCVA) IS 20/28 AND UNCORRECTED VISUAL ACUITY (UCVA) IS 20/28. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6962654&pc=MTA | 2023826-2017-01609 | 2017-04-24 | Injury | STAAR SURGICAL COMPANY | 2017-10-19 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/073 DIOPTER, IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE PATIENT EXPERIENCED HIGH EYE PRESSURE AND PAIN. THE LENS WAS DISCARDED BY THE FACILITY. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/30. Manufacturer Narrative: THE REPORTER INDICATED THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2017, FOR A SHORTER LENS. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING WELL, THE VAULT IS SUFFICIENT AND THE INTRAOCULAR PRESSURE IS WITHIN NORMAL LIMITS. (B)(4). REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6832461&pc=MTA | 2023826-2017-01383 | 2017-08-02 | Injury | STAAR SURGICAL COMPANY | 2017-08-30 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, -10.5 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE REPORTER STATED THE PATIENT WAS IN PAIN AND HAD ELEVATED INTRAOCULAR PRESSURE (PRESSURE SPIKE). THE PERIPHERAL IRIDOTOMIES WERE CLOSED, WITH A SHALLOW CHAMBER AND ANGLE CLOSURE. THE SURGEON DECIDED TO EXPLANT THE LENS THE SAME DAY. A THIRD LENS WAS NOT IMPLANTED. THE LENS WAS DISCARDED. THE PATIENT IS DOING FINE. THIS LENS WAS THE EXCHANGE LENS FOR A LONGER LENS IMPLANTED - SEE MDR #2023826-2017-01382. Manufacturer Narrative: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611920&pc=MTA | 2023826-2019-00780 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI, WITH A ROUND OF DIAMOX, COMBIGAN, AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PI, YAG, AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, AND ADDITIONAL SUTURES WERE DONE ON (B)(6) 2018, PER REPORTER AS OF (B)(6) 2019 THE PATIENT HAS PAIN AND PHOTOPHOBIA. PER REPORTER AS OF (B)(6) 2019 THE PATIENT'S ACUITY IS CORRECTABLE TO 20/20 OU. REPORTEDLY "PATIENT IS NOT INTERESTED IN REPLACEMENT AT THIS TIME." Manufacturer Narrative: ENLARGEMENT OF THE PERIPHERAL IRIDOTOMIES WAS PERFORMED BY YAG LASER. ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS WAS PERFORMED. AN ADDITIONAL SUTURE WAS REQUIRED. THE ICL WAS EXPLANTED ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: 1 SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611924&pc=MTA | 2023826-2019-00781 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, MICL13.2, -9.0 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35 MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI, WITH A ROUND OF DIAMOX, COMBIGAN, AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PI, YAG, AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS, AND ADDITIONAL SUTURES WERE DONE ON (B)(6) 2018, PER REPORTER AS OF (B)(6) 2019 THE PATIENT HAS PAIN AND PHOTOPHOBIA. PER REPORTER AS OF (B)(6) 2019 THE PATIENT'S ACUITY IS CORRECTABLE TO 20/20 OU. REPORTEDLY "PATIENT IS NOT INTERESTED IN REPLACEMENT AT THIS TIME." Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM MICL 13.2 OF -9.0 DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE(OS) ON (B)(6) 2019. ANGLE CLOSURE WITH ELEVATED IOP (35MMHG) AND EXCESSIVE VAULT WAS OBSERVED. "AFTER AN ADDITIONAL LPI WITH A ROUND OF DIAMOX; COMBIGAN AND SIMBRINZA THE IOP DID RESPOND." ENLARGEMENT OF PERIPHERAL IRIDOTOMIES; YAG AND THE ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO/FLUIDS AND ADDITIONAL SUTURES WERE DONE AND LENS WAS EXPLANTED ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782966&pc=MTA | 2023826-2020-00405 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, HIGH PRESSURE SPIKE AND EYE PAIN. A SHORTER LENS WILL BE IMPLANTED WITHIN THE NEXT MONTH. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9782969&pc=MTA | 2023826-2020-00406 |  | Injury | STAAR SURGICAL COMPANY | 2020-03-03 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, - 8.50 DIOPTER, IN THE PATIENTS EYE ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS WAS TOO LARGE, HIGH PRESSURE SPIKE AND EYE PAIN. A SHORTER LENS WILL BE IMPLANTED WITHIN THE NEXT MONTH. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5596421&pc=MTA | 2023826-2016-00507 |  | Injury | STAAR SURGICAL COMPANY | 2016-04-21 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -14.0 DIOPTER, IN THE PATIENT'S EYE ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A PRESSURE SPIKE. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: ADDITIONAL INFORMATION RECEIVED - THE LENS WAS IMPLANTED AND REMOVED FROM THE RIGHT EYE (OD) DUE TO EXCESSIVE VAULTING. THE PATIENT EXPERIENCED NARROWING OF THE ANGLE, SHALLOWING OF THE ANTERIOR CHAMBER DEPTH AND A LITTLE PAIN. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY WAS 20/20. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS NOT THE DEVICE BUT WAS DUE TO LENS SIZE SELECTION (TOO LARGE). Manufacturer Narrative: METHOD: (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULT: (NO FAILURE DETECTED): BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CONCLUSION: (UNABLE TO CONFIRM COMPLAINT): BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, PRODUCT EVALUATION, MEDICAL REVIEW AND DEVICE HISTORY RECORD REVIEW, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: (B)(4). RESULT: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. CONCLUSION: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE PRODUCT EVALUATION, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7636103&pc=MTA | 2023826-2018-00953 | 2018-06-04 | Injury | STAAR SURGICAL COMPANY | 2018-06-25 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; No Known Impact Or Consequence To Patient; No Code Available | NA | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE REPORTER INDICATED THE LENS HAD AN EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2018 FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED INTRAOCULAR PRESSURE INCREASE (ASSOCIATED WITH PAIN) DUE TO A HYPERVAULT OF THE ICL LEADING TO A NARROW CHAMBER ANGLE AND A FLATTENING OF THE ANTERIOR CHAMBER. PIGMENT DISPERSION AND IRIS ATROPHY DUE TO A HYPERVAULT OF THE ICL. NARROW CHAMBER ANGLE, FLATTENING OF ANTERIOR CHAMBER. PIGMENT DISPERSION, IRIS ATROPHY. DEVICE HISTORY RECORD REVIEW - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5004906&pc=MTA | 2023826-2015-01083 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block | P030016 | 1 | Event Description: RECEIVED A COPY OF AN ARTICLE PUBLISHED IN "REVIEW OF OPTOMETRY", 07/15/2015 AND WRITTEN BY A PHYSICIAN. THE TOPIC OF THE ARTICLE WAS - "ARE SEVERE EYE PAIN AND ELEVATED INTRAOCULAR PRESSURE (IOP) COMMON SIDE EFFECTS OF A VISIAN IMPLANTABLE COLLAMER LENS (ICL) PROCEDURE?". THE PHYSICIAN REFERRED TO A RECENT CASE - "A (B)(6) CAUCASIAN FEMALE PATIENT PRESENTED COMPLAINING OF SEVERE EYE PAIN, WITH A HISTORY OF A VISIAN ICL IMPLANTED NINE DAYS PRIOR. THE PATIENT'S IOP AT THE TIME OF THE VISIT WAS 46MM HG. BOTH EYES WERE IMPLANTED, WITH IOLS DISPLAYING EXCESSIVE ANTERIOR VAULTING (300% OD, 500% OS). THE PATIENT HAD PATENT PIS AT 11:00 IN THE RIGHT EYE AND AT 1:00 IN THE LEFT. OCT IMAGING CONFIRMED THAT BOTH EYES SHOWED PUPILLARY BLOCK, AND THE LEFT EYE DEMONSTRATED SEVERE ANGLE NARROWING. THE PATIENT WAS TREATED WITH MEDICATION TO REDUCE IOP TO 38MM HG. ATTENDING DOCTORS ACHIEVED FURTHER REDUCTION IN IOP TO 15MM HG USING COMPRESSION PRESSURE WITH A FOUR-MIRROR GONIO LENS. THE POSTERIOR FORCE OF THE GONIO LENS MECHANICALLY OPENED THE IRIDOCORNEAL ANGLE, RESULTING IN RAPID LOWERING OF IOP. THE PATIENT WAS INSTRUCTED TO CONTINUE WITH THE MEDICATION UNTIL A FOLLOW-UP COULD BE SCHEDULED IN FOUR DAYS WITH HER SURGEON. THE SURGEON AGREED WITH THE PROPOSED TREATMENT PLAN, AND SCHEDULED AN ICL EXCHANGE WITH A SHORTER LENS UPON HER RETURN". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. THIS COMPLAINT IS FOR THE PATIENT'S LEFT EYE - SEE MFR. REPORT # 2023826-2015-01084 FOR THE RIGHT EYE. Manufacturer Narrative: LENS NOT RETURNED. (B)(4)
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5004907&pc=MTA | 2023826-2015-01084 |  | Injury | STAAR SURGICAL COMPANY | 2015-08-14 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block | P030016 | 1 | Event Description: RECEIVED A COPY OF AN ARTICLE PUBLISHED IN "REVIEW OF OPTOMETRY", 07/15/2015 AND WRITTEN BY A PHYSICIAN. THE TOPIC OF THE ARTICLE WAS - "ARE SEVERE EYE PAIN AND ELEVATED INTRAOCULAR PRESSURE (IOP) COMMON SIDE EFFECTS OF A VISIAN IMPLANTABLE COLLAMER LENS (ICL) PROCEDURE?". THE PHYSICIAN REFERRED TO A RECENT CASE - "A (B)(6) CAUCASIAN FEMALE PATIENT PRESENTED COMPLAINING OF SEVERE EYE PAIN, WITH A HISTORY OF A VISIAN ICL IMPLANTED NINE DAYS PRIOR. THE PATIENT'S IOP AT THE TIME OF THE VISIT WAS 46MM HG. BOTH EYES WERE IMPLANTED, WITH IOLS DISPLAYING EXCESSIVE ANTERIOR VAULTING (300% OD, 500% OS). THE PATIENT HAD PATENT PIS AT 11:00 IN THE RIGHT EYE AND AT 1:00 IN THE LEFT. OCT IMAGING CONFIRMED THAT BOTH EYES SHOWED PUPILLARY BLOCK, AND THE LEFT EYE DEMONSTRATED SEVERE ANGLE NARROWING. THE PATIENT WAS TREATED WITH MEDICATION TO REDUCE IOP TO 38MM HG. ATTENDING DOCTORS ACHIEVED FURTHER REDUCTION IN IOP TO 15MM HG USING COMPRESSION PRESSURE WITH A FOUR-MIRROR GONIO LENS. THE POSTERIOR FORCE OF THE GONIO LENS MECHANICALLY OPENED THE IRIDOCORNEAL ANGLE, RESULTING IN RAPID LOWERING OF IOP. THE PATIENT WAS INSTRUCTED TO CONTINUE WITH THE MEDICATION UNTIL A FOLLOW-UP COULD BE SCHEDULED IN FOUR DAYS WITH HER SURGEON. THE SURGEON AGREED WITH THE PROPOSED TREATMENT PLAN, AND SCHEDULED AN ICL EXCHANGE WITH A SHORTER LENS UPON HER RETURN." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. THIS COMPLAINT IS FOR THE PATIENT'S RIGHT EYE - SEE MFR. REPORT # 2023826-2015-01083 FOR THE LEFT EYE. (B)(4) Manufacturer Narrative: DATE OF BIRTH - UNK. EXPIRATION DATE - UNK. DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVENT PROBLEM AND EVALUATION: (NEW INCISION); (SECONDARY SURGERY). (VAULTING); (EXPLANTED). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762431&pc=MTA | 2023826-2017-01260 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, DIOPTER -11.0, INTO THE PATIENT'S EYE. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULTING, ELEVATED IOP, AND ANGLE CLOSURE. ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: ADDITIONAL SYMPTOMS INCLUDE: IRIDECTOMY, BLURRED VISION, PAIN, PUPIL BLOCK. LENS EXCHANGE WAS PERFORMED ON THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PROBLEM WAS RESOLVED WITH LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOM: IRIDECTOMY. _CORRECTED: IRIDECTOMY IS NOT A SYMPTOM, IT WAS PERFORMED PRIOR ICL IMPLANTATION SURGERY. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762437&pc=MTA | 2023826-2017-01262 | 2017-07-07 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO HIGH VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. PREVIOUS IN MDR FOLLOW-UP#1: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. ADDITIONAL SYMPTOMS: NARROWING AND CLOSURE OF ANGLES, PAIN, BLURRED VISION, PUPILLARY BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND PI WAS REVISED PRIOR TO EXPLANTATION. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE LEFT EYE (OS). (B)(4). Manufacturer Narrative: ADDITIONAL SYMPTOMS: PI REVISION, NARROWING AND CLOSURE OF ANGLES, IRIDECTOMY, PAIN, BLURRED VISION, PUPILLARY BLOCK. THE LENS WAS EXCHANGED FOR A SHORTER LENS, SURGERY OCCURRED IN THE RIGHT (OD) EYE. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND A PIECE OF THE HAPTIC MISSING. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2017. THE PROBLEM WAS RESOLVED. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6762432&pc=MTA | 2023826-2017-01261 | 2017-07-08 | Injury | STAAR SURGICAL COMPANY | 2017-08-02 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2017 THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS, DIOPTER -14.5, INTO THE PATIENT'S EYE. IT IS PLANNED TO REPLACE THE LENS WITH A SHORTER LENS DUE TO "A LOT" OF VAULTING AND PRESSURE SPIKE . ATTEMPTS TO OBTAIN FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. Manufacturer Narrative: UPDATED INFORMATION: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PROBLEM WAS RESOLVED WITH THE LENS EXCHANGE. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPES WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: PREVIOUS IN MDR FOLLOW-UP#1: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, IRIDECTOMY, PAIN , BLURRED VISION, YAG REVISION, AND PUPIL BLOCK. CORRECTED: THE LENS WAS EXCHANGED IN THE RIGHT (OD) EYE ON (B)(6) 2017 FOR A SHORTER LENS. AFTER THE LENS EXCHANGE THE IOP DECREASED TO 13 (PREVIOUSLY SPIKED AT 68). ADDITIONAL SYMPTOMS INCLUDE NARROWING AND ANGLE CLOSURE, PAIN, BLURRED VISION, AND PUPIL BLOCK. IRIDECTOMY WAS PERFORMED PRIOR LENS IMPLANT SURGERY AND THE SURGEON ALSO PERFORMED AC DECOMPRESSION AND YAG REVISION PRIOR TO EXPLANTATION. (B)(4). Manufacturer Narrative: (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4743979&pc=MTA | 2023826-2015-00474 | 2015-04-09 | Injury | STAAR SURGICAL COMPANY | 2015-05-01 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block; Blurred Vision; No Code Available | P030016 | 1 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM MICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING, NARROWING OF THE ANGLE, ANGLE CLOSURE AND SHALLOWING OF THE ANTERIOR CHAMBER. THE PATIENT EXPERIENCED ELEVATED INTRAOCULAR PRESSURE - 50MM HG (PUPILLARY BLOCK), PAIN AND BLURRY VISION. ANOTHER PERIPHERAL IRIDECTOMY WAS PERFORMED. A SHORTER LENS WAS IMPLANTED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/25, THE CHAMBER WAS SHALLOW AND THE EYE IRRITATED. Manufacturer Narrative: (B)(4). DEVICE EVALUATED BY MANUFACTURER? NO. LENS NOT RETURNED. EVALUATION METHOD: WORK ORDER SEARCH. RESULTS: A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS: (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: METHOD: MEDICAL REVIEW. RESULTS: PER MEDICAL REVIEW - GIVEN THE REPORTED INFORMATION, THE EVENT WAS CAUSED BY EXCESSIVE VAULTING OF THE ICL LENS. THE HIGH IOP WAS THE CONSEQUENCE OF PUPILLARY BLOCK CAUSED BY EXCESSIVE VAULTING OF THE LENS. THE LENS EXPLANTATION, EXCHANGE AND ADDITIONAL IRIDOCTOMY WERE NECESSARY TO PREVENT FURTHER DAMAGE. ACCORDING TO FMEA (FAILURE MODES AND EFFECT ANALYSIS) IT HAS BEEN DETERMINED THAT EXCESSIVE VAULTING IS A CONSEQUENCE OF A WRONG LENS USE FAILURE MODE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF THE WHITE TO WHITE MEASUREMENTS BASED UPON DIFFERENT TECHNIQUES UTILIZED, IMPROPER SULCUS TO SULCUS MEASUREMENT (IF UBM USED), AND PATIENT CONDITION; POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS IN AN INDIVIDUAL CASE; IRREGULAR CILIARY SULCUS OR CILIARY SULCUS CYST). CONCLUSIONS: BASED ON THE COMPLAINT HISTORY, WORK ORDER SEARCH, MEDICAL REVIEW AND THE EVALUATION OF THE RETURNED PRODUCT, A PROBABLE ROOT CAUSE OF EXCESSIVE VAULTING HAS BEEN DETERMINED TO BE RELATED TO THE INACCURACY OF THE WHITE TO WHITE MEASUREMENT OR A MISMATCH BETWEEN WHITE TO WHITE AND THE SULCUS TO SULCUS DIAMETER. (B)(4). Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS WAS RETURNED IN LIQUID. (B)(4). Manufacturer Narrative: METHOD (PROCESS EVALUATION): DEVICE HISTORY RECORD REVIEW. RESULTS (EVALUATION RESULT): A REVIEW OF THE DEVICE HISTORY RECORD WAS PERFORMED AND NOTHING WAS FOUND IN THE MANUFACTURING AND PACKAGING PROCESSES OF THIS LENS THAT COULD BE IDENTIFIED AS THE CAUSE OF THE COMPLAINT. AS REPORTED ON THE MEDICAL REVIEW, EXCESSIVE VAULTING IS A CONSEQUENCE OF WRONG LENS USE (I.E. IMPROPER WHITE TO WHITE MEASUREMENT, VARIABILITY OF WHITE TO WHITE MEASUREMENTS BASED ON TECHNIQUES, IMPROPER SULCUS TO SULCUS MEASUREMENT, AND PATIENT CONDITION, POOR CORRELATION OF WHITE TO WHITE MEASUREMENT AND LENGTH OF CILIARY SULCUS, IRREGULAR SULCUS OR CILIARY SULCUS CYST). (B)(4).
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=5804108&pc=MTA | 2023826-2016-00958 | 2016-06-02 | Injury | STAAR SURGICAL COMPANY | 2016-07-19 | MTA | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | Intraocular Pressure Increased; Pain; Pupillary Block; No Code Available | P030016 | 1 | Event Description: THE REPORTER STATED THE SURGEON IMPLANTED A 13.7MM MICL13.7 IMPLANTABLE COLLAMER LENS, -8.0 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO AN EXCESSIVE VAULT. THE PATIENT EXPERIENCED AN ELEVATED IOP, PUPILLARY BLOCK, POST-OP PAIN, AND CLOUDY VISUAL ACUITY (VA). THE REPORTER INDICATED THE EVENT WAS PRODUCT-RELATED AND THE DEVICE CONTRIBUTED TO THE EVENT. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE REPORTER STATED THE CAUSE OF THE EVENT WAS WHITE-TO-WHITE MISMEASUREMENT. THE PATIENT'S POST-OP BEST-CORRECTED VISUAL ACUITY (BCVA) WAS 20/25. Manufacturer Narrative: VISUAL INSPECTION OF THE RETURNED PRODUCT FOUND A PIECE OF ONE HAPTIC TORN OFF AND MISSING. THE LENS WAS RETURNED IN VIAL IN LIQUID. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: PT WEIGHT: UNK. DEVICE EVALUATED BY MANUFACTURER? NO - 81 (OTHER): LENS NOT RETURNED. (B)(4). LENS NOT RETURNED.
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7119246&pc=MTA | 2023826-2017-01886 | 2017-11-16 | Injury | STAAR SURGICAL COMPANY | 2017-12-14 | MTA | IMPLANTABLE COLLAMER LENS (ICL) | Device Operates Differently Than Expected; Adverse Event Without Identified Device or Use Problem | Intraocular Pressure Increased; Pain; Retinal Detachment; Vitreous Hemorrhage; Halo; Vitrectomy; No Code Available | P030016 | 1 | Event Description: IT WAS REPORTED THAT ON (B)(6) 2017 AN ICL WAS REMOVED DUE TO RETINAL DETACHMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BROUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. VITRECTOMY WAS ALSO PERFORMED. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). METHOD: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM MICL12.6 OF -12.0 DIOPTER IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE ON (B)(6) 2017. ON (B)(6) 2017 THE LENS (AND NATURAL LENS) WAS EXPLANTED DUE TO RETINAL TEAR LEADING TO ITS SUBSEQUENT DETACHMENT. LASER RETINOPEXY WAS PERFORMED FOLLOWED BY PNEUMATIC CRYOPEXY ONLY TO HAVE THE CURTAIN IN PATIENT'S VISION RETURN. VITRECTOMY PERFORMED LED TO VITREOUS HEMORRHAGE. A SECOND RETINAL TEAR WAS NOTICED LEADING TO A SECOND VITRECTOMY, SCLERAL BUCKLE AND "LENSECTOMY" OR EXPLANT OF THE MICL AND NATURAL LENS.THE PATIENT ALSO HAD ELEVATED IOP, NARROWING OF ANGLE, PAIN AND EXPERIENCED "SOME HALOS AND FLOATING LIGHTS". ADDITIONAL PERIPHERAL IRIDECTOMY BOUGHT THE ELEVATED IOP BACK TO PRE-OP LEVEL AND THE ENLARGEMENT OF PERIPHERAL IRIDOTOMY PERFORMED INCREASED DEPTH. THE SURGEON REPORTED THAT THE PATIENT IS "CURRENTLY APHAKIC WILL NEED LENS IMPLANT". (B)(4). Manufacturer Narrative: (B)(4) Manufacturer Narrative: (B)(4).